## Sheet1
MDR_REPORT_KEY | MDR_TEXT_KEY | TEXT_TYPE_CODE | FOI_TEXT_TRIM | strNode | strtaxonomy
297087 | 180907 | D | MFR RECEIVED A REPORT FROM A USER ALLEGING THAT THE SEAT PAN ON THE FRONT PORTION OF THE WHEELCHAIR CAUSED THE LARGE PRESSURE SORE ON THE POSTERIOR ASPECT OF THE RIGHT LEG JUST BELOW THE KNEE | Pain | Patient problem code
297087 | 180907 | D | MFR RECEIVED A REPORT FROM A USER ALLEGING THAT THE SEAT PAN ON THE FRONT PORTION OF THE WHEELCHAIR CAUSED THE LARGE PRESSURE SORE ON THE POSTERIOR ASPECT OF THE RIGHT LEG JUST BELOW THE KNEE | Pressure Sores/Ulcers | Patient problem code
297087 | 180907 | D | MFR RECEIVED A REPORT FROM A USER ALLEGING THAT THE SEAT PAN ON THE FRONT PORTION OF THE WHEELCHAIR CAUSED THE LARGE PRESSURE SORE ON THE POSTERIOR ASPECT OF THE RIGHT LEG JUST BELOW THE KNEE | Insufficient Information | Device problem code
303005 | 199853 | D | CONSUMER CALLED MFR TO COMPLAIN THAT THEIR SCOOTER TIPPED OVER ON THEM WHILE RIDING SIDWAYS ON A HILL CONSUMER WENT TO ER AND AFTER TESTS AND XRAYS WAS SENT HOME WITH NO INJURIES OR ABNORMALITIES | Insufficient Information | Device problem code
303005 | 199853 | D | CONSUMER CALLED MFR TO COMPLAIN THAT THEIR SCOOTER TIPPED OVER ON THEM WHILE RIDING SIDWAYS ON A HILL CONSUMER WENT TO ER AND AFTER TESTS AND XRAYS WAS SENT HOME WITH NO INJURIES OR ABNORMALITIES | No Consequences Or Impact To Patient | Patient problem code
422154 | 251870 | D | HOSPITAL REPORTED DURING GRAFTING OF CABG FOAMING AND FRANK AIR OBSERVED IN ARTERIAL LINE RESUSCITATIVE MEASURES WERE UNDERTAKEN BUT PT DIED MINUTES INTO THE PROCEDURE HOSPITAL ALSO REPORTED THAT ALARMS WERE NOT HEARD BY OR STAFF MEDICAL EXAMINER LISTED CAUSE OF DEATHAIR EMBOLISM TERUMO SERVICE AND PROFESSIONAL SERVICE PERSONNEL EVALUATED THE UNIT AND FOUND THE SYSTEM WORKED ACCORDING TO SPECIFICATION INCLUDING ALL SAFETY SYSTEMS AND THERE WAS NO EVIDENCE TO INDICATE THERE MAY HAVE BEEN A PREVIOUS MALFUNCTION DURING TERUMO EVALUATIONS THE SAFETY SYSTEMS AIR DETECTION AND LOW LEVEL DETECTION WERE CONNECTED PROPERLY TURNED ON AND FUCNTIONED ACCORDING TO SPECIFICATION DURING THE INITIAL TERUMO EVALUATION THE SERVICE REP WAS INFORMED BY THE HOSPITAL THAT THE LOW LEVEL ALARM SENSOR WAS NOT IS USE AT THE TIME OF THE INCIDENT AUDIBLE VOLUME SETTING FOR THE SAFETY SYSTEMS WAS FOUND IN MEDIUM SETTING THE USER FACILITY HAS NOT PROVIDED SPECIFIC DETAILS ABOUT SETUP OF SYSTEM INCLUDING AUDIBLE SETTING OR WHETHER SYSTEMS WERE TURNED ON AND TESTED OR ABOUT ANY CONCLUSION AS TO THE CUASE OF AIR IN ARTERIAL LINE | Adverse Event Without Identified Device or Use Problem | Device Problem Code
422154 | 251870 | D | HOSPITAL REPORTED DURING GRAFTING OF CABG FOAMING AND FRANK AIR OBSERVED IN ARTERIAL LINE RESUSCITATIVE MEASURES WERE UNDERTAKEN BUT PT DIED MINUTES INTO THE PROCEDURE HOSPITAL ALSO REPORTED THAT ALARMS WERE NOT HEARD BY OR STAFF MEDICAL EXAMINER LISTED CAUSE OF DEATHAIR EMBOLISM TERUMO SERVICE AND PROFESSIONAL SERVICE PERSONNEL EVALUATED THE UNIT AND FOUND THE SYSTEM WORKED ACCORDING TO SPECIFICATION INCLUDING ALL SAFETY SYSTEMS AND THERE WAS NO EVIDENCE TO INDICATE THERE MAY HAVE BEEN A PREVIOUS MALFUNCTION DURING TERUMO EVALUATIONS THE SAFETY SYSTEMS AIR DETECTION AND LOW LEVEL DETECTION WERE CONNECTED PROPERLY TURNED ON AND FUCNTIONED ACCORDING TO SPECIFICATION DURING THE INITIAL TERUMO EVALUATION THE SERVICE REP WAS INFORMED BY THE HOSPITAL THAT THE LOW LEVEL ALARM SENSOR WAS NOT IS USE AT THE TIME OF THE INCIDENT AUDIBLE VOLUME SETTING FOR THE SAFETY SYSTEMS WAS FOUND IN MEDIUM SETTING THE USER FACILITY HAS NOT PROVIDED SPECIFIC DETAILS ABOUT SETUP OF SYSTEM INCLUDING AUDIBLE SETTING OR WHETHER SYSTEMS WERE TURNED ON AND TESTED OR ABOUT ANY CONCLUSION AS TO THE CUASE OF AIR IN ARTERIAL LINE | Death | Patient problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Fever | Patient problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Hypotension | Patient problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Pain, Abdominal | Patient problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Pain, Neck | Patient problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Insufficient Information | Device problem code
405762 | 270790 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Headache | Patient problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Fever | Patient problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Pain, Neck | Patient problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Pain, Abdominal | Patient problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Headache | Patient problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Insufficient Information | Device problem code
405755 | 275889 | D | IT WAS REPORTED THAT A PT DIAGNOSED WITH THALASSEMIA MAJOR EXPERIENCED A SEVERE TRANSFUSION REACTION FEVER HEADACHE NECK PAIN BACK PAIN ABDOMINAL PAIN AND HYPOTENSION WHILE RECEIVING A RED CELL TRANSFUSION THROUGH A PALL TRANSFUSION FILTER IN THE THALASSEMIA CENTER THE PT WAS TRANSFERRED TO THE ICU WHERE AFTER APPROX SIX HOURS THE PT HAD STABILIZED AND WAS THEN DISCHARGED FROM THE HOSPITAL NO FURTHER SEQUELAE HAVE BEEN REPORTED THIS REPORT IS ONE OF SIMILAR REPORTS RECEIVED RECENTLY FROM THIS HEALTH CARE FACILITY | Hypotension | Patient problem code
485254 | 293998 | D | AN AORTIC CONNECTOR WAS UTILIZED TO ANASTOMOSE THE SVG TO RCA DURING CABG INTRAOPERATIVE TEE SHOWED NO ABNORMALITIES THE PT PRESENTED SIX DAYS POSTOPERATIVELY WITH CHEST PAIN TIA AND HORNERS SYNDROME REDO CABG WAS PERFORMED SECONDARY TO A DISSECTING AORTA THE SURGEON INDICATED THE DISSECTION APPEARED TO ORIGINATE FROM THE CONNECTOR IT WAS ALSO REPORTED THE PT HAD A THINWALLED DILATED AORTA THE PT DID WELL AND WAS DISCHARGED TO REHAB | Chest Pain | Patient problem code
485254 | 293998 | D | AN AORTIC CONNECTOR WAS UTILIZED TO ANASTOMOSE THE SVG TO RCA DURING CABG INTRAOPERATIVE TEE SHOWED NO ABNORMALITIES THE PT PRESENTED SIX DAYS POSTOPERATIVELY WITH CHEST PAIN TIA AND HORNERS SYNDROME REDO CABG WAS PERFORMED SECONDARY TO A DISSECTING AORTA THE SURGEON INDICATED THE DISSECTION APPEARED TO ORIGINATE FROM THE CONNECTOR IT WAS ALSO REPORTED THE PT HAD A THINWALLED DILATED AORTA THE PT DID WELL AND WAS DISCHARGED TO REHAB | Insufficient Information | Device problem code
485254 | 293998 | D | AN AORTIC CONNECTOR WAS UTILIZED TO ANASTOMOSE THE SVG TO RCA DURING CABG INTRAOPERATIVE TEE SHOWED NO ABNORMALITIES THE PT PRESENTED SIX DAYS POSTOPERATIVELY WITH CHEST PAIN TIA AND HORNERS SYNDROME REDO CABG WAS PERFORMED SECONDARY TO A DISSECTING AORTA THE SURGEON INDICATED THE DISSECTION APPEARED TO ORIGINATE FROM THE CONNECTOR IT WAS ALSO REPORTED THE PT HAD A THINWALLED DILATED AORTA THE PT DID WELL AND WAS DISCHARGED TO REHAB | Dissection | Patient problem code
485254 | 293998 | D | AN AORTIC CONNECTOR WAS UTILIZED TO ANASTOMOSE THE SVG TO RCA DURING CABG INTRAOPERATIVE TEE SHOWED NO ABNORMALITIES THE PT PRESENTED SIX DAYS POSTOPERATIVELY WITH CHEST PAIN TIA AND HORNERS SYNDROME REDO CABG WAS PERFORMED SECONDARY TO A DISSECTING AORTA THE SURGEON INDICATED THE DISSECTION APPEARED TO ORIGINATE FROM THE CONNECTOR IT WAS ALSO REPORTED THE PT HAD A THINWALLED DILATED AORTA THE PT DID WELL AND WAS DISCHARGED TO REHAB | No Code Available | Patient problem code
628725 | 434726 | D | A DIAPHORETIC PATIENT SUSTAINED BURNS AT FOUR SITES UNDER FOUR ECG ELECTRODES DURING AN MRI PROCEDURE THE ECG ELECTRODES WERE USED TO MONITOR THE PATIENT DURING THE MRI IT WAS REPORTED THAT THE SEVERITY OF THE BURNS VARIED FROM RD TO TH DEGREE NO TREATMENT INFORMATION HAS BEEN REPORTED TO M THE PATIENT HAD A HIGH FEVER WAS SWEATING AND UNCONCIOUS AT THE TIME OF THE MRI PROCEDURE IT WAS REPORTED THAT THE MRI READING WAS NORMAL AND THAT THE MRI LASTED APPROXIMATELY MINUTES | Sweating | Patient problem code
628725 | 434726 | D | A DIAPHORETIC PATIENT SUSTAINED BURNS AT FOUR SITES UNDER FOUR ECG ELECTRODES DURING AN MRI PROCEDURE THE ECG ELECTRODES WERE USED TO MONITOR THE PATIENT DURING THE MRI IT WAS REPORTED THAT THE SEVERITY OF THE BURNS VARIED FROM RD TO TH DEGREE NO TREATMENT INFORMATION HAS BEEN REPORTED TO M THE PATIENT HAD A HIGH FEVER WAS SWEATING AND UNCONCIOUS AT THE TIME OF THE MRI PROCEDURE IT WAS REPORTED THAT THE MRI READING WAS NORMAL AND THAT THE MRI LASTED APPROXIMATELY MINUTES | Adverse Event Without Identified Device or Use Problem | Device problem code
628725 | 434726 | D | A DIAPHORETIC PATIENT SUSTAINED BURNS AT FOUR SITES UNDER FOUR ECG ELECTRODES DURING AN MRI PROCEDURE THE ECG ELECTRODES WERE USED TO MONITOR THE PATIENT DURING THE MRI IT WAS REPORTED THAT THE SEVERITY OF THE BURNS VARIED FROM RD TO TH DEGREE NO TREATMENT INFORMATION HAS BEEN REPORTED TO M THE PATIENT HAD A HIGH FEVER WAS SWEATING AND UNCONCIOUS AT THE TIME OF THE MRI PROCEDURE IT WAS REPORTED THAT THE MRI READING WAS NORMAL AND THAT THE MRI LASTED APPROXIMATELY MINUTES | Fever | Patient problem code
628725 | 434726 | D | A DIAPHORETIC PATIENT SUSTAINED BURNS AT FOUR SITES UNDER FOUR ECG ELECTRODES DURING AN MRI PROCEDURE THE ECG ELECTRODES WERE USED TO MONITOR THE PATIENT DURING THE MRI IT WAS REPORTED THAT THE SEVERITY OF THE BURNS VARIED FROM RD TO TH DEGREE NO TREATMENT INFORMATION HAS BEEN REPORTED TO M THE PATIENT HAD A HIGH FEVER WAS SWEATING AND UNCONCIOUS AT THE TIME OF THE MRI PROCEDURE IT WAS REPORTED THAT THE MRI READING WAS NORMAL AND THAT THE MRI LASTED APPROXIMATELY MINUTES | Third-Degree Burn | Patient problem code
816854 | 568362 | D | EVENT DESCRIPTION BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE AND SUSTAINED A FALL A HOLTER STUDY REVEALED TWO PACING PAUSES WHICH WERE THE RESULT OF A THRESHOLD TEST BEING PERFORMED WHILE THE HOLTER WAS COLLECTING DATA HOWEVER THE CAUSE OF THE SYNCOPAL EVENT IS UNKNOWN AT THIS TIME DEVICE TESTING NOTED NORMAL SENSING AND PACING AND AN XRAY WAS UNREMARKABLE FOR ANY LEAD ANOMALIES THE PHYSICIAN WILL CONTINUE TO PERFORM REGULAR DEVICE FOLLOW UPS | Adverse Event Without Identified Device or Use Problem | Device problem code
816854 | 568362 | D | EVENT DESCRIPTION BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE AND SUSTAINED A FALL A HOLTER STUDY REVEALED TWO PACING PAUSES WHICH WERE THE RESULT OF A THRESHOLD TEST BEING PERFORMED WHILE THE HOLTER WAS COLLECTING DATA HOWEVER THE CAUSE OF THE SYNCOPAL EVENT IS UNKNOWN AT THIS TIME DEVICE TESTING NOTED NORMAL SENSING AND PACING AND AN XRAY WAS UNREMARKABLE FOR ANY LEAD ANOMALIES THE PHYSICIAN WILL CONTINUE TO PERFORM REGULAR DEVICE FOLLOW UPS | Syncope | Patient problem code
816854 | 568362 | D | EVENT DESCRIPTION BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE AND SUSTAINED A FALL A HOLTER STUDY REVEALED TWO PACING PAUSES WHICH WERE THE RESULT OF A THRESHOLD TEST BEING PERFORMED WHILE THE HOLTER WAS COLLECTING DATA HOWEVER THE CAUSE OF THE SYNCOPAL EVENT IS UNKNOWN AT THIS TIME DEVICE TESTING NOTED NORMAL SENSING AND PACING AND AN XRAY WAS UNREMARKABLE FOR ANY LEAD ANOMALIES THE PHYSICIAN WILL CONTINUE TO PERFORM REGULAR DEVICE FOLLOW UPS | Fall | Patient problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Inflammation | Patient problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Insufficient Information | Device problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Adhesion(s) | Patient problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Obstruction | Patient problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Hernia | Patient problem code
954978 | 771453 | D | HERNIA REPAIR SURGERY IN WITH BARD MESH MONOFILAMENT KNITTED POLYPROPYLENE LOT IJD AND LOT ELD HAVE EXPERIENCED CONSTANT PAIN COMPLICATION SUCH AS BOWEL OBSTRUCTIONS ADHESIONS AND SERIOUS INFLAMMATIONS HAVE HAD MULTIPLE SURGERIES TO REMOVE INGUINAL NERVE GALLBLADDER AND NERVE RESECTIONS NOW IN SERIOUS NEED OF ADDITIONAL SURGERY FOR ADDITIONAL CHRONIC LOWER RIGHT ABDOMINAL PAIN AND MORE ADHESIONS IN THE GALLBLADDER AREA I NEED TO KNOW IF THERE ARE ANY STUDIES IN THE USA REGARDING THIS TYPE OF MESH IF THERE ARE ANY RECALLS IN PROGRESS ALL HELP AND INFO IS APPRECIATED I AM SERIOUSLY CHRONICALLY ILL AND HAVE NEVER EXPERIENCED IN PAIN RELIEF FROM ANY OF THE ABOVE MENTIONED SURGERIES I TAKE MULTIPLE DIFFERENT PAIN MEDICATIONS AND HAVE BECAME INDIGENT FROM ALL THE HOSP AND DOCTORS BILLS AND AM UNABLE TO WORK IN MY CAREER AS A PROFESSIONAL IN THE MORTGAGE INDUSTRY I AM IN SERIOUS NEED OF HELP DATES OF USE DIAGNOSIS OR REASON FOR USE HERNIA REPAIR HERNIA REPAIR EVENT REAPPEARED AFTER REINTRODUCTION YES YES | Pain | Patient problem code
1476065 | 1148655 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER IN B TO SUBMIT THIS EVENT THAT HAPPENED IN B PROBLEM PT HAD A MR PROCEDURE THE PT WAS POSITIONED SUPINE WITH THE HEAD FIRST INTO THE MAGNET SHE WAS PLANNED FOR A NECK EXAMINATION WITH THE SYNERGY HEADNECK COIL IMMEDIATELY AFTER THE EXAMINATION TWO SECOND DEGREE BURNS APPEARED AT THE INSIDE OF THE UPPER ARM AND THE BREAST | Second-Degree Burn | Patient problem code
1476065 | 1148655 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER IN B TO SUBMIT THIS EVENT THAT HAPPENED IN B PROBLEM PT HAD A MR PROCEDURE THE PT WAS POSITIONED SUPINE WITH THE HEAD FIRST INTO THE MAGNET SHE WAS PLANNED FOR A NECK EXAMINATION WITH THE SYNERGY HEADNECK COIL IMMEDIATELY AFTER THE EXAMINATION TWO SECOND DEGREE BURNS APPEARED AT THE INSIDE OF THE UPPER ARM AND THE BREAST | Insufficient Information | Device problem code
1440032 | 1190234 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER TO SUBMIT THIS EVENT THAT HAPPENED IN OTHER COUNTRY PROBLEM PT HAD A MR PROCEDURE THE PT WAS SCANNED ON AN INTERA T MR SYSTEM WITH NO ADDITIONAL RF COIL ATTACHED IMMEDIATELY AFTER THE EXAMINATION SECOND TO THIRD DEGREE BURNS APPEARED ON THE INSIDE OF HIS CALVES | Insufficient Information | Device problem code
1440032 | 1190234 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER TO SUBMIT THIS EVENT THAT HAPPENED IN OTHER COUNTRY PROBLEM PT HAD A MR PROCEDURE THE PT WAS SCANNED ON AN INTERA T MR SYSTEM WITH NO ADDITIONAL RF COIL ATTACHED IMMEDIATELY AFTER THE EXAMINATION SECOND TO THIRD DEGREE BURNS APPEARED ON THE INSIDE OF HIS CALVES | Third-Degree Burn | Patient problem code
1582690 | 1310201 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER IN B TO SUBMIT THIS EVENT THAT HAPPENED IN B PROBLEM THIS MALE PT HAD A MR PROCEDURE HE WAS LYING ON THE SYNERGY SPINE COIL WHICH WAS COVERED WITH A TWO CENTIMETER THICK COIL MAT HE WAS HEAVILY SWEATING WHILE WEARING A SHIRT AND BRIEFS DURING THE EXAMINATION THIS THUMB WAS CLOSE TO HIS HIP SKIN THE COIL CABLE WAS PROPERLY POSITIONED WITH NO LOOPS UNDER THE MAT THE PT DID NOT HAVE DIRECT CONTACT TO THE COIL OR CABLE IMMEDIATELY AFTER THE EXAMINATION THE PT WAS FOUND TO HAVE TWO SECOND DEGREE BURNS ONE WAS ON THE LEFT THUMB AND THE OTHER WAS LOCATED ON THE LEFT HIPBACKSIDE WITH A BLISTER SIZE OF APPROXIMATELY CM | Insufficient Information | Device problem code
1582690 | 1310201 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER IN B TO SUBMIT THIS EVENT THAT HAPPENED IN B PROBLEM THIS MALE PT HAD A MR PROCEDURE HE WAS LYING ON THE SYNERGY SPINE COIL WHICH WAS COVERED WITH A TWO CENTIMETER THICK COIL MAT HE WAS HEAVILY SWEATING WHILE WEARING A SHIRT AND BRIEFS DURING THE EXAMINATION THIS THUMB WAS CLOSE TO HIS HIP SKIN THE COIL CABLE WAS PROPERLY POSITIONED WITH NO LOOPS UNDER THE MAT THE PT DID NOT HAVE DIRECT CONTACT TO THE COIL OR CABLE IMMEDIATELY AFTER THE EXAMINATION THE PT WAS FOUND TO HAVE TWO SECOND DEGREE BURNS ONE WAS ON THE LEFT THUMB AND THE OTHER WAS LOCATED ON THE LEFT HIPBACKSIDE WITH A BLISTER SIZE OF APPROXIMATELY CM | Second-Degree Burn | Patient problem code
1582690 | 1310201 | D | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR MR MANUFACTURER IN B TO SUBMIT THIS EVENT THAT HAPPENED IN B PROBLEM THIS MALE PT HAD A MR PROCEDURE HE WAS LYING ON THE SYNERGY SPINE COIL WHICH WAS COVERED WITH A TWO CENTIMETER THICK COIL MAT HE WAS HEAVILY SWEATING WHILE WEARING A SHIRT AND BRIEFS DURING THE EXAMINATION THIS THUMB WAS CLOSE TO HIS HIP SKIN THE COIL CABLE WAS PROPERLY POSITIONED WITH NO LOOPS UNDER THE MAT THE PT DID NOT HAVE DIRECT CONTACT TO THE COIL OR CABLE IMMEDIATELY AFTER THE EXAMINATION THE PT WAS FOUND TO HAVE TWO SECOND DEGREE BURNS ONE WAS ON THE LEFT THUMB AND THE OTHER WAS LOCATED ON THE LEFT HIPBACKSIDE WITH A BLISTER SIZE OF APPROXIMATELY CM | Sweating | Patient problem code
1568197 | 1316033 | D | IT WAS REPORTED THAT A PT SUSTAINED SECOND DEGREE AND THIRD DEGREE BURNS ON HIS LEFT PALM AND LEFT HANDS FINGERS RESPECTIVELY AFTER AN MRI EXAM OF THE HEAD AND NECK AT THE START OF THE EXAM A PULSE OXIMETER WAS PLACED ON PTS LEFT INDEX FINGER WHILE BLOOD PRESSURE MONITOR WAS PLACED ON THE RIGHT ARM THE RISK MANAGER INDICATED THAT BOTH OF THE PTS ARMS WERE SECURED AT ELBOW WITH INCH WIDE STRAP DUE TO HIS FLACCID LEFT ARM THE STRAP WRAPPED ACROSS THE PTS CHEST UPON COMPLETION OF THE EXAM AND REMOVAL OF THE PULSE OXIMETER AND BLOOD PRESSURE CUFF THE PT WAS TRANSPORTED BACK TO THE CRITICAL CARE UNIT AT THIS TIME THE ATTENDING HEALTHCARE PROFESSIONAL DID NOT NOTICE ANY VISIBLE INJURY TO THE LEFT HAND APPROX MINUTES LATER THE MRI UNIT WAS NOTIFIED THAT THE PT HAD DEVELOPED MULTIPLE BLISTERS ON HIS LEFT HANDS PALM AND THE LEFT HANDS FINGERTIPS WITH AN EXCEPTION TO THE THUMB THE HOSPITAL ASSESSED THE BLISTER ON PALM AS SECOND DEGREE WHILE BLISTERS ON THE FINGERTIPS AS THIRD DEGREE THE THIRD DEGREE BURNS RESULTED IN PARTIAL AMPUTATION OF THE PTS FINGERS THE HOSPITAL CONTINUED TO USE THE MRI SCANNER WITH NO FURTHER INCIDENTS | Second-Degree Burn | Patient problem code
1568197 | 1316033 | D | IT WAS REPORTED THAT A PT SUSTAINED SECOND DEGREE AND THIRD DEGREE BURNS ON HIS LEFT PALM AND LEFT HANDS FINGERS RESPECTIVELY AFTER AN MRI EXAM OF THE HEAD AND NECK AT THE START OF THE EXAM A PULSE OXIMETER WAS PLACED ON PTS LEFT INDEX FINGER WHILE BLOOD PRESSURE MONITOR WAS PLACED ON THE RIGHT ARM THE RISK MANAGER INDICATED THAT BOTH OF THE PTS ARMS WERE SECURED AT ELBOW WITH INCH WIDE STRAP DUE TO HIS FLACCID LEFT ARM THE STRAP WRAPPED ACROSS THE PTS CHEST UPON COMPLETION OF THE EXAM AND REMOVAL OF THE PULSE OXIMETER AND BLOOD PRESSURE CUFF THE PT WAS TRANSPORTED BACK TO THE CRITICAL CARE UNIT AT THIS TIME THE ATTENDING HEALTHCARE PROFESSIONAL DID NOT NOTICE ANY VISIBLE INJURY TO THE LEFT HAND APPROX MINUTES LATER THE MRI UNIT WAS NOTIFIED THAT THE PT HAD DEVELOPED MULTIPLE BLISTERS ON HIS LEFT HANDS PALM AND THE LEFT HANDS FINGERTIPS WITH AN EXCEPTION TO THE THUMB THE HOSPITAL ASSESSED THE BLISTER ON PALM AS SECOND DEGREE WHILE BLISTERS ON THE FINGERTIPS AS THIRD DEGREE THE THIRD DEGREE BURNS RESULTED IN PARTIAL AMPUTATION OF THE PTS FINGERS THE HOSPITAL CONTINUED TO USE THE MRI SCANNER WITH NO FURTHER INCIDENTS | Adverse Event Without Identified Device or Use Problem | Device problem code
1873359 | 1831762 | D | IT WAS REPORTED THAT DURING A RIGHT SALPINGOOOPHORECTOMY PROCEDURE WHILE TRYING TO REMOVE A VERY LARGE OVARY THE POUCH TORE BY THE METAL RING A SECOND POUCH WAS PULLED AND WHILE TRYING TO PULL THE LARGE OVARY THROUGH THE HOLE THE POUCH TORETHE CONTENTS DID NOT FALL OUT OF THE POUCH THE SURGEON IS NOT SURE IF ANY OF THE POUCH FELL INTO THE PATIENT BUT WAS CALLING TO SEE WHAT WOULD HAPPEN IF THERE WAS A PIECE OF THE POUCH LEFT INTO THE PATIENT THEY CUT THE LARGE OVARY INTO SMALLER PIECES AND REMOVED THE OVARY WITH NO FURTHER PROBLEMS THE PATIENT WAS A VERY LARGE PERSON WITH A BMI OF OVER B THE CASE WAS COMPLETED WITH NO PATIENT CONSEQUENCE THE FIRST TWO POUCHES WERE DISCARDED | Device Handling Problem | Device problem code
1873359 | 1831762 | D | IT WAS REPORTED THAT DURING A RIGHT SALPINGOOOPHORECTOMY PROCEDURE WHILE TRYING TO REMOVE A VERY LARGE OVARY THE POUCH TORE BY THE METAL RING A SECOND POUCH WAS PULLED AND WHILE TRYING TO PULL THE LARGE OVARY THROUGH THE HOLE THE POUCH TORE THE CONTENTS DID NOT FALL OUT OF THE POUCH THE SURGEON IS NOT SURE IF ANY OF THE POUCH FELL INTO THE PATIENT BUT WAS CALLING TO SEE WHAT WOULD HAPPEN IF THERE WAS A PIECE OF THE POUCH LEFT INTO THE PATIENT THEY CUT THE LARGE OVARY INTO SMALLER PIECES AND REMOVED THE OVARY WITH NO FURTHER PROBLEMS THE PATIENT WAS A VERY LARGE PERSON WITH A BMI OF OVER B THE CASE WAS COMPLETED WITH NO PATIENT CONSEQUENCE THE FIRST TWO POUCHES WERE DISCARDED | No Consequences Or Impact To Patient | Patient problem code
2375151 | 2446917 | D | THE PATIENT WAS REVISED TO ADDRESS ELEVATED METAL ION LEVELS THE CUP WAS IN VERTICAL POSITION DEGREES PLUS | Toxicity | Patient problem code
2375151 | 2446917 | D | THE PATIENT WAS REVISED TO ADDRESS ELEVATED METAL ION LEVELS THE CUP WAS IN VERTICAL POSITION DEGREES PLUS | Insufficient Information | Device Problem Code
2504389 | 2600688 | D | LITIGATION PAPERS ALLEGE PATIENT BEGAN EXPERIENCING PAIN AND DISCOMFORT IN HIS HIP AND LEG IT IS FURTHER ALLEGED PATIENTS BLOOD TESTS REVEALED ELEVATED LEVELS OF CHROMIUM AND COBALT IN HIS BLOOD | Discomfort | Patient problem code
2504389 | 2600688 | D | LITIGATION PAPERS ALLEGE PATIENT BEGAN EXPERIENCING PAIN AND DISCOMFORT IN HIS HIP AND LEG IT IS FURTHER ALLEGED PATIENTS BLOOD TESTS REVEALED ELEVATED LEVELS OF CHROMIUM AND COBALT IN HIS BLOOD | Pain | Patient problem code
2504389 | 2600688 | D | LITIGATION PAPERS ALLEGE PATIENT BEGAN EXPERIENCING PAIN AND DISCOMFORT IN HIS HIP AND LEG IT IS FURTHER ALLEGED PATIENTS BLOOD TESTS REVEALED ELEVATED LEVELS OF CHROMIUM AND COBALT IN HIS BLOOD | Toxicity | Patient problem code
2504389 | 2600688 | D | LITIGATION PAPERS ALLEGE PATIENT BEGAN EXPERIENCING PAIN AND DISCOMFORT IN HIS HIP AND LEG IT IS FURTHER ALLEGED PATIENTS BLOOD TESTS REVEALED ELEVATED LEVELS OF CHROMIUM AND COBALT IN HIS BLOOD | Insufficient Information | Device Problem Code
2740660 | 2895870 | D | PATIENT WAS REVISED TO ADDRESS PAIN SWELLING AND LARGE QUANTITIES OF PURULENT FLUID IT WAS REPORTED THAT THE PATIENT WAS NOT INFECTED | Pain | Patient problem code
2740660 | 2895870 | D | PATIENT WAS REVISED TO ADDRESS PAIN SWELLING AND LARGE QUANTITIES OF PURULENT FLUID IT WAS REPORTED THAT THE PATIENT WAS NOT INFECTED | Purulent Discharge | Patient problem code
2740660 | 2895870 | D | PATIENT WAS REVISED TO ADDRESS PAIN SWELLING AND LARGE QUANTITIES OF PURULENT FLUID IT WAS REPORTED THAT THE PATIENT WAS NOT INFECTED | Swelling | Patient problem code
2740660 | 2895870 | D | PATIENT WAS REVISED TO ADDRESS PAIN SWELLING AND LARGE QUANTITIES OF PURULENT FLUID IT WAS REPORTED THAT THE PATIENT WAS NOT INFECTED | Insufficient Information | Device Problem Code
2802758 | 3088427 | D | THE PATIENT WAS REVISED BECAUSE OF PAIN LEFT HIP | Pain | Patient problem code
2802758 | 3088427 | D | THE PATIENT WAS REVISED BECAUSE OF PAIN LEFT HIP | Insufficient Information | Device Problem Code
3052046 | 3454003 | D | UPDATE B PATIENTS MEDICAL RECORDS WERE RECEIVED RECORDS INDICATE THE PATIENT WAS ALSO REVISED TO ADDRESS INCREASED METAL ION LEVELS UPON REVISION LARGE VOLUMES OF TISSUE CAPSULE AND ABDUCTOR MUSCULATURE WAS REMOVED HAVING BEEN INVOLVED WITH METAL ION DISEASE | Toxicity | Patient problem code
3052046 | 3454003 | D | UPDATE B PATIENTS MEDICAL RECORDS WERE RECEIVED RECORDS INDICATE THE PATIENT WAS ALSO REVISED TO ADDRESS INCREASED METAL ION LEVELS UPON REVISION LARGE VOLUMES OF TISSUE CAPSULE AND ABDUCTOR MUSCULATURE WAS REMOVED HAVING BEEN INVOLVED WITH METAL ION DISEASE | Insufficient Information | Device Problem Code
3251076 | 3666513 | D | ON B PHYSICIAN DID RECOLLECT THE INCIDENT HE DID NOT RECALL SPECIFICS HE DID NOT RECALL THE SURROUNDING HEALTH OR THE LOCATION OF THE WOUND HE DID RECALL THE PT DID NOT HAVE A LASTING INJURY PT EXPERIENCED BURNING PAIN DURING MRI BUT STOPPED WHEN SCAN WAS DONE THE LEVEL OF RESONANCE BEING USED IS NOT DOCUMENTED THE FACILITY THAT INITIATED REPORT DID NOT DESCRIBE THE PROTOCOL FOR TESTING SURROUNDING THERAPYIES WERE NOT DESCRIBED B | Insufficient Information | Device problem code
3251076 | 3666513 | D | ON B PHYSICIAN DID RECOLLECT THE INCIDENT HE DID NOT RECALL SPECIFICS HE DID NOT RECALL THE SURROUNDING HEALTH OR THE LOCATION OF THE WOUND HE DID RECALL THE PT DID NOT HAVE A LASTING INJURY PT EXPERIENCED BURNING PAIN DURING MRI BUT STOPPED WHEN SCAN WAS DONE THE LEVEL OF RESONANCE BEING USED IS NOT DOCUMENTED THE FACILITY THAT INITIATED REPORT DID NOT DESCRIBE THE PROTOCOL FOR TESTING SURROUNDING THERAPYIES WERE NOT DESCRIBED B | Burning Sensation | Patient problem code
3251076 | 3666513 | D | ON B PHYSICIAN DID RECOLLECT THE INCIDENT HE DID NOT RECALL SPECIFICS HE DID NOT RECALL THE SURROUNDING HEALTH OR THE LOCATION OF THE WOUND HE DID RECALL THE PT DID NOT HAVE A LASTING INJURY PT EXPERIENCED BURNING PAIN DURING MRI BUT STOPPED WHEN SCAN WAS DONE THE LEVEL OF RESONANCE BEING USED IS NOT DOCUMENTED THE FACILITY THAT INITIATED REPORT DID NOT DESCRIBE THE PROTOCOL FOR TESTING SURROUNDING THERAPYIES WERE NOT DESCRIBED B | Pain | Patient problem code
3386238 | 3798068 | D | LITIGATION ALLEGES PAIN AND DIFFICULTY AMBULATING LITIGATION ALSO ALLEGES PARAESTHESIA OF THE LOWER EXTREMITIES | Ambulation Difficulties | Patient problem code
3386238 | 3798068 | D | LITIGATION ALLEGES PAIN AND DIFFICULTY AMBULATING LITIGATION ALSO ALLEGES PARAESTHESIA OF THE LOWER EXTREMITIES | Pain | Patient problem code
3386238 | 3798068 | D | LITIGATION ALLEGES PAIN AND DIFFICULTY AMBULATING LITIGATION ALSO ALLEGES PARAESTHESIA OF THE LOWER EXTREMITIES | Insufficient Information | Device Problem Code
3808519 | 5302217 | D | UPDATE RECEIVED THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND ELEVATED CHROMIUM LEVELSCOMPLAINT WAS UPDATED ON | Pain | Patient problem code
3808519 | 5302217 | D | UPDATE RECEIVED THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND ELEVATED CHROMIUM LEVELSCOMPLAINT WAS UPDATED ON | Toxicity | Patient problem code
3808519 | 5302217 | D | UPDATE RECEIVED THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND ELEVATED CHROMIUM LEVELSCOMPLAINT WAS UPDATED ON | Insufficient Information | Device problem code
4870943 | 6060199 | D | UPDATE PFS AND MEDICAL RECORDS RECEIVED AFTER REVIEW OF THE MEDICAL RECORDS FOR MDR REPORTABILITY THE REVISION OPERATIVE NOTE INDICATED INCREASED METAL IONS NO LABS PAIN AND METAL ON METAL REACTION THE FEMORAL STEM IS BEING ADDED THE COMPLAINT WAS UPDATED ON | Pain | Patient problem code
4870943 | 6060199 | D | UPDATE PFS AND MEDICAL RECORDS RECEIVED AFTER REVIEW OF THE MEDICAL RECORDS FOR MDR REPORTABILITY THE REVISION OPERATIVE NOTE INDICATED INCREASED METAL IONS NO LABS PAIN AND METAL ON METAL REACTION THE FEMORAL STEM IS BEING ADDED THE COMPLAINT WAS UPDATED ON | Insufficient Information | Device Problem Code
926821 | 7861432 | N | DISCUSSIONS BETWEEN A SIEMENS REPRESENTATIVE AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM THE SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Difficult to Remove | Device problem code
926821 | 7861432 | N | DISCUSSIONS BETWEEN A SIEMENS REPRESENTATIVE AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM THE SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Mechanical Jam | Device problem code
926821 | 7861432 | N | DISCUSSIONS BETWEEN A SIEMENS REPRESENTATIVE AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM THE SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Mechanical Problem Identified | Medical Evaluation result code
864219 | 7871232 | N | NARRATIVE DISCUSSIONS BETWEEN A SIEMENS REP AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Mechanical Jam | Device problem code
864219 | 7871232 | N | NARRATIVE DISCUSSIONS BETWEEN A SIEMENS REP AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Difficult to Remove | Device problem code
864219 | 7871232 | N | NARRATIVE DISCUSSIONS BETWEEN A SIEMENS REP AND THE CUSTOMER INDICATED THAT SHE WAS USING THE PROPER TECHNIQUE TO REMOVE THE REAGENT CARTRIDGE BUT IT WAS STUCK IN THE INSTRUMENT THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS DIAGNOSTICS FOR FURTHER EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Mechanical Problem Identified | Medical Evaluation result code
1027255 | 7994274 | N | NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED ANALYSIS OF A SIMILAR DEVICE FROM THE SAME LOT SHOWED IMPRECISION THAT MAY CAUSE FALSELY ELEVATED RESULTS ELEVATIONS UP TO NGML UGL HAVE BEEN CONFIRMED BY INTERNAL TESTING AT CONCENTRATIONS BETWEEN AND NGML UGL WHICH HAVE POTENTIAL IMPACT ON THE MONITORING OF POSTPROSTATECTOMY PTS THE AFFECTED LOT DEMONSTRATES SPORADIC VARIABILITY IN RECOVERY OF CALIBRATOR AND QUALITY CONTROL QC THEREFORE THIS ISSUE MAY NOT BE READILY DETECTED UPON CALIBRATION OR DURING ROUTINE PROCESSING OF QC CUSTOMERS HAVE BEEN INSTRUCTED TO DISCONTINUE USE OF THIS LOT | Calibration Problem | Device problem code
1027255 | 7994274 | N | NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED ANALYSIS OF A SIMILAR DEVICE FROM THE SAME LOT SHOWED IMPRECISION THAT MAY CAUSE FALSELY ELEVATED RESULTS ELEVATIONS UP TO NGML UGL HAVE BEEN CONFIRMED BY INTERNAL TESTING AT CONCENTRATIONS BETWEEN AND NGML UGL WHICH HAVE POTENTIAL IMPACT ON THE MONITORING OF POSTPROSTATECTOMY PTS THE AFFECTED LOT DEMONSTRATES SPORADIC VARIABILITY IN RECOVERY OF CALIBRATOR AND QUALITY CONTROL QC THEREFORE THIS ISSUE MAY NOT BE READILY DETECTED UPON CALIBRATION OR DURING ROUTINE PROCESSING OF QC CUSTOMERS HAVE BEEN INSTRUCTED TO DISCONTINUE USE OF THIS LOT | Imprecision | Device problem code
1027255 | 7994274 | N | NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED ANALYSIS OF A SIMILAR DEVICE FROM THE SAME LOT SHOWED IMPRECISION THAT MAY CAUSE FALSELY ELEVATED RESULTS ELEVATIONS UP TO NGML UGL HAVE BEEN CONFIRMED BY INTERNAL TESTING AT CONCENTRATIONS BETWEEN AND NGML UGL WHICH HAVE POTENTIAL IMPACT ON THE MONITORING OF POSTPROSTATECTOMY PTS THE AFFECTED LOT DEMONSTRATES SPORADIC VARIABILITY IN RECOVERY OF CALIBRATOR AND QUALITY CONTROL QC THEREFORE THIS ISSUE MAY NOT BE READILY DETECTED UPON CALIBRATION OR DURING ROUTINE PROCESSING OF QC CUSTOMERS HAVE BEEN INSTRUCTED TO DISCONTINUE USE OF THIS LOT | Failure to Calibrate | Medical Evaluation result code
1084906 | 8106260 | N | THERE WAS NO EVIDENCE OF SYSTEM MALFUNCTION ACCORDING TO THE FACILITY HEARING PROTECTION WAS NOT USED DURING THE SCAN THE SYSTEM IS DESIGNED TO COMPLY WITH IEC AND OSHA STANDARDS FOR ACOUSTIC LIMITS SINCE ACOUSTIC LEVELS OF THE SYSTEM MAY EXCEED DBA IN THE BORE HEARING PROTECTION IS REQUIRED FOR ALL PEOPLE IN THE MAGNET ROOM DURING A SCAN THE SYSTEMS OPERATOR MANUAL PROVIDES A WARNING OF THE ACOUSTIC LEVELS AS WELL AS INSTRUCTION TO USE ADEQUATE HEARING PROTECTION DURING A SCAN | No Device Problem Found | Medical Evaluation result code
1118753 | 8106487 | N | ON ADDITIONAL INFO REGARDING THE EVENT WAS PROVIDED BY THE PTS SIBLING IT WAS REPORTED THAT AT THE END OF THE DA VINCI S HYSTERECTOMY SURGICAL PROCEDURE THE SURGICAL STAFF INSERTED A GYNECARE XTRACT MORCELLATOR INTO THE PT HOWEVER IT WAS NOT SECURELY ATTACHED WHEN THE SURGEON TRIED TO ATTACH THE MORCELLATOR INTRAABDOMINALLY THE SURGEON LOST CONTROL OF THE MORCELLATOR AND INJURED THE PTS AORTA INFERIOR VENA CAVA AND COMMON ILIAC ARTERY CIA AS A RESULT THE PT NEEDED A VEIN GRAFT FOR THE CIA REPAIR ON ADDITIONAL INFO WAS PROVIDED BY AN ISI REP WHO REPORTED THAT DURING THE TIME OF THE EVENT A SYSTEM ARM WAS UNDOCKED IN ORDER FOR THE SURGICAL STAFF TO INSERT THE MORCELLATOR THROUGH THAT PORT SITE AND ONLY THE SYSTEM CAMERA WAS BEING USED WHEN THE INJURY OCCURRED BASED ON THE INFO PROVIDED IT WAS DETERMINED THAT THE DA VINCI S SURGICAL SYSTEM DID NOT MALFUNCTION IN A WAY THAT WOULD CAUSE NOR CONTRIBUTE TO THE PT INJURY | No Device Problem Found | Medical Evaluation result code
1084907 | 8107405 | N | THE GE FIELD ENGINEER FE INSPECTED THE SYSTEM AND FOUND IT TO BE OPERATIONAL ACCORDING TO THE SITE HEARING PROTECTION WAS NOT USED DURING THE EXAM THE SYSTEM IS DESIGNED TO COMPLY WITH IEC AND OSHA STANDARDS FOR ACOUSTIC LIMITS SINCE ACOUSTIC LEVELS OF THE SYSTEM MAY EXCEED DBA IN THE BORE HEARING PROTECTION IS REQUIRED FOR ALL PEOPLE IN THE MAGNET ROOM DURING A SCAN THE SYSTEMS OPERATOR MANUAL PROVIDES A WARNING OF THE ACOUSTIC LEVELS AS WELL AS INSTRUCTION TO USE ADEQUATE HEARING PROTECTION DURING A SCAN ATTEMPTS WERE MADE TO OBTAIN INFORMATION REGARDING THE STATUS OF THE PATIENT HOWEVER THE SITE INDICATED THAT THEY HAD NOT HEARD FROM THE PATIENT SINCE RECEIVING THE REPORT ADDITIONALLY GE MADE SEVERAL ATTEMPTS TO OBTAIN FURTHER DETAILS OF THE EVENT SUCH AS THE SCAN LENGTH AT THIS TIME NO INFORMATION HAVE BEEN RECEIVED FROM THE SITE | No Findings Available | Medical Evaluation result code
1334527 | 8184660 | N | SINCE THE INCIDENT OCCURRED IN AN ACOUSTIC TEST WILL NOT BE PERFORMED ON THE SYSTEM NOW SINCE THE RESULTS WOULD NOT BE INDICATIVE OF THE STATUS OF THE SYSTEM AT THE TIME OF THE INCIDENT A REVIEW OF THE COMPLAINTS FROM THE MONTHS SURROUNDING THE DATE OF THIS INCIDENT WAS DONE AND THERE WERE NO OTHER TINNITUS COMPLAINTS REPORTED FOR THIS SYSTEM A REVIEW OF THE SERVICE DISPATCHES FROM THE MONTHS SURROUNDING THE DATE OF THE INCIDENT WAS ALSO DONE THERE WERE NO OTHER ACOUSTIC NOISE TINNITUS OR HEARING LOSS COMPLAINTS NOTED IN THE SERVICE DISPATCHES DURING THIS TIME PERIOD AS WELL | No Device Problem Found | Medical Evaluation result code
1244788 | 8246334 | N | RESULTS OBSTRUCTION FINAL DEVICE ANALYSIS REVEALED A RELIABILITY NONCONFORMANCE THERE WAS AN OBSTRUCTION IN THE CONNECTOR PORT ADHESIVE WAS PRESENT AROUND THE CONNECTOR IN THE BORE AT THE CONNECTOR AND AT THE FIRST INNER SEALING RING A KNOWN GOOD LEAD COULD ONLY INSERT INTO THE CONNECTOR PORT UP TO THE POINT WHERE THE PROXIMAL END OF THE LEAD REACHED THE CONNECTOR | Obstruction of Flow | Device problem code
1244788 | 8246334 | N | RESULTS OBSTRUCTION FINAL DEVICE ANALYSIS REVEALED A RELIABILITY NONCONFORMANCE THERE WAS AN OBSTRUCTION IN THE CONNECTOR PORT ADHESIVE WAS PRESENT AROUND THE CONNECTOR IN THE BORE AT THE CONNECTOR AND AT THE FIRST INNER SEALING RING A KNOWN GOOD LEAD COULD ONLY INSERT INTO THE CONNECTOR PORT UP TO THE POINT WHERE THE PROXIMAL END OF THE LEAD REACHED THE CONNECTOR | Device Difficult to Operate | Medical Evaluation result code
1525020 | 8359487 | N | DEVICE EVALUATION CUSTOMER REPORT WAS CONFIRMED RECD TRIPLE DPT KIT IV TUBING WAS COMPLETELY DETACHED FROM SOLVENT BOND JOINT WITH MALE CONNECTOR ATTACHED TO ARTERIAL DPT INDICATION OF BONDING SOLVENT WAS VISIBLE ON MOST PART OF BOND AREAS B | Detachment of Device or device Component | Device problem code
1525020 | 8359487 | N | DEVICE EVALUATION CUSTOMER REPORT WAS CONFIRMED RECD TRIPLE DPT KIT IV TUBING WAS COMPLETELY DETACHED FROM SOLVENT BOND JOINT WITH MALE CONNECTOR ATTACHED TO ARTERIAL DPT INDICATION OF BONDING SOLVENT WAS VISIBLE ON MOST PART OF BOND AREAS B | Incompatible Component/ Accessory | Medical Evaluation result code
1450178 | 8413026 | N | THE CUSTOMER REPORTED GETTING POWER LOST ERRORS THE UNIT WAS EVALUATED AT PHILIPS AND THE SYSTEM WAS VERIFIED REPLACING THE BATTERY PCA RESOLVED THE ISSUE | Loss of Power | Device Problem Code
1450178 | 8413026 | N | THE CUSTOMER REPORTED GETTING POWER LOST ERRORS THE UNIT WAS EVALUATED AT PHILIPS AND THE SYSTEM WAS VERIFIED REPLACING THE BATTERY PCA RESOLVED THE ISSUE | Energy Storage System Problem | Medical Evaluation result code
1399864 | 8458060 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Unintended power up | Device Problem Code
1399864 | 8458060 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Power Source Problem Identified | Medical Evaluation result code
1914584 | 8971746 | N | THE PARTS RETURNED ARE SHOWING TRACES OF IMPLANTATION AND REMOVAL NO DAMAGES OR DEFECTS HAVE BEEN IDENTIFIED ON THE CONNECTION HEAD AND THE SETSCREWS WITH THE AVAILABLE INFORMATION THE REASON FOR THE SETSCREW BACKOUT HAS NOT BEEN DETERMINED POTENTIAL REASONS FOR SETSCREW LOOSENING WITH FIXED ANGLE SCREWS INCLUDES WRONG APPROXIMATION OF THE RODFIXED ANGLE SCREW ORTHOGONALITY NONUNION OF THE INSTRUMENTED LEVEL | Connection Problem | Device problem code
1914584 | 8971746 | N | THE PARTS RETURNED ARE SHOWING TRACES OF IMPLANTATION AND REMOVAL NO DAMAGES OR DEFECTS HAVE BEEN IDENTIFIED ON THE CONNECTION HEAD AND THE SETSCREWS WITH THE AVAILABLE INFORMATION THE REASON FOR THE SETSCREW BACKOUT HAS NOT BEEN DETERMINED POTENTIAL REASONS FOR SETSCREW LOOSENING WITH FIXED ANGLE SCREWS INCLUDES WRONG APPROXIMATION OF THE RODFIXED ANGLE SCREW ORTHOGONALITY NONUNION OF THE INSTRUMENTED LEVEL | Incompatible Component/ Accessory | Medical Evaluation result code
1910149 | 9201360 | N | RADIOGRAPHIC REVIEW OF MULTIPLE FILMS MRI AND XRAYS OF LUMBAR SPINE THERE IS EVIDENCE OF LL DEGENERATIVE SPONDYLOLISTHESIS WITH INSTABILITY SUPPORTED BY MOVEMENT ON DYNAMIC XRAYS AND FLUID WITHIN THE LL FACET JOINTS EFFUSION NO EVIDENCE OF SURGICAL INTERVENTION OR ADVERSE EVENT IS NOTED | No Device Problem Found | Medical Evaluation result code
4753517 | 13451269 | N | A REVIEW OF MEDTRONICS DATABASES DID NOT SHOW ANY COMPLAINTS HAD PREVIOUSLY BEEN RECEIVED FROM THIS HEALTH CARE FACILITY FOR THAT MODEL FAMILY NO INFORMATION HAS BEEN RECEIVED THAT WOULD INDICATE THE DEVICE DOES NOT REMAIN IMPLANTED B ARTICLE BILATERAL PERCUTANEOUS STENT IMPLANTATION IN THE PULMONARY VASCULATURE A MODALITY NOW AVAILABLE IN LEBANON AUTHORS RAMZI SOUKI AGHIAD AL KOUTOUBI M SAMIR ARNAOUT MOUNIR OBEID FADI BITAR THE LEBANESE MEDICAL JOURNAL | No Device Problem Found | Medical Evaluation result code
2762792 | 15021222 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Discomfort | Patient problem code
2762792 | 15021222 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Pain | Patient problem code
2762792 | 15021222 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Noise, Audible | Device problem code
816591 | 15219936 | D | EVENT DESCRIPTION BOSTON SCIENTIFIC CARDIAC RHYTHM MANAGEMENT CRM RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSVENOUS DEFIBRILLATION LEAD COMPLETE HEART BLOCK WAS INDUCED NO ADVERSE PATIENT EFFECTS HAVE BEEN REPORTED | No Consequences Or Impact To Patient | Patient problem code
816591 | 15219936 | D | EVENT DESCRIPTION BOSTON SCIENTIFIC CARDIAC RHYTHM MANAGEMENT CRM RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSVENOUS DEFIBRILLATION LEAD COMPLETE HEART BLOCK WAS INDUCED NO ADVERSE PATIENT EFFECTS HAVE BEEN REPORTED | Adverse Event Without Identified Device or Use Problem | Device problem code
170739 | 15616711 | D | ACCORDING TO A LETTER RECD FROM AN ATTORNEY A PT SUSTAINED SECOND DEGREE BURNS WHEN A CRYOSURGICAL INSTRUMENT MALFUNCTIONED DURING A CRYOSURGICAL PROCEDURE IN CORRESPONDENCE FROM THE ATTORNEY IT WAS STATED THAT THE DOCTOR HAD SOME PROBLEMS WITH THE EQUIPMENT BUT DECIDED TO PERFORM THE PROCEDURE MIDWAY THROUGH THE PROCEDURE THE PT SUFFERED BURNS IN THE VAGINAL AREA | Component or Accessory Incompatibility | Device problem code
170739 | 15616711 | D | ACCORDING TO A LETTER RECD FROM AN ATTORNEY A PT SUSTAINED SECOND DEGREE BURNS WHEN A CRYOSURGICAL INSTRUMENT MALFUNCTIONED DURING A CRYOSURGICAL PROCEDURE IN CORRESPONDENCE FROM THE ATTORNEY IT WAS STATED THAT THE DOCTOR HAD SOME PROBLEMS WITH THE EQUIPMENT BUT DECIDED TO PERFORM THE PROCEDURE MIDWAY THROUGH THE PROCEDURE THE PT SUFFERED BURNS IN THE VAGINAL AREA | Burn(s) | Patient problem code
471007 | 15687300 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROX CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Deformation Due to Compressive Stress | Device problem code
471007 | 15687300 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROX CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Occlusion | Patient problem code
471007 | 15687300 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROX CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Stenosis | Patient problem code
1399861 | 15988697 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Unintended power up | Device Problem Code
1399861 | 15988697 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Power Source Problem Identified | Medical Evaluation result code
471001 | 16035403 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROXIMATELY CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Occlusion | Patient problem code
471001 | 16035403 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROXIMATELY CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Stenosis | Patient problem code
471001 | 16035403 | D | THREE GRAFTS WERE ANASTOMOSED WITH AORTIC CONNECTORS DIAGONAL ARTERY OM PDA APPROXIMATELY SIX WEEKS POSTOPERATIVELY ANGIOGRAM REVEALED THE GRAFTS TO THE DIAGONAL ARTERY AND OM WERE STENOSED THE STENOSED AREAS WERE APPROXIMATELY CM DISTAL TO THE CONNECTORS AND NO KINKING WAS OBSERVED STENTS WERE PLACED IN THE MIDDLE OF BOTH GRAFTS IN ANGIOGRAM REVEALED THE GRAFT TO THE PDA WAS STENOSED NEAR THE CONNECTOR DUE TO A KINK IN THE VEIN AND THE OM GRAFT WAS OCCLUDED THE GRAFTS WERE STENTED AND ALL THREE CONNECTORS REMAIN IMPLANTED | Deformation Due to Compressive Stress | Device problem code
561588 | 16140073 | D | REPORTER INDICATED THAT VNS PT HAD PASSED AWAY CAUSE OF DEATH IS NOT KNOWN AT THIS TIME TREATING NEUROLOGIST INDICATED THAT THE PT HAD BEEN FOUND DEAD LAYING ON A HEATING PAD AND WAS TAKING METHADONE AND VALIUM FOR CHRONIC PAIN THE PT SUSTAINED SECOND DEGREE BURNS FROM LAYING ON THE HEATING PAD NEUROLOGIST REQUESTED AN AUTOPSY FOR TOXICOLOGY RESULTS BUT THE NURSE AT THE CORONERS OFFICE INDICATED THAT THEIR FACILITY HAD ELECTED NOT TO PERFORM AN AUTOPSY AND IT WAS BELIEVED THAT THE PTS DEATH MAY HAVE BEEN SEIZURERELATED THE PTS SIGNIFICANT OTHER REPORTEDLY INDICATED THAT THE PT WAS NOT TAKING ANY PAIN MEDICATIONS REVIEW OF MANUFACTURING RECORDS FOR BOTH THE PULSE GENERATOR AND THE BIPOLAR LEAD REVEALED NO ANOMALIES THAT WOULD ADVERSELY EFFECT DEVICE PERFORMANCE DEVICE DIAGNOSTIC TESTING PERFORMED ON THE DAY OF IMPLANT SURGERY WAS WITHIN NORMAL LIMITS INDICATING PROPER DEVICE FUNCTION AT THAT TIME | Death | Patient problem code
561588 | 16140073 | D | REPORTER INDICATED THAT VNS PT HAD PASSED AWAY CAUSE OF DEATH IS NOT KNOWN AT THIS TIME TREATING NEUROLOGIST INDICATED THAT THE PT HAD BEEN FOUND DEAD LAYING ON A HEATING PAD AND WAS TAKING METHADONE AND VALIUM FOR CHRONIC PAIN THE PT SUSTAINED SECOND DEGREE BURNS FROM LAYING ON THE HEATING PAD NEUROLOGIST REQUESTED AN AUTOPSY FOR TOXICOLOGY RESULTS BUT THE NURSE AT THE CORONERS OFFICE INDICATED THAT THEIR FACILITY HAD ELECTED NOT TO PERFORM AN AUTOPSY AND IT WAS BELIEVED THAT THE PTS DEATH MAY HAVE BEEN SEIZURERELATED THE PTS SIGNIFICANT OTHER REPORTEDLY INDICATED THAT THE PT WAS NOT TAKING ANY PAIN MEDICATIONS REVIEW OF MANUFACTURING RECORDS FOR BOTH THE PULSE GENERATOR AND THE BIPOLAR LEAD REVEALED NO ANOMALIES THAT WOULD ADVERSELY EFFECT DEVICE PERFORMANCE DEVICE DIAGNOSTIC TESTING PERFORMED ON THE DAY OF IMPLANT SURGERY WAS WITHIN NORMAL LIMITS INDICATING PROPER DEVICE FUNCTION AT THAT TIME | Pain | Patient problem code
561588 | 16140073 | D | REPORTER INDICATED THAT VNS PT HAD PASSED AWAY CAUSE OF DEATH IS NOT KNOWN AT THIS TIME TREATING NEUROLOGIST INDICATED THAT THE PT HAD BEEN FOUND DEAD LAYING ON A HEATING PAD AND WAS TAKING METHADONE AND VALIUM FOR CHRONIC PAIN THE PT SUSTAINED SECOND DEGREE BURNS FROM LAYING ON THE HEATING PAD NEUROLOGIST REQUESTED AN AUTOPSY FOR TOXICOLOGY RESULTS BUT THE NURSE AT THE CORONERS OFFICE INDICATED THAT THEIR FACILITY HAD ELECTED NOT TO PERFORM AN AUTOPSY AND IT WAS BELIEVED THAT THE PTS DEATH MAY HAVE BEEN SEIZURERELATED THE PTS SIGNIFICANT OTHER REPORTEDLY INDICATED THAT THE PT WAS NOT TAKING ANY PAIN MEDICATIONS REVIEW OF MANUFACTURING RECORDS FOR BOTH THE PULSE GENERATOR AND THE BIPOLAR LEAD REVEALED NO ANOMALIES THAT WOULD ADVERSELY EFFECT DEVICE PERFORMANCE DEVICE DIAGNOSTIC TESTING PERFORMED ON THE DAY OF IMPLANT SURGERY WAS WITHIN NORMAL LIMITS INDICATING PROPER DEVICE FUNCTION AT THAT TIME | Adverse Event Without Identified Device or Use Problem | Device problem code
1253283 | 16170476 | D | THE CUSTOMER REPORTED THE VENTILATOR ALARMED AND THE SAFETY VALVE OPENED SVO DUE TO AN OCCLUSION WHEN AN OCCLUSION IS DETECTED DURING NORMAL VENTILATION THE ESPRIT WILL OPEN THE SAFETY VALVE SAFETY VALVE IS A COMPONENT UTILIZED DURING SAFETY VALVE OPEN SVO AN EMERGENCY MODE OF VENTILATION THAT ALLOWS THE PATIENT TO BREATHE THROUGH THE SYSTEM THE DEVICE WAS IN USE ON A PATIENT AND THE CUSTOMER REPORTED THERE WAS NO PATIENT HARM THE MANUFACTURERS SERVICE TECHNICIAN WAS UNABLE TO DUPLICATE THE REPORTED OCCLUSION EVENT THE SERVICE TECHNICIAN CONFIRMED A DIAGNOSTIC CODE IN THE VENTILATOR LOG HISTORY THAT INDICATED THE VENTILATOR HAD ALARMED SVO DUE TO AN OCCLUSION THE SERVICE TECHNICIAN TESTED THE VENTILATOR AND REPORTED THE EXHALATION PRESSURE WAS OUT OF SPECIFICATION THE SERVICE TECHNICIAN REPLACED THE SENSOR BOARD TO CORRECT THE FINDING PERFORMANCE VERIFICATION TESTING WAS COMPLETED AND THE VENTILATOR PASSED ALL TESTS TO OPERATING SPECIFICATIONS | Appropriate Term/Code Not Available | Device Problem Code
1253283 | 16170476 | D | THE CUSTOMER REPORTED THE VENTILATOR ALARMED AND THE SAFETY VALVE OPENED SVO DUE TO AN OCCLUSION WHEN AN OCCLUSION IS DETECTED DURING NORMAL VENTILATION THE ESPRIT WILL OPEN THE SAFETY VALVE SAFETY VALVE IS A COMPONENT UTILIZED DURING SAFETY VALVE OPEN SVO AN EMERGENCY MODE OF VENTILATION THAT ALLOWS THE PATIENT TO BREATHE THROUGH THE SYSTEM THE DEVICE WAS IN USE ON A PATIENT AND THE CUSTOMER REPORTED THERE WAS NO PATIENT HARM THE MANUFACTURERS SERVICE TECHNICIAN WAS UNABLE TO DUPLICATE THE REPORTED OCCLUSION EVENT THE SERVICE TECHNICIAN CONFIRMED A DIAGNOSTIC CODE IN THE VENTILATOR LOG HISTORY THAT INDICATED THE VENTILATOR HAD ALARMED SVO DUE TO AN OCCLUSION THE SERVICE TECHNICIAN TESTED THE VENTILATOR AND REPORTED THE EXHALATION PRESSURE WAS OUT OF SPECIFICATION THE SERVICE TECHNICIAN REPLACED THE SENSOR BOARD TO CORRECT THE FINDING PERFORMANCE VERIFICATION TESTING WAS COMPLETED AND THE VENTILATOR PASSED ALL TESTS TO OPERATING SPECIFICATIONS | No Consequences Or Impact To Patient | Patient problem code
1479938 | 17231268 | N | A REVIEW OF THE CHROMATOGRAMS FROM THE CUSTOMER SITE REVEALED THAT THE PEAKS WERE BROAD AND ROUNDED IN THE RESULTS PRODUCED ON B THE CHROMATOGRAM PEAKS PRODUCED ON B AFTER THE CHANGE OF THE GUARD CARTRIDGE AND CALIBRATION WERE SHARP AND WELL DEFINED THIS IS INDICATIVE OF A FAILED GUARD CARTRIDGE HOWEVER THE CUSTOMER HAD ALREADY DISPOSED OFF THE GUARD CARTRIDGE PRIOR TO REPORTING THIS ISSUE WE HAVE NOT BEEN ABLE TO REPRODUCE THIS ISSUE INTERNALLY | Component or Accessory Incompatibility | Device problem code
1479938 | 17231268 | N | A REVIEW OF THE CHROMATOGRAMS FROM THE CUSTOMER SITE REVEALED THAT THE PEAKS WERE BROAD AND ROUNDED IN THE RESULTS PRODUCED ON B THE CHROMATOGRAM PEAKS PRODUCED ON B AFTER THE CHANGE OF THE GUARD CARTRIDGE AND CALIBRATION WERE SHARP AND WELL DEFINED THIS IS INDICATIVE OF A FAILED GUARD CARTRIDGE HOWEVER THE CUSTOMER HAD ALREADY DISPOSED OFF THE GUARD CARTRIDGE PRIOR TO REPORTING THIS ISSUE WE HAVE NOT BEEN ABLE TO REPRODUCE THIS ISSUE INTERNALLY | Incompatible Component/ Accessory | Medical Evaluation result code
527188 | 17439394 | D | THE SALUTE FIXATION DEVICE IS INTENDED FOR FIXATION OF PROSTHETIC MATERIAL THE SYSTEM WAS BEING UTILIZED FOR A LAPAROSCOPIC VENTRAL HERNIA REPAIR THE SURGEON SUCCESSFULLY DEPLOYED ONE DISPOSABLE IMPLANT CARTRIDGE AND REQUIRED MORE Q CONSTRUCTS A SECOND DISPOSABLE IMPLANT CARTRIDGE WAS LOADED A STRAIGHT WIRE WAS DEPLOYED FROM THE SALUTE TIP WHICH PUNCTURED THE SURGEONS GLOVE AND FINGER THE PT HAD HEP A THERE WAS NO ADVERSE AFFECT TO THE PT THERE WAS NO FURTHER INFO AVAILABLE FROM THE HOSP | Material Puncture / Hole | Device problem code
527188 | 17439394 | D | THE SALUTE FIXATION DEVICE IS INTENDED FOR FIXATION OF PROSTHETIC MATERIAL THE SYSTEM WAS BEING UTILIZED FOR A LAPAROSCOPIC VENTRAL HERNIA REPAIR THE SURGEON SUCCESSFULLY DEPLOYED ONE DISPOSABLE IMPLANT CARTRIDGE AND REQUIRED MORE Q CONSTRUCTS A SECOND DISPOSABLE IMPLANT CARTRIDGE WAS LOADED A STRAIGHT WIRE WAS DEPLOYED FROM THE SALUTE TIP WHICH PUNCTURED THE SURGEONS GLOVE AND FINGER THE PT HAD HEP A THERE WAS NO ADVERSE AFFECT TO THE PT THERE WAS NO FURTHER INFO AVAILABLE FROM THE HOSP | No Consequences Or Impact To Patient | Patient problem code
2762816 | 17973969 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Discomfort | Patient problem code
2762816 | 17973969 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Pain | Patient problem code
2762816 | 17973969 | D | LITIGATION ALLEGES PATIENT EXPERIENCED SEVERE PAIN DISCOMFORT GRINDING AND POPPING | Noise, Audible | Device problem code
488442 | 18481779 | D | CUSTOMER REPORTS THE PT WAS ADMITTED TO THE MRI WITH ELECTRODES IN PLACE THE PT SUFFERED THIRD DEGREE BURNS | Insufficient Information | Device problem code
488442 | 18481779 | D | CUSTOMER REPORTS THE PT WAS ADMITTED TO THE MRI WITH ELECTRODES IN PLACE THE PT SUFFERED THIRD DEGREE BURNS | Third-Degree Burn | Patient problem code
1909424 | 18991182 | N | PREOP FILMS SHOW SIGNIFICANT DISK NARROWING AT LS WITH A DEGENERATIVE SLIP AND STENOSIS AT LL POSTOP FILMS SHOW PEDICLE SCREWS WITH SEGMENTAL BILATERAL RODS LS A FENESTRATED CIRCULAR IMPLANT OCCUPIES ROUGHLY OF THE DISK SPACE INSTRUMENTATION AND ANTERIOR IMPLANT DOES NOT EXTEND BEYOND ITS REASONABLE POSITIONING THE ANTERIOR IMPLANT HAS SETTLED INTO THE SUPRA AND SUB JACENT ENDPLATES CAUSING THE FORMER DISK SPACE TO BECOME COMPLETELY NARROWED MORE IMMEDIATE POSTOP FILMS SHOW MM OR MORE INTERBODY SPACE OSSEOUS HEALING ANTERIORLY HAS LIKELY OCCURRED IN THE MONTH POSTOP FILM AT LL | Material Deformation | Device problem code
1909424 | 18991182 | N | PREOP FILMS SHOW SIGNIFICANT DISK NARROWING AT LS WITH A DEGENERATIVE SLIP AND STENOSIS AT LL POSTOP FILMS SHOW PEDICLE SCREWS WITH SEGMENTAL BILATERAL RODS LS A FENESTRATED CIRCULAR IMPLANT OCCUPIES ROUGHLY OF THE DISK SPACE INSTRUMENTATION AND ANTERIOR IMPLANT DOES NOT EXTEND BEYOND ITS REASONABLE POSITIONING THE ANTERIOR IMPLANT HAS SETTLED INTO THE SUPRA AND SUB JACENT ENDPLATES CAUSING THE FORMER DISK SPACE TO BECOME COMPLETELY NARROWED MORE IMMEDIATE POSTOP FILMS SHOW MM OR MORE INTERBODY SPACE OSSEOUS HEALING ANTERIORLY HAS LIKELY OCCURRED IN THE MONTH POSTOP FILM AT LL | Deformation Problem | Medical Evaluation result code
1399862 | 19428231 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Unintended power up | Device Problem Code
1399862 | 19428231 | N | THE PHILIPS REPRESENTATIVE REPORTED THAT THE DEVICE WOULD INTERMITTENTLY POWER ON WHEN THE SWITCH WAS IN THE OFF POSITION IF THE DEVICE WAS TO SPONTANEOUSLY POWER UP AND GO UNNOTICED BY THE USERS THE BATTERY COULD DRAIN WHICH COULD PREVENT THE DELIVERY OF THERAPY VIA BATTERY POWER THE PHILIPS DISTRIBUTOR EVALUATED THE DEVICE THE SYMPTOM WAS VERIFIED THE ISSUE WAS LOCALIZED TO THE ENERGY SELECT SWITCH ASSEMBLY THIS EVENT WAS REPORTED LATE PURSUANT TO FDA AUDIT FINDINGS | Power Source Problem Identified | Medical Evaluation result code
490004 | 19961749 | D | IN PT WAS ADMITTED TO HOSP AFTER EXPERIENCING FOCAL SEIZURE TEGRETOL LEVELS WERE MONITORED AND PT WAS GIVEN IV DECADRON PT WAS OBSERVED OVERNIGHT AND DISCHARGED RELATIONSHIP TO TREATMENT IS POSSIBLE | Seizures, Focal | Patient problem code
490004 | 19961749 | D | IN PT WAS ADMITTED TO HOSP AFTER EXPERIENCING FOCAL SEIZURE TEGRETOL LEVELS WERE MONITORED AND PT WAS GIVEN IV DECADRON PT WAS OBSERVED OVERNIGHT AND DISCHARGED RELATIONSHIP TO TREATMENT IS POSSIBLE | Insufficient Information | Device problem code
916489 | 21819847 | N | NARRATIVE THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS MEDICAL SOLUTIONS DIAGNOSTICS FOR FURTHER EVALUATION EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM THE SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Difficult to Remove | Device problem code
916489 | 21819847 | N | NARRATIVE THE CARTRIDGE AND INSTRUMENT ARE BEING SENT BACK TO SIEMENS MEDICAL SOLUTIONS DIAGNOSTICS FOR FURTHER EVALUATION EVALUATION THE REAGENT CARTRIDGE WAS DIFFICULT TO REMOVE FROM THE SYSTEM THIS EVENT IS BEING REPORTED BECAUSE IT OCCURRED IN A POTENTIALLY BIOHAZARDOUS ENVIRONMENT | Operational Problem Identified | Medical Evaluation result code
1525767 | 21861732 | N | DEVICE EVALUATION CUSTOMER REPORT WAS NOT CONFIRMED DPT WITH SALINE PRIMED VAMP PLUS KIT WAS RETURNED IT WAS NOT ABLE TO FIND ANY CONTAMINANTS IN THE VAMP SYRINGE BEFORE DECONTAMINATION FURTHER VISUAL INSPECTION WAS PERFORMED BUT ANY CONTAMINANT WAS NOT FOUND NO VISIBLE DEFECT WAS OBSERVED FROM ENTIRE KIT | No Device Problem Found | Medical Evaluation result code
491928 | 21951572 | D | VNS PT PASSED AWAY IT WAS REPORTED THAT BEFORE DEATH THE PATIENT HAD FALLEN AND WAS TAKEN TO THE HOSPITAL BUT WAS DISCHARGED TO HOME IT IS UNKNOWN WHETHER THE PATIENT PASSED AWAY AT HOME OR WAS TAKEN BACK TO THE HOSPITAL CAUSE OF DEATH IS NOT KNOWN AT THIS TIME THERE IS NO EVIDENCE AT THIS TIME THAT THE NCP SYSTEM CAUSED OR CONTRIBUTED TO THE REPORTED EVENT | Death | Patient problem code
491928 | 21951572 | D | VNS PT PASSED AWAY IT WAS REPORTED THAT BEFORE DEATH THE PATIENT HAD FALLEN AND WAS TAKEN TO THE HOSPITAL BUT WAS DISCHARGED TO HOME IT IS UNKNOWN WHETHER THE PATIENT PASSED AWAY AT HOME OR WAS TAKEN BACK TO THE HOSPITAL CAUSE OF DEATH IS NOT KNOWN AT THIS TIME THERE IS NO EVIDENCE AT THIS TIME THAT THE NCP SYSTEM CAUSED OR CONTRIBUTED TO THE REPORTED EVENT | Fall | Patient problem code
491928 | 21951572 | D | VNS PT PASSED AWAY IT WAS REPORTED THAT BEFORE DEATH THE PATIENT HAD FALLEN AND WAS TAKEN TO THE HOSPITAL BUT WAS DISCHARGED TO HOME IT IS UNKNOWN WHETHER THE PATIENT PASSED AWAY AT HOME OR WAS TAKEN BACK TO THE HOSPITAL CAUSE OF DEATH IS NOT KNOWN AT THIS TIME THERE IS NO EVIDENCE AT THIS TIME THAT THE NCP SYSTEM CAUSED OR CONTRIBUTED TO THE REPORTED EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
2801977 | 22043332 | D | PATIENT WAS REVISED TO ADDRESS PAIN HIP | Pain | Patient problem code
2801977 | 22043332 | D | PATIENT WAS REVISED TO ADDRESS PAIN HIP | Insufficient Information | Device Problem Code
4116204 | 22266919 | D | THE CUSTOMER STATED THAT SHE WAS HOSPITALIZED WITH A LOW BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED LOW BLOOD GLUCOSE FOR AN HOUR AFTER SHE GAVE HERSELF TOO MUCH INSULIN DURING BOLUS DELIVERY THE CUSTOMER COULD NOT RECALL IF A SET CHANGE HAD BEEN PERFORMED PRIOR TO THIS THE CUSTOMER STATED THE INSULIN PUMP WAS NOT EXPOSED TO ANY STRONG MAGNETIC FIELDS SHE COULD NOT CHECK THE PROGRAMMING OF THE INSULIN PUMP AS IT WAS NO LONGER IN THE CUSTOMERS POSSESSION AT THE TIME OF REPORTING THE EVENT WHICH SHE STATED OCCURRED BETWEEN AND NOTHING FURTHER WAS REPORTED | Insufficient Information | Device problem code
4116204 | 22266919 | D | THE CUSTOMER STATED THAT SHE WAS HOSPITALIZED WITH A LOW BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED LOW BLOOD GLUCOSE FOR AN HOUR AFTER SHE GAVE HERSELF TOO MUCH INSULIN DURING BOLUS DELIVERY THE CUSTOMER COULD NOT RECALL IF A SET CHANGE HAD BEEN PERFORMED PRIOR TO THIS THE CUSTOMER STATED THE INSULIN PUMP WAS NOT EXPOSED TO ANY STRONG MAGNETIC FIELDS SHE COULD NOT CHECK THE PROGRAMMING OF THE INSULIN PUMP AS IT WAS NO LONGER IN THE CUSTOMERS POSSESSION AT THE TIME OF REPORTING THE EVENT WHICH SHE STATED OCCURRED BETWEEN AND NOTHING FURTHER WAS REPORTED | Hypoglycemia | Patient Problem Code
4846864 | 22530575 | D | THE CUSTOMER RECEIVED QUESTIONABLE TROPONIN I STAT RESULTS FOR QC MATERIAL AND ONE PATIENT SAMPLE ONLY THE RESULTS FOR THE PATIENT SAMPLE WERE A REPORTABLE MALFUNCTION THE INITIAL RESULT WAS UGL THE REPEAT RESULT WAS UGL WITH A DATA FLAG AND THE REPEAT RESULT ON B WAS UGL WITH A DATA FLAG ALL OF THE RESULTS WERE REPORTED OUTSIDE THE LABORATORY THE PATIENT WAS NOT ADVERSELY AFFECTED THE REAGENT LOT NUMBER WAS WITH AN EXPIRATION DATE OF A SPECIFIC ROOT CAUSE COULD NOT BE IDENTIFIED AN ERRONEOUS PIPETTING WAS MOST LIKELY THE CAUSE OF THE ISSUE IT WAS NOTED THE CUSTOMER WAS NOT FOLLOWING THE TUBE MANUFACTURERS RECOMMENDATION FOR CENTRIFUGATION CONDITIONS THIS MAY HAVE CONTRIBUTED TO THE ISSUE THE PROVIDED ANALYZER ALARM DATA DID NOT INDICATE ANY ISSUES AND THE ANALYZER PERFORMANCE DATA WAS WITHIN SPECIFICATIONS | No Consequences Or Impact To Patient | Patient problem code
4846864 | 22530575 | D | THE CUSTOMER RECEIVED QUESTIONABLE TROPONIN I STAT RESULTS FOR QC MATERIAL AND ONE PATIENT SAMPLE ONLY THE RESULTS FOR THE PATIENT SAMPLE WERE A REPORTABLE MALFUNCTION THE INITIAL RESULT WAS UGL THE REPEAT RESULT WAS UGL WITH A DATA FLAG AND THE REPEAT RESULT ON B WAS UGL WITH A DATA FLAG ALL OF THE RESULTS WERE REPORTED OUTSIDE THE LABORATORY THE PATIENT WAS NOT ADVERSELY AFFECTED THE REAGENT LOT NUMBER WAS WITH AN EXPIRATION DATE OF A SPECIFIC ROOT CAUSE COULD NOT BE IDENTIFIED AN ERRONEOUS PIPETTING WAS MOST LIKELY THE CAUSE OF THE ISSUE IT WAS NOTED THE CUSTOMER WAS NOT FOLLOWING THE TUBE MANUFACTURERS RECOMMENDATION FOR CENTRIFUGATION CONDITIONS THIS MAY HAVE CONTRIBUTED TO THE ISSUE THE PROVIDED ANALYZER ALARM DATA DID NOT INDICATE ANY ISSUES AND THE ANALYZER PERFORMANCE DATA WAS WITHIN SPECIFICATIONS | Incorrect, Inadequate or Imprecise Result or Readings | Device Problem Code
2232643 | 26967385 | D | UPDATE RECD PFS AND MEDICAL RECORDS RECEIVED CORRECT DOI AND DOR WAS GIVEN AFTER REVIEW OF THE OPEATIVE NOTE METALLOSIS WAS CONFIRMED THERE IS NO NEW ADDITIONAL INFORMATION THAT WOULD AFFECT THE EXISTING MDR DECISION | Toxicity | Patient problem code
2232643 | 26967385 | D | UPDATE RECD PFS AND MEDICAL RECORDS RECEIVED CORRECT DOI AND DOR WAS GIVEN AFTER REVIEW OF THE OPEATIVE NOTE METALLOSIS WAS CONFIRMED THERE IS NO NEW ADDITIONAL INFORMATION THAT WOULD AFFECT THE EXISTING MDR DECISION | Appropriate Term/Code Not Available | Device Problem Code
2779427 | 27259884 | D | SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN ELEVATED METAL ION LEVELS AND AN UNDERANTEVERTED CUP POSITION | Pain | Patient problem code
2779427 | 27259884 | D | SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN ELEVATED METAL ION LEVELS AND AN UNDERANTEVERTED CUP POSITION | Toxicity | Patient problem code
2779427 | 27259884 | D | SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN ELEVATED METAL ION LEVELS AND AN UNDERANTEVERTED CUP POSITION | Insufficient Information | Device Problem Code
2760996 | 27379492 | D | THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND METALLOSIS | Pain | Patient problem code
2760996 | 27379492 | D | THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND METALLOSIS | Toxicity | Patient problem code
2760996 | 27379492 | D | THE SALES REP HAS REPORTED THE REVISION SURGERY PATIENT WAS REVISED TO ADDRESS PAIN AND METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
2316977 | 32904651 | D | UPDATE RECD THE PATIENTS MEDICAL RECORDS WERE RECEIVED MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY ACCORDING TO THE MEDICAL RECORDS THE PATIENT ALSO HAD METALLOSIS THERE WAS METAL DEBRIS THROUGHOUT THE CAPSULE AT THIS TIME THE PATIENTS FEMORAL STEM IS BEING REPORTED FOR THE ELEVATED METAL ION LEVELS THE COMPLAINT WAS UPDATED ON | Toxicity | Patient problem code
2316977 | 32904651 | D | UPDATE RECD THE PATIENTS MEDICAL RECORDS WERE RECEIVED MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY ACCORDING TO THE MEDICAL RECORDS THE PATIENT ALSO HAD METALLOSIS THERE WAS METAL DEBRIS THROUGHOUT THE CAPSULE AT THIS TIME THE PATIENTS FEMORAL STEM IS BEING REPORTED FOR THE ELEVATED METAL ION LEVELS THE COMPLAINT WAS UPDATED ON | Appropriate Term/Code Not Available | Device Problem Code
2317011 | 32905027 | D | UPDATE RECD THE PATIENTS MEDICAL RECORDS WERE RECEIVED MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY ACCORDING TO THE MEDICAL RECORDS THE PATIENT ALSO HAD METALLOSIS THERE WAS METAL DEBRIS THROUGHOUT THE CAPSULE AT THIS TIME THE PATIENTS FEMORAL STEM IS BEING REPORTED FOR THE ELEVATED METAL ION LEVELS THE COMPLAINT WAS UPDATED ON | Toxicity | Patient problem code
2317011 | 32905027 | D | UPDATE RECD THE PATIENTS MEDICAL RECORDS WERE RECEIVED MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY ACCORDING TO THE MEDICAL RECORDS THE PATIENT ALSO HAD METALLOSIS THERE WAS METAL DEBRIS THROUGHOUT THE CAPSULE AT THIS TIME THE PATIENTS FEMORAL STEM IS BEING REPORTED FOR THE ELEVATED METAL ION LEVELS THE COMPLAINT WAS UPDATED ON | Appropriate Term/Code Not Available | Device Problem Code
5471062 | 42778159 | N | THE CUSTOMER PROVIDED ADDITIONAL ERRONEOUS RESULTS FOR AN ADDITIONAL PATIENT WITH B ON B THE INITIAL TROPONIN T HS STAT RESULT WAS PGML WITH A DATA FLAG ON THE E WITH SERIAL NUMBER B THE SAMPLE WAS REPEATED ON THE SAME ANALYZER WITH A RESULT OF PGML THE SAMPLE WAS ALSO REPEATED ON THE E ANALYZER WITH SERIAL NUMBER B AND THE RESULT WAS PGML NO ERRONEOUS RESULTS WERE REPORTED OUTSIDE OF THE LABORATORY AND THIS PATIENT WAS NOT ADVERSELY AFFECTED | No Device Problem Found | Medical Evaluation result code
5890924 | 52737123 | N | A CUSTOMER CONTACTED A SIEMENS CUSTOMER CARE CENTER CCC SPECIALIST THE CCC INQUIRED ABOUT CALIBRATION AND QUALITY CONTROL QC THE CUSTOMER STATED THAT THE CALIBRATION AND QC WERE VALID AND IN RANGE SIEMENS IS INVESTIGATING THE ISSUE | Results Pending Completion of Investigation | Medical Evaluation result code
5934551 | 54139479 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B AT AN UNNAMED HOSPITAL IN B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
5934551 | 54139479 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B AT AN UNNAMED HOSPITAL IN B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
5934551 | 61628523 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B AT AN UNNAMED HOSPITAL IN B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
5934551 | 61628523 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B AT AN UNNAMED HOSPITAL IN B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6188510 | 62797184 | D | IT IS ALLEGED THAT THE PATIENT RECEIVED A COOK GUNTHER TULIP ON B AT B MEDICAL CENTER IN B IT IS ALLEGED THAT PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6188510 | 62797184 | D | IT IS ALLEGED THAT THE PATIENT RECEIVED A COOK GUNTHER TULIP ON B AT B MEDICAL CENTER IN B IT IS ALLEGED THAT PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION PATIENT IS SEEKING PUNITIVE DAMAGES HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6208393 | 63704550 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR WE ARE UNAWARE ABOUT PATIENT INJURIES | Component Incompatible | Device problem code
6208393 | 63704550 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR WE ARE UNAWARE ABOUT PATIENT INJURIES | Incompatible Component/ Accessory | Medical Evaluation result code
6208402 | 63704720 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURIES DEVICE EVALUATED BY DISTRIBUTOR | Component Incompatible | Device problem code
6208402 | 63704720 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURIES DEVICE EVALUATED BY DISTRIBUTOR | Incompatible Component/ Accessory | Medical Evaluation result code
6209885 | 63707683 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RECOGNIZED THE ATTACHED ADULT ELECTRODES AS PEDIATRIC ELECTRODES COMPLAINANT ALLEGED THE CLINICIAN OBTAINED ANOTHER SET OF PADS TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6209885 | 63707683 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RECOGNIZED THE ATTACHED ADULT ELECTRODES AS PEDIATRIC ELECTRODES COMPLAINANT ALLEGED THE CLINICIAN OBTAINED ANOTHER SET OF PADS TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6208410 | 63709784 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURIES DEVICE EVALUATED BY DISTRIBUTOR | Component Incompatible | Device problem code
6208410 | 63709784 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURIES DEVICE EVALUATED BY DISTRIBUTOR | Incompatible Component/ Accessory | Medical Evaluation result code
6213425 | 63807602 | D | CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP WAS TURNING OFF AND ON WITHOUT ANY LOW BATTERY ALERT CUSTOMERS BLOOD GLUCOSE WAS MMOLL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND CUSTOMER STATED IT WAS THE SECOND OCCURRENCE OF OFF NOW POWER WITHOUT A LOW BATTERY WARNING PRIOR TO SHUTTING OFF CUSTOMER WAS ADVISED THE DEVICE WILL NEED TO BE REPLACED AND LOCAL OFFICE WILL DO THAT DEVICE IS CURRENTLY NOT RETURNING FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6213425 | 63807602 | D | CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP WAS TURNING OFF AND ON WITHOUT ANY LOW BATTERY ALERT CUSTOMERS BLOOD GLUCOSE WAS MMOLL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND CUSTOMER STATED IT WAS THE SECOND OCCURRENCE OF OFF NOW POWER WITHOUT A LOW BATTERY WARNING PRIOR TO SHUTTING OFF CUSTOMER WAS ADVISED THE DEVICE WILL NEED TO BE REPLACED AND LOCAL OFFICE WILL DO THAT DEVICE IS CURRENTLY NOT RETURNING FOR ANALYSIS | Battery Problem | Device problem code
6215514 | 63829499 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE FAILED TO OPERATE NO FURTHER INFORMATION WAS AVAILABLE REGARDING THE MALFUNCTION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6215514 | 63829499 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE FAILED TO OPERATE NO FURTHER INFORMATION WAS AVAILABLE REGARDING THE MALFUNCTION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6095414 | 64151387 | D | DESCRIPTION ACCORDING TO SHORT FORM COMPLAINT FILED IT IS ALLEGED THAT PT RECEIVED A COOK CELECT FILTER ON B PATIENT OUTCOME IT IS ALLEGED THAT PT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6095414 | 64151387 | D | DESCRIPTION ACCORDING TO SHORT FORM COMPLAINT FILED IT IS ALLEGED THAT PT RECEIVED A COOK CELECT FILTER ON B PATIENT OUTCOME IT IS ALLEGED THAT PT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6215514 | 65693688 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICES SPO WOULD NOT FUNCTION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6215514 | 65693688 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICES SPO WOULD NOT FUNCTION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6373205 | 68842326 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6373205 | 68842326 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6380772 | 69364261 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFER TO MEDWATCH REPORT WHICH IS A SECOND DEVICE FOR THE SAME PATIENT | Failure to Analyze Signal | Device Problem Code
6380772 | 69364261 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFER TO MEDWATCH REPORT WHICH IS A SECOND DEVICE FOR THE SAME PATIENT | No Known Impact Or Consequence To Patient | Patient problem code
6380763 | 69391446 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFER TO MEDWATCH REPORT WHICH IS A SECOND DEVICE FOR THE SAME PATIENT | No Known Impact Or Consequence To Patient | Patient problem code
6380763 | 69391446 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFER TO MEDWATCH REPORT WHICH IS A SECOND DEVICE FOR THE SAME PATIENT | Failure to Analyze Signal | Device Problem Code
6397122 | 69733710 | D | IT WAS REPORTED A PHYSICIAN PERFORMED A NOVASURE ENDOMETRIAL ABLATION ON B AND SECONDS INTO THE PROCEDURE THE PATIENTS BLOOD PRESSURE DROPPED AND THE PATIENT CODED CHEST COMPRESSIONS WERE PERFORMED AND THE PATIENT WAS REVIVED THE PATIENT WAS TRANSFERRED TO THE EMERGENCY ROOM ER ON B IT WAS REPORTED THE PATIENT WAS DISCHARGED FROM THE ER WITH NO DIAGNOSIS AND IS CURRENTLY DOING FINE | Insufficient Information | Device problem code
6397122 | 69733710 | D | IT WAS REPORTED A PHYSICIAN PERFORMED A NOVASURE ENDOMETRIAL ABLATION ON B AND SECONDS INTO THE PROCEDURE THE PATIENTS BLOOD PRESSURE DROPPED AND THE PATIENT CODED CHEST COMPRESSIONS WERE PERFORMED AND THE PATIENT WAS REVIVED THE PATIENT WAS TRANSFERRED TO THE EMERGENCY ROOM ER ON B IT WAS REPORTED THE PATIENT WAS DISCHARGED FROM THE ER WITH NO DIAGNOSIS AND IS CURRENTLY DOING FINE | No Known Impact Or Consequence To Patient | Patient Problem Code
6391224 | 69810018 | D | COMPLAINANT ALLEGED THAT WHILE ANALYZING THE HEART RHYTHM OF A PATIENT AGE AND GENDER UNKNOWN THE DEVICE RETURNED A SHOCK ADVISED DETERMINATION FOR WHAT APPEARED TO BE A NONSHOCKABLE HEART RHYTHM AND SUBSEQUENTLY DELIVERED THE SHOCK COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defibrillation Problem | Device Problem Code
6391224 | 69810018 | D | COMPLAINANT ALLEGED THAT WHILE ANALYZING THE HEART RHYTHM OF A PATIENT AGE AND GENDER UNKNOWN THE DEVICE RETURNED A SHOCK ADVISED DETERMINATION FOR WHAT APPEARED TO BE A NONSHOCKABLE HEART RHYTHM AND SUBSEQUENTLY DELIVERED THE SHOCK COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6401857 | 70104670 | D | OLYMPUS MEDICAL SYSTEMS CORP OMSC WAS INFORMED THAT DURING A ROUTINE SURVEILLANCE CULTURING CONDUCTED BY THE USER FACILITY ALL CHANNELS OF THE SUBJECT DEVICE TESTED POSITIVE FOR STAPHYLOCOCCUS NON AUREUSTHERE WAS NO REPORT OF INFECTION ASSOCIATED WITH THIS REPORT | Infection | Patient problem code
6401857 | 70104670 | D | OLYMPUS MEDICAL SYSTEMS CORP OMSC WAS INFORMED THAT DURING A ROUTINE SURVEILLANCE CULTURING CONDUCTED BY THE USER FACILITY ALL CHANNELS OF THE SUBJECT DEVICE TESTED POSITIVE FOR STAPHYLOCOCCUS NON AUREUSTHERE WAS NO REPORT OF INFECTION ASSOCIATED WITH THIS REPORT | Insufficient Information | Device Problem Code
6435427 | 71205971 | D | THE CUSTOMER REPORTED THAT THE DISCHARGED SIGN IS ON THE DISPLAY THERE WAS NO REPORTED PATIENT INVOLVEMENT | No Patient Involvement | Patient problem code
6435427 | 71205971 | D | THE CUSTOMER REPORTED THAT THE DISCHARGED SIGN IS ON THE DISPLAY THERE WAS NO REPORTED PATIENT INVOLVEMENT | Unintended Ejection | Device problem code
6471450 | 72365459 | N | B THE PUMP ICON APPEARED AFTER THE CUSTOMER TURNED ON AND OFF THE POWER SOURCE BUT AFTER A PERIOD OF TIME IT DISAPPEARED AGAIN THEY TRIED TO TURN ON POWER SOURCE AGAIN THE ICON APPEARED BUT THE PUMP STOPPED | Pumping Stopped | Device problem code
6471450 | 72365459 | N | B THE PUMP ICON APPEARED AFTER THE CUSTOMER TURNED ON AND OFF THE POWER SOURCE BUT AFTER A PERIOD OF TIME IT DISAPPEARED AGAIN THEY TRIED TO TURN ON POWER SOURCE AGAIN THE ICON APPEARED BUT THE PUMP STOPPED | Operational Problem Identified | Medical Evaluation result code
6413376 | 72704083 | N | INVESTIGATION THE RUN DATA FILE RDF WAS ANALYZED FOR THIS EVENT SIGNALS IN THE RDF CONFIRMED THAT THE SPECTRA OPTIA OPERATED AS INTENDED BY DETECTING THE REPORTED HEMOLYSIS THAT WAS OBSERVED IN THE PLASMA LINE APPROXIMATELY MINUTES INTO THE PROCEDURE WAS THE FIRST OCCURRENCE OF THE CELLS WERE DETECTED IN PLASMA LINE FROM CENTRIFUGE ALARM THERE WERE MORE OCCURRENCES OF THIS ALARM DURING THE FIRST MINUTES OF THE PROCEDURE AFTER MINUTES THE ALARMS SUBSIDED AND THE PROCEDURE CONTINUED WITHOUT ANY FURTHER ALARMS INVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | Results Pending Completion of Investigation | Medical Evaluation result code
6499447 | 72990297 | D | AFTER THE UNSPECIFIED PROCEDURE WITH THE SUBJECT UES MICROPERFORATION WAS FOUND IN THE PATIENTS BLADDER WE HAD A MEETING WITH THE USER FACILITY AND ASKED FOR INFORMATION OF THE SUBJECT EVENT THE USER FACILITY REPORTED THAT THERE WAS NO PROBLEM IN THE PROCEDURES PERFORMED AFTER THE SUBJECT EVENT THERE WAS NO REPORT OF THE PATIENT INJURY OTHER THAN ABOVE | Adverse Event Without Identified Device or Use Problem | Device problem code
6499447 | 72990297 | D | AFTER THE UNSPECIFIED PROCEDURE WITH THE SUBJECT UES MICROPERFORATION WAS FOUND IN THE PATIENTS BLADDER WE HAD A MEETING WITH THE USER FACILITY AND ASKED FOR INFORMATION OF THE SUBJECT EVENT THE USER FACILITY REPORTED THAT THERE WAS NO PROBLEM IN THE PROCEDURES PERFORMED AFTER THE SUBJECT EVENT THERE WAS NO REPORT OF THE PATIENT INJURY OTHER THAN ABOVE | Perforation | Patient problem code
6499527 | 73254607 | N | THE INVESTIGATION DETERMINED THAT A LOWER THAN EXPECTED CKMB RESULT WAS OBTAINED FROM A PATIENT SAMPLE PROCESSED USING VITROS CKMB SLIDES ON A VITROS INTEGRATED SYSTEM THE MOST LIKELY ASSIGNABLE CAUSE FOR THE EVENT IS INSTRUMENT RELATED AS A WITHINRUN PRECISION TEST WAS OUTSIDE OF ORTHO CLINICAL DIAGNOSTIC GUIDELINES INDICATING THE VITROS SYSTEM WAS NOT PERFORMING AS INTENDED THE ORTHO FIELD ENGINEER PERFORMED SERVICE ACTIONS AND THE POST SERVICE WITHINRUN PRECISION WAS ACCEPTABLE INDICATING THE VITROS SYSTEM WAS PERFORMING AS EXPECTED FOLLOWING SERVICE ACTIONS ACCEPTABLE VITROS CKMB PERFORMANCE WAS OBSERVED | Results Pending Completion of Investigation | Medical Evaluation result code
6511216 | 73374515 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO DETECT THE ATTACHED ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Under-Sensing | Device Problem Code
6511216 | 73374515 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO DETECT THE ATTACHED ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6518701 | 73735168 | D | THE USER IS REPORTING THE DEVICE INAPPROPRIATELY DISCHARGED DURING A PATIENT USE EVENT | No Information | Patient problem code
6518701 | 73735168 | D | THE USER IS REPORTING THE DEVICE INAPPROPRIATELY DISCHARGED DURING A PATIENT USE EVENT | Insufficient Information | Device problem code
6526963 | 74021211 | N | B THE CUSTOMER OBSERVED NO GAS LEAKS THEY DID ATTEMPT A SECOND CALIBRATION AND THE CALIBRATIONS WERE PERFORMED IMMEDIATELY PRIOR TO CASE | Calibration Problem | Device problem code
6526963 | 74021211 | N | B THE CUSTOMER OBSERVED NO GAS LEAKS THEY DID ATTEMPT A SECOND CALIBRATION AND THE CALIBRATIONS WERE PERFORMED IMMEDIATELY PRIOR TO CASE | Failure to Calibrate | Medical Evaluation result code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Occlusion | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Urticaria | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Rash | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Itching | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Collapse | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Skin Inflammation | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Hypersensitivity | Patient problem code
6533410 | 74039477 | D | THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF URTICARIA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD AND ECZEMA PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IN A FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW IN B THE PATIENT HAD ESSURE INSERTED IN B THE PATIENT EXPERIENCED URTICARIA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED WITH PRURITUS ECZEMA SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED AGITATION EXCITABILITY AND INSOMNIA INSOMNIA THE PATIENT WAS TREATED WITH CORTICOSTEROIDS DERMATOLOGICAL PREPARATIONS PARAFFIN MOISTURIZER ANTIHISTAMINES PREDNISONE SURGERY SURGICAL REMOVAL OF IMPLANTS ESSURE WAS REMOVED IN B AT THE TIME OF THE REPORT THE URTICARIA ECZEMA AGITATION AND INSOMNIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR AGITATION ECZEMA INSOMNIA AND URTICARIA WITH ESSURE FURTHER COMPANY FOLLOWUP WITH THE CONSUMER IS NOT POSSIBLE COMPANY CAUSALITY COMMENT THIS SPONTANEOUS CASE REPORT REFERS TO A FEMALE PATIENT WHO HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND APPROXIMATELY TWO MONTHS LATER HAD PROFUSE URTICARIAL AND ECZEMATIFORM ERUPTION POSSIBLY MORE MARKED WHICH PROGRESSIVELY SPREAD IMPLANTS WERE SURGICALLY REMOVED MONTHS AFTER INSERTION AND THE EVENT WAS RESOLVING THIS EVENT IS ANTICIPATED IN THE REFERENCE SAFETY INFORMATION FOR ESSURE THE ESSURE INSERT CONSISTS OF A NICKELTITANIUM ALLOY ALLERGIC REACTIONS TO ESSURE ARE MORE COMMONLY RELATED TO NICKEL SENSITIVITY AND ITS MANIFESTATIONS INCLUDE SKIN ITCHING RASH ECZEMATOUS DERMATITIS AND HIVES THAT RESOLVE AFTER DEVICE REMOVAL IN THIS PARTICULAR CASE THE EVENT OCCURRED TWO MONTHS AFTER DEVICE INSERTION THEREFORE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED NONSERIOUS EVENTS WERE ALSO REPORTED THIS CASE WAS REGARDED AS INCIDENT SINCE DEVICE REMOVAL WAS REQUIRED A PRODUCT TECHNICAL ANALYSIS IS EXPECTED | Adverse Event Without Identified Device or Use Problem | Device problem code
6534313 | 74088775 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH SUGAR THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6534313 | 74088775 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH SUGAR THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Display or Visual Feedback Problem | Device problem code
6534367 | 74089745 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED VOMITING THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS IN THE HOSPITAL FOR DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6534367 | 74089745 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED VOMITING THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS IN THE HOSPITAL FOR DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Vomiting | Patient problem code
6534367 | 74089745 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED VOMITING THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS IN THE HOSPITAL FOR DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6534367 | 74089745 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED VOMITING THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS IN THE HOSPITAL FOR DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Device Handling Problem | Device problem code
6537662 | 74204467 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF INCIDENT THE CUSTOMER STATED THAT THE EMERGENCY MEDICAL SERVICES WERE CALLED FOR THE LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY WERE GIVEN GLUCOSE TABLETS TO TREAT THE BLOOD GLUCOSE THE CUSTOMER STATED THAT THEY WERE WEARING THE INSULIN PUMP DURING HOSPITALIZATION THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6537662 | 74204467 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF INCIDENT THE CUSTOMER STATED THAT THE EMERGENCY MEDICAL SERVICES WERE CALLED FOR THE LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY WERE GIVEN GLUCOSE TABLETS TO TREAT THE BLOOD GLUCOSE THE CUSTOMER STATED THAT THEY WERE WEARING THE INSULIN PUMP DURING HOSPITALIZATION THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Display or Visual Feedback Problem | Device problem code
6537814 | 74206005 | D | THE CUSTOMER REPORTED THAT SHE RECENTLY HAD A BLOOD GLUCOSE LEVEL OF MGDL WHICH REQUIRED PARAMEDICS TO BE DISPATCHED CUSTOMER STATED THAT SHE FELL SEVERAL TIMES AND GLUCAGON DID NOT RAISE HER BLOOD GLUCOSE LEVEL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE INCIDENT BUT HAD BEEN OFF FOR LESS THAN HOURS BLOOD GLUCOSE LEVEL AT THE TIME OF THE CALL WAS MGDL THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | Device Handling Problem | Device problem code
6537814 | 74206005 | D | THE CUSTOMER REPORTED THAT SHE RECENTLY HAD A BLOOD GLUCOSE LEVEL OF MGDL WHICH REQUIRED PARAMEDICS TO BE DISPATCHED CUSTOMER STATED THAT SHE FELL SEVERAL TIMES AND GLUCAGON DID NOT RAISE HER BLOOD GLUCOSE LEVEL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE INCIDENT BUT HAD BEEN OFF FOR LESS THAN HOURS BLOOD GLUCOSE LEVEL AT THE TIME OF THE CALL WAS MGDL THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | Fall | Patient Problem Code
6539119 | 74253469 | D | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE NONTORTUOUS STENOSED BELOW THE BIFURCATION OF THE INTERNAL CAROTID ARTERY A X MM XACT STENT DELIVERY SYSTEM SDS WAS SELECTED FOR THE PROCEDURE THE STENT IMPLANT WAS SUCCESSFULLY DEPLOYED WITH NO ISSUES NOTED ONE HOUR AFTER THE COMPLETION OF THE PROCEDURE THE PATIENT SUFFERED A MINOR STROKE WAS TREATED WITH MEDICATION AND WAS DOING FINE IN THE OPINION OF THE ATTENDING PHYSICIAN THE PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE STROKE EXPERIENCED BY THE PATIENT NO ADDITIONAL INFORMATION WAS PROVIDED | Stenosis | Patient problem code
6539119 | 74253469 | D | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE NONTORTUOUS STENOSED BELOW THE BIFURCATION OF THE INTERNAL CAROTID ARTERY A X MM XACT STENT DELIVERY SYSTEM SDS WAS SELECTED FOR THE PROCEDURE THE STENT IMPLANT WAS SUCCESSFULLY DEPLOYED WITH NO ISSUES NOTED ONE HOUR AFTER THE COMPLETION OF THE PROCEDURE THE PATIENT SUFFERED A MINOR STROKE WAS TREATED WITH MEDICATION AND WAS DOING FINE IN THE OPINION OF THE ATTENDING PHYSICIAN THE PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE STROKE EXPERIENCED BY THE PATIENT NO ADDITIONAL INFORMATION WAS PROVIDED | Cerebrovascular Accident | Patient problem code
6539119 | 74253469 | D | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE NONTORTUOUS STENOSED BELOW THE BIFURCATION OF THE INTERNAL CAROTID ARTERY A X MM XACT STENT DELIVERY SYSTEM SDS WAS SELECTED FOR THE PROCEDURE THE STENT IMPLANT WAS SUCCESSFULLY DEPLOYED WITH NO ISSUES NOTED ONE HOUR AFTER THE COMPLETION OF THE PROCEDURE THE PATIENT SUFFERED A MINOR STROKE WAS TREATED WITH MEDICATION AND WAS DOING FINE IN THE OPINION OF THE ATTENDING PHYSICIAN THE PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE STROKE EXPERIENCED BY THE PATIENT NO ADDITIONAL INFORMATION WAS PROVIDED | Adverse Event Without Identified Device or Use Problem | Device problem code
6534444 | 74259246 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIME TEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO COMPLETED THE FUNCTIONAL TESTING DUE TO MOTOR ERROR ALARM UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6534444 | 74259246 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIME TEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO COMPLETED THE FUNCTIONAL TESTING DUE TO MOTOR ERROR ALARM UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6534444 | 74259246 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIME TEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO COMPLETED THE FUNCTIONAL TESTING DUE TO MOTOR ERROR ALARM UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6534444 | 74259246 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIME TEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO COMPLETED THE FUNCTIONAL TESTING DUE TO MOTOR ERROR ALARM UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Operational Problem Identified | Medical Evaluation result code
6542192 | 74357750 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL AND DROPPED TO MGDL DUE TO OVER BOLUSING THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH FOOD THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6542192 | 74357750 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL AND DROPPED TO MGDL DUE TO OVER BOLUSING THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH FOOD THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Device Unsafe to Use in Environment | Device problem code
6542214 | 74358655 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE AROUND MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Device Unsafe to Use in Environment | Device problem code
6542214 | 74358655 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE AROUND MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6538523 | 74362709 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6538523 | 74362709 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6538523 | 74362709 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6538523 | 74362709 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Operational Problem Identified | Medical Evaluation result code
6537885 | 74362841 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLD MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED NO MOTOR ERROR ALARM NOTED SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Failure to Prime | Device problem code
6537885 | 74362841 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLD MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED NO MOTOR ERROR ALARM NOTED SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6537885 | 74362841 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLD MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED NO MOTOR ERROR ALARM NOTED SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Device Alarm System | Device problem code
6537885 | 74362841 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLD MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED NO MOTOR ERROR ALARM NOTED SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6537885 | 74362841 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLD MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED NO MOTOR ERROR ALARM NOTED SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Operational Problem Identified | Medical Evaluation result code
6537851 | 74365364 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6537851 | 74365364 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6537851 | 74365364 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6537778 | 74394060 | N | UNIT UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLDUNIT PASSED THE BASIC OCCLUSION TEST DISPLACEMENT TEST AND U BOLUSDELIVERY NO MOTOR ERROR ALARMS OCCURRED DURING TEST MOTOR TESTED OKSCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Failure to Prime | Device problem code
6537778 | 74394060 | N | UNIT UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLDUNIT PASSED THE BASIC OCCLUSION TEST DISPLACEMENT TEST AND U BOLUSDELIVERY NO MOTOR ERROR ALARMS OCCURRED DURING TEST MOTOR TESTED OKSCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6537778 | 74394060 | N | UNIT UNABLE TO PRIME DURING PRIMEA TEST DUE TO FAULTY FSRGOLDUNIT PASSED THE BASIC OCCLUSION TEST DISPLACEMENT TEST AND U BOLUSDELIVERY NO MOTOR ERROR ALARMS OCCURRED DURING TEST MOTOR TESTED OKSCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6538239 | 74406861 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMATHERM PUMP WAS NOT WORKING THE USER FACILITY HAD A BACKUP DEVICE THAT ALLOWED THE PROCEDURE TO CONTINUE SUCCESSFULLY CSZ RD PARTY TECHNICIAN B FOUND A LOOSE CONNECTION BETWEEN THE F FUSE AND THE FUSE HOLDER THE TECHNICIAN SECURED THE FUSE TO RESOLVE THE ISSUE | Loose or Intermittent Connection | Device problem code
6538239 | 74406861 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMATHERM PUMP WAS NOT WORKING THE USER FACILITY HAD A BACKUP DEVICE THAT ALLOWED THE PROCEDURE TO CONTINUE SUCCESSFULLY CSZ RD PARTY TECHNICIAN B FOUND A LOOSE CONNECTION BETWEEN THE F FUSE AND THE FUSE HOLDER THE TECHNICIAN SECURED THE FUSE TO RESOLVE THE ISSUE | Detachment of Device or device Component | Device problem code
6538239 | 74406861 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMATHERM PUMP WAS NOT WORKING THE USER FACILITY HAD A BACKUP DEVICE THAT ALLOWED THE PROCEDURE TO CONTINUE SUCCESSFULLY CSZ RD PARTY TECHNICIAN B FOUND A LOOSE CONNECTION BETWEEN THE F FUSE AND THE FUSE HOLDER THE TECHNICIAN SECURED THE FUSE TO RESOLVE THE ISSUE | Connection Problem | Device problem code
6538239 | 74406861 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMATHERM PUMP WAS NOT WORKING THE USER FACILITY HAD A BACKUP DEVICE THAT ALLOWED THE PROCEDURE TO CONTINUE SUCCESSFULLY CSZ RD PARTY TECHNICIAN B FOUND A LOOSE CONNECTION BETWEEN THE F FUSE AND THE FUSE HOLDER THE TECHNICIAN SECURED THE FUSE TO RESOLVE THE ISSUE | Interoperability Problem Identified | Medical Evaluation result code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6537643 | 74409133 | N | UNIT PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST UNIT RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING SCRATCHED LCD WINDOW NOTED DURING VISUAL INSPECTION NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6544995 | 74438232 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B DUE TO DIABETIC KETOACIDOSIS AND HIGH BLOOD GLUCOSE THE CUSTOMERS OLD INSULIN PUMP WAS MALFUNCTIONING AND THEIR BLOOD GLUCOSE LEVELS WERE HIGH THE CUSTOMER TREATED THEIR HIGH BLOOD GLUCOSE WITH A SLIDING SCALE AND SYRINGES THE CUSTOMERS BLOOD GLUCOSE LEVEL AT THE TIME OF ADMISSION WAS IN THE RANGE OF MGDL THE CUSTOMER ALSO HAD AT LEAST TWO BLOOD GLUCOSE READINGS THAT WERE IN THE RANGE OF MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE HOSPITALIZATION THE CUSTOMER WAS OFF THE INSULIN PUMP FOR LESS THAN HOURS PRIOR TO THE HOSPITALIZATION THE INSULIN PUMP WAS ALREADY BEING REPLACED THE CUSTOMER DID NOT WANT TO CONTINUE TROUBLESHOOTING THE CUSTOMER WAS ADVISED TO CALL IS THEY NEEDED ANY ASSISTANCE | Diabetic Ketoacidosis | Patient problem code
6544995 | 74438232 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B DUE TO DIABETIC KETOACIDOSIS AND HIGH BLOOD GLUCOSE THE CUSTOMERS OLD INSULIN PUMP WAS MALFUNCTIONING AND THEIR BLOOD GLUCOSE LEVELS WERE HIGH THE CUSTOMER TREATED THEIR HIGH BLOOD GLUCOSE WITH A SLIDING SCALE AND SYRINGES THE CUSTOMERS BLOOD GLUCOSE LEVEL AT THE TIME OF ADMISSION WAS IN THE RANGE OF MGDL THE CUSTOMER ALSO HAD AT LEAST TWO BLOOD GLUCOSE READINGS THAT WERE IN THE RANGE OF MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE HOSPITALIZATION THE CUSTOMER WAS OFF THE INSULIN PUMP FOR LESS THAN HOURS PRIOR TO THE HOSPITALIZATION THE INSULIN PUMP WAS ALREADY BEING REPLACED THE CUSTOMER DID NOT WANT TO CONTINUE TROUBLESHOOTING THE CUSTOMER WAS ADVISED TO CALL IS THEY NEEDED ANY ASSISTANCE | Hyperglycemia | Patient problem code
6544995 | 74438232 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B DUE TO DIABETIC KETOACIDOSIS AND HIGH BLOOD GLUCOSE THE CUSTOMERS OLD INSULIN PUMP WAS MALFUNCTIONING AND THEIR BLOOD GLUCOSE LEVELS WERE HIGH THE CUSTOMER TREATED THEIR HIGH BLOOD GLUCOSE WITH A SLIDING SCALE AND SYRINGES THE CUSTOMERS BLOOD GLUCOSE LEVEL AT THE TIME OF ADMISSION WAS IN THE RANGE OF MGDL THE CUSTOMER ALSO HAD AT LEAST TWO BLOOD GLUCOSE READINGS THAT WERE IN THE RANGE OF MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE HOSPITALIZATION THE CUSTOMER WAS OFF THE INSULIN PUMP FOR LESS THAN HOURS PRIOR TO THE HOSPITALIZATION THE INSULIN PUMP WAS ALREADY BEING REPLACED THE CUSTOMER DID NOT WANT TO CONTINUE TROUBLESHOOTING THE CUSTOMER WAS ADVISED TO CALL IS THEY NEEDED ANY ASSISTANCE | Component or Accessory Incompatibility | Device problem code
6545306 | 74439638 | D | THE CUSTOMER REPORTED THAT HE WAS RECENTLY HOSPITALIZED DUE TO A BLOOD GLUCOSE LEVEL OF MGDL CUSTOMER STATED THAT HE WAS TRANSPORTED VIA AMBULANCE AND ADMITTED THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE HOSPITALIZATION BUT HAD BEEN OFF FOR LESS THAN HOURS THE CUSTOMER STATED THAT HE BELIEVED THE DEVICE WAS OVERDELIVERING INSULIN DURING TROUBLESHOOTING THE INSULIN PUMP DID NOT SHOW ANY SIGN OF MALFUNCTION THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | Excess Flow or Over-Infusion | Device problem code
6545306 | 74439638 | D | THE CUSTOMER REPORTED THAT HE WAS RECENTLY HOSPITALIZED DUE TO A BLOOD GLUCOSE LEVEL OF MGDL CUSTOMER STATED THAT HE WAS TRANSPORTED VIA AMBULANCE AND ADMITTED THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE HOSPITALIZATION BUT HAD BEEN OFF FOR LESS THAN HOURS THE CUSTOMER STATED THAT HE BELIEVED THE DEVICE WAS OVERDELIVERING INSULIN DURING TROUBLESHOOTING THE INSULIN PUMP DID NOT SHOW ANY SIGN OF MALFUNCTION THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6545563 | 74440349 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT SHE RECENTLY WENT TO THE EMERGENCY ROOM DUE TO A BLOOD GLUCOSE LEVEL OF MGDL BLOOD GLUCOSE LEVEL AT THE TIME OF THE CALL WAS MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE EMERGENCY ROOM VISIT BUT HAD BEEN OFF FOR LESS THAN HOURS DURING TROUBLESHOOTING THE INSULIN PUMP DID NOT SHOW ANY MALFUNCTION BUT CUSTOMER STATED THAT SHE THOUGHT IT WAS OVERDELIVERING THE CUSTOMER WAS ADVISED THAT THE INSULIN PUMP WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | Excess Flow or Over-Infusion | Device problem code
6545563 | 74440349 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT SHE RECENTLY WENT TO THE EMERGENCY ROOM DUE TO A BLOOD GLUCOSE LEVEL OF MGDL BLOOD GLUCOSE LEVEL AT THE TIME OF THE CALL WAS MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF THE EMERGENCY ROOM VISIT BUT HAD BEEN OFF FOR LESS THAN HOURS DURING TROUBLESHOOTING THE INSULIN PUMP DID NOT SHOW ANY MALFUNCTION BUT CUSTOMER STATED THAT SHE THOUGHT IT WAS OVERDELIVERING THE CUSTOMER WAS ADVISED THAT THE INSULIN PUMP WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6542194 | 74477035 | N | THE INSULIN PUMP WAS RECEIVED WITH NO BUTTON RESPONSE DUE TO UNLOCKED CONNECTOR ON LCD BOARD NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6542194 | 74477035 | N | THE INSULIN PUMP WAS RECEIVED WITH NO BUTTON RESPONSE DUE TO UNLOCKED CONNECTOR ON LCD BOARD NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6542095 | 74477261 | N | RELIABILITY ANALYSIS INSPECTED ONE OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Component or Accessory Incompatibility | Device problem code
6542095 | 74477261 | N | RELIABILITY ANALYSIS INSPECTED ONE OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Interoperability Problem Identified | Medical Evaluation result code
6542257 | 74487047 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME TEST AND DISPLACEMENT TESTS THE DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DEVICE HAD SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6542257 | 74487047 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME TEST AND DISPLACEMENT TESTS THE DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DEVICE HAD SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6542257 | 74487047 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME TEST AND DISPLACEMENT TESTS THE DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DEVICE HAD SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6542143 | 74490077 | N | THE INSULIN PUMP WAS RECEIVED WITH DOG CHEW MARKS ON CASE NEAR KEYPAD OVERLAY AREA AND MISSING RESERVOIR TUBE RETAINER | Component Missing | Device problem code
6542143 | 74490077 | N | THE INSULIN PUMP WAS RECEIVED WITH DOG CHEW MARKS ON CASE NEAR KEYPAD OVERLAY AREA AND MISSING RESERVOIR TUBE RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
6542143 | 74490077 | N | THE INSULIN PUMP WAS RECEIVED WITH DOG CHEW MARKS ON CASE NEAR KEYPAD OVERLAY AREA AND MISSING RESERVOIR TUBE RETAINER | Stress Problem Identified | Medical Evaluation result code
6542505 | 74498497 | N | PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION AND PRIMEA TESTS UNIT WAS RECEIVED WITH STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUSBASAL DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6542505 | 74498497 | N | PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION AND PRIMEA TESTS UNIT WAS RECEIVED WITH STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUSBASAL DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6542505 | 74498497 | N | PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION AND PRIMEA TESTS UNIT WAS RECEIVED WITH STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUSBASAL DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6542505 | 74498497 | N | PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION AND PRIMEA TESTS UNIT WAS RECEIVED WITH STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUSBASAL DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6216412 | 74512840 | N | THE PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE FORMATTED HISTORY FILE CONFIRMED TWO PUMP ERROR ALARMS LINE NUMBER FILE NUMBER AND ONE PUMP ERROR ALARM LINE NUMBER FILE NUMBER DUE TO A SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A SCRATCHED CASE A FADING SERIAL NUMBER LABEL A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Scratched Material | Device problem code
6216412 | 74512840 | N | THE PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE FORMATTED HISTORY FILE CONFIRMED TWO PUMP ERROR ALARMS LINE NUMBER FILE NUMBER AND ONE PUMP ERROR ALARM LINE NUMBER FILE NUMBER DUE TO A SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A SCRATCHED CASE A FADING SERIAL NUMBER LABEL A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
6545312 | 74513703 | N | THE INSULIN PUMP HAD A MOTOR ERROR ALARM DURING REWIND DUE TO A MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO A MOTOR ERROR ALARM THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6545312 | 74513703 | N | THE INSULIN PUMP HAD A MOTOR ERROR ALARM DURING REWIND DUE TO A MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO A MOTOR ERROR ALARM THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6548418 | 74544139 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER WAS TREATED WITH IV AND SHOTS CUSTOMERS BLOOD GLUCOSE AT THE TIME OF CALL WAS MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN FORTY EIGHT HOURS CUSTOMER DECLINED TO TROUBLESHOOT FOR HIGH BLOOD GLUCOSE CUSTOMER STATED THAT WHEN THEY TOOK RESERVOIR OUT IT WAS WET SO THEY WERE ASSISTED WITH PUMP EXPOSED TO FLUID TROUBLESHOOTING CUSTOMER WAS ADVISED TO REPLACE RESERVOIR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Fluid Leak | Device problem code
6548418 | 74544139 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER WAS TREATED WITH IV AND SHOTS CUSTOMERS BLOOD GLUCOSE AT THE TIME OF CALL WAS MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN FORTY EIGHT HOURS CUSTOMER DECLINED TO TROUBLESHOOT FOR HIGH BLOOD GLUCOSE CUSTOMER STATED THAT WHEN THEY TOOK RESERVOIR OUT IT WAS WET SO THEY WERE ASSISTED WITH PUMP EXPOSED TO FLUID TROUBLESHOOTING CUSTOMER WAS ADVISED TO REPLACE RESERVOIR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6548425 | 74546275 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THE CUSTOMER WAS HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN FORTY EIGHT HOURS THE CUSTOMER CALLED BACK THE FOLLOWING DAY TO PROVIDE FURTHER INFORMATION REGARDING HOSPITALIZATION AND WAS ASSISTED WITH HIGH BLOOD GLUCOSE TROUBLESHOOTING CUSTOMER WAS ADVISED TO CHANGE THE INFUSION SET RESERVOIR AND INSULIN AND TO TREAT PER HEALTHCARE PROFESSIONALS RECOMMENDATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6548425 | 74546275 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THE CUSTOMER WAS HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN FORTY EIGHT HOURS THE CUSTOMER CALLED BACK THE FOLLOWING DAY TO PROVIDE FURTHER INFORMATION REGARDING HOSPITALIZATION AND WAS ASSISTED WITH HIGH BLOOD GLUCOSE TROUBLESHOOTING CUSTOMER WAS ADVISED TO CHANGE THE INFUSION SET RESERVOIR AND INSULIN AND TO TREAT PER HEALTHCARE PROFESSIONALS RECOMMENDATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6548425 | 74546275 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THE CUSTOMER WAS HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL THE CUSTOMER EXPERIENCED DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN FORTY EIGHT HOURS THE CUSTOMER CALLED BACK THE FOLLOWING DAY TO PROVIDE FURTHER INFORMATION REGARDING HOSPITALIZATION AND WAS ASSISTED WITH HIGH BLOOD GLUCOSE TROUBLESHOOTING CUSTOMER WAS ADVISED TO CHANGE THE INFUSION SET RESERVOIR AND INSULIN AND TO TREAT PER HEALTHCARE PROFESSIONALS RECOMMENDATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6549641 | 74580079 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMER STATED THAT THE EMERGENCY MEDICAL SERVICES CAME FOR THE LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY WERE GIVEN FOOD TO TREAT THE BLOOD GLUCOSE AND WENT UP TO MGDL THE CUSTOMER STATED THAT THEY WERE WEARING THE INSULIN PUMP AT THE TIME OF EVENT THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6549641 | 74580079 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMER STATED THAT THE EMERGENCY MEDICAL SERVICES CAME FOR THE LOW BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY WERE GIVEN FOOD TO TREAT THE BLOOD GLUCOSE AND WENT UP TO MGDL THE CUSTOMER STATED THAT THEY WERE WEARING THE INSULIN PUMP AT THE TIME OF EVENT THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Display or Visual Feedback Problem | Device problem code
6549534 | 74580418 | D | THE CUSTOMERS SON REPORTED VIA PHONE CALL THAT THE CUSTOMER WAS HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS DRIVING WHEN HE EXPERIENCED THE LOW HE RAN OFF THE ROAD AND HIT A CONCRETE BOX CONVERTER CUSTOMER WAS THEN TAKEN TO THE HOSPITAL AND HAS BEEN EXPERIENCING LOW BLOOD GLUCOSE FOR THE PAST HOUR THE CUSTOMERS SON DID NOT KNOW WHEN THE LAST SET CHANGE WAS PERFORMED THE CUSTOMER WAS POSSIBLY TREATED WITH FLUIDS DURING THE AMBULANCE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION AND HAD BEEN OF THE DEVICE FOR GREATER THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6549534 | 74580418 | D | THE CUSTOMERS SON REPORTED VIA PHONE CALL THAT THE CUSTOMER WAS HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS DRIVING WHEN HE EXPERIENCED THE LOW HE RAN OFF THE ROAD AND HIT A CONCRETE BOX CONVERTER CUSTOMER WAS THEN TAKEN TO THE HOSPITAL AND HAS BEEN EXPERIENCING LOW BLOOD GLUCOSE FOR THE PAST HOUR THE CUSTOMERS SON DID NOT KNOW WHEN THE LAST SET CHANGE WAS PERFORMED THE CUSTOMER WAS POSSIBLY TREATED WITH FLUIDS DURING THE AMBULANCE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION AND HAD BEEN OF THE DEVICE FOR GREATER THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6549770 | 74582586 | D | COMPANY REPRESENTATIVE REPORTED VIA EMAIL THE CUSTOMER HAD RELAYED SHE WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE MGDL CUSTOMER STATED HER BLOOD GLUCOSE WERE FINE AND SUDDENLY HER BLOOD GLUCOSE STARTED RISING CUSTOMER WAS WITH HER MOTHER WHO CALLED EMERGENCY MEDICAL SERVICES AS SHE HAD HIGH BLOOD GLUCOSE AND WAS VOMITING CUSTOMER DID NOT RECALL IF THE INSULIN PUMP HAD ALARMED OR IS SHE HAD ISSUES WITH THE INFUSION SET OR RESERVOIR CUSTOMER WAS HOSPITALIZED FOR A MONTH AND THREE WEEKS IN INTENSIVE CARE UNIT ICU UPON RELEASE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND START MANUAL INJECTIONS AS DOCTOR BELIEVED THE INSULIN PUMP MIGHT BE MALFUNCTIONING DURING THE HOSPITALIZATION CUSTOMER WENT INTO COMA AND WAS ON LIFE SUPPORT FOR OVER A WEEK THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT SHE HAD BEEN OFF LESS THAN HOURS CUSTOMER WAS UNABLE TO PERFORM TROUBLESHOOTING ON THE INSULIN PUMP AS THE CUSTOMER DID NOT HAVE A BATTERY FOR THE DEVICE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Component or Accessory Incompatibility | Device problem code
6549770 | 74582586 | D | COMPANY REPRESENTATIVE REPORTED VIA EMAIL THE CUSTOMER HAD RELAYED SHE WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE MGDL CUSTOMER STATED HER BLOOD GLUCOSE WERE FINE AND SUDDENLY HER BLOOD GLUCOSE STARTED RISING CUSTOMER WAS WITH HER MOTHER WHO CALLED EMERGENCY MEDICAL SERVICES AS SHE HAD HIGH BLOOD GLUCOSE AND WAS VOMITING CUSTOMER DID NOT RECALL IF THE INSULIN PUMP HAD ALARMED OR IS SHE HAD ISSUES WITH THE INFUSION SET OR RESERVOIR CUSTOMER WAS HOSPITALIZED FOR A MONTH AND THREE WEEKS IN INTENSIVE CARE UNIT ICU UPON RELEASE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND START MANUAL INJECTIONS AS DOCTOR BELIEVED THE INSULIN PUMP MIGHT BE MALFUNCTIONING DURING THE HOSPITALIZATION CUSTOMER WENT INTO COMA AND WAS ON LIFE SUPPORT FOR OVER A WEEK THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT SHE HAD BEEN OFF LESS THAN HOURS CUSTOMER WAS UNABLE TO PERFORM TROUBLESHOOTING ON THE INSULIN PUMP AS THE CUSTOMER DID NOT HAVE A BATTERY FOR THE DEVICE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Vomiting | Patient problem code
6549770 | 74582586 | D | COMPANY REPRESENTATIVE REPORTED VIA EMAIL THE CUSTOMER HAD RELAYED SHE WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE MGDL CUSTOMER STATED HER BLOOD GLUCOSE WERE FINE AND SUDDENLY HER BLOOD GLUCOSE STARTED RISING CUSTOMER WAS WITH HER MOTHER WHO CALLED EMERGENCY MEDICAL SERVICES AS SHE HAD HIGH BLOOD GLUCOSE AND WAS VOMITING CUSTOMER DID NOT RECALL IF THE INSULIN PUMP HAD ALARMED OR IS SHE HAD ISSUES WITH THE INFUSION SET OR RESERVOIR CUSTOMER WAS HOSPITALIZED FOR A MONTH AND THREE WEEKS IN INTENSIVE CARE UNIT ICU UPON RELEASE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND START MANUAL INJECTIONS AS DOCTOR BELIEVED THE INSULIN PUMP MIGHT BE MALFUNCTIONING DURING THE HOSPITALIZATION CUSTOMER WENT INTO COMA AND WAS ON LIFE SUPPORT FOR OVER A WEEK THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT SHE HAD BEEN OFF LESS THAN HOURS CUSTOMER WAS UNABLE TO PERFORM TROUBLESHOOTING ON THE INSULIN PUMP AS THE CUSTOMER DID NOT HAVE A BATTERY FOR THE DEVICE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Coma | Patient problem code
6549770 | 74582586 | D | COMPANY REPRESENTATIVE REPORTED VIA EMAIL THE CUSTOMER HAD RELAYED SHE WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE MGDL CUSTOMER STATED HER BLOOD GLUCOSE WERE FINE AND SUDDENLY HER BLOOD GLUCOSE STARTED RISING CUSTOMER WAS WITH HER MOTHER WHO CALLED EMERGENCY MEDICAL SERVICES AS SHE HAD HIGH BLOOD GLUCOSE AND WAS VOMITING CUSTOMER DID NOT RECALL IF THE INSULIN PUMP HAD ALARMED OR IS SHE HAD ISSUES WITH THE INFUSION SET OR RESERVOIR CUSTOMER WAS HOSPITALIZED FOR A MONTH AND THREE WEEKS IN INTENSIVE CARE UNIT ICU UPON RELEASE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND START MANUAL INJECTIONS AS DOCTOR BELIEVED THE INSULIN PUMP MIGHT BE MALFUNCTIONING DURING THE HOSPITALIZATION CUSTOMER WENT INTO COMA AND WAS ON LIFE SUPPORT FOR OVER A WEEK THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT SHE HAD BEEN OFF LESS THAN HOURS CUSTOMER WAS UNABLE TO PERFORM TROUBLESHOOTING ON THE INSULIN PUMP AS THE CUSTOMER DID NOT HAVE A BATTERY FOR THE DEVICE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6550116 | 74582926 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER HAD TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE CALL THE CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE THE CUSTOMER ALSO REPORTED GETTING A SECOND OCCURENCE OF REPLACE BATTERY ALARM WITHOUT A LOW BATTERY WARNING CUSTOMER WAS ADVISED THAT PUMP WILL NEED TO BE REPLACED THE PRODUCT WILL BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6550116 | 74582926 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER HAD TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE CALL THE CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE THE CUSTOMER ALSO REPORTED GETTING A SECOND OCCURENCE OF REPLACE BATTERY ALARM WITHOUT A LOW BATTERY WARNING CUSTOMER WAS ADVISED THAT PUMP WILL NEED TO BE REPLACED THE PRODUCT WILL BE RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6550240 | 74583442 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY HAD BEEN EXPERIENCING UNEXPLAINABLE HIGH BLOOD GLUCOSE THE CUSTOMER REPORTED THAT THEY HAD GONE THROUGH A FULL BODY SCANNER WITH THEIR INSULIN PUMP ON NUMEROUS TIMES THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL THEY ALSO REPORTED THAT THEY EXPERIENCED A HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER HAD BEEN TRYING TO CORRECT THEIR HIGH BLOOD GLUCOSE WITH THE INSULIN PUMP TROUBLESHOOTING WAS PERFORMED AND INSULIN WAS ABLE TO EXIT WITHOUT INCIDENT THE CUSTOMER WILL BE SENT A TUBING CLAMP TO PERFORM THE OCCLUSION TEST THEY WILL CALL BACK ONCE THEY RECEIVE THE TUBING CLAMP THE DEVICE WILL NOT RETURN FOR ANALYSIS | Hyperglycemia | Patient problem code
6550240 | 74583442 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY HAD BEEN EXPERIENCING UNEXPLAINABLE HIGH BLOOD GLUCOSE THE CUSTOMER REPORTED THAT THEY HAD GONE THROUGH A FULL BODY SCANNER WITH THEIR INSULIN PUMP ON NUMEROUS TIMES THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL THEY ALSO REPORTED THAT THEY EXPERIENCED A HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER HAD BEEN TRYING TO CORRECT THEIR HIGH BLOOD GLUCOSE WITH THE INSULIN PUMP TROUBLESHOOTING WAS PERFORMED AND INSULIN WAS ABLE TO EXIT WITHOUT INCIDENT THE CUSTOMER WILL BE SENT A TUBING CLAMP TO PERFORM THE OCCLUSION TEST THEY WILL CALL BACK ONCE THEY RECEIVE THE TUBING CLAMP THE DEVICE WILL NOT RETURN FOR ANALYSIS | Component or Accessory Incompatibility | Device problem code
6545278 | 74589289 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6545278 | 74589289 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6545278 | 74589289 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6545278 | 74589289 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6545875 | 74591351 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURNT MARKS ON THE CIRCUIT BOARD WHICH RESULTED IN THE DEVICE POWER ANOMALY | Intermittent loss of power | Device problem code
6545875 | 74591351 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURNT MARKS ON THE CIRCUIT BOARD WHICH RESULTED IN THE DEVICE POWER ANOMALY | Power Source Problem Identified | Medical Evaluation result code
6544927 | 74606598 | D | IT WAS REPORTED OF REPLACE BATTERY NOW ALARM THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE PUMP ASKED HER TO REPLACE THE BATTERY EVERY HOURS THE CUSTOMER MENTIONED THAT SHE DID NOT GET A LOW BATTERY WARNING THE CUSTOMER DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THE PRODUCT WAS RETURNED FOR ANALYSIS | Battery Problem: Low Impedance | Device problem code
6544927 | 74606598 | D | IT WAS REPORTED OF REPLACE BATTERY NOW ALARM THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE PUMP ASKED HER TO REPLACE THE BATTERY EVERY HOURS THE CUSTOMER MENTIONED THAT SHE DID NOT GET A LOW BATTERY WARNING THE CUSTOMER DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THE PRODUCT WAS RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6544761 | 74632943 | N | MOTOR ERROR ALARMED DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6544761 | 74632943 | N | MOTOR ERROR ALARMED DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6544761 | 74632943 | N | MOTOR ERROR ALARMED DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6544889 | 74633315 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNIT RETURNED WITH MISSING END CAP STICKER AND SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6545440 | 74633318 | N | THE INSUILN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE WAS RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS BASAL DELIVERY AND ERROR ALARM CONFIRMED IN THEHISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED THE DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6545440 | 74633318 | N | THE INSUILN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE WAS RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS BASAL DELIVERY AND ERROR ALARM CONFIRMED IN THEHISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED THE DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6545440 | 74633318 | N | THE INSUILN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE WAS RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS BASAL DELIVERY AND ERROR ALARM CONFIRMED IN THEHISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED THE DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6545440 | 74633318 | N | THE INSUILN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE WAS RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS BASAL DELIVERY AND ERROR ALARM CONFIRMED IN THEHISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED THE DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6545440 | 74633318 | N | THE INSUILN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE WAS RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS BASAL DELIVERY AND ERROR ALARM CONFIRMED IN THEHISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED THE DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Inaccurate Delivery | Device Problem Code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Component Missing | Device problem code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Failure to Prime | Device problem code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Failure to Deliver | Device problem code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Operational Problem Identified | Medical Evaluation result code
6547984 | 74633393 | N | UNABLE TO PRIME DURING PRIMEDELIVERY ALARM TEST DUE TO FAULTY BACK PRESSURE SENSOR GOLD UNABLE TO PERFORM BASIC OCCLUSION OCCLUSION AND EXCESSIVE NO DELIVERY ALARM TESTS MISSING END CAP STICKER NOTED NO A ALARMS NOTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
6547941 | 74634696 | N | UNIT WAS RECEIVED WITH CURRENTS IN SPEC MOTOR ERROR ALARM DURING THE PRIMEDELIVERY TEST DUE TO GEARBOX JAMMED | Component or Accessory Incompatibility | Device Problem Code
6547941 | 74634696 | N | UNIT WAS RECEIVED WITH CURRENTS IN SPEC MOTOR ERROR ALARM DURING THE PRIMEDELIVERY TEST DUE TO GEARBOX JAMMED | Mechanical Jam | Device problem code
6547941 | 74634696 | N | UNIT WAS RECEIVED WITH CURRENTS IN SPEC MOTOR ERROR ALARM DURING THE PRIMEDELIVERY TEST DUE TO GEARBOX JAMMED | Mechanical Problem Identified | Medical Evaluation result code
6547941 | 74634696 | N | UNIT WAS RECEIVED WITH CURRENTS IN SPEC MOTOR ERROR ALARM DURING THE PRIMEDELIVERY TEST DUE TO GEARBOX JAMMED | Interoperability Problem Identified | Medical Evaluation result code
6544898 | 74634922 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIMEATEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6544898 | 74634922 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIMEATEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6544898 | 74634922 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIMEATEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6544898 | 74634922 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIMEATEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6544898 | 74634922 | N | UNIT PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION AND PRIMEATEST UNIT HAD MOTOR ERROR STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNIT RECEIVED WITH SCRATCHED ON DISPLAY WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6548226 | 74639327 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548226 | 74639327 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6548226 | 74639327 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6548226 | 74639327 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND AND BASIC OCCLUSION TEST DEVICE WAS RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE NOT DETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST DEVICE WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6551824 | 74641113 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH FOOD THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6551824 | 74641113 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH FOOD THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Device Unsafe to Use in Environment | Device problem code
6547979 | 74641812 | N | PUMP RECEIVED STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND E ALARM NOTED IN ALARM HISTORY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED MOTOR MAY HAVE HAD INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING HOWEVER PUMP PASSED DISPLACEMENT TEST AND BASIC OCCLUSION TESTS | Component or Accessory Incompatibility | Device problem code
6547979 | 74641812 | N | PUMP RECEIVED STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND E ALARM NOTED IN ALARM HISTORY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED MOTOR MAY HAVE HAD INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING HOWEVER PUMP PASSED DISPLACEMENT TEST AND BASIC OCCLUSION TESTS | Mechanical Jam | Device problem code
6547979 | 74641812 | N | PUMP RECEIVED STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND E ALARM NOTED IN ALARM HISTORY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED MOTOR MAY HAVE HAD INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING HOWEVER PUMP PASSED DISPLACEMENT TEST AND BASIC OCCLUSION TESTS | Mechanical Problem Identified | Medical Evaluation result code
6547979 | 74641812 | N | PUMP RECEIVED STUCK IN MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND E ALARM NOTED IN ALARM HISTORY MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED MOTOR MAY HAVE HAD INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING HOWEVER PUMP PASSED DISPLACEMENT TEST AND BASIC OCCLUSION TESTS | Interoperability Problem Identified | Medical Evaluation result code
6548237 | 74643229 | N | UNIT ALARMED MOTOR ERROR DURING BASIC OCCLUSION TEST DUE TO FAULTY FSRGOLD THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED UNIT RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548237 | 74643229 | N | UNIT ALARMED MOTOR ERROR DURING BASIC OCCLUSION TEST DUE TO FAULTY FSRGOLD THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED UNIT RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6548237 | 74643229 | N | UNIT ALARMED MOTOR ERROR DURING BASIC OCCLUSION TEST DUE TO FAULTY FSRGOLD THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED UNIT RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6547981 | 74647095 | N | MOTOR ERROR ALARMED DURING REWIND DUE TO OUT OF PHASE MOTOR ENCODER SIGNAL UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS | Component or Accessory Incompatibility | Device Problem Code
6547981 | 74647095 | N | MOTOR ERROR ALARMED DURING REWIND DUE TO OUT OF PHASE MOTOR ENCODER SIGNAL UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MOTOR ERROR ALARMS | Interoperability Problem Identified | Medical Evaluation result code
6548379 | 74649974 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO EXCESSIVE MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6548379 | 74649974 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO EXCESSIVE MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548379 | 74649974 | N | UNIT RECEIVED MOTOR ERROR ALARMS DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM DISPLACEMENT TEST DUE TO EXCESSIVE MOTOR ERROR ALARMS SCRATCHED ON DISPLAY WINDOW NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6547936 | 74651269 | N | UNIT FAILED DISPLACEMENT TEST UNITS REMAINING ON STATUS SCREEN AND MOTOR ERROR ALARM DURING REWIND DUE TO MOTOR ENCODER OUT OF PHASE | Component or Accessory Incompatibility | Device problem code
6547936 | 74651269 | N | UNIT FAILED DISPLACEMENT TEST UNITS REMAINING ON STATUS SCREEN AND MOTOR ERROR ALARM DURING REWIND DUE TO MOTOR ENCODER OUT OF PHASE | Interoperability Problem Identified | Medical Evaluation result code
6548850 | 74652881 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENTTESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUSDELIVERY THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOTDETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST UNIT HAD CRACKEDBATTERY TUBE THREADS RESERVOIR TUBE LIP SCRATCHED LCD WINDOW AND CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6548850 | 74652881 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENTTESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUSDELIVERY THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOTDETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST UNIT HAD CRACKEDBATTERY TUBE THREADS RESERVOIR TUBE LIP SCRATCHED LCD WINDOW AND CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6548850 | 74652881 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENTTESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUSDELIVERY THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOTDETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST UNIT HAD CRACKEDBATTERY TUBE THREADS RESERVOIR TUBE LIP SCRATCHED LCD WINDOW AND CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548850 | 74652881 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENTTESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUSDELIVERY THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOTDETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST UNIT HAD CRACKEDBATTERY TUBE THREADS RESERVOIR TUBE LIP SCRATCHED LCD WINDOW AND CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6548850 | 74652881 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENTTESTS PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUSDELIVERY THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOTDETECTED DURING OUR TESTING MOTOR PASSED MOTOR TEST UNIT HAD CRACKEDBATTERY TUBE THREADS RESERVOIR TUBE LIP SCRATCHED LCD WINDOW AND CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6548830 | 74657125 | N | PUMP WAS RECEIVED WITH MOTOR ERROR ALARM DURING OCCLUSION TEST AND UNABLE TO PRIME AT LBS DURING PRIMEA TEST DUE TO FAULTY FSRGOLD UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Failure to Prime | Device problem code
6548830 | 74657125 | N | PUMP WAS RECEIVED WITH MOTOR ERROR ALARM DURING OCCLUSION TEST AND UNABLE TO PRIME AT LBS DURING PRIMEA TEST DUE TO FAULTY FSRGOLD UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6548830 | 74657125 | N | PUMP WAS RECEIVED WITH MOTOR ERROR ALARM DURING OCCLUSION TEST AND UNABLE TO PRIME AT LBS DURING PRIMEA TEST DUE TO FAULTY FSRGOLD UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548830 | 74657125 | N | PUMP WAS RECEIVED WITH MOTOR ERROR ALARM DURING OCCLUSION TEST AND UNABLE TO PRIME AT LBS DURING PRIMEA TEST DUE TO FAULTY FSRGOLD UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6548055 | 74660267 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE DEVICE HAD A STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING THE DEVICE HAD A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Mechanical Jam | Device problem code
6548055 | 74660267 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE DEVICE HAD A STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING THE DEVICE HAD A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6548055 | 74660267 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE DEVICE HAD A STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING THE DEVICE HAD A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6548055 | 74660267 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE DEVICE HAD A STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING THE DEVICE HAD A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6548079 | 74673701 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT A REPLACE BATTERY NOW ALERT WAS NOT RECEIVED PRIOR TO A LOW BATTERY WARNING ON THE INSULIN PUMP THE CUSTOMERS BLOOD GLUCOSE READING WAS UNKNOWN AT THE TIME OF INCIDENT TROUBLESHOOTING FOUND THAT THE BATTERIESS HAVE TO BE REPLACED IN THE PUMP TO TIMES A DAY THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE DEVICE AND REVERT TO A BACKUP PLAN PER THEIR DOCTORS INSTRUCTION THE CUSTOMER WAS ALSO ADVISED THAT THE DEVICE WOULD BE REPLACED AND TO RETURN THE PRODUCT FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6548079 | 74673701 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT A REPLACE BATTERY NOW ALERT WAS NOT RECEIVED PRIOR TO A LOW BATTERY WARNING ON THE INSULIN PUMP THE CUSTOMERS BLOOD GLUCOSE READING WAS UNKNOWN AT THE TIME OF INCIDENT TROUBLESHOOTING FOUND THAT THE BATTERIESS HAVE TO BE REPLACED IN THE PUMP TO TIMES A DAY THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE DEVICE AND REVERT TO A BACKUP PLAN PER THEIR DOCTORS INSTRUCTION THE CUSTOMER WAS ALSO ADVISED THAT THE DEVICE WOULD BE REPLACED AND TO RETURN THE PRODUCT FOR ANALYSIS | Battery Problem | Device problem code
6550268 | 74705192 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6550268 | 74705192 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6550106 | 74705946 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TEST PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | Mechanical Jam | Device problem code
6550106 | 74705946 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TEST PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | Component or Accessory Incompatibility | Device problem code
6550106 | 74705946 | N | PUMP PASSED THE REWIND BASIC OCCLUSION PRIMEA AND DISPLACEMENT TEST PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | Interoperability Problem Identified | Medical Evaluation result code
6550335 | 74706300 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6550335 | 74706300 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6537709 | 74713830 | N | THE CUSTOMER STATED THAT THEY HAD HIGH BLOOD GLUCOSE OF MGDL SO THEY TOOK SEVERAL INJECTIONS WHICH RESULTED IN LOW BLOOD GLUCOSE OF MGDL THE PARAMEDICS WERE CALLED AND ADMINISTERED TREATMENT TROUBLESHOOTING WAS PERFORMED FOR THE REPORTED NO DELIVERY ALARM THE TIMEDATE WAS CHECKED AND IT WAS NOT CORRECT ASSISTANCE WAS PROVIDED TO CORRECT THE TIMEDATE SETTING THE INSULIN PUMP REMAINS IN USE | Date/Time-Related Software Problem | Device problem code
6537709 | 74713830 | N | THE CUSTOMER STATED THAT THEY HAD HIGH BLOOD GLUCOSE OF MGDL SO THEY TOOK SEVERAL INJECTIONS WHICH RESULTED IN LOW BLOOD GLUCOSE OF MGDL THE PARAMEDICS WERE CALLED AND ADMINISTERED TREATMENT TROUBLESHOOTING WAS PERFORMED FOR THE REPORTED NO DELIVERY ALARM THE TIMEDATE WAS CHECKED AND IT WAS NOT CORRECT ASSISTANCE WAS PROVIDED TO CORRECT THE TIMEDATE SETTING THE INSULIN PUMP REMAINS IN USE | Incorrect Dimension | Medical Evaluation result code
6537709 | 74713830 | N | THE CUSTOMER STATED THAT THEY HAD HIGH BLOOD GLUCOSE OF MGDL SO THEY TOOK SEVERAL INJECTIONS WHICH RESULTED IN LOW BLOOD GLUCOSE OF MGDL THE PARAMEDICS WERE CALLED AND ADMINISTERED TREATMENT TROUBLESHOOTING WAS PERFORMED FOR THE REPORTED NO DELIVERY ALARM THE TIMEDATE WAS CHECKED AND IT WAS NOT CORRECT ASSISTANCE WAS PROVIDED TO CORRECT THE TIMEDATE SETTING THE INSULIN PUMP REMAINS IN USE | Inaccurate Delivery | Device Problem Code
6550612 | 74752905 | N | THE INSULIN PUMP WAS RECEIVED MOTOR ERROR ALARMS DURING REWIND THE PROBLEM ISOLATED TO MOTOR | Mechanical Problem | Device problem code
6550612 | 74752905 | N | THE INSULIN PUMP WAS RECEIVED MOTOR ERROR ALARMS DURING REWIND THE PROBLEM ISOLATED TO MOTOR | Mechanical Problem Identified | Medical Evaluation result code
6550061 | 74754891 | N | UNIT HAD MISSING SEGMENTS DUE TO BLEEDING LCD GLASS UNIT HAD SCRATCHED DISPLAY WINDOW | Component Missing | Device problem code
6550061 | 74754891 | N | UNIT HAD MISSING SEGMENTS DUE TO BLEEDING LCD GLASS UNIT HAD SCRATCHED DISPLAY WINDOW | Scratched Material | Device problem code
6550061 | 74754891 | N | UNIT HAD MISSING SEGMENTS DUE TO BLEEDING LCD GLASS UNIT HAD SCRATCHED DISPLAY WINDOW | Stress Problem Identified | Medical Evaluation result code
6550061 | 74754891 | N | UNIT HAD MISSING SEGMENTS DUE TO BLEEDING LCD GLASS UNIT HAD SCRATCHED DISPLAY WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
6550551 | 74789958 | N | THE INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE NO MOTION SENSOR TEST FAILURE ALARM NOTED ANALYSIS WAS UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN ALARM HISTORY SCREEN DUE TO MOTOR ERROR ALARMS NO AUDIOBEEP ANOMALY NOTED DURING TESTING | Component or Accessory Incompatibility | Device problem code
6550551 | 74789958 | N | THE INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE NO MOTION SENSOR TEST FAILURE ALARM NOTED ANALYSIS WAS UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN ALARM HISTORY SCREEN DUE TO MOTOR ERROR ALARMS NO AUDIOBEEP ANOMALY NOTED DURING TESTING | Interoperability Problem Identified | Medical Evaluation result code
6550427 | 74790065 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Device Alarm System | Device problem code
6550427 | 74790065 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6550427 | 74790065 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Heavier Menses | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Fatigue | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Blood Loss | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Pain, Abdominal | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Anemia | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Hemorrhage | Patient Problem Code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Hematoma | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Itching | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Arthralgia | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Numbness | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Hair Loss | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Insufficient Information | Device problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Malaise | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Pain | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Occlusion | Patient problem code
6556562 | 74828981 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY B B ON APR THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN LOWER PAIN LOWER ABDOMINAL PAIN IN A FEMALE PATIENT WHO HAD ESSURE INSERTED FOR CONTRACEPTION THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW THE PATIENTS PAST MEDICAL HISTORY INCLUDED CESAREAN SECTION CESAREAN SECTIONS ON AN UNKNOWN DATE THE PATIENT HAD ESSURE INSERTED IN THE PATIENT EXPERIENCED ABDOMINAL PAIN LOWER SERIOUSNESS CRITERIA HOSPITALIZATION DISABILITY AND MEDICALLY SIGNIFICANT FATIGUE CHRONIC FATIGUE MENORRHAGIA HEAVY MENSTRUAL BLEEDING LARGE BLOOD CLOTS DURING MENSTRUAL CYCLE MALAISE SICKNESS BACK PAIN LOWER BACK PAIN ARTHRALGIA JOINT PAINACHES ANAEMIA ANEMIA ALOPECIA HAIR LOSS WEIGHT INCREASED WEIGHT GAIN ABDOMINAL DISTENSION STOMACH BLOATING PRURITUS ITCHING AND HYPOAESTHESIA NUMBNESS AT THE TOP OF LEFT LEG THE PATIENT WAS TREATED WITH FERROUS SULFATE FERROUS SULPHATE PARACETAMOL AND IBUPROFEN IPROFEN AT THE TIME OF THE REPORT THE ABDOMINAL PAIN LOWER FATIGUE MENORRHAGIA MALAISE BACK PAIN ARTHRALGIA ANAEMIA ALOPECIA WEIGHT INCREASED ABDOMINAL DISTENSION AND PRURITUS HAD NOT RESOLVED AND THE HYPOAESTHESIA WAS RESOLVING THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISTENSION ABDOMINAL PAIN LOWER ALOPECIA ANAEMIA ARTHRALGIA BACK PAIN FATIGUE HYPOAESTHESIA MALAISE MENORRHAGIA PRURITUS AND WEIGHT INCREASED WITH ESSURE REPORTERS COMMENT THE REACTION WAS SERIOUS INVOLVED OR PROLONGED INPATIENT HOSPITALISATION INVOLVED PERSISTENT OR SIGNIFICANT DISABILITY OR INCAPACITY MEDICALLY SIGNIFICANT ADMITTED INTO HOSPITAL WITH LOWER ABDOMINAL PAIN CAUSE UNKNOWN PRESCRIBED FERROUS SULPHATE DUE TO BLOOD LOSS DURING MENSTRUAL CYCLE REFERRED BACK TO GYNAECOLOGIST AFTER EXPERIENCING CHRONIC FATIGUE COMPANY CAUSALITY COMMENT A FEMALE CONSUMER HAD ESSURE FALLOPIAN TUBE OCCLUSION INSERT INSERTED AND EXPERIENCED PAIN LOWER ABDOMINAL PAIN THIS EVENT IS REGARDED AS ANTICIPATED ACCORDING TO THE REFERENCE SAFETY INFORMATION FOR ESSURE ABDOMINAL PAIN MAY OCCUR WITH ESSURE THERAPY AND IS HIGHLY PREVALENT IN WOMEN AND MAY HAVE GYNAECOLOGICAL AND NONGYNAECOLOGICAL CAUSES IN THIS CASE NO ALTERNATIVE EXPLANATION WAS PROVIDED BASED ON ITS NATURE A CAUSAL RELATIONSHIP WITH ESSURE CANNOT BE EXCLUDED THIS CASE WAS REGARDED AS INCIDENT DUE TO HOSPITALIZATION AND DISABILITY REPORTED ADDITIONALLY NONSERIOUS EVENTS WERE REPORTED FURTHER INFORMATION AND PRODUCT TECHNICAL ANALYSIS ARE BEING SOUGHT | Disability | Patient problem code
6551767 | 74829034 | N | THE INSULIN PUMP WAS RECEIVED WITH A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6551767 | 74829034 | N | THE INSULIN PUMP WAS RECEIVED WITH A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6554923 | 74891738 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY HAD ISSUES WITH THE BATTERY LIFE ON THEIR INSULIN PUMP THEY RECEIVED MULTIPLE CHANGE BATTERY NOW ALERTS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE ALARM HISTORY WAS REVIEWED AND IT WAS NOTED THAT THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE CHANGE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF THE REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6554923 | 74891738 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY HAD ISSUES WITH THE BATTERY LIFE ON THEIR INSULIN PUMP THEY RECEIVED MULTIPLE CHANGE BATTERY NOW ALERTS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE ALARM HISTORY WAS REVIEWED AND IT WAS NOTED THAT THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE CHANGE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF THE REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6554877 | 74896042 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6554877 | 74896042 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6554960 | 74903052 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY HAD REPLACED THE INSULIN PUMPS BATTERIES TWICE BUT THEY HAD ONLY LAST HOURS THE ALARM HISTORY WAS REVIEWED AND IT WAS NOTED THAT THERE WERE REPLACE BATTERY ERRORS THAT HAD OCCURRED IN THE LAST MINUTES THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL IT WAS ALSO NOTED IN TROUBLESHOOTING THAT THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6554960 | 74903052 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY HAD REPLACED THE INSULIN PUMPS BATTERIES TWICE BUT THEY HAD ONLY LAST HOURS THE ALARM HISTORY WAS REVIEWED AND IT WAS NOTED THAT THERE WERE REPLACE BATTERY ERRORS THAT HAD OCCURRED IN THE LAST MINUTES THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL IT WAS ALSO NOTED IN TROUBLESHOOTING THAT THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6557395 | 74910630 | N | ONE OPENEDUSED SENSOR WAS INSPECTED AND CONTINUITY RESISTANCE TEST WAS PERFORMED AND THE SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6557395 | 74910630 | N | ONE OPENEDUSED SENSOR WAS INSPECTED AND CONTINUITY RESISTANCE TEST WAS PERFORMED AND THE SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6557933 | 74914388 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6557933 | 74914388 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6561200 | 74960213 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL AND INQUIRED ABOUT SITE AND INSULIN USAGE THE CUSTOMER HAD SYMPTOMS SUCH AS NAUSEA AND SHORTNESS OF BREATH THE CUSTOMER TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS MGDL AT THE TIME OF THE CALL TROUBLESHOOTING WAS PERFORMED AND CUSTOMER STATED THAT THERE WERE NO AIR BUBBLES THE CUSTOMER DECLINED TO CONTINUE WITH THE TROUBLESHOOTING AND STATED THAT THEY WILL WAIT TO GET HOME AND DO A SET CHANGE THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Nausea | Patient problem code
6561200 | 74960213 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL AND INQUIRED ABOUT SITE AND INSULIN USAGE THE CUSTOMER HAD SYMPTOMS SUCH AS NAUSEA AND SHORTNESS OF BREATH THE CUSTOMER TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS MGDL AT THE TIME OF THE CALL TROUBLESHOOTING WAS PERFORMED AND CUSTOMER STATED THAT THERE WERE NO AIR BUBBLES THE CUSTOMER DECLINED TO CONTINUE WITH THE TROUBLESHOOTING AND STATED THAT THEY WILL WAIT TO GET HOME AND DO A SET CHANGE THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6561200 | 74960213 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL AND INQUIRED ABOUT SITE AND INSULIN USAGE THE CUSTOMER HAD SYMPTOMS SUCH AS NAUSEA AND SHORTNESS OF BREATH THE CUSTOMER TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS MGDL AT THE TIME OF THE CALL TROUBLESHOOTING WAS PERFORMED AND CUSTOMER STATED THAT THERE WERE NO AIR BUBBLES THE CUSTOMER DECLINED TO CONTINUE WITH THE TROUBLESHOOTING AND STATED THAT THEY WILL WAIT TO GET HOME AND DO A SET CHANGE THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Dyspnea | Patient problem code
6561200 | 74960213 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL AND INQUIRED ABOUT SITE AND INSULIN USAGE THE CUSTOMER HAD SYMPTOMS SUCH AS NAUSEA AND SHORTNESS OF BREATH THE CUSTOMER TREATED WITH THE INSULIN PUMP CUSTOMERS BLOOD GLUCOSE VALUE WAS MGDL AT THE TIME OF THE CALL TROUBLESHOOTING WAS PERFORMED AND CUSTOMER STATED THAT THERE WERE NO AIR BUBBLES THE CUSTOMER DECLINED TO CONTINUE WITH THE TROUBLESHOOTING AND STATED THAT THEY WILL WAIT TO GET HOME AND DO A SET CHANGE THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6558218 | 74991369 | D | IT WAS REPORTED VIA PHONE CALL THAT THEIR INSULIN PUMP HAD ALARMED REPLACE BATTERY NOW CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN AT THE TIME OF THE INCIDENT IT WAS REPORTED THAT TROUBLESHOOTING FOR REPLACE BATTERY NOW WAS PERFORMED IT WAS REPORTED THAT THE INSULIN PUMP WAS NEITHER DROPPED NOR BUMPED IT WAS ALSO STATED THAT THERE WAS NO UNATTENDED ALARMS IT WAS REPORTED THAT THE INSULIN PUMP RECEIVED REPLACE BATTERY NOW WITHOUT RECEIVING LOW BATTERY WARNING TWICE IT WAS REPORTED THAT THE CUSTOMER WAS ADVISED THAT THE INSULIN PUMP NEEDED TO BE REPLACED AND WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6558218 | 74991369 | D | IT WAS REPORTED VIA PHONE CALL THAT THEIR INSULIN PUMP HAD ALARMED REPLACE BATTERY NOW CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN AT THE TIME OF THE INCIDENT IT WAS REPORTED THAT TROUBLESHOOTING FOR REPLACE BATTERY NOW WAS PERFORMED IT WAS REPORTED THAT THE INSULIN PUMP WAS NEITHER DROPPED NOR BUMPED IT WAS ALSO STATED THAT THERE WAS NO UNATTENDED ALARMS IT WAS REPORTED THAT THE INSULIN PUMP RECEIVED REPLACE BATTERY NOW WITHOUT RECEIVING LOW BATTERY WARNING TWICE IT WAS REPORTED THAT THE CUSTOMER WAS ADVISED THAT THE INSULIN PUMP NEEDED TO BE REPLACED AND WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6558085 | 74996868 | N | THE INSULIN PUMP HAD UNRESPONSIVE BUTTON THEN BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6558085 | 74996868 | N | THE INSULIN PUMP HAD UNRESPONSIVE BUTTON THEN BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6558082 | 74997079 | N | THE INSULIN PUMP WAS RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6558082 | 74997079 | N | THE INSULIN PUMP WAS RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6558075 | 75000744 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED UP BUTTON DOME SWITCH UNABLE TO VERIFY EXCESSIVE NO DELIVERY ALARM DUE TO BUTTON ERROR ALARM THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6558075 | 75000744 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED UP BUTTON DOME SWITCH UNABLE TO VERIFY EXCESSIVE NO DELIVERY ALARM DUE TO BUTTON ERROR ALARM THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6558048 | 75002306 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM AND NO ESC AND ACT BUTTON RESPONSE DUE TO FLATTENED ESC AND ACT BUTTON DOME SWITCH THE INSULIN PUMP HAD A SCRATCHES ON LCD WINDOW AND KEYPAD OVERLAY NEAR ESC AND ACT BUTTONS NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6558048 | 75002306 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM AND NO ESC AND ACT BUTTON RESPONSE DUE TO FLATTENED ESC AND ACT BUTTON DOME SWITCH THE INSULIN PUMP HAD A SCRATCHES ON LCD WINDOW AND KEYPAD OVERLAY NEAR ESC AND ACT BUTTONS NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6558048 | 75002306 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM AND NO ESC AND ACT BUTTON RESPONSE DUE TO FLATTENED ESC AND ACT BUTTON DOME SWITCH THE INSULIN PUMP HAD A SCRATCHES ON LCD WINDOW AND KEYPAD OVERLAY NEAR ESC AND ACT BUTTONS NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6557884 | 75006008 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON UNIT RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6557884 | 75006008 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON UNIT RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6557884 | 75006008 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON UNIT RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6557884 | 75006008 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON UNIT RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6560868 | 75037203 | N | PUMP RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED UP BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6560868 | 75037203 | N | PUMP RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED UP BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6564402 | 75060345 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMERS NURSE STATED THAT THE CUSTOMER RAN OUT OF INSULIN DID NOT TELL ANYONE THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS NAUSEA AND VOMITING THE CUSTOMER WAS TREATED WITH IV INSULIN THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Nausea | Patient problem code
6564402 | 75060345 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMERS NURSE STATED THAT THE CUSTOMER RAN OUT OF INSULIN DID NOT TELL ANYONE THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS NAUSEA AND VOMITING THE CUSTOMER WAS TREATED WITH IV INSULIN THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Vomiting | Patient problem code
6564402 | 75060345 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMERS NURSE STATED THAT THE CUSTOMER RAN OUT OF INSULIN DID NOT TELL ANYONE THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS NAUSEA AND VOMITING THE CUSTOMER WAS TREATED WITH IV INSULIN THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6564402 | 75060345 | D | THE CUSTOMERS NURSE REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMERS NURSE STATED THAT THE CUSTOMER RAN OUT OF INSULIN DID NOT TELL ANYONE THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS NAUSEA AND VOMITING THE CUSTOMER WAS TREATED WITH IV INSULIN THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6564511 | 75061604 | N | UNIT PASSED THE FUNCTIONAL TEST INCLUDING THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DATA TEST AT INCHES NO BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING UNIT HAD MINOR SCRATCHED DISPLAY WINDOW | Scratched Material | Device problem code
6564511 | 75061604 | N | UNIT PASSED THE FUNCTIONAL TEST INCLUDING THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DATA TEST AT INCHES NO BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING UNIT HAD MINOR SCRATCHED DISPLAY WINDOW | Stress Problem Identified | Medical Evaluation result code
6561203 | 75094595 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6561203 | 75094595 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6561196 | 75101289 | N | THE INSULIN PUMP HAD INTERMITTENT BUTTON RESPONSE DUE TO A FLATTENED DOME SWITCH ON ESC BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6561196 | 75101289 | N | THE INSULIN PUMP HAD INTERMITTENT BUTTON RESPONSE DUE TO A FLATTENED DOME SWITCH ON ESC BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6559886 | 75104143 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY DEVICE EVALUATED BY DISTRIBUTOR | Component Incompatible | Device problem code
6559886 | 75104143 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY DEVICE EVALUATED BY DISTRIBUTOR | Interoperability Problem Identified | Medical Evaluation result code
6561012 | 75109177 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY RECEIVED MULTIPLE REPLACE BATTERY NOW ALARMS THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MMOLL THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM TROUBLESHOOTING WAS PERFORMED AND IT WAS NOTED THAT IT WAS THE SECOND OCCURRENCE OF A REPLACE BATTERY NOW ALARM WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6561012 | 75109177 | D | IT WAS REPORTED VIA PHONE CALL THAT THEY RECEIVED MULTIPLE REPLACE BATTERY NOW ALARMS THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MMOLL THEY DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM TROUBLESHOOTING WAS PERFORMED AND IT WAS NOTED THAT IT WAS THE SECOND OCCURRENCE OF A REPLACE BATTERY NOW ALARM WITHOUT A LOW BATTERY WARNING THE DEVICE WAS RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6561150 | 75111201 | N | THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED ON KEYPAD OVERLAY AND MISSING END CAP STICKER THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ALL BUTTONS ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6561150 | 75111201 | N | THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED ON KEYPAD OVERLAY AND MISSING END CAP STICKER THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ALL BUTTONS ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6561150 | 75111201 | N | THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED ON KEYPAD OVERLAY AND MISSING END CAP STICKER THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ALL BUTTONS ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6561150 | 75111201 | N | THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED ON KEYPAD OVERLAY AND MISSING END CAP STICKER THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ALL BUTTONS ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6561150 | 75111201 | N | THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED ON KEYPAD OVERLAY AND MISSING END CAP STICKER THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH ALL BUTTONS ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6537413 | 75111311 | N | AFTER FURTHER REVIEW OF THE COMPLAINT IT WAS DETERMINED THIS COMPLAINT WAS SUBMITTED IN ERROR THE CUSTOMER DID EXPRESS THE DEVICE HAD WAS MISSING THE ELECTRODE PRIOR TO INSERTION THE CUSTOMER DID NOT REPORT ANY PAIN DISCOMFORT OR ANY IRREGULARITIES AT THE INSERTION SITE DUE TO THE ELECTRODE BEING STUCK IN THE BODY THE DIVIDE DID NOT CONTRIBUTE TO A REPORTABLE MALFUNCTION | Component Missing | Device problem code
6537413 | 75111311 | N | AFTER FURTHER REVIEW OF THE COMPLAINT IT WAS DETERMINED THIS COMPLAINT WAS SUBMITTED IN ERROR THE CUSTOMER DID EXPRESS THE DEVICE HAD WAS MISSING THE ELECTRODE PRIOR TO INSERTION THE CUSTOMER DID NOT REPORT ANY PAIN DISCOMFORT OR ANY IRREGULARITIES AT THE INSERTION SITE DUE TO THE ELECTRODE BEING STUCK IN THE BODY THE DIVIDE DID NOT CONTRIBUTE TO A REPORTABLE MALFUNCTION | Mechanical Jam | Device problem code
6537413 | 75111311 | N | AFTER FURTHER REVIEW OF THE COMPLAINT IT WAS DETERMINED THIS COMPLAINT WAS SUBMITTED IN ERROR THE CUSTOMER DID EXPRESS THE DEVICE HAD WAS MISSING THE ELECTRODE PRIOR TO INSERTION THE CUSTOMER DID NOT REPORT ANY PAIN DISCOMFORT OR ANY IRREGULARITIES AT THE INSERTION SITE DUE TO THE ELECTRODE BEING STUCK IN THE BODY THE DIVIDE DID NOT CONTRIBUTE TO A REPORTABLE MALFUNCTION | Mechanical Problem | Device Problem Code
6537413 | 75111311 | N | AFTER FURTHER REVIEW OF THE COMPLAINT IT WAS DETERMINED THIS COMPLAINT WAS SUBMITTED IN ERROR THE CUSTOMER DID EXPRESS THE DEVICE HAD WAS MISSING THE ELECTRODE PRIOR TO INSERTION THE CUSTOMER DID NOT REPORT ANY PAIN DISCOMFORT OR ANY IRREGULARITIES AT THE INSERTION SITE DUE TO THE ELECTRODE BEING STUCK IN THE BODY THE DIVIDE DID NOT CONTRIBUTE TO A REPORTABLE MALFUNCTION | Mechanical Problem Identified | Medical Evaluation result code
6565765 | 75147441 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE THE DEVICE HAD A SCRATCHED CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6565765 | 75147441 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE THE DEVICE HAD A SCRATCHED CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6565765 | 75147441 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE THE DEVICE HAD A SCRATCHED CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6565765 | 75147441 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE THE DEVICE HAD A SCRATCHED CASE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6561285 | 75152661 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED EXPRESS BOLUS BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6561285 | 75152661 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED EXPRESS BOLUS BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6564057 | 75161431 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON DOWN ARROW BUTTON THE INSULIN PUMP WAS RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6564057 | 75161431 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON DOWN ARROW BUTTON THE INSULIN PUMP WAS RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6564057 | 75161431 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON DOWN ARROW BUTTON THE INSULIN PUMP WAS RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6564057 | 75161431 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON DOWN ARROW BUTTON THE INSULIN PUMP WAS RETURNED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6563894 | 75162189 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6563894 | 75162189 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6563798 | 75162670 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH ALSO RECEIVED A WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6563798 | 75162670 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH ALSO RECEIVED A WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6563798 | 75162670 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH ALSO RECEIVED A WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6563798 | 75162670 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH ALSO RECEIVED A WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6565758 | 75163076 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED A MOTOR POWER SUPPLY VOLTAGE ERROR DETECTED THE DELIVERY WAS STOPPED AND THEY WERE PROMPTED FOR A REWIND SO THEY DID BUT IT KEPT HAPPENING THE CUSTOMER WAS IN TRAINING FOR THE NEW INSULIN PUMP AND WAS DRIVING HOME WHEN THEY RECEIVED THE ERROR THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MGDL THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD IT WAS NOTED IN TROUBLESHOOTING THAT THEY WERE UNABLE TO CHECK THE ALARM HISTORY THE CUSTOMER STATED THAT THEY WERE UNABLE TO REWIND THE INSULIN PUMP THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE DEVICE AND REVERT TO A BACKUP PLAN THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6565758 | 75163076 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED A MOTOR POWER SUPPLY VOLTAGE ERROR DETECTED THE DELIVERY WAS STOPPED AND THEY WERE PROMPTED FOR A REWIND SO THEY DID BUT IT KEPT HAPPENING THE CUSTOMER WAS IN TRAINING FOR THE NEW INSULIN PUMP AND WAS DRIVING HOME WHEN THEY RECEIVED THE ERROR THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MGDL THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD IT WAS NOTED IN TROUBLESHOOTING THAT THEY WERE UNABLE TO CHECK THE ALARM HISTORY THE CUSTOMER STATED THAT THEY WERE UNABLE TO REWIND THE INSULIN PUMP THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE DEVICE AND REVERT TO A BACKUP PLAN THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | Power Problem | Device problem code
6564100 | 75169266 | N | THE INSULIN PUMP HAD UNRESPONSIVE BUTTON THEN BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6564100 | 75169266 | N | THE INSULIN PUMP HAD UNRESPONSIVE BUTTON THEN BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6564121 | 75173417 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED ALL BUTTONS DOME SWITCH THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION PRIME AND NO DELIVERY TESTS NO EXCESSIVE NO DELIVERY ALARM NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Failure to Deliver | Device problem code
6564121 | 75173417 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED ALL BUTTONS DOME SWITCH THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION PRIME AND NO DELIVERY TESTS NO EXCESSIVE NO DELIVERY ALARM NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6564121 | 75173417 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED ALL BUTTONS DOME SWITCH THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION PRIME AND NO DELIVERY TESTS NO EXCESSIVE NO DELIVERY ALARM NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Inaccurate Delivery | Device Problem Code
6564121 | 75173417 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED ALL BUTTONS DOME SWITCH THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND BASIC OCCLUSION PRIME AND NO DELIVERY TESTS NO EXCESSIVE NO DELIVERY ALARM NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6564438 | 75175955 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED REPLACE BATTERY NOW THE CUSTOMER WAS ASSISTED WITH TROUBLESHOOTING FOR REPLACE BATTERY NOW CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THIS WAS THE SECOND OCCURRENCE OF INSULIN PUMP ALARMING REPLACE BATTERY NOW WITHOUT RECEIVING LOW BATTERY WARNING FIRST CUSTOMER STATED THAT THE BATTERY WAS NOT PREVIOUSLY USED THE DRIVE SUPPORT CAP WAS NORMAL AND THAT THERE WERE NO UNATTENDED ALARMS THE REPORTED WAS ADVISED THAT THE INSULIN PUMP NEEDED TO BE REPLACED AND WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP WAS RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6564438 | 75175955 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED REPLACE BATTERY NOW THE CUSTOMER WAS ASSISTED WITH TROUBLESHOOTING FOR REPLACE BATTERY NOW CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THIS WAS THE SECOND OCCURRENCE OF INSULIN PUMP ALARMING REPLACE BATTERY NOW WITHOUT RECEIVING LOW BATTERY WARNING FIRST CUSTOMER STATED THAT THE BATTERY WAS NOT PREVIOUSLY USED THE DRIVE SUPPORT CAP WAS NORMAL AND THAT THERE WERE NO UNATTENDED ALARMS THE REPORTED WAS ADVISED THAT THE INSULIN PUMP NEEDED TO BE REPLACED AND WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP WAS RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6563623 | 75212888 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP WOULD START BUZZING BETWEEN TO HOURS THAT THE BATTERY NEEDS TO BE REPLACED THEY REPORTED THAT THE POWER OUT OCCURRED THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN THE CUSTOMER DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THE INSULIN PUMP WAS NOT DROPPED THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW ALARM WITHOUT A LOW BATTERY WARNING THE DEVICE WILL BE RETURNED FOR ANALYSIS | Battery Problem | Device problem code
6563623 | 75212888 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP WOULD START BUZZING BETWEEN TO HOURS THAT THE BATTERY NEEDS TO BE REPLACED THEY REPORTED THAT THE POWER OUT OCCURRED THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN THE CUSTOMER DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THE INSULIN PUMP WAS NOT DROPPED THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW ALARM WITHOUT A LOW BATTERY WARNING THE DEVICE WILL BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6565494 | 75253058 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORM CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6565494 | 75253058 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORM CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6565679 | 75312820 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6565679 | 75312820 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6565818 | 75314403 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6565818 | 75314403 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6572069 | 75360934 | D | THERE WAS NO PATIENT INVOLVED IN THIS EVENT LOW BATTERY WARNING | Battery Problem: Low Impedance | Device problem code
6572069 | 75360934 | D | THERE WAS NO PATIENT INVOLVED IN THIS EVENT LOW BATTERY WARNING | No Patient Involvement | Patient problem code
6555480 | 75361025 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Hypotension | Patient problem code
6555480 | 75361025 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Adverse Event Without Identified Device or Use Problem | Device problem code
6555480 | 75361025 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Tingling | Patient problem code
6555479 | 75361354 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Hypotension | Patient problem code
6555479 | 75361354 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Tingling | Patient problem code
6555479 | 75361354 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Adverse Event Without Identified Device or Use Problem | Device problem code
6555481 | 75361592 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Adverse Event Without Identified Device or Use Problem | Device problem code
6555481 | 75361592 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Hypotension | Patient problem code
6555481 | 75361592 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SYMPTOMS WERE TINGLING TO HANDS AND FACE SHE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS ON B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Tingling | Patient problem code
6555477 | 75361691 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS WERE RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Hypotension | Patient problem code
6555477 | 75361691 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS WERE RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Adverse Event Without Identified Device or Use Problem | Device problem code
6555477 | 75361691 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS WERE RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENTS SIGNS AND SYMPTOMS WERE TINGLING TO HANDS AND FACE HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Tingling | Patient problem code
6555478 | 75361746 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENT WAS THE FIRST ONE WITH SIGNS AND SYMPTOMS HE HAD TINGLING TO HANDS AND FACE BAD TASTE IN MOUTH HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Hypotension | Patient problem code
6555478 | 75361746 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENT WAS THE FIRST ONE WITH SIGNS AND SYMPTOMS HE HAD TINGLING TO HANDS AND FACE BAD TASTE IN MOUTH HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Adverse Event Without Identified Device or Use Problem | Device problem code
6555478 | 75361746 | D | DURING THE RD SHIFT ON BSIDE OF THE OP DIALYSIS UNIT PATIENTS WERE FINISHED AND REMAINED THE WERE HALFWAY DONE WITH THEIR TREATMENT WHEN THEIR MACHINES BEGAN ALARMING SPORADICALLY AND PATIENTS BEGAN TO COMPLAIN OF TINGLING IN HANDS AND FACE SOME EXPERIENCED MODERATE HYPOTENSION TREATMENT STOPPED AND BLOOD WAS RETURNED RENAL PANELS AND BLOOD CULTURES WERE OBTAINED PATIENTS WERE GIVEN NOVASOURCE AFTER LABS WERE RESULTED THE PATIENTS WERE MOVED TO A SIDE FOR ONE HOUR HEMODIALYSIS WITH A K BATH SYMPTOMS RESOLVED AND PATIENTS WERE DISCHARGED THIS PATIENT WAS THE FIRST ONE WITH SIGNS AND SYMPTOMS HE HAD TINGLING TO HANDS AND FACE BAD TASTE IN MOUTH HE WAS IN CHAIR K BATH UNITS A AND B ARE ON DIFFERENT LOOPS DAILY BICARB AND ACID ADDED AT AM FOLLOWING ROUTINE PROCESS AND USUAL STAFF RO WATER TESTED NORMAL STORAGE TANKS FOR BICARB CHECKED AFTER INCIDENT NO SLUDGE SEEN CONFERENCE CALL TO FRESENIUS B TO DISCUSS INCIDENT AND ASK FOR ASSISTANCE IN DETERMINING POSSIBLE CAUSE | Tingling | Patient problem code
6573834 | 75401155 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST AT INCHES THE INSULIN PUMP WAS PROGRAMMED WITH A UNIT BOLUS THE BOLUS DELIVERED PROPERLY NO BOLUS DELIVERY ANOMALY OR AIR BUBBLE ANOMALY NOTED THE MOTOR FUNCTIONS PROPERLY DURING TESTING NO DRIVEMOTOR ANOMALY NOTED THE MOTOR WAS TESTED OUTSIDE OF THE UNIT AND PASSED THE INSULIN PUMP HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY | Scratched Material | Device problem code
6573834 | 75401155 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST AT INCHES THE INSULIN PUMP WAS PROGRAMMED WITH A UNIT BOLUS THE BOLUS DELIVERED PROPERLY NO BOLUS DELIVERY ANOMALY OR AIR BUBBLE ANOMALY NOTED THE MOTOR FUNCTIONS PROPERLY DURING TESTING NO DRIVEMOTOR ANOMALY NOTED THE MOTOR WAS TESTED OUTSIDE OF THE UNIT AND PASSED THE INSULIN PUMP HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY | Fracture Problem | Medical Evaluation result code
6573772 | 75403568 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH DRINKS THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Display or Visual Feedback Problem | Device problem code
6573772 | 75403568 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER STATED THAT THEY TREATED THE LOW BLOOD GLUCOSE WITH DRINKS THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP APPEARED NORMAL THE CUSTOMER STATED THAT THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6569813 | 75404714 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED REPLACE BATTERY NOW IF THE CUSTOMER DID NOT REPLACE THE BATTERY THE INSULIN PUMP WOULD SHUT OFF IT WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6569813 | 75404714 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP ALARMED REPLACE BATTERY NOW IF THE CUSTOMER DID NOT REPLACE THE BATTERY THE INSULIN PUMP WOULD SHUT OFF IT WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | Battery Problem | Device problem code
6569802 | 75408167 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6569802 | 75408167 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6570087 | 75408640 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6570087 | 75408640 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6570179 | 75409122 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6570179 | 75409122 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6569536 | 75416146 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Component or Accessory Incompatibility | Device problem code
6569536 | 75416146 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Interoperability Problem Identified | Medical Evaluation result code
6572217 | 75425241 | D | ENDOLEAK AORTIC ANEURISM GROWTH AFTER ENDOLOGIX AFX IMPLANTED IN | Hemorrhage | Patient problem code
6572217 | 75425241 | D | ENDOLEAK AORTIC ANEURISM GROWTH AFTER ENDOLOGIX AFX IMPLANTED IN | Insufficient Information | Device problem code
6574986 | 75429629 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SELECT ENERGY LEVEL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6574986 | 75429629 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SELECT ENERGY LEVEL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Energy Output Problem | Device Problem Code
6573871 | 75505957 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6573871 | 75505957 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6573914 | 75509482 | N | UNABLE TO PRIME DURING THE PRIMEA TEST DUE TO FAULTY FSR GOLD UNABLE TO PERFORM THE EXCESSIVE NO DELIVERY TEST DUE TO PRIME ANOMALY PUMP PASSED SELF TEST A ERROR TEST DISPLACEMENT REWIND BASIC OCCLUSION AND OCCLUSION TESTS NO MOTOR ERROR ALARM DURING TESTING NOTED MOTOR PASSED MOTOR TEST PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Component or Accessory Incompatibility | Device problem code
6573914 | 75509482 | N | UNABLE TO PRIME DURING THE PRIMEA TEST DUE TO FAULTY FSR GOLD UNABLE TO PERFORM THE EXCESSIVE NO DELIVERY TEST DUE TO PRIME ANOMALY PUMP PASSED SELF TEST A ERROR TEST DISPLACEMENT REWIND BASIC OCCLUSION AND OCCLUSION TESTS NO MOTOR ERROR ALARM DURING TESTING NOTED MOTOR PASSED MOTOR TEST PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Scratched Material | Device problem code
6573914 | 75509482 | N | UNABLE TO PRIME DURING THE PRIMEA TEST DUE TO FAULTY FSR GOLD UNABLE TO PERFORM THE EXCESSIVE NO DELIVERY TEST DUE TO PRIME ANOMALY PUMP PASSED SELF TEST A ERROR TEST DISPLACEMENT REWIND BASIC OCCLUSION AND OCCLUSION TESTS NO MOTOR ERROR ALARM DURING TESTING NOTED MOTOR PASSED MOTOR TEST PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Interoperability Problem Identified | Medical Evaluation result code
6573914 | 75509482 | N | UNABLE TO PRIME DURING THE PRIMEA TEST DUE TO FAULTY FSR GOLD UNABLE TO PERFORM THE EXCESSIVE NO DELIVERY TEST DUE TO PRIME ANOMALY PUMP PASSED SELF TEST A ERROR TEST DISPLACEMENT REWIND BASIC OCCLUSION AND OCCLUSION TESTS NO MOTOR ERROR ALARM DURING TESTING NOTED MOTOR PASSED MOTOR TEST PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Stress Problem Identified | Medical Evaluation result code
6573713 | 75515154 | N | THE INSULIN PUMP HAD NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESCAPE BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6573713 | 75515154 | N | THE INSULIN PUMP HAD NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESCAPE BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6573527 | 75517053 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6573527 | 75517053 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6534375 | 75530817 | N | THE INSULIN PUMP PASSED FUNCTIONAL TESTING INCLUDING THE OPERATING CURRENTS MEASUREMENT SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST THE BACKLIGHT FUNCTIONED PROPERLY NO MISSING SEGMENTS OR PARTIAL DISPLAY OR DISPLAY ANOMALY NOTED THE INSULIN PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Stress Problem Identified | Medical Evaluation result code
6534375 | 75530817 | N | THE INSULIN PUMP PASSED FUNCTIONAL TESTING INCLUDING THE OPERATING CURRENTS MEASUREMENT SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST THE BACKLIGHT FUNCTIONED PROPERLY NO MISSING SEGMENTS OR PARTIAL DISPLAY OR DISPLAY ANOMALY NOTED THE INSULIN PUMP HAD MINOR SCRATCHED DISPLAY WINDOW | Scratched Material | Device problem code
6576341 | 75550805 | N | NO BUTTON RESPONSE DUE TO FLATTENED BUTTON DOME SWITCH ON ESC BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6576341 | 75550805 | N | NO BUTTON RESPONSE DUE TO FLATTENED BUTTON DOME SWITCH ON ESC BUTTON NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6574870 | 75555304 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6579742 | 75593671 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE LEVELS THE CUSTOMER WAS OFF THE INSULIN PUMP FOR HOURS AND TREATED WITH SYRINGES THE CUSTOMER WENT TO THE EMERGENCY ROOM AS A RESULT OF HER BLOOD GLUCOSE LEVELS THE TUBING LENGTH HAD ONE SMALL CHAMPAGNE BUBBLE THE CUSTOMERS MOTHER REPORTED THAT ON B SHE HAD A HIGH BLOOD GLUCOSE LEVEL OF OVER MGDL SHE WAS TAKEN BY AN AMBULANCE TO FURTHER TREAT HER DIABETIC KETOACIDOSIS HER BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION SHE WAS OFF THE INSULIN PUMP LESS THAN HOURS THE DEVICE WAS NOT RETURNED | Hyperglycemia | Patient problem code
6579742 | 75593671 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE LEVELS THE CUSTOMER WAS OFF THE INSULIN PUMP FOR HOURS AND TREATED WITH SYRINGES THE CUSTOMER WENT TO THE EMERGENCY ROOM AS A RESULT OF HER BLOOD GLUCOSE LEVELS THE TUBING LENGTH HAD ONE SMALL CHAMPAGNE BUBBLE THE CUSTOMERS MOTHER REPORTED THAT ON B SHE HAD A HIGH BLOOD GLUCOSE LEVEL OF OVER MGDL SHE WAS TAKEN BY AN AMBULANCE TO FURTHER TREAT HER DIABETIC KETOACIDOSIS HER BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION SHE WAS OFF THE INSULIN PUMP LESS THAN HOURS THE DEVICE WAS NOT RETURNED | Diabetic Ketoacidosis | Patient problem code
6579742 | 75593671 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE LEVELS THE CUSTOMER WAS OFF THE INSULIN PUMP FOR HOURS AND TREATED WITH SYRINGES THE CUSTOMER WENT TO THE EMERGENCY ROOM AS A RESULT OF HER BLOOD GLUCOSE LEVELS THE TUBING LENGTH HAD ONE SMALL CHAMPAGNE BUBBLE THE CUSTOMERS MOTHER REPORTED THAT ON B SHE HAD A HIGH BLOOD GLUCOSE LEVEL OF OVER MGDL SHE WAS TAKEN BY AN AMBULANCE TO FURTHER TREAT HER DIABETIC KETOACIDOSIS HER BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION SHE WAS OFF THE INSULIN PUMP LESS THAN HOURS THE DEVICE WAS NOT RETURNED | Vomiting | Patient problem code
6579742 | 75593671 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE LEVELS THE CUSTOMER WAS OFF THE INSULIN PUMP FOR HOURS AND TREATED WITH SYRINGES THE CUSTOMER WENT TO THE EMERGENCY ROOM AS A RESULT OF HER BLOOD GLUCOSE LEVELS THE TUBING LENGTH HAD ONE SMALL CHAMPAGNE BUBBLE THE CUSTOMERS MOTHER REPORTED THAT ON B SHE HAD A HIGH BLOOD GLUCOSE LEVEL OF OVER MGDL SHE WAS TAKEN BY AN AMBULANCE TO FURTHER TREAT HER DIABETIC KETOACIDOSIS HER BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER WAS VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION SHE WAS OFF THE INSULIN PUMP LESS THAN HOURS THE DEVICE WAS NOT RETURNED | Suction Problem | Device problem code
6579470 | 75596024 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS AND DEHYDRATION AND WAS TREATED WITH INSULIN DRIP AND FLUIDS FOR DEHYDRATION THE CUSTOMERS MOTHER STATED THAT THE CUSTOMER LOST THE INSULIN PUMP AND WAS OFF OF IT FOR APPROXIMATELY TWELVE HOURS AND BUT THE TIME IT WAS FOUND THE CUSTOMER ALREADY HAD HIGH BLOOD GLUCOSE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6579470 | 75596024 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS AND DEHYDRATION AND WAS TREATED WITH INSULIN DRIP AND FLUIDS FOR DEHYDRATION THE CUSTOMERS MOTHER STATED THAT THE CUSTOMER LOST THE INSULIN PUMP AND WAS OFF OF IT FOR APPROXIMATELY TWELVE HOURS AND BUT THE TIME IT WAS FOUND THE CUSTOMER ALREADY HAD HIGH BLOOD GLUCOSE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6579470 | 75596024 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS AND DEHYDRATION AND WAS TREATED WITH INSULIN DRIP AND FLUIDS FOR DEHYDRATION THE CUSTOMERS MOTHER STATED THAT THE CUSTOMER LOST THE INSULIN PUMP AND WAS OFF OF IT FOR APPROXIMATELY TWELVE HOURS AND BUT THE TIME IT WAS FOUND THE CUSTOMER ALREADY HAD HIGH BLOOD GLUCOSE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6579470 | 75596024 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS AND DEHYDRATION AND WAS TREATED WITH INSULIN DRIP AND FLUIDS FOR DEHYDRATION THE CUSTOMERS MOTHER STATED THAT THE CUSTOMER LOST THE INSULIN PUMP AND WAS OFF OF IT FOR APPROXIMATELY TWELVE HOURS AND BUT THE TIME IT WAS FOUND THE CUSTOMER ALREADY HAD HIGH BLOOD GLUCOSE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT FOR LESS THAN HOURS THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Dehydrated | Patient problem code
6579773 | 75596552 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THE INFUSION SET CANNULA TIPPED OVER CUSTOMERS BLOOD GLUCOSE WAS MGDL AFTER TROUBLESHOOTING CUSTOMER WAS ADVISED TO CHANGE THE INFUSION SET CUSTOMER WILL NOT BE RETURNING THE DEVICE FOR ANALYSIS | Device Tipped Over | Device problem code
6579773 | 75596552 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THE INFUSION SET CANNULA TIPPED OVER CUSTOMERS BLOOD GLUCOSE WAS MGDL AFTER TROUBLESHOOTING CUSTOMER WAS ADVISED TO CHANGE THE INFUSION SET CUSTOMER WILL NOT BE RETURNING THE DEVICE FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6580571 | 75600804 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED IN THE INTENSIVE CARE UNIT DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL CUSTOMER REPORTED THAT BLOOD GLUCOSE WAS HIGH AT MGDL UPON WAKING AND DID NOT LOWER THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION AS IT WAS REMOVED LESS THAN HOURS PRIOR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS AS CUSTOMER DECLINED TROUBLESHOOTING | Hyperglycemia | Patient problem code
6580571 | 75600804 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED IN THE INTENSIVE CARE UNIT DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL CUSTOMER REPORTED THAT BLOOD GLUCOSE WAS HIGH AT MGDL UPON WAKING AND DID NOT LOWER THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION AS IT WAS REMOVED LESS THAN HOURS PRIOR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS AS CUSTOMER DECLINED TROUBLESHOOTING | Insufficient Information | Device problem code
6580571 | 75600804 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED IN THE INTENSIVE CARE UNIT DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL CUSTOMER REPORTED THAT BLOOD GLUCOSE WAS HIGH AT MGDL UPON WAKING AND DID NOT LOWER THE CUSTOMER WAS TREATED WITH INSULIN DRIP THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF HOSPITALIZATION AS IT WAS REMOVED LESS THAN HOURS PRIOR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS AS CUSTOMER DECLINED TROUBLESHOOTING | Diabetic Ketoacidosis | Patient problem code
6575167 | 75603375 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6575167 | 75603375 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
6575937 | 75608503 | N | THE INSULIN PUMP PASSED BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE INSULIN PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ERROR ALARM CONFIRMED IN THE HISTORY FILE THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED IN OUR TESTING | Mechanical Jam | Device Problem Code
6575937 | 75608503 | N | THE INSULIN PUMP PASSED BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE INSULIN PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ERROR ALARM CONFIRMED IN THE HISTORY FILE THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED IN OUR TESTING | Component or Accessory Incompatibility | Device Problem Code
6575937 | 75608503 | N | THE INSULIN PUMP PASSED BASIC OCCLUSION PRIME AND DISPLACEMENT TEST THE INSULIN PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ERROR ALARM CONFIRMED IN THE HISTORY FILE THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED IN OUR TESTING | Interoperability Problem Identified | Medical Evaluation result code
6575293 | 75608580 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6576095 | 75613064 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6576095 | 75613064 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6574831 | 75613713 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY DEVICE EVALUATED BY DISTRIBUTOR | Component Incompatible | Device problem code
6574831 | 75613713 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY DEVICE EVALUATED BY DISTRIBUTOR | Interoperability Problem Identified | Medical Evaluation result code
6576583 | 75620939 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND NO BUTTON RESPONSE DUE TO A FLATTENED UP AND DOWN BUTTON DOME SWITCH THE DEVICE HAD A MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6576583 | 75620939 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND NO BUTTON RESPONSE DUE TO A FLATTENED UP AND DOWN BUTTON DOME SWITCH THE DEVICE HAD A MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6576583 | 75620939 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND NO BUTTON RESPONSE DUE TO A FLATTENED UP AND DOWN BUTTON DOME SWITCH THE DEVICE HAD A MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6576583 | 75620939 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND NO BUTTON RESPONSE DUE TO A FLATTENED UP AND DOWN BUTTON DOME SWITCH THE DEVICE HAD A MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6579363 | 75661663 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6579363 | 75661663 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESC BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6579576 | 75674375 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6579576 | 75674375 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6579576 | 75674375 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6579576 | 75674375 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH UNIT HAD SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6579665 | 75683550 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6579665 | 75683550 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6579652 | 75685246 | N | THE INSULIN PUMP HAD A HIGH STOP CURRENT DUE TO A FAULTY SCREEN DRIVER THE INSULIN PUMP HAD A MINOR SCRATCHED DISPLAY WINDOW | Component or Accessory Incompatibility | Device problem code
6579652 | 75685246 | N | THE INSULIN PUMP HAD A HIGH STOP CURRENT DUE TO A FAULTY SCREEN DRIVER THE INSULIN PUMP HAD A MINOR SCRATCHED DISPLAY WINDOW | Scratched Material | Device problem code
6579652 | 75685246 | N | THE INSULIN PUMP HAD A HIGH STOP CURRENT DUE TO A FAULTY SCREEN DRIVER THE INSULIN PUMP HAD A MINOR SCRATCHED DISPLAY WINDOW | Stress Problem Identified | Medical Evaluation result code
6579652 | 75685246 | N | THE INSULIN PUMP HAD A HIGH STOP CURRENT DUE TO A FAULTY SCREEN DRIVER THE INSULIN PUMP HAD A MINOR SCRATCHED DISPLAY WINDOW | Interoperability Problem Identified | Medical Evaluation result code
6579482 | 75686109 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6579482 | 75686109 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6579591 | 75719307 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6579591 | 75719307 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ACT BUTTON ON KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6578508 | 75721905 | N | B THE USER FACILITY LEFT THE SYSTEM TURNED ON OVERNIGHT AND WHEN THEY CAME IN THE NEXT MORNING AN ERROR MESSAGE ON THE CENTRAL CONTROL MONITOR CCM WAS OBSERVED THE FIELD SERVICE REPRESENTATIVE FSR WAS ONSITE AND WAS NOT ABLE TO VERIFY THE ISSUE DATA LOGS WERE RETURNED ON MAY | Device Displays Incorrect Message | Device problem code
6578508 | 75721905 | N | B THE USER FACILITY LEFT THE SYSTEM TURNED ON OVERNIGHT AND WHEN THEY CAME IN THE NEXT MORNING AN ERROR MESSAGE ON THE CENTRAL CONTROL MONITOR CCM WAS OBSERVED THE FIELD SERVICE REPRESENTATIVE FSR WAS ONSITE AND WAS NOT ABLE TO VERIFY THE ISSUE DATA LOGS WERE RETURNED ON MAY | Interoperability Problem Identified | Medical Evaluation result code
6580467 | 75731680 | N | THE INSULIN PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARMS DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6580467 | 75731680 | N | THE INSULIN PUMP WAS RECEIVED WITH UNRESPONSIVE BUTTONS AND BUTTON ERROR ALARMS DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6580621 | 75737457 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO UNIT HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6580621 | 75737457 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO UNIT HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Moisture or Humidity Problem | Device problem code
6580621 | 75737457 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO UNIT HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Storage Problem Identified | Medical Evaluation result code
6580621 | 75737457 | N | UNIT HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO UNIT HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Electrical Power Problem | Device problem code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Component Missing | Device problem code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Missing or Inadequate Safety Measures | Medical Evaluation result code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Failure to Deliver | Device problem code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Inappropriate or Unexpected Reset | Device Problem Code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Inaccurate Delivery | Device Problem Code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Scratched Material | Device problem code
6582233 | 75769190 | N | THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW MISSING RESERVOIR TUBE ORING AND DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP THE TEST RESERVOIR WOULD NOT LOCKCLICK IN PLACE THE INSULIN PUMP PASSED THE IDLE CURRENT TEST RUN CURRENT TEST SELFTEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST UNABLE TO PERFORM BASIC OCCLUSION TEST AND OCCLUSION TEST DUE TO DOG CHEW MARKS SEVERELY ON RESERVOIR TUBE LIP | Operational Problem Identified | Medical Evaluation result code
6582323 | 75769274 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6582323 | 75769274 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARMS DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6582263 | 75791822 | N | THE INSULIN PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER B FILE NUMBER B AND PUMP ERROR ALARM DUE TO SOFTWARE ERROR THE PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Component or Accessory Incompatibility | Device problem code
6582263 | 75791822 | N | THE INSULIN PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER B FILE NUMBER B AND PUMP ERROR ALARM DUE TO SOFTWARE ERROR THE PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Scratched Material | Device problem code
6582263 | 75791822 | N | THE INSULIN PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER B FILE NUMBER B AND PUMP ERROR ALARM DUE TO SOFTWARE ERROR THE PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Stress Problem Identified | Medical Evaluation result code
6582263 | 75791822 | N | THE INSULIN PUMP PASSED THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER B FILE NUMBER B AND PUMP ERROR ALARM DUE TO SOFTWARE ERROR THE PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Interoperability Problem Identified | Medical Evaluation result code
6545069 | 75830052 | N | IT WAS REPORTED BY THE CUSTOMERS FATHER THAT HE PERFORMED THE HIGHPRESSURE TEST AND IT DID NOT PASS THE INSULIN PUMP FAILED TO ALARM NO DELIVERY THE CUSTOMERS FATHER ALSO STATED CUSTOMER IS HAVING HIGH BLOOD GLUCOSE DUE TO THE INFUSION SET CANNULA IS KINKING | Failure to Deliver | Device problem code
6545069 | 75830052 | N | IT WAS REPORTED BY THE CUSTOMERS FATHER THAT HE PERFORMED THE HIGHPRESSURE TEST AND IT DID NOT PASS THE INSULIN PUMP FAILED TO ALARM NO DELIVERY THE CUSTOMERS FATHER ALSO STATED CUSTOMER IS HAVING HIGH BLOOD GLUCOSE DUE TO THE INFUSION SET CANNULA IS KINKING | Inaccurate Delivery | Device Problem Code
6545069 | 75830052 | N | IT WAS REPORTED BY THE CUSTOMERS FATHER THAT HE PERFORMED THE HIGHPRESSURE TEST AND IT DID NOT PASS THE INSULIN PUMP FAILED TO ALARM NO DELIVERY THE CUSTOMERS FATHER ALSO STATED CUSTOMER IS HAVING HIGH BLOOD GLUCOSE DUE TO THE INFUSION SET CANNULA IS KINKING | Device Alarm System | Device problem code
6545069 | 75830052 | N | IT WAS REPORTED BY THE CUSTOMERS FATHER THAT HE PERFORMED THE HIGHPRESSURE TEST AND IT DID NOT PASS THE INSULIN PUMP FAILED TO ALARM NO DELIVERY THE CUSTOMERS FATHER ALSO STATED CUSTOMER IS HAVING HIGH BLOOD GLUCOSE DUE TO THE INFUSION SET CANNULA IS KINKING | Interoperability Problem Identified | Medical Evaluation result code
6585016 | 75856085 | N | OPEN BOOK CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION THE CRITICALPUMP ERROR WAS GENERATED BY PUMP ERROR PER BOOTLOADER TRACES AND PUMPERROR AT THE LONG TRACE PROBLEM WAS ISOLATED TO PCBA THE MOTOR WASTESTES OUTSIDE THE DEVICE AND PASSED UNABLE TO PERFORM FUNCTIONAL TESTINCLUDING SELFTEST REWIND PRIMESEATING TEST BASIC OCCLUSION TESTOCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DUE TO CRITICALPUMP ERROR UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE ANDKEYPAD OVERLAY | Scratched Material | Device problem code
6585016 | 75856085 | N | OPEN BOOK CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION THE CRITICALPUMP ERROR WAS GENERATED BY PUMP ERROR PER BOOTLOADER TRACES AND PUMPERROR AT THE LONG TRACE PROBLEM WAS ISOLATED TO PCBA THE MOTOR WASTESTES OUTSIDE THE DEVICE AND PASSED UNABLE TO PERFORM FUNCTIONAL TESTINCLUDING SELFTEST REWIND PRIMESEATING TEST BASIC OCCLUSION TESTOCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DUE TO CRITICALPUMP ERROR UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE ANDKEYPAD OVERLAY | Interoperability Problem Identified | Medical Evaluation result code
6585016 | 75856085 | N | OPEN BOOK CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION THE CRITICALPUMP ERROR WAS GENERATED BY PUMP ERROR PER BOOTLOADER TRACES AND PUMPERROR AT THE LONG TRACE PROBLEM WAS ISOLATED TO PCBA THE MOTOR WASTESTES OUTSIDE THE DEVICE AND PASSED UNABLE TO PERFORM FUNCTIONAL TESTINCLUDING SELFTEST REWIND PRIMESEATING TEST BASIC OCCLUSION TESTOCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DUE TO CRITICALPUMP ERROR UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE ANDKEYPAD OVERLAY | Component or Accessory Incompatibility | Device problem code
6588670 | 75902224 | D | IT WAS REPORTED VIA PHONE CALL THAT THE INSULIN PUMP HAD ALARMED MULTIPLE PUMP ERRORS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL AT THE TIME OF THE INCIDENT AND TREATED WITH ORANGE JUICE IT WAS REPORTED THAT THE INSULIN PUMP WAS NOT EXPOSED TO MAGNETIC FIELDS THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Pumping Problem | Device problem code
6588670 | 75902224 | D | IT WAS REPORTED VIA PHONE CALL THAT THE INSULIN PUMP HAD ALARMED MULTIPLE PUMP ERRORS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL AT THE TIME OF THE INCIDENT AND TREATED WITH ORANGE JUICE IT WAS REPORTED THAT THE INSULIN PUMP WAS NOT EXPOSED TO MAGNETIC FIELDS THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6589263 | 75918141 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO ENTER DEFIB MODE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6589263 | 75918141 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO ENTER DEFIB MODE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defibrillation Problem | Device Problem Code
6586408 | 75928305 | D | IT WAS REPORTED THAT THE CATHETER DID NOT HAVE A COATING ON IT CAUSING SOME RESISTANCE WHEN THE PATIENT WAS ATTEMPTING TO INSERT THE CATHETER HOWEVER THE PATIENT WAS ABLE TO INSERT THE CATHETER AND DRAIN URINE THE PATIENT ALLEGEDLY EXPERIENCED SOME PAIN AND A BURNING SENSATION UPON REMOVAL OF THE CATHETER | Burning Sensation | Patient problem code
6586408 | 75928305 | D | IT WAS REPORTED THAT THE CATHETER DID NOT HAVE A COATING ON IT CAUSING SOME RESISTANCE WHEN THE PATIENT WAS ATTEMPTING TO INSERT THE CATHETER HOWEVER THE PATIENT WAS ABLE TO INSERT THE CATHETER AND DRAIN URINE THE PATIENT ALLEGEDLY EXPERIENCED SOME PAIN AND A BURNING SENSATION UPON REMOVAL OF THE CATHETER | Pain | Patient problem code
6586408 | 75928305 | D | IT WAS REPORTED THAT THE CATHETER DID NOT HAVE A COATING ON IT CAUSING SOME RESISTANCE WHEN THE PATIENT WAS ATTEMPTING TO INSERT THE CATHETER HOWEVER THE PATIENT WAS ABLE TO INSERT THE CATHETER AND DRAIN URINE THE PATIENT ALLEGEDLY EXPERIENCED SOME PAIN AND A BURNING SENSATION UPON REMOVAL OF THE CATHETER | Difficult to Insert | Device problem code
6545563 | 75932881 | N | THE INSULIN PUMP PASSED FUNCTIONAL TESTING INCLUDING SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DISPLACEMENT ACCURACY TEST NO BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING OUR TESTING THE INSULIN PUMP HAD SCRATCHED CASE | Scratched Material | Device problem code
6545563 | 75932881 | N | THE INSULIN PUMP PASSED FUNCTIONAL TESTING INCLUDING SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DISPLACEMENT ACCURACY TEST NO BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING OUR TESTING THE INSULIN PUMP HAD SCRATCHED CASE | Stress Problem Identified | Medical Evaluation result code
6592522 | 76011394 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED FUNCTIONAL TESTING INCLUDING THE SELFTEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DAT TEST AT INCHES THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Scratched Material | Device problem code
6592522 | 76011394 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED FUNCTIONAL TESTING INCLUDING THE SELFTEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DAT TEST AT INCHES THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Stress Problem Identified | Medical Evaluation result code
6588984 | 76023551 | N | THE INSULIN PUMP RECEIVED WITH NO BUTTON RESPONSE DUE TO LCD KEYPAD CONNECTOR UNLOCKED NO BUTTON ERROR ALARM NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Device Alarm System | Device problem code
6588984 | 76023551 | N | THE INSULIN PUMP RECEIVED WITH NO BUTTON RESPONSE DUE TO LCD KEYPAD CONNECTOR UNLOCKED NO BUTTON ERROR ALARM NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6588984 | 76023551 | N | THE INSULIN PUMP RECEIVED WITH NO BUTTON RESPONSE DUE TO LCD KEYPAD CONNECTOR UNLOCKED NO BUTTON ERROR ALARM NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Stress Problem Identified | Medical Evaluation result code
6588537 | 76029308 | N | RELIABILITY ANALYSIS INSPECTED ONE OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Component or Accessory Incompatibility | Device problem code
6588537 | 76029308 | N | RELIABILITY ANALYSIS INSPECTED ONE OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Interoperability Problem Identified | Medical Evaluation result code
6588948 | 76030075 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6588948 | 76030075 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM DUE TO A FLATTENED DOME SWITCH AT THE ACT BUTTON ON THE KEYPAD TRACE NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6592053 | 76039425 | N | THE INSULIN PUMP ALARMED BUTTON ERROR FOLLOWED BY UNRESPONSIVE BUTTONS DUE TO A FLATTENED ESCAPE DOME SWITCH UNABLE TO VERIFY THE RESET ALARMS DUE TO BUTTON ANOMALY THE DEVICE HAD A MISSING END CAP STICKER AND A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6592053 | 76039425 | N | THE INSULIN PUMP ALARMED BUTTON ERROR FOLLOWED BY UNRESPONSIVE BUTTONS DUE TO A FLATTENED ESCAPE DOME SWITCH UNABLE TO VERIFY THE RESET ALARMS DUE TO BUTTON ANOMALY THE DEVICE HAD A MISSING END CAP STICKER AND A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6592053 | 76039425 | N | THE INSULIN PUMP ALARMED BUTTON ERROR FOLLOWED BY UNRESPONSIVE BUTTONS DUE TO A FLATTENED ESCAPE DOME SWITCH UNABLE TO VERIFY THE RESET ALARMS DUE TO BUTTON ANOMALY THE DEVICE HAD A MISSING END CAP STICKER AND A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6592053 | 76039425 | N | THE INSULIN PUMP ALARMED BUTTON ERROR FOLLOWED BY UNRESPONSIVE BUTTONS DUE TO A FLATTENED ESCAPE DOME SWITCH UNABLE TO VERIFY THE RESET ALARMS DUE TO BUTTON ANOMALY THE DEVICE HAD A MISSING END CAP STICKER AND A SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6592870 | 76061304 | D | IT WAS REPORTED THAT A NURSE USE A G X IN BD ML BD LUERLOK BLUNT FILL NEEDLE TO DRAW UP NORMAL SALINE AS THE NURSE RECAPPED THE NEEDLE THE NEEDLE PUNCTURED ITS CAP AND THE NURSE SUFFERED A NEEDLE STICK INJURY THERE WAS NO REPORT OF PATIENT CONTACT WITH THE USED DEVICE OR ANY MEDICAL INTERVENTION | No Known Impact Or Consequence To Patient | Patient problem code
6592870 | 76061304 | D | IT WAS REPORTED THAT A NURSE USE A G X IN BD ML BD LUERLOK BLUNT FILL NEEDLE TO DRAW UP NORMAL SALINE AS THE NURSE RECAPPED THE NEEDLE THE NEEDLE PUNCTURED ITS CAP AND THE NURSE SUFFERED A NEEDLE STICK INJURY THERE WAS NO REPORT OF PATIENT CONTACT WITH THE USED DEVICE OR ANY MEDICAL INTERVENTION | Dull, Blunt | Device problem code
6591967 | 76132184 | N | RELIABILITY ANALYSIS INSPECTED ONE RANDOM OPENEDUSED PARTIAL ENLITE SENSOR AND PERFORM CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS ALSO CHECKED ONE OF TWO NEEDLES FOR BURRSHOOKS AND DULL NEEDLE TIP DEFECTS AND NONE WAS FOUND INTRODUCER NEEDLE PASSED PER SPECIFICATIONS MISSING ONE NEEDLE | Missing or Inadequate Safety Measures | Medical Evaluation result code
6592573 | 76139778 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM AND NO ACT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6592573 | 76139778 | N | THE INSULIN PUMP HAD A BUTTON ERROR ALARM AND NO ACT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6593702 | 76149908 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER RESERVOIR TUBE ORING MISSING SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Scratched Material | Device problem code
6593702 | 76149908 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER RESERVOIR TUBE ORING MISSING SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
6593702 | 76149908 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER RESERVOIR TUBE ORING MISSING SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
6594813 | 76175541 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTONCREASED NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Device Alarm System | Device problem code
6594813 | 76175541 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTONCREASED NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6594813 | 76175541 | N | NO BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC AND ACT BUTTONCREASED NO BUTTON ERROR ALARMS NOTED NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6595878 | 76177811 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT SHE RECEIVED MULTIPLE BATTERY AND POWER ALARMS THE CUSTOMER RECEIVED A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW IN LESS THAN HOURS AFTER LOW BATTERY WARNING CUSTOMERS BLOOD GLUCOSE LEVEL WAS NOT REPORTED THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6595878 | 76177811 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT SHE RECEIVED MULTIPLE BATTERY AND POWER ALARMS THE CUSTOMER RECEIVED A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY NOW ALARM THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY NOW IN LESS THAN HOURS AFTER LOW BATTERY WARNING CUSTOMERS BLOOD GLUCOSE LEVEL WAS NOT REPORTED THE CUSTOMER WAS ADVISED THAT THE DEVICE WOULD BE REPLACED AND AGREED TO RETURN THE PRODUCT FOR ANALYSIS | Battery Problem | Device problem code
6216516 | 76182076 | N | RESULTS OF INVESTIGATION CAREFUSION FAILURE ANALYSIS LAB TECHNICIAN WAS ABLE TO VERIFY AND DUPLICATE THE CUSTOMERS REPORTED ISSUE THE EXHALATION DIFFERENTIAL PRESSURE TRANSDUCER FAILED CALIBRATION THE ODV WAS OUTSIDE OF SPECIFICATIONS | Calibration Problem | Device problem code
6216516 | 76182076 | N | RESULTS OF INVESTIGATION CAREFUSION FAILURE ANALYSIS LAB TECHNICIAN WAS ABLE TO VERIFY AND DUPLICATE THE CUSTOMERS REPORTED ISSUE THE EXHALATION DIFFERENTIAL PRESSURE TRANSDUCER FAILED CALIBRATION THE ODV WAS OUTSIDE OF SPECIFICATIONS | Failure to Calibrate | Medical Evaluation result code
6594980 | 76182213 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO A FLATTENED DOWN ARROW BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6594980 | 76182213 | N | THE INSULIN PUMP RECEIVED WITH BUTTON ERROR ALARMS DUE TO A FLATTENED DOWN ARROW BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6594726 | 76188707 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESCAPE BUTTON ON KEYPAD TRACE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6594726 | 76188707 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO FLATTENED DOME SWITCH AT ESCAPE BUTTON ON KEYPAD TRACE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596718 | 76239378 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6596718 | 76239378 | N | THE INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ACT BUTTON DOME SWITCH NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596536 | 76242789 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6596536 | 76242789 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ACT BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596727 | 76244579 | N | NO BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED DOME SWITCH ON ACT AND DOWN ARROW BUTTON THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6596727 | 76244579 | N | NO BUTTON RESPONSE AND BUTTON ERROR ALARM DUE TO FLATTENED DOME SWITCH ON ACT AND DOWN ARROW BUTTON THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596574 | 76245325 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON B BUTTON UNIT HAD SCRATCHED ON DISPLAY WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6596574 | 76245325 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON B BUTTON UNIT HAD SCRATCHED ON DISPLAY WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6596574 | 76245325 | N | UNIT RECEIVED BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON B BUTTON UNIT HAD SCRATCHED ON DISPLAY WINDOW THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596545 | 76246582 | N | INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON CREASED DOME SWITCH UNIT WAS RECEIVED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6596545 | 76246582 | N | INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON CREASED DOME SWITCH UNIT WAS RECEIVED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Missing or Inadequate Safety Measures | Medical Evaluation result code
6596545 | 76246582 | N | INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON CREASED DOME SWITCH UNIT WAS RECEIVED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6596545 | 76246582 | N | INSULIN PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ESC BUTTON CREASED DOME SWITCH UNIT WAS RECEIVED WITH MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6597030 | 76249341 | N | THE INSULIN PUMP WAS RECEIVED WITH A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6597030 | 76249341 | N | THE INSULIN PUMP WAS RECEIVED WITH A BUTTON ERROR ALARMS DUE TO FLATTENED DOME ON ESC BUTTON NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596746 | 76251971 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO THE DEVICE HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6596746 | 76251971 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO THE DEVICE HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component Missing | Device problem code
6596746 | 76251971 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARMED DUE TO MOISTURE ON KEYPAD TRACE ALSO THE DEVICE HAD MISSING END CAP STICKER NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6596930 | 76255222 | N | PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NO BLANK DISPLAY NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device problem code
6596930 | 76255222 | N | PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NO BLANK DISPLAY NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | No Display / Image | Device Problem Code
6596930 | 76255222 | N | PUMP WAS RECEIVED WITH BUTTON ERROR ALARM DUE TO FLATTENED ALL THE BUTTONS DOME SWITCH NO BLANK DISPLAY NOTED THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6599562 | 76264701 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND INTERMITTENT BUTTON RESPONSE NOTED DUE TO FLATTENED DOME SWITCH ON THE UP ARROW BUTTON AND DOWN ARROW BUTTON THE DEVICE WAS RECEIVED WITH DOG CHEW MARKS AROUND CASING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6599562 | 76264701 | N | THE INSULIN PUMP HAD BUTTON ERROR ALARM AND INTERMITTENT BUTTON RESPONSE NOTED DUE TO FLATTENED DOME SWITCH ON THE UP ARROW BUTTON AND DOWN ARROW BUTTON THE DEVICE WAS RECEIVED WITH DOG CHEW MARKS AROUND CASING NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON B SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6600591 | 76266126 | D | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING THE DEVICE SELF DISCHARGED COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6600591 | 76266126 | D | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING THE DEVICE SELF DISCHARGED COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Unintended Ejection | Device problem code
6392290 | 76525919 | D | THIS COMPLAINT WAS CREATED BY ACCIDENTALLY THERE WAS NO DEVICE ISSUE THERE WAS NO REPORTED PATIENT INVOLVEMENT | No Patient Involvement | Patient problem code
6392290 | 76525919 | D | THIS COMPLAINT WAS CREATED BY ACCIDENTALLY THERE WAS NO DEVICE ISSUE THERE WAS NO REPORTED PATIENT INVOLVEMENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6541996 | 77370464 | N | THE EXTERNAL VISUAL INSPECTION REVEALED A MISSING ELECTRODE GROUND RING AND A CUT ELECTRODE PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED ONE OF THE ELECTRICAL TESTS FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED | Component Missing | Device problem code
6541996 | 77370464 | N | THE EXTERNAL VISUAL INSPECTION REVEALED A MISSING ELECTRODE GROUND RING AND A CUT ELECTRODE PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED ONE OF THE ELECTRICAL TESTS FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED | Mechanical Jam | Device problem code
6541996 | 77370464 | N | THE EXTERNAL VISUAL INSPECTION REVEALED A MISSING ELECTRODE GROUND RING AND A CUT ELECTRODE PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED ONE OF THE ELECTRICAL TESTS FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED | Mechanical Problem Identified | Medical Evaluation result code
6541996 | 77370464 | N | THE EXTERNAL VISUAL INSPECTION REVEALED A MISSING ELECTRODE GROUND RING AND A CUT ELECTRODE PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED ONE OF THE ELECTRICAL TESTS FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED | Missing or Inadequate Safety Measures | Medical Evaluation result code
6662841 | 78461129 | N | THE INVESTIGATION DETERMINED THAT LOWER THAN EXPECTED MYOGLOBIN RESULTS WERE OBTAINED FROM BIORAD LEVEL QUALITY CONTROL FLUID USING VITROS MYOGLOBIN REAGENT ON A VITROS INTEGRATED SYSTEM THE ASSIGNABLE CAUSE WAS DETERMINED TO BE INSTRUMENT RELATED REAGENT AND SAMPLE METERING CONDITION CODES WERE BEING POSTED BY THE INSTRUMENT AT THE TIME THE LOWER THAN EXPECTED VITROS MYOG RESULTS WERE GENERATED ORTHO TECHNICAL SUPPORT REVIEWED HISTORICAL BIORAD QC FLUID PERFORMANCE ACROSS THE TWO VITROS SYSTEMS IN THE LABORATORY AND ONLY ONE VITROS SYSTEM SHOWED MYOGLOBIN IMPRECISION SUGGESTING THAT THE ISSUE IS MOST LIKELY INSTRUMENT AND NOT REAGENT RELATED THE CUSTOMER PERFORMED MAINTENANCE ACTIONS OF FLUSHING THE REAGENT AND SAMPLE METERING PROBES FOLLOWING THESE ACTIONS ACCEPTABLE WITHIN RUN PRECISION RESULTS WERE OBTAINED ON THE AFFECTED VITROS SYSTEM IN ADDITION ACCEPTABLE BIORAD QC PERFORMANCE HAS MAINTAINED FOLLOWING MAINTENANCE ACTIONS | Low Test Results | Device problem code
6662841 | 78461129 | N | THE INVESTIGATION DETERMINED THAT LOWER THAN EXPECTED MYOGLOBIN RESULTS WERE OBTAINED FROM BIORAD LEVEL QUALITY CONTROL FLUID USING VITROS MYOGLOBIN REAGENT ON A VITROS INTEGRATED SYSTEM THE ASSIGNABLE CAUSE WAS DETERMINED TO BE INSTRUMENT RELATED REAGENT AND SAMPLE METERING CONDITION CODES WERE BEING POSTED BY THE INSTRUMENT AT THE TIME THE LOWER THAN EXPECTED VITROS MYOG RESULTS WERE GENERATED ORTHO TECHNICAL SUPPORT REVIEWED HISTORICAL BIORAD QC FLUID PERFORMANCE ACROSS THE TWO VITROS SYSTEMS IN THE LABORATORY AND ONLY ONE VITROS SYSTEM SHOWED MYOGLOBIN IMPRECISION SUGGESTING THAT THE ISSUE IS MOST LIKELY INSTRUMENT AND NOT REAGENT RELATED THE CUSTOMER PERFORMED MAINTENANCE ACTIONS OF FLUSHING THE REAGENT AND SAMPLE METERING PROBES FOLLOWING THESE ACTIONS ACCEPTABLE WITHIN RUN PRECISION RESULTS WERE OBTAINED ON THE AFFECTED VITROS SYSTEM IN ADDITION ACCEPTABLE BIORAD QC PERFORMANCE HAS MAINTAINED FOLLOWING MAINTENANCE ACTIONS | Imprecision | Device problem code
6662841 | 78461129 | N | THE INVESTIGATION DETERMINED THAT LOWER THAN EXPECTED MYOGLOBIN RESULTS WERE OBTAINED FROM BIORAD LEVEL QUALITY CONTROL FLUID USING VITROS MYOGLOBIN REAGENT ON A VITROS INTEGRATED SYSTEM THE ASSIGNABLE CAUSE WAS DETERMINED TO BE INSTRUMENT RELATED REAGENT AND SAMPLE METERING CONDITION CODES WERE BEING POSTED BY THE INSTRUMENT AT THE TIME THE LOWER THAN EXPECTED VITROS MYOG RESULTS WERE GENERATED ORTHO TECHNICAL SUPPORT REVIEWED HISTORICAL BIORAD QC FLUID PERFORMANCE ACROSS THE TWO VITROS SYSTEMS IN THE LABORATORY AND ONLY ONE VITROS SYSTEM SHOWED MYOGLOBIN IMPRECISION SUGGESTING THAT THE ISSUE IS MOST LIKELY INSTRUMENT AND NOT REAGENT RELATED THE CUSTOMER PERFORMED MAINTENANCE ACTIONS OF FLUSHING THE REAGENT AND SAMPLE METERING PROBES FOLLOWING THESE ACTIONS ACCEPTABLE WITHIN RUN PRECISION RESULTS WERE OBTAINED ON THE AFFECTED VITROS SYSTEM IN ADDITION ACCEPTABLE BIORAD QC PERFORMANCE HAS MAINTAINED FOLLOWING MAINTENANCE ACTIONS | Interoperability Problem Identified | Medical Evaluation result code
6477257 | 78606955 | N | PHILIPS RECEIVED UPDATED INFORMATION STATING THAT THIS REPORT WAS A FAILURE TO DISCHARGE THAT DID NOT OCCUR WHILE IN USE ON A PATIENT | Leak / Splash | Device problem code
6477257 | 78606955 | N | PHILIPS RECEIVED UPDATED INFORMATION STATING THAT THIS REPORT WAS A FAILURE TO DISCHARGE THAT DID NOT OCCUR WHILE IN USE ON A PATIENT | Failure to Discharge | Device Problem Code
6671494 | 78957281 | N | THE INVESTIGATION DETERMINED THAT A HIGHER THAN EXPECTED VALP RESULT WAS OBTAINED FROM A VITROS PERFORMANCE VERIFIER QUALITY CONTROL FLUID PROCESSED USING VITROS VALP REAGENT WITH A VITROS CHEMISTRY SYSTEM THERE IS NO INDICATION A REAGENT ISSUE CONTRIBUTED TO THE EVENT THE MOST LIKELY ASSIGNABLE CAUSE IS INSTRUMENT RELATED THE RESULTS OF PRESERVICE WITHINRUN PRECISION TESTING DEMONSTRATED THAT THE VITROS SYSTEM WAS NOT OPERATING AS EXPECTED AT THE TIME OF THE EVENT AN ORTHO FIELD ENGINEER PERFORMED SERVICE ACTIONS TO ADJUST THE CUVETTE INCUBATOR METERING TRUCKS AND CLEAN RESIDUE FROM THE READ CHANNEL THE VALP QUALITY CONTROL RESULTS OBTAINED AFTER SERVICE INDICATE ACCEPTABLE PERFORMANCE HAS BEEN MAINTAINED | Residue After Decontamination | Device problem code
6671494 | 78957281 | N | THE INVESTIGATION DETERMINED THAT A HIGHER THAN EXPECTED VALP RESULT WAS OBTAINED FROM A VITROS PERFORMANCE VERIFIER QUALITY CONTROL FLUID PROCESSED USING VITROS VALP REAGENT WITH A VITROS CHEMISTRY SYSTEM THERE IS NO INDICATION A REAGENT ISSUE CONTRIBUTED TO THE EVENT THE MOST LIKELY ASSIGNABLE CAUSE IS INSTRUMENT RELATED THE RESULTS OF PRESERVICE WITHINRUN PRECISION TESTING DEMONSTRATED THAT THE VITROS SYSTEM WAS NOT OPERATING AS EXPECTED AT THE TIME OF THE EVENT AN ORTHO FIELD ENGINEER PERFORMED SERVICE ACTIONS TO ADJUST THE CUVETTE INCUBATOR METERING TRUCKS AND CLEAN RESIDUE FROM THE READ CHANNEL THE VALP QUALITY CONTROL RESULTS OBTAINED AFTER SERVICE INDICATE ACCEPTABLE PERFORMANCE HAS BEEN MAINTAINED | Contamination of Environment by Device | Medical Evaluation result code
6699429 | 79572490 | D | ON B THE PATIENT UNDERWENT TREATMENT OF A TYPE B AORTIC DISSECTION WITH TWO GORE TAG THORACIC ENDOPROSTHESES ON B FOLLOW UP IMAGING REPORTEDLY SHOWED A PROXIMAL TYPE I ENDOLEAK PERSISTENT FALSE LUMEN PERFUSION EXPANSION OF THE FALSE LUMEN AMOUNT UNKNOWN AND EXTENSION OF THE DISSECTION PROXIMALLY INTO THE ASCENDING AORTA IT WAS REPORTED THE PROXIMAL TYPE I ENDOLEAK WAS CAUSED BY EXPANSION OF THE EXISTING DISSECTION RESULTING IN A LOSS OF SEAL PROXIMALLY THE PATIENT WAS REPORTED TO BE ASYMPTOMATIC AND DOING WELL ON AN UNKNOWN DATE A DEBRANCHING PROCEDURE WAS PERFORMED WHEREBY A SURGICAL GRAFT WAS IMPLANTED IN THE ASCENDING AORTA TO BYPASS THE INNOMINATE AND CAROTID ARTERIES ON B AN ADDITIONAL PROCEDURE WAS PERFORMED WHEREBY A COOK ALPHA GRAFT WAS IMPLANTED FOR TREATMENT OF THE ENDOLEAK AND PERSISTENT FALSE LUMEN PERFUSION IT WAS REPORTED A SMALL PROXIMAL TYPE I ENDOLEAK REMAINED AT THE END OF THE PROCEDURE THE PHYSICIAN REPORTEDLY PLANS TO MONITOR THE ENDOLEAK AT THIS TIME NO FURTHER ADVERSE EVENTS WERE REPORTED AND THE PATIENT TOLERATED THE PROCEDURE ADDITIONAL INFORMATION HAS BEEN REQUESTED | Insufficient Information | Device problem code
6699429 | 79572490 | D | ON B THE PATIENT UNDERWENT TREATMENT OF A TYPE B AORTIC DISSECTION WITH TWO GORE TAG THORACIC ENDOPROSTHESES ON B FOLLOW UP IMAGING REPORTEDLY SHOWED A PROXIMAL TYPE I ENDOLEAK PERSISTENT FALSE LUMEN PERFUSION EXPANSION OF THE FALSE LUMEN AMOUNT UNKNOWN AND EXTENSION OF THE DISSECTION PROXIMALLY INTO THE ASCENDING AORTA IT WAS REPORTED THE PROXIMAL TYPE I ENDOLEAK WAS CAUSED BY EXPANSION OF THE EXISTING DISSECTION RESULTING IN A LOSS OF SEAL PROXIMALLY THE PATIENT WAS REPORTED TO BE ASYMPTOMATIC AND DOING WELL ON AN UNKNOWN DATE A DEBRANCHING PROCEDURE WAS PERFORMED WHEREBY A SURGICAL GRAFT WAS IMPLANTED IN THE ASCENDING AORTA TO BYPASS THE INNOMINATE AND CAROTID ARTERIES ON B AN ADDITIONAL PROCEDURE WAS PERFORMED WHEREBY A COOK ALPHA GRAFT WAS IMPLANTED FOR TREATMENT OF THE ENDOLEAK AND PERSISTENT FALSE LUMEN PERFUSION IT WAS REPORTED A SMALL PROXIMAL TYPE I ENDOLEAK REMAINED AT THE END OF THE PROCEDURE THE PHYSICIAN REPORTEDLY PLANS TO MONITOR THE ENDOLEAK AT THIS TIME NO FURTHER ADVERSE EVENTS WERE REPORTED AND THE PATIENT TOLERATED THE PROCEDURE ADDITIONAL INFORMATION HAS BEEN REQUESTED | Hemorrhage | Patient problem code
6691087 | 79740465 | N | THE INVESTIGATION HAS DETERMINED THAT NONREPRODUCIBLE LOWER THAN EXPECTED VITROS GLU RESULT WAS OBTAINED FROM ONE PATIENT SAMPLE TESTED ON A VITROS INTEGRATED SYSTEM THE ISSUE WAS DESCRIBED BY THE CUSTOMER AS RECURRING AT RANDOM TIMES FOR THE PAST SEVERAL WEEKS AND WAS OBSERVED ACROSS MULTIPLE ASSAYS AND ONLY ON THIS VITROS SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT WAS INSTRUMENT RELATED AN ORTHO FIELD ENGINEER FE PERFORMED SERVICE ACTIONS TO THE SAMPLE METERING UNIT ALONG WITH ALL NECESSARY MAINTENANCE ADJUSTMENTS RELATED SPECIFICALLY TO THE ZAXIS PINIONS OF THE SAMPLE METERING PUMP AS WELL AS ADJUSTMENTS TO THE TIP LOCATOR SYSTEM FOLLOWING THIS ACTIVITY THE CUSTOMER WAS ABLE TO PERFORM VITROS TESTING WITHOUT ANY ADDITIONAL RECURRENCES OF NONREPRODUCIBLE RESULTS AND CONDITION CODES | Component or Accessory Incompatibility | Device Problem Code
6691087 | 79740465 | N | THE INVESTIGATION HAS DETERMINED THAT NONREPRODUCIBLE LOWER THAN EXPECTED VITROS GLU RESULT WAS OBTAINED FROM ONE PATIENT SAMPLE TESTED ON A VITROS INTEGRATED SYSTEM THE ISSUE WAS DESCRIBED BY THE CUSTOMER AS RECURRING AT RANDOM TIMES FOR THE PAST SEVERAL WEEKS AND WAS OBSERVED ACROSS MULTIPLE ASSAYS AND ONLY ON THIS VITROS SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT WAS INSTRUMENT RELATED AN ORTHO FIELD ENGINEER FE PERFORMED SERVICE ACTIONS TO THE SAMPLE METERING UNIT ALONG WITH ALL NECESSARY MAINTENANCE ADJUSTMENTS RELATED SPECIFICALLY TO THE ZAXIS PINIONS OF THE SAMPLE METERING PUMP AS WELL AS ADJUSTMENTS TO THE TIP LOCATOR SYSTEM FOLLOWING THIS ACTIVITY THE CUSTOMER WAS ABLE TO PERFORM VITROS TESTING WITHOUT ANY ADDITIONAL RECURRENCES OF NONREPRODUCIBLE RESULTS AND CONDITION CODES | Interoperability Problem Identified | Medical Evaluation result code
6691087 | 79740465 | N | THE INVESTIGATION HAS DETERMINED THAT NONREPRODUCIBLE LOWER THAN EXPECTED VITROS GLU RESULT WAS OBTAINED FROM ONE PATIENT SAMPLE TESTED ON A VITROS INTEGRATED SYSTEM THE ISSUE WAS DESCRIBED BY THE CUSTOMER AS RECURRING AT RANDOM TIMES FOR THE PAST SEVERAL WEEKS AND WAS OBSERVED ACROSS MULTIPLE ASSAYS AND ONLY ON THIS VITROS SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT WAS INSTRUMENT RELATED AN ORTHO FIELD ENGINEER FE PERFORMED SERVICE ACTIONS TO THE SAMPLE METERING UNIT ALONG WITH ALL NECESSARY MAINTENANCE ADJUSTMENTS RELATED SPECIFICALLY TO THE ZAXIS PINIONS OF THE SAMPLE METERING PUMP AS WELL AS ADJUSTMENTS TO THE TIP LOCATOR SYSTEM FOLLOWING THIS ACTIVITY THE CUSTOMER WAS ABLE TO PERFORM VITROS TESTING WITHOUT ANY ADDITIONAL RECURRENCES OF NONREPRODUCIBLE RESULTS AND CONDITION CODES | Low Test Results | Device problem code
6691087 | 79740465 | N | THE INVESTIGATION HAS DETERMINED THAT NONREPRODUCIBLE LOWER THAN EXPECTED VITROS GLU RESULT WAS OBTAINED FROM ONE PATIENT SAMPLE TESTED ON A VITROS INTEGRATED SYSTEM THE ISSUE WAS DESCRIBED BY THE CUSTOMER AS RECURRING AT RANDOM TIMES FOR THE PAST SEVERAL WEEKS AND WAS OBSERVED ACROSS MULTIPLE ASSAYS AND ONLY ON THIS VITROS SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT WAS INSTRUMENT RELATED AN ORTHO FIELD ENGINEER FE PERFORMED SERVICE ACTIONS TO THE SAMPLE METERING UNIT ALONG WITH ALL NECESSARY MAINTENANCE ADJUSTMENTS RELATED SPECIFICALLY TO THE ZAXIS PINIONS OF THE SAMPLE METERING PUMP AS WELL AS ADJUSTMENTS TO THE TIP LOCATOR SYSTEM FOLLOWING THIS ACTIVITY THE CUSTOMER WAS ABLE TO PERFORM VITROS TESTING WITHOUT ANY ADDITIONAL RECURRENCES OF NONREPRODUCIBLE RESULTS AND CONDITION CODES | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6708676 | 79884580 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6708676 | 79884580 | D | IT IS ALLEGED THAT PT RECEIVED A COOK GUNTHER TULIP ON B IT IS ALLEGED THAT THE PATIENT WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6713252 | 80008657 | N | B RESULTS OF INVESTIGATION A CAREFUSION FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR CHECKED THE MEAN AIRWAY PRESSURE MAP TRANSDUCER FOR DRIFT AND CALIBRATED THE PNEUMATICS CORNER AND THE DRIVER DISPLACEMENT INDICATOR DDI BOARD THE FSR FOUND THE DDI WAS OUT OF SPECIFICATION THE FSR ADJUSTED THE DDI TO TIGHT TOLERANCES AND COMPLETED THE ALARMS TEST THE FSR COMPLETED THE PATIENT CIRCUIT CALIBRATION AND THE PERFORMANCE CHECK THE UNIT MEETS MANUFACTURER SPECIFICATIONS | Calibration Problem | Device problem code
6713252 | 80008657 | N | B RESULTS OF INVESTIGATION A CAREFUSION FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR CHECKED THE MEAN AIRWAY PRESSURE MAP TRANSDUCER FOR DRIFT AND CALIBRATED THE PNEUMATICS CORNER AND THE DRIVER DISPLACEMENT INDICATOR DDI BOARD THE FSR FOUND THE DDI WAS OUT OF SPECIFICATION THE FSR ADJUSTED THE DDI TO TIGHT TOLERANCES AND COMPLETED THE ALARMS TEST THE FSR COMPLETED THE PATIENT CIRCUIT CALIBRATION AND THE PERFORMANCE CHECK THE UNIT MEETS MANUFACTURER SPECIFICATIONS | Failure to Calibrate | Medical Evaluation result code
6713857 | 80044808 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE FAILED TO DETECT THE ATTACHED EXTERNAL PADDLES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Device Difficult to Maintain | Device Problem Code
6713857 | 80044808 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE FAILED TO DETECT THE ATTACHED EXTERNAL PADDLES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6710293 | 80047601 | N | B THE ISSUE WAS DISCOVERED WHEN THE BIOMEDICAL TECHNICIAN BIOMED WAS TESTING THE PUMP THE BIOMED SHUT DOWN THE MACHINE SWAPPED THE PUMP CABLE AND AFTER THAT THE UNIT WORKED FINE | Component or Accessory Incompatibility | Device Problem Code
6710293 | 80047601 | N | B THE ISSUE WAS DISCOVERED WHEN THE BIOMEDICAL TECHNICIAN BIOMED WAS TESTING THE PUMP THE BIOMED SHUT DOWN THE MACHINE SWAPPED THE PUMP CABLE AND AFTER THAT THE UNIT WORKED FINE | Interoperability Problem Identified | Medical Evaluation result code
6724088 | 80382795 | N | B THE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE TO EVALUATE THE SUSPECT DEVICE THE FSR CONFIRMED THE REPORTED EVENT AND DETERMINED THE MOST LIKELY CAUSE OF THE ISSUE IS THE EXHALATION VALVE AFTER REPLACING THE EXHALATION VALVE AND PERFORMING CALIBRATION THE ISSUE WAS RESOLVED THE FSR PERFORMED THE OPERATIONAL VERIFICATION PROCEDURE AND REPORTED THE UNIT MEETS SERVICE SPECIFICATIONS AT THIS TIME THE SUSPECT COMPONENT IS NOT AVAILABLE FOR RETURN FOR FURTHER EVALUATION DUE TO THE PART NOT BEING AVAILABLE TO BE RETURNED NO FURTHER EVALUATION CAN BE COMPLETED AT THIS TIME | Calibration Problem | Device problem code
6724088 | 80382795 | N | B THE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE TO EVALUATE THE SUSPECT DEVICE THE FSR CONFIRMED THE REPORTED EVENT AND DETERMINED THE MOST LIKELY CAUSE OF THE ISSUE IS THE EXHALATION VALVE AFTER REPLACING THE EXHALATION VALVE AND PERFORMING CALIBRATION THE ISSUE WAS RESOLVED THE FSR PERFORMED THE OPERATIONAL VERIFICATION PROCEDURE AND REPORTED THE UNIT MEETS SERVICE SPECIFICATIONS AT THIS TIME THE SUSPECT COMPONENT IS NOT AVAILABLE FOR RETURN FOR FURTHER EVALUATION DUE TO THE PART NOT BEING AVAILABLE TO BE RETURNED NO FURTHER EVALUATION CAN BE COMPLETED AT THIS TIME | Failure to Calibrate | Medical Evaluation result code
6719320 | 80400807 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE ESCAPE BUTTON DUE TO UNLOCKED J CONNECTOR AT LCD BOARD PUMP WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Component or Accessory Incompatibility | Device Problem Code
6719320 | 80400807 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE ESCAPE BUTTON DUE TO UNLOCKED J CONNECTOR AT LCD BOARD PUMP WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Scratched Material | Device problem code
6719320 | 80400807 | N | PUMP WAS RECEIVED WITH UNRESPONSIVE ESCAPE BUTTON DUE TO UNLOCKED J CONNECTOR AT LCD BOARD PUMP WAS RECEIVED WITH SCRATCHED LCD WINDOW NOTE THIS IS A REMEDIATION MDR MEDTRONIC DIABETES IMPLEMENTED REVISED MDR REPORTABILITY CRITERIA EFFECTIVE ON JULY SUBSEQUENTLY MEDTRONIC DIABETES CONDUCTED A ONE YEAR RETROSPECTIVE REVIEW OF COMPLAINTS THIS EVENT WAS RETROSPECTIVELY IDENTIFIED TO BE REPORTABLE BASED ON THE REVISED MDR REPORTABILITY CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6725077 | 80440264 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM WITH A GORE EXCLUDER BIFURCATED ENDOPROSTHESIS TRUNKIPSILATERAL LEG COMPONENT AND CONTRALATERAL LEG COMPONENT ON AN UNKNOWN DATE FOLLOWUP IMAGING SHOWED THAT THE ANEURYSM HAD GROWN AN UNKNOWN AMOUNT IT WAS REPORTED THAT THE PHYSICIAN BELIEVES THE ANEURYSM ENLARGEMENT WAS CAUSED BY ENDOTENSION DUE TO NOT SEEING ANY EVIDENCE OF ENDOLEAK ON B AN ADDITIONAL PROCEDURE WAS PERFORMED WHEREBY THE PREVIOUSLY IMPLANTED DEVICES WERE RELINED WITH TWO ADDITIONAL GORE EXCLUDER AAA ENDOPROSTHESES TO TREAT THE SUSPECTED ENDOTENSION THE PATIENT TOLERATED THE PROCEDURE | Aneurysm | Patient problem code
6725077 | 80440264 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM WITH A GORE EXCLUDER BIFURCATED ENDOPROSTHESIS TRUNKIPSILATERAL LEG COMPONENT AND CONTRALATERAL LEG COMPONENT ON AN UNKNOWN DATE FOLLOWUP IMAGING SHOWED THAT THE ANEURYSM HAD GROWN AN UNKNOWN AMOUNT IT WAS REPORTED THAT THE PHYSICIAN BELIEVES THE ANEURYSM ENLARGEMENT WAS CAUSED BY ENDOTENSION DUE TO NOT SEEING ANY EVIDENCE OF ENDOLEAK ON B AN ADDITIONAL PROCEDURE WAS PERFORMED WHEREBY THE PREVIOUSLY IMPLANTED DEVICES WERE RELINED WITH TWO ADDITIONAL GORE EXCLUDER AAA ENDOPROSTHESES TO TREAT THE SUSPECTED ENDOTENSION THE PATIENT TOLERATED THE PROCEDURE | Insufficient Information | Device problem code
6734156 | 80748966 | N | B MEDICATIONS AMLODIPINE ASPIRIN ATORVASTATIN CITALOPRAM LISINOPRIL METOPROLOL RANITIDINE DICYCLOMINE PROBIOTIC THE REVIEW OF THE MANUFACTURING PAPERWORK VERIFIED THAT THIS LOT MET ALL PRERELEASE SPECIFICATIONSACCORDING TO THE GORE EXCLUDER AAA ENDOPROSTHESES INSTRUCTIONS FOR USE ADVERSE EVENTS WHICH MAY OCCUR AND REQUIRE INTERVENTION INCLUDE ENDOLEAK | No Findings Available | Medical Evaluation result code
6737122 | 80940933 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE SERVICE TERRITORY MANAGER STM REPORTED THAT THE CUSTOMER HAD ISSUES WITH THE DISPLAY IT WAS UNREADABLE AND SCRAMBLED THE IABP CONTINUED TO OPERATE WITHOUT LEGIBLE VIDEO THE STM COULD NOT RECREATE THE REPORTED ISSUE WITH DISPLAY DID REPLACE THE PCBCOLOR VIDEO RECEIVER AND CABLECOILED ASSOCIATED WITH DISPLAY TO TRY AND PREVENT FURTHER ISSUES THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Component or Accessory Incompatibility | Device Problem Code
6737122 | 80940933 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE SERVICE TERRITORY MANAGER STM REPORTED THAT THE CUSTOMER HAD ISSUES WITH THE DISPLAY IT WAS UNREADABLE AND SCRAMBLED THE IABP CONTINUED TO OPERATE WITHOUT LEGIBLE VIDEO THE STM COULD NOT RECREATE THE REPORTED ISSUE WITH DISPLAY DID REPLACE THE PCBCOLOR VIDEO RECEIVER AND CABLECOILED ASSOCIATED WITH DISPLAY TO TRY AND PREVENT FURTHER ISSUES THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Interoperability Problem Identified | Medical Evaluation result code
6489935 | 81142921 | N | THE INITIAL EVALUATION OF THE DEVICE WAS DONE BY THE HOSPITAL BIOMEDICAL ENGINEER WHO NOTED AN ERROR CODE IN THE DEVICE LOGS THIS ERROR CODE INDICATES A POTENTIAL ISSUE WITH THE POWER PCATHE BIOMEDICAL WAS INFORMED TO REPLACE THE POWER PCA THE BIOMED STATED THAT HE WILL REPLACE THE POWER PCA AND IF THAT DOES NOT FIX THE PROBLEM HE WILL SEND THE UNIT TO THE BENCH THE DEVICE REMAINS AT THE CUSTOMER SITE NO FURTHER ACTION WAS TAKEN AND NONE IS WARRANTED A SEARCH OF SUBSEQUENT RECORDS WAS CONDUCTED AND YIELDED NO RESULTS RELATED TO THE CUSTOMERS REPORTED PROBLEMTHIS WILL BE CONSIDERED A MALFUNCTION OF POWER PCA BASED ON THE LIMITED AVAILABLE INFORMATION THERE IS NO INDICATION OF A SYSTEMIC PROBLEM NO FURTHER INVESTIGATION OR ACTION IS WARRANTED PATIENT INFORMATION HAS BEEN REQUESTED UNAVAILABLE AT THE TIME OF THIS REPORT | Device Displays Incorrect Message | Device problem code
6489935 | 81142921 | N | THE INITIAL EVALUATION OF THE DEVICE WAS DONE BY THE HOSPITAL BIOMEDICAL ENGINEER WHO NOTED AN ERROR CODE IN THE DEVICE LOGS THIS ERROR CODE INDICATES A POTENTIAL ISSUE WITH THE POWER PCATHE BIOMEDICAL WAS INFORMED TO REPLACE THE POWER PCA THE BIOMED STATED THAT HE WILL REPLACE THE POWER PCA AND IF THAT DOES NOT FIX THE PROBLEM HE WILL SEND THE UNIT TO THE BENCH THE DEVICE REMAINS AT THE CUSTOMER SITE NO FURTHER ACTION WAS TAKEN AND NONE IS WARRANTED A SEARCH OF SUBSEQUENT RECORDS WAS CONDUCTED AND YIELDED NO RESULTS RELATED TO THE CUSTOMERS REPORTED PROBLEMTHIS WILL BE CONSIDERED A MALFUNCTION OF POWER PCA BASED ON THE LIMITED AVAILABLE INFORMATION THERE IS NO INDICATION OF A SYSTEMIC PROBLEM NO FURTHER INVESTIGATION OR ACTION IS WARRANTED PATIENT INFORMATION HAS BEEN REQUESTED UNAVAILABLE AT THE TIME OF THIS REPORT | Results Pending Completion of Investigation | Medical Evaluation result code
6744979 | 81253346 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE TEST WELL IMAGES IN QUESTION ON THE INSTRUMENT IN QUESTION WHICH APPEARED VISUALLY VERY WEAKLY POSITIVE | Connection Problem | Device problem code
6744979 | 81253346 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE TEST WELL IMAGES IN QUESTION ON THE INSTRUMENT IN QUESTION WHICH APPEARED VISUALLY VERY WEAKLY POSITIVE | Communications Problem Identified | Medical Evaluation result code
6746926 | 81295648 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE IABP WAS EVALUATED BY THE FACILITYS BIOMEDICAL ENGINEER AS PER THE FACILITYS BIOMEDICAL ENGINEER A NEW FIBER OPTIC ASSEMBLY WAS PURCHASED AND THE REPORTED ISSUE WAS RESOLVED | Component or Accessory Incompatibility | Device Problem Code
6746926 | 81295648 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE IABP WAS EVALUATED BY THE FACILITYS BIOMEDICAL ENGINEER AS PER THE FACILITYS BIOMEDICAL ENGINEER A NEW FIBER OPTIC ASSEMBLY WAS PURCHASED AND THE REPORTED ISSUE WAS RESOLVED | Interoperability Problem Identified | Medical Evaluation result code
6753866 | 81393918 | N | THE CUSTOMER REPORTED THAT PATIENT RESPONDED WITH BURN ON TREATED AREA A FIELD SERVICE ENGINEER VISITED THIS ACCOUNT ON JULY RD AND REPORTED THAT THIS SYSTEM WAS MISSING CORRECT CALIBRATION OUR ADDITIONAL ATTEMPTS TO COLLECT MORE INFORMATION WERE NOT SUCCESSFUL IT IS DIFFICULT TO CONNECT AND COLLECT ANY INFORMATION BECAUSE THE COUNTRY IS IN PROTEST FOR THE LAST WEEKS | Calibration Problem | Device problem code
6753866 | 81393918 | N | THE CUSTOMER REPORTED THAT PATIENT RESPONDED WITH BURN ON TREATED AREA A FIELD SERVICE ENGINEER VISITED THIS ACCOUNT ON JULY RD AND REPORTED THAT THIS SYSTEM WAS MISSING CORRECT CALIBRATION OUR ADDITIONAL ATTEMPTS TO COLLECT MORE INFORMATION WERE NOT SUCCESSFUL IT IS DIFFICULT TO CONNECT AND COLLECT ANY INFORMATION BECAUSE THE COUNTRY IS IN PROTEST FOR THE LAST WEEKS | Component Missing | Device problem code
6753866 | 81393918 | N | THE CUSTOMER REPORTED THAT PATIENT RESPONDED WITH BURN ON TREATED AREA A FIELD SERVICE ENGINEER VISITED THIS ACCOUNT ON JULY RD AND REPORTED THAT THIS SYSTEM WAS MISSING CORRECT CALIBRATION OUR ADDITIONAL ATTEMPTS TO COLLECT MORE INFORMATION WERE NOT SUCCESSFUL IT IS DIFFICULT TO CONNECT AND COLLECT ANY INFORMATION BECAUSE THE COUNTRY IS IN PROTEST FOR THE LAST WEEKS | Failure to Calibrate | Medical Evaluation result code
6753866 | 81393918 | N | THE CUSTOMER REPORTED THAT PATIENT RESPONDED WITH BURN ON TREATED AREA A FIELD SERVICE ENGINEER VISITED THIS ACCOUNT ON JULY RD AND REPORTED THAT THIS SYSTEM WAS MISSING CORRECT CALIBRATION OUR ADDITIONAL ATTEMPTS TO COLLECT MORE INFORMATION WERE NOT SUCCESSFUL IT IS DIFFICULT TO CONNECT AND COLLECT ANY INFORMATION BECAUSE THE COUNTRY IS IN PROTEST FOR THE LAST WEEKS | Missing or Inadequate Safety Measures | Medical Evaluation result code
6755563 | 81465196 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HER PRODIGY DIABETES METER THE END USER WAS SWEATING AND NONRESPONSIVE ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL ADDITIONAL BLOOD GLUCOSE TESTS WERE PERFORMED WITH RESULTS RANGING FROM MGDL MGDL THE END USER WAS TAKEN TO THE ER AND UPON ARRIVAL HER BLOOD GLUCOSE WAS MGDL SHE WAS GIVEN HAM SANDWICHES AND ORANGE JUICE AFTER HOURS AT THE ER THE END USER WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL SHE WAS INSTRUCTED TO FOLLOWUP WITH HER PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Sweating | Patient problem code
6755563 | 81465196 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HER PRODIGY DIABETES METER THE END USER WAS SWEATING AND NONRESPONSIVE ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL ADDITIONAL BLOOD GLUCOSE TESTS WERE PERFORMED WITH RESULTS RANGING FROM MGDL MGDL THE END USER WAS TAKEN TO THE ER AND UPON ARRIVAL HER BLOOD GLUCOSE WAS MGDL SHE WAS GIVEN HAM SANDWICHES AND ORANGE JUICE AFTER HOURS AT THE ER THE END USER WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL SHE WAS INSTRUCTED TO FOLLOWUP WITH HER PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Hyperglycemia | Patient problem code
6755563 | 81465196 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HER PRODIGY DIABETES METER THE END USER WAS SWEATING AND NONRESPONSIVE ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL ADDITIONAL BLOOD GLUCOSE TESTS WERE PERFORMED WITH RESULTS RANGING FROM MGDL MGDL THE END USER WAS TAKEN TO THE ER AND UPON ARRIVAL HER BLOOD GLUCOSE WAS MGDL SHE WAS GIVEN HAM SANDWICHES AND ORANGE JUICE AFTER HOURS AT THE ER THE END USER WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL SHE WAS INSTRUCTED TO FOLLOWUP WITH HER PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Insufficient Information | Device problem code
6750745 | 81472728 | D | A US DISTRIBUTOR CONTACTED ZOLL AND REPORTED THAT THE PATIENT PASSED AWAY WHILE WEARING THE LIFEVEST DEVICE DOWNLOAD DATA REVEALED THAT THE PATIENT WAS TREATED ONE TIME DURING THE EVENT AT THE PATIENT WAS TREATED INAPPROPRIATELY THE PATIENT WAS IN SUPRAVENTRICULAR TACHYCARDIA SVT AT THE TIME OF THE TREATMENT SVT RATE ABOVE THE PROGRAMMED TREATMENT THRESHOLD CONTRIBUTED TO THE FALSE DETECTION IT WAS REPORTED THAT WHILE IN THE HOSPITAL THE LIFEVEST WAS REMOVED AND THAT THE PATIENT CODED THREE TIMES PRIOR TO PASSING THE PATIENT PASSED AWAY ON B WHILE IN THE HOSPITAL | High Sensing Threshold | Device problem code
6750745 | 81472728 | D | A US DISTRIBUTOR CONTACTED ZOLL AND REPORTED THAT THE PATIENT PASSED AWAY WHILE WEARING THE LIFEVEST DEVICE DOWNLOAD DATA REVEALED THAT THE PATIENT WAS TREATED ONE TIME DURING THE EVENT AT THE PATIENT WAS TREATED INAPPROPRIATELY THE PATIENT WAS IN SUPRAVENTRICULAR TACHYCARDIA SVT AT THE TIME OF THE TREATMENT SVT RATE ABOVE THE PROGRAMMED TREATMENT THRESHOLD CONTRIBUTED TO THE FALSE DETECTION IT WAS REPORTED THAT WHILE IN THE HOSPITAL THE LIFEVEST WAS REMOVED AND THAT THE PATIENT CODED THREE TIMES PRIOR TO PASSING THE PATIENT PASSED AWAY ON B WHILE IN THE HOSPITAL | Death | Patient problem code
6751216 | 81474353 | N | DATE OF EVENT UNKNOWN THE DATE RECEIVED BY MANUFACTURER HAS BEEN USED FOR THIS FIELD RESULTS THE COMPLAINTS LAB RECEIVED ONE SAMPLE THE SAMPLE WAS VISUALLY INSPECTED NO FLAKES WERE OBSERVED ON THE BARREL OF THE SYRINGE NO OTHER ISSUES OR ABNORMALITIES WERE OBSERVED WHILE THIS COMPLAINT IS UNCONFIRMED FROM THE VERBATIM THE LIKELY ISSUE IS BURNT PLASTIC CONCLUSION BD WAS NOT ABLE TO CONFIRM THE CUSTOMERS INDICATED FAILURE COMPLAINTS RECEIVED FOR THIS DEVICE AND REPORTED CONDITION WILL CONTINUE TO BE TRACKED AND TRENDED INFORMATION WILL BE CAPTURED ON TREND REPORTS AND MONITORED MONTHLY B | Failure to Calibrate | Medical Evaluation result code
6752036 | 81474795 | D | THE HOSPITAL REPORTED THAT DURING AN ENDOSCOPIC VEIN HARVESTING PROCEDURE A METAL RIVET AT THE END OF THE BISECTOR OF THE VASOVIEW HEMOPRO WAS MAKING IT HARD TO MOVE BACK AND FORTH ALSO THE PLASTIC COATING NORMALLY COVERING THE RIVET RUBBED OFF AND THERE WAS A PLASTIC BIT HANGING FROM THE DEVICE THE PIECE DID NOT FALL IN THE LEG THE HOSPITAL DID NOT REPORT ANY PATIENT EFFECTS | No Consequences Or Impact To Patient | Patient problem code
6752036 | 81474795 | D | THE HOSPITAL REPORTED THAT DURING AN ENDOSCOPIC VEIN HARVESTING PROCEDURE A METAL RIVET AT THE END OF THE BISECTOR OF THE VASOVIEW HEMOPRO WAS MAKING IT HARD TO MOVE BACK AND FORTH ALSO THE PLASTIC COATING NORMALLY COVERING THE RIVET RUBBED OFF AND THERE WAS A PLASTIC BIT HANGING FROM THE DEVICE THE PIECE DID NOT FALL IN THE LEG THE HOSPITAL DID NOT REPORT ANY PATIENT EFFECTS | Physical Resistance / Sticking | Device problem code
6754958 | 81484112 | N | B ANY ADDITIONAL INFORMATION PROVIDED BY THE CUSTOMER WILL BE INCLUDED IN A FOLLOW UP REPORT B RESULTS OF INVESTIGATION CAREFUSION FIELD SERVICE REPRESENTATIVE WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE THE SERVICE TECHNICIAN INSTALLED A NEW DIAPHRAGM ADJUSTED THE FLOW SENSOR AND TIGHTENED THE TILT STAND THE CUSTOMERS REPORTED ISSUE WAS RESOLVED AND THE UNIT PASSED ALL TESTING | No Device Problem Found | Medical Evaluation result code
6754719 | 81499491 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR OF THE INTRAAORTIC BALLOON PUMP IABP INVOLVED IN THE EVENT WAS REVIEWED THERE WERE NO NONCONFORMANCES IN THE PRODUCTION DHR RELATED TO THE REPORTED EVENT THE FULL EVENT SITE NAME LISTED IS B THE IABP WAS EVALUATED BY THE FACILITYS BIOMEDICAL ENGINEER A GETINGE SERVICE TERRITORY MANAGER STM CONTACTED THE CUSTOMER BY PHONE AND WAS ABLE TO HELP DETERMINE THAT THIS ISSUE WAS CAUSED BY A LEAKY TEST BALLOON THIS IABP DEVICE WAS SUCCESSFUL WHEN TESTED BY THE CUSTOMER WITH A FUNCTIONAL TEST BALLOON AND THEN RETURNED TO CLINICAL SERVICE | Component or Accessory Incompatibility | Device Problem Code
6754719 | 81499491 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR OF THE INTRAAORTIC BALLOON PUMP IABP INVOLVED IN THE EVENT WAS REVIEWED THERE WERE NO NONCONFORMANCES IN THE PRODUCTION DHR RELATED TO THE REPORTED EVENT THE FULL EVENT SITE NAME LISTED IS B THE IABP WAS EVALUATED BY THE FACILITYS BIOMEDICAL ENGINEER A GETINGE SERVICE TERRITORY MANAGER STM CONTACTED THE CUSTOMER BY PHONE AND WAS ABLE TO HELP DETERMINE THAT THIS ISSUE WAS CAUSED BY A LEAKY TEST BALLOON THIS IABP DEVICE WAS SUCCESSFUL WHEN TESTED BY THE CUSTOMER WITH A FUNCTIONAL TEST BALLOON AND THEN RETURNED TO CLINICAL SERVICE | Interoperability Problem Identified | Medical Evaluation result code
6756018 | 81530023 | D | THE CUSTOMERS NEXT TO KIN REPORTED THAT THE INSULIN PUMP HAD AN ERROR ALARM THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED ON EXTERNAL TEMPERATURE THE CUSTOMER WAS ADVISED TO RESTART AN INSULIN PUMP BY REMOVING BATTERY FOR MINUTES AND TO WAIT UNTIL PUMP ALARM THE CUSTOMER WAS ADVISED TO CALL BACK IF PROBLEM CONTINUES TO OCCUR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6756018 | 81530023 | D | THE CUSTOMERS NEXT TO KIN REPORTED THAT THE INSULIN PUMP HAD AN ERROR ALARM THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE INSULIN PUMP WAS NOT EXPOSED ON EXTERNAL TEMPERATURE THE CUSTOMER WAS ADVISED TO RESTART AN INSULIN PUMP BY REMOVING BATTERY FOR MINUTES AND TO WAIT UNTIL PUMP ALARM THE CUSTOMER WAS ADVISED TO CALL BACK IF PROBLEM CONTINUES TO OCCUR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Pumping Problem | Device problem code
6754781 | 81564953 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Component or Accessory Incompatibility | Device problem code
6754781 | 81564953 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Interoperability Problem Identified | Medical Evaluation result code
6754883 | 81572187 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER THE PUMP WAS UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING MINOR SCRATCHES ON LCD WINDOW AND PILLOWING KEYPAD OVERLAY | Scratched Material | Device problem code
6754883 | 81572187 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER THE PUMP WAS UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING MINOR SCRATCHES ON LCD WINDOW AND PILLOWING KEYPAD OVERLAY | Component Missing | Device problem code
6758777 | 81578684 | D | IT WAS REPORTED THAT THE DEVICE ARCTIC SUN SN B EXPERIENCED AN ERRATIC TEMPERATURE ALARM WHILE THE TEMPERATURE SENSING CATHETER WAS IN PLACE THERE WAS ALSO AN ESOPHAGEAL PROBE IN PLACE AS A SECOND TEMPERATURE SOURCE HOWEVER THE TEMPERATURE READINGS WERE NOT CORRELATING A CALL WAS PLACED TO THE MSS TROUBLESHOOTING LINE BY THE ICU NURSE DURING THE CALL THE PATIENTS TEMPERATURE WENT FROM C TO C THE CURRENT PATIENT TEMPERATURE WAS C THE WATER TEMPERATURE WAS C AND THE TREND WAS INCREASING THERE WAS GOOD PAD COVERAGE WITH OPTIMAL FLOW THERE HAD NOT BEEN ANY PROBLEMS COOLING THE PATIENT AND THE NURSE STATED THAT THE PROBLEM COULD BE THE PATIENTS CONDITION AND NOT THE DEVICE THE NURSE STATED THAT HE WOULD KEEP AN EYE ON THE PATIENT AND CALL BACK IF THE DEVICE RECEIVED ANY ADDITIONAL ALARMS AS THE PATIENT STABILIZED DURING THE CALL PER THE NURSE THE PATIENT REACHED THE REWARM TARGET THE PATIENT WAS IN CARDIAC ARREST WITH AN ANOXIC BRAIN INJURY DURING THE FOLLOW UP CALLS ON THE NURSE STATED THAT HE BELIEVES IT WAS THE TEMPERATURE SENSING CATHETER DEFAULTING AS THE PROBLEM WAS RESOLVED AFTER THE CATHETER WAS REPLACED WITH AN ESOPHAGEAL PROBE IT WAS ALSO REPORTED THAT THE PATIENT LATER EXPIRED DUE TO COMPLICATIONS OF THE ANOXIC BRAIN INJURY THE ICU NURSE STATED THAT THE PATIENT S DEATH WAS DUE TO THEIR CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Brain Injury | Patient problem code
6758777 | 81578684 | D | IT WAS REPORTED THAT THE DEVICE ARCTIC SUN SN B EXPERIENCED AN ERRATIC TEMPERATURE ALARM WHILE THE TEMPERATURE SENSING CATHETER WAS IN PLACE THERE WAS ALSO AN ESOPHAGEAL PROBE IN PLACE AS A SECOND TEMPERATURE SOURCE HOWEVER THE TEMPERATURE READINGS WERE NOT CORRELATING A CALL WAS PLACED TO THE MSS TROUBLESHOOTING LINE BY THE ICU NURSE DURING THE CALL THE PATIENTS TEMPERATURE WENT FROM C TO C THE CURRENT PATIENT TEMPERATURE WAS C THE WATER TEMPERATURE WAS C AND THE TREND WAS INCREASING THERE WAS GOOD PAD COVERAGE WITH OPTIMAL FLOW THERE HAD NOT BEEN ANY PROBLEMS COOLING THE PATIENT AND THE NURSE STATED THAT THE PROBLEM COULD BE THE PATIENTS CONDITION AND NOT THE DEVICE THE NURSE STATED THAT HE WOULD KEEP AN EYE ON THE PATIENT AND CALL BACK IF THE DEVICE RECEIVED ANY ADDITIONAL ALARMS AS THE PATIENT STABILIZED DURING THE CALL PER THE NURSE THE PATIENT REACHED THE REWARM TARGET THE PATIENT WAS IN CARDIAC ARREST WITH AN ANOXIC BRAIN INJURY DURING THE FOLLOW UP CALLS ON THE NURSE STATED THAT HE BELIEVES IT WAS THE TEMPERATURE SENSING CATHETER DEFAULTING AS THE PROBLEM WAS RESOLVED AFTER THE CATHETER WAS REPLACED WITH AN ESOPHAGEAL PROBE IT WAS ALSO REPORTED THAT THE PATIENT LATER EXPIRED DUE TO COMPLICATIONS OF THE ANOXIC BRAIN INJURY THE ICU NURSE STATED THAT THE PATIENT S DEATH WAS DUE TO THEIR CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Death | Patient problem code
6758777 | 81578684 | D | IT WAS REPORTED THAT THE DEVICE ARCTIC SUN SN B EXPERIENCED AN ERRATIC TEMPERATURE ALARM WHILE THE TEMPERATURE SENSING CATHETER WAS IN PLACE THERE WAS ALSO AN ESOPHAGEAL PROBE IN PLACE AS A SECOND TEMPERATURE SOURCE HOWEVER THE TEMPERATURE READINGS WERE NOT CORRELATING A CALL WAS PLACED TO THE MSS TROUBLESHOOTING LINE BY THE ICU NURSE DURING THE CALL THE PATIENTS TEMPERATURE WENT FROM C TO C THE CURRENT PATIENT TEMPERATURE WAS C THE WATER TEMPERATURE WAS C AND THE TREND WAS INCREASING THERE WAS GOOD PAD COVERAGE WITH OPTIMAL FLOW THERE HAD NOT BEEN ANY PROBLEMS COOLING THE PATIENT AND THE NURSE STATED THAT THE PROBLEM COULD BE THE PATIENTS CONDITION AND NOT THE DEVICE THE NURSE STATED THAT HE WOULD KEEP AN EYE ON THE PATIENT AND CALL BACK IF THE DEVICE RECEIVED ANY ADDITIONAL ALARMS AS THE PATIENT STABILIZED DURING THE CALL PER THE NURSE THE PATIENT REACHED THE REWARM TARGET THE PATIENT WAS IN CARDIAC ARREST WITH AN ANOXIC BRAIN INJURY DURING THE FOLLOW UP CALLS ON THE NURSE STATED THAT HE BELIEVES IT WAS THE TEMPERATURE SENSING CATHETER DEFAULTING AS THE PROBLEM WAS RESOLVED AFTER THE CATHETER WAS REPLACED WITH AN ESOPHAGEAL PROBE IT WAS ALSO REPORTED THAT THE PATIENT LATER EXPIRED DUE TO COMPLICATIONS OF THE ANOXIC BRAIN INJURY THE ICU NURSE STATED THAT THE PATIENT S DEATH WAS DUE TO THEIR CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6759801 | 81592361 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HIS PRODIGY DIABETES METER THE END USER PERFORMED SEVERAL BLOOD GLUCOSE TEST WITH RESULTS IN THE S HE DID NOT EXPERIENCE ANY SIGNIFICANT SYMPTOMS BUT OUT OF CONCERN HE WENT TO THE ER UPON ARRIVAL TO THE ER AT PM HIS BLOOD GLUCOSE WAS MGDL NO TREATMENT WAS WARRANTED DUE TO THE FACT THAT HIS BLOOD GLUCOSE WAS WITHIN NORMAL RANGE AFTER ONE HOUR IN THE ER HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO PURCHASE A NEW METER NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | No Consequences Or Impact To Patient | Patient problem code
6759801 | 81592361 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HIS PRODIGY DIABETES METER THE END USER PERFORMED SEVERAL BLOOD GLUCOSE TEST WITH RESULTS IN THE S HE DID NOT EXPERIENCE ANY SIGNIFICANT SYMPTOMS BUT OUT OF CONCERN HE WENT TO THE ER UPON ARRIVAL TO THE ER AT PM HIS BLOOD GLUCOSE WAS MGDL NO TREATMENT WAS WARRANTED DUE TO THE FACT THAT HIS BLOOD GLUCOSE WAS WITHIN NORMAL RANGE AFTER ONE HOUR IN THE ER HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO PURCHASE A NEW METER NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Insufficient Information | Device problem code
6760478 | 81611777 | D | IT WAS REPORTED AT TWO DIFFERENT COLLECTION SITES USING DIFFERENT NEEDLE SETS REUSABLE NEEDLE HOLDER AND SINGLE USE TWO PATIENTS DEVELOPED INFECTIONS AFTER VENIPUNCTURE PROCEDURES WITH A G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER PATIENTS WERE PLACED ON ANTIBIOTICS THE SPECIFIC DETAILS FOR EACH PATIENTINCIDENT WERE NOT PROVIDED | Insufficient Information | Device problem code
6760478 | 81611777 | D | IT WAS REPORTED AT TWO DIFFERENT COLLECTION SITES USING DIFFERENT NEEDLE SETS REUSABLE NEEDLE HOLDER AND SINGLE USE TWO PATIENTS DEVELOPED INFECTIONS AFTER VENIPUNCTURE PROCEDURES WITH A G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER PATIENTS WERE PLACED ON ANTIBIOTICS THE SPECIFIC DETAILS FOR EACH PATIENTINCIDENT WERE NOT PROVIDED | Infection | Patient problem code
6760460 | 81619752 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HIS PRODIGY DIABETES METER THE END USER COULD NOT RECALL THE ACTUAL DATE OF THE MEDICAL INTERVENTION THE END USER FELT WEAK COLD WAS SHAKING AND THEN FELT HOT ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL HE WAS TAKEN TO THE ER AND UPON ARRIVAL HIS BLOOD GLUCOSE WAS MGDL WHILE AT THE ER HE RECEIVED CRACKERS AND ORANGE JUICE AFTER HOURS HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO FOLLOWUP WITH HIS PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Chills | Patient problem code
6760460 | 81619752 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HIS PRODIGY DIABETES METER THE END USER COULD NOT RECALL THE ACTUAL DATE OF THE MEDICAL INTERVENTION THE END USER FELT WEAK COLD WAS SHAKING AND THEN FELT HOT ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL HE WAS TAKEN TO THE ER AND UPON ARRIVAL HIS BLOOD GLUCOSE WAS MGDL WHILE AT THE ER HE RECEIVED CRACKERS AND ORANGE JUICE AFTER HOURS HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO FOLLOWUP WITH HIS PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6760460 | 81619752 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT AFTER THE END USER RECEIVED HIGHER THAN NORMAL BLOOD GLUCOSE RESULTS FROM HIS PRODIGY DIABETES METER THE END USER COULD NOT RECALL THE ACTUAL DATE OF THE MEDICAL INTERVENTION THE END USER FELT WEAK COLD WAS SHAKING AND THEN FELT HOT ACCOMPANIED WITH A BLOOD GLUCOSE READING OF MGDL HE WAS TAKEN TO THE ER AND UPON ARRIVAL HIS BLOOD GLUCOSE WAS MGDL WHILE AT THE ER HE RECEIVED CRACKERS AND ORANGE JUICE AFTER HOURS HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO FOLLOWUP WITH HIS PCP NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Weakness | Patient problem code
6761088 | 81627292 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEIR BLOOD GLUCOSE WAS MGDL THE CUSTOMER TREATED WITH FOOD AND SUGAR WATER DURING TROUBLESHOOTING THE CUSTOMER CONFIRMED THAT THEIR INSULIN PUMP PROGRAMMING WAS ACCURATE THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THERE WAS NO MRI EXPOSURE EITHER NO APPARENT ISSUES WERE NOTED ON THE DEVICE THE INSULIN PUMP WAS NOT RETURNED FOR ANALYSIS | Adverse Event Without Identified Device or Use Problem | Device problem code
6761088 | 81627292 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEIR BLOOD GLUCOSE WAS MGDL THE CUSTOMER TREATED WITH FOOD AND SUGAR WATER DURING TROUBLESHOOTING THE CUSTOMER CONFIRMED THAT THEIR INSULIN PUMP PROGRAMMING WAS ACCURATE THE RESERVOIR WAS SHOWING THE SAME AMOUNT OF INSULIN AS SHOWN ON THE STATUS SCREEN THERE WAS NO MRI EXPOSURE EITHER NO APPARENT ISSUES WERE NOTED ON THE DEVICE THE INSULIN PUMP WAS NOT RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6761220 | 81630702 | N | RESULTS CODE THE LOT NUMBER HAS BEEN REQUESTED BUT REMAINS UNKNOWN THEREFORE A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED CONCLUSION CODE ACCORDING TO THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE IFU ADVERSE EVENTS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ANEURYSM ENLARGEMENT AND ENDOLEAK | No Findings Available | Medical Evaluation result code
6761220 | 81630703 | D | THE FOLLOWING INFORMATION WAS REPORTED TO GORE ON B A PATIENT WAS TREATED FOR AN UNKNOWN PATHOLOGY WITH A GORE EXCLUDER AAA ENDOPROSTHESIS ON B FOLLOWUP IMAGING TAKEN FOR COMPARISON DUE TO SAC GROWTH SHOWED AN ENDOLEAK OF UNKNOWN ORIGIN | Insufficient Information | Device problem code
6761220 | 81630703 | D | THE FOLLOWING INFORMATION WAS REPORTED TO GORE ON B A PATIENT WAS TREATED FOR AN UNKNOWN PATHOLOGY WITH A GORE EXCLUDER AAA ENDOPROSTHESIS ON B FOLLOWUP IMAGING TAKEN FOR COMPARISON DUE TO SAC GROWTH SHOWED AN ENDOLEAK OF UNKNOWN ORIGIN | Hemorrhage | Patient problem code
6762080 | 81666988 | D | ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC DISSECTION USING A GORE TAG THORACIC ENDOPROSTHESIS TG PREIMPLANT MEASUREMENT OF MAXIMUM AORTIC DIAMETER WAS MM INTRAPROCEDURE IMAGING REVEALED A TYPE II ENDOLEAK AND THE PROCEDURE WAS CONCLUDED WITH A WAITANDWATCH APPROACH TAKEN TO THE ENDOLEAK ON B A FOURYEAR FOLLOWUP STUDY REPORTEDLY REVEALED DISTAL TYPE I AND PERSISTENT TYPE II ENDOLEAKS THE MAXIMUM AORTIC DIAMETER HAD ENLARGED TO MM ON B THE PATIENT WAS IMPLANTED WITH AN ADDITIONAL GORE TAG THORACIC ENDOPROSTHESIS TO TREAT THE ENDOLEAKS AND FALSE LUMEN EXPANSION ON B A FIVEYEAR FOLLOWUP STUDY REPORTEDLY SHOWED THAT THE ENDOLEAKS WERE RESOLVED THE MAXIMUM AORTIC DIAMETER MEASURED MM | Hemorrhage | Patient problem code
6762080 | 81666988 | D | ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC DISSECTION USING A GORE TAG THORACIC ENDOPROSTHESIS TG PREIMPLANT MEASUREMENT OF MAXIMUM AORTIC DIAMETER WAS MM INTRAPROCEDURE IMAGING REVEALED A TYPE II ENDOLEAK AND THE PROCEDURE WAS CONCLUDED WITH A WAITANDWATCH APPROACH TAKEN TO THE ENDOLEAK ON B A FOURYEAR FOLLOWUP STUDY REPORTEDLY REVEALED DISTAL TYPE I AND PERSISTENT TYPE II ENDOLEAKS THE MAXIMUM AORTIC DIAMETER HAD ENLARGED TO MM ON B THE PATIENT WAS IMPLANTED WITH AN ADDITIONAL GORE TAG THORACIC ENDOPROSTHESIS TO TREAT THE ENDOLEAKS AND FALSE LUMEN EXPANSION ON B A FIVEYEAR FOLLOWUP STUDY REPORTEDLY SHOWED THAT THE ENDOLEAKS WERE RESOLVED THE MAXIMUM AORTIC DIAMETER MEASURED MM | Insufficient Information | Device problem code
6762706 | 81668919 | D | IN A LITERATURE REVIEW THE FOLLOWING INFORMATION WAS OBTAINED E ITO ET AL ENDOVASCULAR REPAIR OF ABDOMINAL AORTIC ANEURYSM WITH HIGHLY ANGULATED PROXIMAL NECK COMPARISON OF TREATMENT RESULTS BETWEEN EXCLUDER AND AORFIX JOURNAL OF THE JAPANESE SOCIETY FOR VASCULAR SURGERY FROM JULY TO JUNE THE TOTAL OF PATIENTS WITH PROXIMAL NECK ANGLE GREATER THAN DEGREES UNDERWENT ENDOVASCULAR REPAIR OF AN ABDOMINAL AORTIC ANEURYSM USING GORE EXCLUDER AAA ENDOPROSTHESES AMONG THOSE PATIENTS THE AVERAGE PROXIMAL NECK ANGLE WAS DEGREES AND THE AVERAGE PROXIMAL NECK LENGTH WAS MM IN ONE PATIENT PROXIMAL TYPE I ENDOLEAK WITH ANEURYSM ENLARGEMENT AMOUNT OF ENLARGEMENT IS UNKNOWN WAS REVEALED POST INITIAL IMPLANT PROCEDURE REINTERVENTION WAS PERFORMED AND THIS PATIENT WAS IMPLANTED WITH AN ADDITIONAL ENDOPROSTHESIS TO REPAIR THE ENDOLEAK | Aneurysm | Patient problem code
6762706 | 81668919 | D | IN A LITERATURE REVIEW THE FOLLOWING INFORMATION WAS OBTAINED E ITO ET AL ENDOVASCULAR REPAIR OF ABDOMINAL AORTIC ANEURYSM WITH HIGHLY ANGULATED PROXIMAL NECK COMPARISON OF TREATMENT RESULTS BETWEEN EXCLUDER AND AORFIX JOURNAL OF THE JAPANESE SOCIETY FOR VASCULAR SURGERY FROM JULY TO JUNE THE TOTAL OF PATIENTS WITH PROXIMAL NECK ANGLE GREATER THAN DEGREES UNDERWENT ENDOVASCULAR REPAIR OF AN ABDOMINAL AORTIC ANEURYSM USING GORE EXCLUDER AAA ENDOPROSTHESES AMONG THOSE PATIENTS THE AVERAGE PROXIMAL NECK ANGLE WAS DEGREES AND THE AVERAGE PROXIMAL NECK LENGTH WAS MM IN ONE PATIENT PROXIMAL TYPE I ENDOLEAK WITH ANEURYSM ENLARGEMENT AMOUNT OF ENLARGEMENT IS UNKNOWN WAS REVEALED POST INITIAL IMPLANT PROCEDURE REINTERVENTION WAS PERFORMED AND THIS PATIENT WAS IMPLANTED WITH AN ADDITIONAL ENDOPROSTHESIS TO REPAIR THE ENDOLEAK | Hemorrhage | Patient problem code
6762706 | 81668919 | D | IN A LITERATURE REVIEW THE FOLLOWING INFORMATION WAS OBTAINED E ITO ET AL ENDOVASCULAR REPAIR OF ABDOMINAL AORTIC ANEURYSM WITH HIGHLY ANGULATED PROXIMAL NECK COMPARISON OF TREATMENT RESULTS BETWEEN EXCLUDER AND AORFIX JOURNAL OF THE JAPANESE SOCIETY FOR VASCULAR SURGERY FROM JULY TO JUNE THE TOTAL OF PATIENTS WITH PROXIMAL NECK ANGLE GREATER THAN DEGREES UNDERWENT ENDOVASCULAR REPAIR OF AN ABDOMINAL AORTIC ANEURYSM USING GORE EXCLUDER AAA ENDOPROSTHESES AMONG THOSE PATIENTS THE AVERAGE PROXIMAL NECK ANGLE WAS DEGREES AND THE AVERAGE PROXIMAL NECK LENGTH WAS MM IN ONE PATIENT PROXIMAL TYPE I ENDOLEAK WITH ANEURYSM ENLARGEMENT AMOUNT OF ENLARGEMENT IS UNKNOWN WAS REVEALED POST INITIAL IMPLANT PROCEDURE REINTERVENTION WAS PERFORMED AND THIS PATIENT WAS IMPLANTED WITH AN ADDITIONAL ENDOPROSTHESIS TO REPAIR THE ENDOLEAK | Insufficient Information | Device problem code
6757883 | 81677344 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER DEVICE WAS RECEIVED WITH SCRATCHED CASING MINOR SCRATCHES ON LCD WINDOW SCRATCHES ON DISPLAY WINDOW COVER AND PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
6757883 | 81677344 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER DEVICE WAS RECEIVED WITH SCRATCHED CASING MINOR SCRATCHES ON LCD WINDOW SCRATCHES ON DISPLAY WINDOW COVER AND PILLOWING KEYPAD OVERLAY B | Component Missing | Device problem code
6757883 | 81677344 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER DEVICE WAS RECEIVED WITH SCRATCHED CASING MINOR SCRATCHES ON LCD WINDOW SCRATCHES ON DISPLAY WINDOW COVER AND PILLOWING KEYPAD OVERLAY B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6757844 | 81684700 | N | RESERVOIR WAS RETURNED INCOMPLETE ONLY TRANSFER GUARD RETURNED MISSING PLUNGER BARREL AND STOPPER PLUNGER PERFORMED VISUAL INSPECTION AND FOUND NEEDLE DETACHED FROM TRANSFER GUARD LOOSE NEEDLE B | Component Missing | Device problem code
6757844 | 81684700 | N | RESERVOIR WAS RETURNED INCOMPLETE ONLY TRANSFER GUARD RETURNED MISSING PLUNGER BARREL AND STOPPER PLUNGER PERFORMED VISUAL INSPECTION AND FOUND NEEDLE DETACHED FROM TRANSFER GUARD LOOSE NEEDLE B | Detachment of Device or device Component | Device problem code
6757844 | 81684700 | N | RESERVOIR WAS RETURNED INCOMPLETE ONLY TRANSFER GUARD RETURNED MISSING PLUNGER BARREL AND STOPPER PLUNGER PERFORMED VISUAL INSPECTION AND FOUND NEEDLE DETACHED FROM TRANSFER GUARD LOOSE NEEDLE B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6763639 | 81714172 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM A RIGHT EXTERNAL ILIAC ANEURYSM AND A LEFT INTERNAL ILIAC ANEURYSM WITH GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE FOLLOWUP IMAGING REPORTEDLY REVEALED A DISTAL TYPE I ENDOLEAK FROM THE CONTRALATERAL LIMB ON THE LEFT SIDE PXCJ AND ENLARGEMENT OF THE LEFT INTERNAL ILIAC ANEURYSM SIZE UNKNOWN ON B THE PATIENT UNDERWENT A REINTERVENTION PROCEDURE WHEREBY THE LEFT INTERNAL ILIAC ANEURYSM WAS COILEMBOLIZED AND TWO ADDITIONAL EXCLUDER COMPONENTS WERE IMPLANTED DISTALLY ON THE LEFT SIDE TO EXTEND THE CONTRALATERAL LIMB AND TREAT THE DISTAL TYPE I ENDOLEAK THE ENDOLEAK WAS RESOLVED AND THE PATIENT TOLERATED THE PROCEDURE | Insufficient Information | Device problem code
6763639 | 81714172 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM A RIGHT EXTERNAL ILIAC ANEURYSM AND A LEFT INTERNAL ILIAC ANEURYSM WITH GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE FOLLOWUP IMAGING REPORTEDLY REVEALED A DISTAL TYPE I ENDOLEAK FROM THE CONTRALATERAL LIMB ON THE LEFT SIDE PXCJ AND ENLARGEMENT OF THE LEFT INTERNAL ILIAC ANEURYSM SIZE UNKNOWN ON B THE PATIENT UNDERWENT A REINTERVENTION PROCEDURE WHEREBY THE LEFT INTERNAL ILIAC ANEURYSM WAS COILEMBOLIZED AND TWO ADDITIONAL EXCLUDER COMPONENTS WERE IMPLANTED DISTALLY ON THE LEFT SIDE TO EXTEND THE CONTRALATERAL LIMB AND TREAT THE DISTAL TYPE I ENDOLEAK THE ENDOLEAK WAS RESOLVED AND THE PATIENT TOLERATED THE PROCEDURE | Hemorrhage | Patient problem code
6764818 | 81744254 | D | COMPLAINANT ALLEGED THAT THE DEVICE WAS UNABLE TO RECOGNIZE AN INSTALLED BATTERY COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6764818 | 81744254 | D | COMPLAINANT ALLEGED THAT THE DEVICE WAS UNABLE TO RECOGNIZE AN INSTALLED BATTERY COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Battery Problem | Device Problem Code
6765109 | 81752348 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED AN IRRITATION WHILE WEARING THE LIFEVEST FOLLOW UP INDICATED THAT THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO END USE OF THE LIFEVEST AND THE IRRITATION RESOLVED | Irritation | Patient problem code
6765109 | 81752348 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED AN IRRITATION WHILE WEARING THE LIFEVEST FOLLOW UP INDICATED THAT THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO END USE OF THE LIFEVEST AND THE IRRITATION RESOLVED | Insufficient Information | Device problem code
6764961 | 81774811 | D | IT WAS REPORTED THAT THE PATIENT DEVELOPED A STAGE PRESSURE SORE AT THE SITE OF A PORTEX BLUE LINE ULTRA TRACHEOSTOMY TUBE IT WAS NOTED THAT THE FLANGE ON THE DEVICE DOES NOT CONTAIN ANY GEL LIKE OR OTHER SOFT BARRIER TO HELP PREVENT PRESSURE INJURIES THE TRACHEOSTOMY TUBE REQUIRES USE OF OTHER MATERIALS SUCH AS GAUZE OR GEL PADS TO DECREASEELIMINATE PRESSURE NO PERMANENT INJURY WAS REPORTED | Injury | Patient problem code
6764961 | 81774811 | D | IT WAS REPORTED THAT THE PATIENT DEVELOPED A STAGE PRESSURE SORE AT THE SITE OF A PORTEX BLUE LINE ULTRA TRACHEOSTOMY TUBE IT WAS NOTED THAT THE FLANGE ON THE DEVICE DOES NOT CONTAIN ANY GEL LIKE OR OTHER SOFT BARRIER TO HELP PREVENT PRESSURE INJURIES THE TRACHEOSTOMY TUBE REQUIRES USE OF OTHER MATERIALS SUCH AS GAUZE OR GEL PADS TO DECREASEELIMINATE PRESSURE NO PERMANENT INJURY WAS REPORTED | Pain | Patient problem code
6764961 | 81774811 | D | IT WAS REPORTED THAT THE PATIENT DEVELOPED A STAGE PRESSURE SORE AT THE SITE OF A PORTEX BLUE LINE ULTRA TRACHEOSTOMY TUBE IT WAS NOTED THAT THE FLANGE ON THE DEVICE DOES NOT CONTAIN ANY GEL LIKE OR OTHER SOFT BARRIER TO HELP PREVENT PRESSURE INJURIES THE TRACHEOSTOMY TUBE REQUIRES USE OF OTHER MATERIALS SUCH AS GAUZE OR GEL PADS TO DECREASEELIMINATE PRESSURE NO PERMANENT INJURY WAS REPORTED | Adverse Event Without Identified Device or Use Problem | Device problem code
6764961 | 81774811 | D | IT WAS REPORTED THAT THE PATIENT DEVELOPED A STAGE PRESSURE SORE AT THE SITE OF A PORTEX BLUE LINE ULTRA TRACHEOSTOMY TUBE IT WAS NOTED THAT THE FLANGE ON THE DEVICE DOES NOT CONTAIN ANY GEL LIKE OR OTHER SOFT BARRIER TO HELP PREVENT PRESSURE INJURIES THE TRACHEOSTOMY TUBE REQUIRES USE OF OTHER MATERIALS SUCH AS GAUZE OR GEL PADS TO DECREASEELIMINATE PRESSURE NO PERMANENT INJURY WAS REPORTED | Pressure Sores/Ulcers | Patient problem code
6764853 | 81775723 | D | ON B THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM MEASURING MM IN DIAMETER AND WAS IMPLANTED WITH GORE EXCLUDER ENDOPROSTHESES IT WAS REPORTED THAT ON B A TYPE II ENDOLEAK PORTION OF THE ANEURYSMAL SAC THAT WAS NOT THROMBOSED NOT A TRUE ENDOLEAK WAS IDENTIFIED FROM B TILL B THE ANEURYSM ENLARGED FROM MM TO MM FOLLOW UP CTA ON B DETERMINED THAT THE MAXIMUM DIAMETER OF AORTIC ANEURYSMLESION WAS MM ON B THE PATIENT UNDERWENT COIL EMBOLIZATION PROCEDURE OF THE ABDOMINAL AORTIC ANEURYSM | Insufficient Information | Device problem code
6764853 | 81775723 | D | ON B THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM MEASURING MM IN DIAMETER AND WAS IMPLANTED WITH GORE EXCLUDER ENDOPROSTHESES IT WAS REPORTED THAT ON B A TYPE II ENDOLEAK PORTION OF THE ANEURYSMAL SAC THAT WAS NOT THROMBOSED NOT A TRUE ENDOLEAK WAS IDENTIFIED FROM B TILL B THE ANEURYSM ENLARGED FROM MM TO MM FOLLOW UP CTA ON B DETERMINED THAT THE MAXIMUM DIAMETER OF AORTIC ANEURYSMLESION WAS MM ON B THE PATIENT UNDERWENT COIL EMBOLIZATION PROCEDURE OF THE ABDOMINAL AORTIC ANEURYSM | Hemorrhage | Patient problem code
6765159 | 81776717 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A YEAST INFECTION WHILE WEARING THE LIFEVEST THE PATIENT ALSO REPORTED THAT SHE MAY ALSO HAVE HAD AN ALLERGIC REACTION THE PATIENTS PHYSICIAN PRESCRIBED A TABLET FOR THE INFECTION FOLLOW UP INDICATED THAT THE INFECTION WAS NOT IMPROVING THE PATIENT WAS GOING TO CONTINUE TO FOLLOW WITH THE PHYSICIAN THE FINAL MEDICAL OUTCOME IS UNKNOWN | Hypersensitivity | Patient problem code
6765159 | 81776717 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A YEAST INFECTION WHILE WEARING THE LIFEVEST THE PATIENT ALSO REPORTED THAT SHE MAY ALSO HAVE HAD AN ALLERGIC REACTION THE PATIENTS PHYSICIAN PRESCRIBED A TABLET FOR THE INFECTION FOLLOW UP INDICATED THAT THE INFECTION WAS NOT IMPROVING THE PATIENT WAS GOING TO CONTINUE TO FOLLOW WITH THE PHYSICIAN THE FINAL MEDICAL OUTCOME IS UNKNOWN | Insufficient Information | Device problem code
6765159 | 81776717 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A YEAST INFECTION WHILE WEARING THE LIFEVEST THE PATIENT ALSO REPORTED THAT SHE MAY ALSO HAVE HAD AN ALLERGIC REACTION THE PATIENTS PHYSICIAN PRESCRIBED A TABLET FOR THE INFECTION FOLLOW UP INDICATED THAT THE INFECTION WAS NOT IMPROVING THE PATIENT WAS GOING TO CONTINUE TO FOLLOW WITH THE PHYSICIAN THE FINAL MEDICAL OUTCOME IS UNKNOWN | Infection, Fungal | Patient problem code
6765745 | 81777212 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE IN THE MGDL RANGE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MGDL AT THE TIME OF ADMISSION THE CUSTOMER USED JUICE TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SEIZURES THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED CUSTOMER ALSO CALLED FOR ASSISTANCE UPLOADING THEIR PUMP THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Seizures | Patient problem code
6765745 | 81777212 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE IN THE MGDL RANGE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MGDL AT THE TIME OF ADMISSION THE CUSTOMER USED JUICE TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SEIZURES THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED CUSTOMER ALSO CALLED FOR ASSISTANCE UPLOADING THEIR PUMP THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6765745 | 81777212 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE IN THE MGDL RANGE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MGDL AT THE TIME OF ADMISSION THE CUSTOMER USED JUICE TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SEIZURES THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED CUSTOMER ALSO CALLED FOR ASSISTANCE UPLOADING THEIR PUMP THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6766315 | 81778742 | N | B RESULTS OF INVESTIGATION A CAREFUSION FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR WAS NOT ABLE TO DUPLICATE THE REPORTED ISSUE THE FSR ATTACHED A KNOWN GOOD PATIENT CIRCUIT TO THE VENTILATOR AND PERFORMED THE PATIENT CIRCUIT CALIBRATION AND THE PERFORMANCE CHECK THE UNIT MEETS MANUFACTURERS SPECIFICATIONS THE FSR PERFORMED A K PREVENTATIVE MAINTENANCE PROCEDURE AND CALIBRATED THE DRIVER DISPLACEMENT INDICATOR THE UNIT RAN PROPERLY WITH NO ALARMS NOR DRIVER SHUT DOWN | Device Alarm System | Device problem code
6766315 | 81778742 | N | B RESULTS OF INVESTIGATION A CAREFUSION FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR WAS NOT ABLE TO DUPLICATE THE REPORTED ISSUE THE FSR ATTACHED A KNOWN GOOD PATIENT CIRCUIT TO THE VENTILATOR AND PERFORMED THE PATIENT CIRCUIT CALIBRATION AND THE PERFORMANCE CHECK THE UNIT MEETS MANUFACTURERS SPECIFICATIONS THE FSR PERFORMED A K PREVENTATIVE MAINTENANCE PROCEDURE AND CALIBRATED THE DRIVER DISPLACEMENT INDICATOR THE UNIT RAN PROPERLY WITH NO ALARMS NOR DRIVER SHUT DOWN | Calibration Problem | Device problem code
6766315 | 81778742 | N | B RESULTS OF INVESTIGATION A CAREFUSION FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR WAS NOT ABLE TO DUPLICATE THE REPORTED ISSUE THE FSR ATTACHED A KNOWN GOOD PATIENT CIRCUIT TO THE VENTILATOR AND PERFORMED THE PATIENT CIRCUIT CALIBRATION AND THE PERFORMANCE CHECK THE UNIT MEETS MANUFACTURERS SPECIFICATIONS THE FSR PERFORMED A K PREVENTATIVE MAINTENANCE PROCEDURE AND CALIBRATED THE DRIVER DISPLACEMENT INDICATOR THE UNIT RAN PROPERLY WITH NO ALARMS NOR DRIVER SHUT DOWN | Failure to Calibrate | Medical Evaluation result code
6763568 | 81803592 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6763568 | 81803592 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
6767174 | 81820689 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS TAKEN OFF OF ONE OF HIS MEDICATIONS AND WAS GIVEN A STEROID FOLLOWUP INDICATED THAT THE IRRITATION HAD IMPROVED | Itching | Patient problem code
6767174 | 81820689 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS TAKEN OFF OF ONE OF HIS MEDICATIONS AND WAS GIVEN A STEROID FOLLOWUP INDICATED THAT THE IRRITATION HAD IMPROVED | Rash | Patient problem code
6767174 | 81820689 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS TAKEN OFF OF ONE OF HIS MEDICATIONS AND WAS GIVEN A STEROID FOLLOWUP INDICATED THAT THE IRRITATION HAD IMPROVED | Irritation | Patient problem code
6767174 | 81820689 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS TAKEN OFF OF ONE OF HIS MEDICATIONS AND WAS GIVEN A STEROID FOLLOWUP INDICATED THAT THE IRRITATION HAD IMPROVED | Insufficient Information | Device problem code
6767163 | 81823584 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SKIN IRRITATION ON HIS BACK UNDER THE THERAPY ELECTRODES THE IRRITATION WAS DESCRIBED AS A HEAT RASH THAT WAS OOZING AND BLEEDING THE RASH DEVELOPED INTO A DEEP RED AREA WITH ABCESS THE PATIENTS PHYSICIAN PRESCRIBED AN ANTIBIOTIC CREAM FOR THE IRRITATION AND THE ABCESS WAS LANCED THE PATIENTS MEDICAL OUTCOME IS UNKNOWN | Skin Irritation | Patient problem code
6767163 | 81823584 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SKIN IRRITATION ON HIS BACK UNDER THE THERAPY ELECTRODES THE IRRITATION WAS DESCRIBED AS A HEAT RASH THAT WAS OOZING AND BLEEDING THE RASH DEVELOPED INTO A DEEP RED AREA WITH ABCESS THE PATIENTS PHYSICIAN PRESCRIBED AN ANTIBIOTIC CREAM FOR THE IRRITATION AND THE ABCESS WAS LANCED THE PATIENTS MEDICAL OUTCOME IS UNKNOWN | Rash | Patient problem code
6767163 | 81823584 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SKIN IRRITATION ON HIS BACK UNDER THE THERAPY ELECTRODES THE IRRITATION WAS DESCRIBED AS A HEAT RASH THAT WAS OOZING AND BLEEDING THE RASH DEVELOPED INTO A DEEP RED AREA WITH ABCESS THE PATIENTS PHYSICIAN PRESCRIBED AN ANTIBIOTIC CREAM FOR THE IRRITATION AND THE ABCESS WAS LANCED THE PATIENTS MEDICAL OUTCOME IS UNKNOWN | Insufficient Information | Device problem code
6767163 | 81823584 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SKIN IRRITATION ON HIS BACK UNDER THE THERAPY ELECTRODES THE IRRITATION WAS DESCRIBED AS A HEAT RASH THAT WAS OOZING AND BLEEDING THE RASH DEVELOPED INTO A DEEP RED AREA WITH ABCESS THE PATIENTS PHYSICIAN PRESCRIBED AN ANTIBIOTIC CREAM FOR THE IRRITATION AND THE ABCESS WAS LANCED THE PATIENTS MEDICAL OUTCOME IS UNKNOWN | Hemorrhage | Patient problem code
6767468 | 81831447 | D | THE CUSTOMER REPORTED THAT A PATIENT WHO WAS B PREGNANT UNDERWENT A RED BLOOD CELL EXCHANGE RBCX PROCEDURE DURING THE PROCEDURE THE PATIENT WAS INTUBATED AND REQUIRED A CAESAREAN C SECTION THE CUSTOMER REPORTS THAT THE PATIENT WAS DISCHARGED FROM THE HOSPITAL HOWEVER HAVE NO INFORMATION ON THE NEWBORN BABY PATIENT IDENTIFIER AGE AND WEIGHT ARE NOT AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
6767468 | 81831447 | D | THE CUSTOMER REPORTED THAT A PATIENT WHO WAS B PREGNANT UNDERWENT A RED BLOOD CELL EXCHANGE RBCX PROCEDURE DURING THE PROCEDURE THE PATIENT WAS INTUBATED AND REQUIRED A CAESAREAN C SECTION THE CUSTOMER REPORTS THAT THE PATIENT WAS DISCHARGED FROM THE HOSPITAL HOWEVER HAVE NO INFORMATION ON THE NEWBORN BABY PATIENT IDENTIFIER AGE AND WEIGHT ARE NOT AVAILABLE AT THIS TIME | No Information | Patient problem code
6767701 | 81839302 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENTS PHYSICIAN INSTRUCTED THE PATIENT TO REMOVE THE LIFEVEST PRESCRIBED AN ANTIBIOTIC AND RECOMMENDED THE PATIENT TO TAKE BENADRYL | Adverse Event Without Identified Device or Use Problem | Device problem code
6767701 | 81839302 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENTS PHYSICIAN INSTRUCTED THE PATIENT TO REMOVE THE LIFEVEST PRESCRIBED AN ANTIBIOTIC AND RECOMMENDED THE PATIENT TO TAKE BENADRYL | Hypersensitivity | Patient problem code
6756876 | 81842922 | N | RESULTS MOLDING PROCEDURES DEFINE THE FREQUENCY OF VISUAL INSPECTIONS AND REACTION PLANS TO EMBEDDED FM WHEN IT IS DETECTED IN THE MOLDING PROCESS LIKEWISE THERE IS A VISUAL INSPECTION OF RESIN AT INCOMING INSPECTION FOR EACH RECEIVING CONCLUSION SINCE EMBEDDED FM IS ISOLATED FROM THE SPECIMEN AND FLUID PATH IT IS SOLELY A COSMETIC ISSUE BASED ON THE PROBABILITY OF OCCURRENCE AND THE SEVERITY OF THE ISSUE NO FURTHER INVESTIGATION IS NECESSARY DATE OF EVENT UNKNOWN THE DATE RECEIVED BY MANUFACTURER HAS BEEN USED FOR THIS FIELD B | Entrapment of Device | Device problem code
6756876 | 81842922 | N | RESULTS MOLDING PROCEDURES DEFINE THE FREQUENCY OF VISUAL INSPECTIONS AND REACTION PLANS TO EMBEDDED FM WHEN IT IS DETECTED IN THE MOLDING PROCESS LIKEWISE THERE IS A VISUAL INSPECTION OF RESIN AT INCOMING INSPECTION FOR EACH RECEIVING CONCLUSION SINCE EMBEDDED FM IS ISOLATED FROM THE SPECIMEN AND FLUID PATH IT IS SOLELY A COSMETIC ISSUE BASED ON THE PROBABILITY OF OCCURRENCE AND THE SEVERITY OF THE ISSUE NO FURTHER INVESTIGATION IS NECESSARY DATE OF EVENT UNKNOWN THE DATE RECEIVED BY MANUFACTURER HAS BEEN USED FOR THIS FIELD B | Results Pending Completion of Investigation | Medical Evaluation result code
6767762 | 81843117 | D | ZOLL MANUFACTURING CORPORATION WAS NOTIFIED ON THAT A B PATIENT HAD PASSED AWAY WHILE WEARING THE LIFEVEST THE PATIENT WAS FOUND ALONE AND DEAD IN HIS FLAT AS REPORTED BY THE CRIMINAL INVESTIGATION DEPARTMENT DUREN KK IT WAS REPORTED BY THE POLICE THAT THE PATIENT WAS FOUND LIFELESS IN HIS BED WEARING THE LIFEVEST BUT THE LIFEVEST BATTERY PACK WAS NOT INSERTED INTO THE MONITOR ONE BATTERY WAS FOUND IN THE KITCHEN AND THE OTHER WAS INSIDE THE LIFEVEST CHARGER IT WAS REPORTED THAT THE PATIENT WOULD BE AUTOPSIED A PHYSICIAN DR B POINTED OUT THAT HE THOUGHT THAT THE BATTERY IS POTENTIALLY TOO EASY TO REMOVE DURING AN ARRHYTHMIA OR ALARM DR B WISHED FOR ZOLL TO INFORM B OF THIS CASE DR B ADDITIONALLY SUGGESTED DEVELOPING A SOLUTION TO AVOID SUCH AN EVENT IN THE FUTURE EVALUATION OF THE DEVICE DOWNLOADED FLAG FILES AND THE ASSOCIATED AUTOMATIC ECG RECORDINGS INDICATED THAT THE LIFEVEST WAS LAST STARTED ONB AT ON B AT AN ARRHYTHMIA WAS DETECTED BY THE LIFEVEST THE ECG STRIPS SHOW THAT THE PATIENT WAS IN SINUS TACHYCARDIA AT BEATS PER MINUTE TRANSITIONING TO VENTRICULAR TACHYCARDIA AT BEATS PER MINUTES THE PATIENT WAS IN VENTRICULAR TACHYCARDIA FOR APPROXIMATELY MINUTES AS CAPTURED BY THE ECG RECORDINGS AT AND THE DOWNLOADED FLAG FILES AND ECG STRIPS SHOW THAT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE ARRHYTHMIA DETECTION THE PATIENT WAS NOT TREATED BY THE LIFEVEST AS THE USE OF THE RESPONSE BUTTONS PREVENTED THE TREATMENT FROM BEING DELIVERED DETECTION OF THE ARRHYTHMIA CONTINUED UNTIL THE LIFEVEST WAS SHUTDOWN BY REMOVAL OF THE BATTERY PACK AT ON B SUBSEQUENT MONITORING OF THE PATIENT WAS NOT POSSIBLE AFTER THE DEVICE SHUTDOWN | Tachycardia | Patient problem code
6767762 | 81843117 | D | ZOLL MANUFACTURING CORPORATION WAS NOTIFIED ON THAT A B PATIENT HAD PASSED AWAY WHILE WEARING THE LIFEVEST THE PATIENT WAS FOUND ALONE AND DEAD IN HIS FLAT AS REPORTED BY THE CRIMINAL INVESTIGATION DEPARTMENT DUREN KK IT WAS REPORTED BY THE POLICE THAT THE PATIENT WAS FOUND LIFELESS IN HIS BED WEARING THE LIFEVEST BUT THE LIFEVEST BATTERY PACK WAS NOT INSERTED INTO THE MONITOR ONE BATTERY WAS FOUND IN THE KITCHEN AND THE OTHER WAS INSIDE THE LIFEVEST CHARGER IT WAS REPORTED THAT THE PATIENT WOULD BE AUTOPSIED A PHYSICIAN DR B POINTED OUT THAT HE THOUGHT THAT THE BATTERY IS POTENTIALLY TOO EASY TO REMOVE DURING AN ARRHYTHMIA OR ALARM DR B WISHED FOR ZOLL TO INFORM B OF THIS CASE DR B ADDITIONALLY SUGGESTED DEVELOPING A SOLUTION TO AVOID SUCH AN EVENT IN THE FUTURE EVALUATION OF THE DEVICE DOWNLOADED FLAG FILES AND THE ASSOCIATED AUTOMATIC ECG RECORDINGS INDICATED THAT THE LIFEVEST WAS LAST STARTED ONB AT ON B AT AN ARRHYTHMIA WAS DETECTED BY THE LIFEVEST THE ECG STRIPS SHOW THAT THE PATIENT WAS IN SINUS TACHYCARDIA AT BEATS PER MINUTE TRANSITIONING TO VENTRICULAR TACHYCARDIA AT BEATS PER MINUTES THE PATIENT WAS IN VENTRICULAR TACHYCARDIA FOR APPROXIMATELY MINUTES AS CAPTURED BY THE ECG RECORDINGS AT AND THE DOWNLOADED FLAG FILES AND ECG STRIPS SHOW THAT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE ARRHYTHMIA DETECTION THE PATIENT WAS NOT TREATED BY THE LIFEVEST AS THE USE OF THE RESPONSE BUTTONS PREVENTED THE TREATMENT FROM BEING DELIVERED DETECTION OF THE ARRHYTHMIA CONTINUED UNTIL THE LIFEVEST WAS SHUTDOWN BY REMOVAL OF THE BATTERY PACK AT ON B SUBSEQUENT MONITORING OF THE PATIENT WAS NOT POSSIBLE AFTER THE DEVICE SHUTDOWN | Arrhythmia | Patient problem code
6767762 | 81843117 | D | ZOLL MANUFACTURING CORPORATION WAS NOTIFIED ON THAT A B PATIENT HAD PASSED AWAY WHILE WEARING THE LIFEVEST THE PATIENT WAS FOUND ALONE AND DEAD IN HIS FLAT AS REPORTED BY THE CRIMINAL INVESTIGATION DEPARTMENT DUREN KK IT WAS REPORTED BY THE POLICE THAT THE PATIENT WAS FOUND LIFELESS IN HIS BED WEARING THE LIFEVEST BUT THE LIFEVEST BATTERY PACK WAS NOT INSERTED INTO THE MONITOR ONE BATTERY WAS FOUND IN THE KITCHEN AND THE OTHER WAS INSIDE THE LIFEVEST CHARGER IT WAS REPORTED THAT THE PATIENT WOULD BE AUTOPSIED A PHYSICIAN DR B POINTED OUT THAT HE THOUGHT THAT THE BATTERY IS POTENTIALLY TOO EASY TO REMOVE DURING AN ARRHYTHMIA OR ALARM DR B WISHED FOR ZOLL TO INFORM B OF THIS CASE DR B ADDITIONALLY SUGGESTED DEVELOPING A SOLUTION TO AVOID SUCH AN EVENT IN THE FUTURE EVALUATION OF THE DEVICE DOWNLOADED FLAG FILES AND THE ASSOCIATED AUTOMATIC ECG RECORDINGS INDICATED THAT THE LIFEVEST WAS LAST STARTED ONB AT ON B AT AN ARRHYTHMIA WAS DETECTED BY THE LIFEVEST THE ECG STRIPS SHOW THAT THE PATIENT WAS IN SINUS TACHYCARDIA AT BEATS PER MINUTE TRANSITIONING TO VENTRICULAR TACHYCARDIA AT BEATS PER MINUTES THE PATIENT WAS IN VENTRICULAR TACHYCARDIA FOR APPROXIMATELY MINUTES AS CAPTURED BY THE ECG RECORDINGS AT AND THE DOWNLOADED FLAG FILES AND ECG STRIPS SHOW THAT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE ARRHYTHMIA DETECTION THE PATIENT WAS NOT TREATED BY THE LIFEVEST AS THE USE OF THE RESPONSE BUTTONS PREVENTED THE TREATMENT FROM BEING DELIVERED DETECTION OF THE ARRHYTHMIA CONTINUED UNTIL THE LIFEVEST WAS SHUTDOWN BY REMOVAL OF THE BATTERY PACK AT ON B SUBSEQUENT MONITORING OF THE PATIENT WAS NOT POSSIBLE AFTER THE DEVICE SHUTDOWN | Death | Patient problem code
6767762 | 81843117 | D | ZOLL MANUFACTURING CORPORATION WAS NOTIFIED ON THAT A B PATIENT HAD PASSED AWAY WHILE WEARING THE LIFEVEST THE PATIENT WAS FOUND ALONE AND DEAD IN HIS FLAT AS REPORTED BY THE CRIMINAL INVESTIGATION DEPARTMENT DUREN KK IT WAS REPORTED BY THE POLICE THAT THE PATIENT WAS FOUND LIFELESS IN HIS BED WEARING THE LIFEVEST BUT THE LIFEVEST BATTERY PACK WAS NOT INSERTED INTO THE MONITOR ONE BATTERY WAS FOUND IN THE KITCHEN AND THE OTHER WAS INSIDE THE LIFEVEST CHARGER IT WAS REPORTED THAT THE PATIENT WOULD BE AUTOPSIED A PHYSICIAN DR B POINTED OUT THAT HE THOUGHT THAT THE BATTERY IS POTENTIALLY TOO EASY TO REMOVE DURING AN ARRHYTHMIA OR ALARM DR B WISHED FOR ZOLL TO INFORM B OF THIS CASE DR B ADDITIONALLY SUGGESTED DEVELOPING A SOLUTION TO AVOID SUCH AN EVENT IN THE FUTURE EVALUATION OF THE DEVICE DOWNLOADED FLAG FILES AND THE ASSOCIATED AUTOMATIC ECG RECORDINGS INDICATED THAT THE LIFEVEST WAS LAST STARTED ONB AT ON B AT AN ARRHYTHMIA WAS DETECTED BY THE LIFEVEST THE ECG STRIPS SHOW THAT THE PATIENT WAS IN SINUS TACHYCARDIA AT BEATS PER MINUTE TRANSITIONING TO VENTRICULAR TACHYCARDIA AT BEATS PER MINUTES THE PATIENT WAS IN VENTRICULAR TACHYCARDIA FOR APPROXIMATELY MINUTES AS CAPTURED BY THE ECG RECORDINGS AT AND THE DOWNLOADED FLAG FILES AND ECG STRIPS SHOW THAT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE ARRHYTHMIA DETECTION THE PATIENT WAS NOT TREATED BY THE LIFEVEST AS THE USE OF THE RESPONSE BUTTONS PREVENTED THE TREATMENT FROM BEING DELIVERED DETECTION OF THE ARRHYTHMIA CONTINUED UNTIL THE LIFEVEST WAS SHUTDOWN BY REMOVAL OF THE BATTERY PACK AT ON B SUBSEQUENT MONITORING OF THE PATIENT WAS NOT POSSIBLE AFTER THE DEVICE SHUTDOWN | Adverse Event Without Identified Device or Use Problem | Device problem code
6767762 | 81843117 | D | ZOLL MANUFACTURING CORPORATION WAS NOTIFIED ON THAT A B PATIENT HAD PASSED AWAY WHILE WEARING THE LIFEVEST THE PATIENT WAS FOUND ALONE AND DEAD IN HIS FLAT AS REPORTED BY THE CRIMINAL INVESTIGATION DEPARTMENT DUREN KK IT WAS REPORTED BY THE POLICE THAT THE PATIENT WAS FOUND LIFELESS IN HIS BED WEARING THE LIFEVEST BUT THE LIFEVEST BATTERY PACK WAS NOT INSERTED INTO THE MONITOR ONE BATTERY WAS FOUND IN THE KITCHEN AND THE OTHER WAS INSIDE THE LIFEVEST CHARGER IT WAS REPORTED THAT THE PATIENT WOULD BE AUTOPSIED A PHYSICIAN DR B POINTED OUT THAT HE THOUGHT THAT THE BATTERY IS POTENTIALLY TOO EASY TO REMOVE DURING AN ARRHYTHMIA OR ALARM DR B WISHED FOR ZOLL TO INFORM B OF THIS CASE DR B ADDITIONALLY SUGGESTED DEVELOPING A SOLUTION TO AVOID SUCH AN EVENT IN THE FUTURE EVALUATION OF THE DEVICE DOWNLOADED FLAG FILES AND THE ASSOCIATED AUTOMATIC ECG RECORDINGS INDICATED THAT THE LIFEVEST WAS LAST STARTED ONB AT ON B AT AN ARRHYTHMIA WAS DETECTED BY THE LIFEVEST THE ECG STRIPS SHOW THAT THE PATIENT WAS IN SINUS TACHYCARDIA AT BEATS PER MINUTE TRANSITIONING TO VENTRICULAR TACHYCARDIA AT BEATS PER MINUTES THE PATIENT WAS IN VENTRICULAR TACHYCARDIA FOR APPROXIMATELY MINUTES AS CAPTURED BY THE ECG RECORDINGS AT AND THE DOWNLOADED FLAG FILES AND ECG STRIPS SHOW THAT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE ARRHYTHMIA DETECTION THE PATIENT WAS NOT TREATED BY THE LIFEVEST AS THE USE OF THE RESPONSE BUTTONS PREVENTED THE TREATMENT FROM BEING DELIVERED DETECTION OF THE ARRHYTHMIA CONTINUED UNTIL THE LIFEVEST WAS SHUTDOWN BY REMOVAL OF THE BATTERY PACK AT ON B SUBSEQUENT MONITORING OF THE PATIENT WAS NOT POSSIBLE AFTER THE DEVICE SHUTDOWN | Ventricular Tachycardia | Patient problem code
6761632 | 81843198 | N | RESULTS TWO OPEN G X MM PEN NEEDLE SAMPLES AND THREE PHOTOS WERE RETURNED FROM AN UNKNOWN LOT NO CAT NO VISUAL EXAMINATION WAS CARRIED OUT ON THE RETURNED SAMPLES AND PHOTOS AND IT WAS OBSERVED THAT THE CANNULA WAS MISSING FROM BOTH SAMPLES CONCLUSION EVALUATION OF THE DEFECT SAMPLE INDICATED THAT THE ADHESIVE BOND HAD SEPARATED FROM THE HUB COMPONENT BASED ON AN EVALUATION OF SEVERITY AND FREQUENCY IT WAS DETERMINED THAT NO CORRECTIVE ACTION IS REQUIRED AT THIS TIME B | Material Separation | Device problem code
6761632 | 81843198 | N | RESULTS TWO OPEN G X MM PEN NEEDLE SAMPLES AND THREE PHOTOS WERE RETURNED FROM AN UNKNOWN LOT NO CAT NO VISUAL EXAMINATION WAS CARRIED OUT ON THE RETURNED SAMPLES AND PHOTOS AND IT WAS OBSERVED THAT THE CANNULA WAS MISSING FROM BOTH SAMPLES CONCLUSION EVALUATION OF THE DEFECT SAMPLE INDICATED THAT THE ADHESIVE BOND HAD SEPARATED FROM THE HUB COMPONENT BASED ON AN EVALUATION OF SEVERITY AND FREQUENCY IT WAS DETERMINED THAT NO CORRECTIVE ACTION IS REQUIRED AT THIS TIME B | Defective Device | Device problem code
6761632 | 81843198 | N | RESULTS TWO OPEN G X MM PEN NEEDLE SAMPLES AND THREE PHOTOS WERE RETURNED FROM AN UNKNOWN LOT NO CAT NO VISUAL EXAMINATION WAS CARRIED OUT ON THE RETURNED SAMPLES AND PHOTOS AND IT WAS OBSERVED THAT THE CANNULA WAS MISSING FROM BOTH SAMPLES CONCLUSION EVALUATION OF THE DEFECT SAMPLE INDICATED THAT THE ADHESIVE BOND HAD SEPARATED FROM THE HUB COMPONENT BASED ON AN EVALUATION OF SEVERITY AND FREQUENCY IT WAS DETERMINED THAT NO CORRECTIVE ACTION IS REQUIRED AT THIS TIME B | Component Missing | Device problem code
6761632 | 81843198 | N | RESULTS TWO OPEN G X MM PEN NEEDLE SAMPLES AND THREE PHOTOS WERE RETURNED FROM AN UNKNOWN LOT NO CAT NO VISUAL EXAMINATION WAS CARRIED OUT ON THE RETURNED SAMPLES AND PHOTOS AND IT WAS OBSERVED THAT THE CANNULA WAS MISSING FROM BOTH SAMPLES CONCLUSION EVALUATION OF THE DEFECT SAMPLE INDICATED THAT THE ADHESIVE BOND HAD SEPARATED FROM THE HUB COMPONENT BASED ON AN EVALUATION OF SEVERITY AND FREQUENCY IT WAS DETERMINED THAT NO CORRECTIVE ACTION IS REQUIRED AT THIS TIME B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6767707 | 81843241 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A HEAT RASH UNDER THE BREASTS THE PATIENTS PHYSICIAN PRESCRIBED A MEDICATION AND THE IRRITATION IMPROVED | Rash | Patient problem code
6767707 | 81843241 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A HEAT RASH UNDER THE BREASTS THE PATIENTS PHYSICIAN PRESCRIBED A MEDICATION AND THE IRRITATION IMPROVED | Irritation | Patient problem code
6767707 | 81843241 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT DEVELOPED A HEAT RASH UNDER THE BREASTS THE PATIENTS PHYSICIAN PRESCRIBED A MEDICATION AND THE IRRITATION IMPROVED | Insufficient Information | Device problem code
6761809 | 81851971 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THE TEST WELL IMAGES WERE CONSISTENT WITH THE ERRONEOUS OUTCOME | Connection Problem | Device problem code
6761809 | 81851971 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THE TEST WELL IMAGES WERE CONSISTENT WITH THE ERRONEOUS OUTCOME | Communication or Transmission Problem | Device Problem Code
6763645 | 81853075 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION AND THAT WELL HAD SLIGHT RED BLOOD CELL ADHERENCE THIS DATA OBTAINED HAD BEEN ARCHIVED THE FULL AND COMPLETE CUSTOMER TELEPHONE EXTENSION NUMBER WAS B | Connection Problem | Device problem code
6763645 | 81853075 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION AND THAT WELL HAD SLIGHT RED BLOOD CELL ADHERENCE THIS DATA OBTAINED HAD BEEN ARCHIVED THE FULL AND COMPLETE CUSTOMER TELEPHONE EXTENSION NUMBER WAS B | No Device Problem Found | Medical Evaluation result code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Component or Accessory Incompatibility | Device problem code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Nausea | Patient problem code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Vomiting | Patient problem code
6768799 | 81895852 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS SHAKING HANDS NAUSEA FALLING AND VOMITING THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO STATED THAT THE DAY BEFORE THE HIGH THEY HAD A LOW AND TREATED WITH GRAPE JUICE AND THEIR PUMP LIGHT IS NO LONGER WORKING THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Fall | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Death | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Cardiac Arrest | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Arrhythmia | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Shock | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Ventricular Fibrillation | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Loss Of Consciousness | Patient problem code
6768151 | 81898995 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST REVIEW OF THE DOWNLOAD DATA INDICATES THAT THE PATIENT RECEIVED AN APPROPRIATE TREATMENT EVENT ON THE DATE OF PASSING THE PATIENT WAS AT HOME AND UNCONSCIOUS AT THE TIME OF THE EVENT IT WAS REPORTED THAT EMS WAS CONTACTED AND THAT PARAMEDICS PRONOUNCED THE PATIENT DEAD ON B A VENTRICULAR FIBRILLATION VF ARRHYTHMIA WAS FIRST DETECTED AT THE PATIENT RECEIVED FOUR NONLIFEVEST DEFIBRILLATION SHOCKS WHILE THE PATIENT REMAINED IN VF AT THE PATIENT RECEIVED A LIFEVEST TREATMENT WHILE IN VF THE PATIENT REMAINED IN VF UNTIL WHEN THE PATIENTS RHYTHM SHOWED ASYSTOLE WITH PACER SPIKES THE PATIENT REMAINED IN ASYSTOLE UNTIL DEVICE REMOVAL AT THE PATIENT PASSED AWAY ON B | Insufficient Information | Device problem code
6767736 | 81903443 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A B PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENT IS REPORTEDLY ALLERGIC TO NICKEL THE PATIENTS PHYSICIAN PRESCRIBED A DERMATOLOGICAL TREATMENT THE PATIENT WAS LATER IMPLANTED WITH AN ICD THE MEDICAL OUTCOME OF THE REACTION IS UNKNOWN | Hypersensitivity | Patient problem code
6767736 | 81903443 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A B PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENT IS REPORTEDLY ALLERGIC TO NICKEL THE PATIENTS PHYSICIAN PRESCRIBED A DERMATOLOGICAL TREATMENT THE PATIENT WAS LATER IMPLANTED WITH AN ICD THE MEDICAL OUTCOME OF THE REACTION IS UNKNOWN | Insufficient Information | Device problem code
6764111 | 81904101 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE INSULIN PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST POWER GRAPH ANALYSIS CONFIRMED LOW BATTERY POWER LOSS REPLACE BATTERY NOW ALARMS AND AN ABNORMAL LOADED VOLTAGE AT THE POWER MANAGEMENT GRAPH DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Scratched Material | Device problem code
6764111 | 81904101 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE INSULIN PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST POWER GRAPH ANALYSIS CONFIRMED LOW BATTERY POWER LOSS REPLACE BATTERY NOW ALARMS AND AN ABNORMAL LOADED VOLTAGE AT THE POWER MANAGEMENT GRAPH DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Disconnection | Device problem code
6764111 | 81904101 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE INSULIN PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST POWER GRAPH ANALYSIS CONFIRMED LOW BATTERY POWER LOSS REPLACE BATTERY NOW ALARMS AND AN ABNORMAL LOADED VOLTAGE AT THE POWER MANAGEMENT GRAPH DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Component or Accessory Incompatibility | Device Problem Code
6764111 | 81904101 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE INSULIN PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST POWER GRAPH ANALYSIS CONFIRMED LOW BATTERY POWER LOSS REPLACE BATTERY NOW ALARMS AND AN ABNORMAL LOADED VOLTAGE AT THE POWER MANAGEMENT GRAPH DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Loss of Power | Device problem code
6764111 | 81904101 | N | THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE INSULIN PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST POWER GRAPH ANALYSIS CONFIRMED LOW BATTERY POWER LOSS REPLACE BATTERY NOW ALARMS AND AN ABNORMAL LOADED VOLTAGE AT THE POWER MANAGEMENT GRAPH DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Operational Problem Identified | Medical Evaluation result code
6764070 | 81915249 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | No Device Problem Found | Medical Evaluation result code
6763455 | 81922127 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND STATED THAT HE CHECKED THE IABPS FAULTERROR LOGS AND THERE WERE NO ERROR MESSAGE LISTED IN THE LOGS THE FSE ALSO STATED THAT HE OBSERVED THE SYSTEM AND COULD NOT REPRODUCE THE REPORTED ERROR THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE WITH ALL TESTS PASSED | Device Displays Incorrect Message | Device problem code
6763455 | 81922127 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND STATED THAT HE CHECKED THE IABPS FAULTERROR LOGS AND THERE WERE NO ERROR MESSAGE LISTED IN THE LOGS THE FSE ALSO STATED THAT HE OBSERVED THE SYSTEM AND COULD NOT REPRODUCE THE REPORTED ERROR THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE WITH ALL TESTS PASSED | Image Artifact | Medical Evaluation result code
6764791 | 81929217 | N | THE CUSTOMER DECLINED TO PROVIDE ANY BLOOD SAMPLE INFORMATION AND THE CUSTOMER ALSO DECLINED TO ALLOW IMMUCOR TECHNICAL SUPPORT ACCESS TO USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE CUSTOMER STATED THAT TECHNOLOGIST ERROR WITH HANDLING THE REAGENT VIALS MAY HAVE BEEN THE CAUSE OF THE CARRYOVER EVENT | Connection Problem | Device problem code
6764791 | 81929217 | N | THE CUSTOMER DECLINED TO PROVIDE ANY BLOOD SAMPLE INFORMATION AND THE CUSTOMER ALSO DECLINED TO ALLOW IMMUCOR TECHNICAL SUPPORT ACCESS TO USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE CUSTOMER STATED THAT TECHNOLOGIST ERROR WITH HANDLING THE REAGENT VIALS MAY HAVE BEEN THE CAUSE OF THE CARRYOVER EVENT | No Device Problem Found | Medical Evaluation result code
6768489 | 82012184 | N | ONE REFURBISHED VULCAN RF GENERATOR WAS RECEIVED COMPLAINT OF GROUND PAD ERRORS WERE CONFIRMED PRODUCT FAILED FUNCTIONAL TESTING WITH NEM CIRCUIT ERROR NEM CIRCUIT TP ON RF PCB PN IS OUT OF SPEC WITH VOLT READING CAUSING ERROR ALARM SPEC IS TO VOLTS THE RF BOARD WAS RECALIBRATED INCLUDING THE NEM CIRCUIT AND UNIT PASSED FUNCTIONAL TESTING | Calibration Problem | Device problem code
6768489 | 82012184 | N | ONE REFURBISHED VULCAN RF GENERATOR WAS RECEIVED COMPLAINT OF GROUND PAD ERRORS WERE CONFIRMED PRODUCT FAILED FUNCTIONAL TESTING WITH NEM CIRCUIT ERROR NEM CIRCUIT TP ON RF PCB PN IS OUT OF SPEC WITH VOLT READING CAUSING ERROR ALARM SPEC IS TO VOLTS THE RF BOARD WAS RECALIBRATED INCLUDING THE NEM CIRCUIT AND UNIT PASSED FUNCTIONAL TESTING | Failure to Calibrate | Device Problem Code
6766864 | 82012963 | D | A CUSTOMER REPORTED A GAUGE VITRECTOMY CUTTER HAD NO CUT FUNCTION WHEN USED IN A VITRECTOMY PROCEDURE IT WAS NOT CUTTING THE VITREOUS ALTHOUGH THE ASPIRATION WAS WORKING THERE WAS NO HARM TO THE PATIENT THE PRODUCT SAMPLE IS AVAILABLE ADDITIONAL INFORMATION WAS REQUESTED HOWEVER NONE HAS BEEN RECEIVED TO DATE | No Consequences Or Impact To Patient | Patient problem code
6766864 | 82012963 | D | A CUSTOMER REPORTED A GAUGE VITRECTOMY CUTTER HAD NO CUT FUNCTION WHEN USED IN A VITRECTOMY PROCEDURE IT WAS NOT CUTTING THE VITREOUS ALTHOUGH THE ASPIRATION WAS WORKING THERE WAS NO HARM TO THE PATIENT THE PRODUCT SAMPLE IS AVAILABLE ADDITIONAL INFORMATION WAS REQUESTED HOWEVER NONE HAS BEEN RECEIVED TO DATE | Failure to Cut | Device problem code
6767086 | 82018132 | D | IT WAS REPORTED WHEN THE DOCTOR CHECKED THE DRESSING DURING NPWT FOR HIP PRESSURE ULCER WITH RENASYS GO PUMP HESHE FOUND THAT EXUDATE FLUID WAS NOT REMOVED PROPERLY AND IT REMAINED UNDER THE DRESSING | Appropriate Term/Code Not Available | Device problem code
6767086 | 82018132 | D | IT WAS REPORTED WHEN THE DOCTOR CHECKED THE DRESSING DURING NPWT FOR HIP PRESSURE ULCER WITH RENASYS GO PUMP HESHE FOUND THAT EXUDATE FLUID WAS NOT REMOVED PROPERLY AND IT REMAINED UNDER THE DRESSING | Pressure Sores/Ulcers | Patient problem code
6773291 | 82026222 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE B ADDITIONAL INFORMATION REQUESTED AND RECEIVED ALL THE ANSWERS AROUND THE PROCEDURE ITSELF ARE UNKNOWN AT THIS TIME AS THE PHYSICIAN HAS YET TO PROVIDE A RESPONSE ANDOR THE VIDEO WAS THE PROCEDURE PERFORMED LAPAROSCOPICALLY OR OPEN LAPAROSCOPIC DOES THE PHYSICAL FEEL THAT THE DEVICE LED TO THE POSTOP COMPLICATIONS PLEASE EXPLAIN YES MY UNDERSTANDING FROM THE SURGEON WHO WAS ONCALL WAS THAT BOTH HE AND DR B PHYSICIAN WHOS PATIENT HAD TO BE BROUGHT BACK THOUGHT IT WAS DUE TO BLEEDING ALONG THE STAPLE LINE WHERE PVS WAS UTILIZED IT WAS REPORTED THAT ON A CASE LAST FRIDAY THERE WAS BLEEDING THRU THE STAPLE LINE FROM THE GONADAL VEIN IF I CAN FIND THE VIDEO I CAN SHOW IT TO YOU WERE YOU REFERRING TO THIS PATIENT OR A DIFFERENT PATIENT IF A DIFFERENT PATIENT HAS A SECOND COMPLAINT BEEN REPORTED SAME PATIENT I COPYPASTED THE SURGEONS EMAIL CONTENT WHEN I FILLED OUT THE REPORT WHAT IS THE CURRENT PATIENT STATUS UNKNOWN ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE WHERE WAS THE TUMOR LOCATED WHAT WAS THE SIZE OF THE TUMOR WAS THERE INVASION OF THE TUMOR INTO THE RENAL VEIN WHAT WAS THE SIZE OF THE GONADAL VEIN PLEASE DESCRIBE THE BLEEDING THROUGH THE STAPLE LINE OOZING ETC HOW WAS THE BLEEDING ADDRESSED IT IS THE SURGEONS NORMAL PRACTICE TO USE A STAPLER ON THE GONADAL VEIN WERE ANY STAPLE FORMATION ISSUES NOTED IN THE PRIMARY PROCEDURE WAS IT CONFIRMED THAT THE ENTIRE VESSEL WAS PROXIMAL TO CUT LINE PRIOR TO FIRING WAS THERE ANY EXTRANEOUS TISSUE BEYOND THE CUT LINE WHEN DEVICE WAS FIRED WAS THE TIP OF DEVICE VISUALIZED PRIOR TO FIRING WERE ANY CLIPS USED IN THE AREA OF THE FIRING WHAT WERE THE POSTOPERATIVE COMPLICATIONS HOW LONG AFTER THE PRIMARY PROCEDURE WAS THE REOPERATION WHAT WAS DONE DURING THE REOPERATION WERE ANY STAPLE FORMATION ISSUES NOTED IN THE REOPERATION IF AVAILABLE PLEASE SEND THE VIDEO WHAT IS THE CURRENT PATIENT STATUS | Failure to Form Staple | Device Problem Code
6773291 | 82026222 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE B ADDITIONAL INFORMATION REQUESTED AND RECEIVED ALL THE ANSWERS AROUND THE PROCEDURE ITSELF ARE UNKNOWN AT THIS TIME AS THE PHYSICIAN HAS YET TO PROVIDE A RESPONSE ANDOR THE VIDEO WAS THE PROCEDURE PERFORMED LAPAROSCOPICALLY OR OPEN LAPAROSCOPIC DOES THE PHYSICAL FEEL THAT THE DEVICE LED TO THE POSTOP COMPLICATIONS PLEASE EXPLAIN YES MY UNDERSTANDING FROM THE SURGEON WHO WAS ONCALL WAS THAT BOTH HE AND DR B PHYSICIAN WHOS PATIENT HAD TO BE BROUGHT BACK THOUGHT IT WAS DUE TO BLEEDING ALONG THE STAPLE LINE WHERE PVS WAS UTILIZED IT WAS REPORTED THAT ON A CASE LAST FRIDAY THERE WAS BLEEDING THRU THE STAPLE LINE FROM THE GONADAL VEIN IF I CAN FIND THE VIDEO I CAN SHOW IT TO YOU WERE YOU REFERRING TO THIS PATIENT OR A DIFFERENT PATIENT IF A DIFFERENT PATIENT HAS A SECOND COMPLAINT BEEN REPORTED SAME PATIENT I COPYPASTED THE SURGEONS EMAIL CONTENT WHEN I FILLED OUT THE REPORT WHAT IS THE CURRENT PATIENT STATUS UNKNOWN ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE WHERE WAS THE TUMOR LOCATED WHAT WAS THE SIZE OF THE TUMOR WAS THERE INVASION OF THE TUMOR INTO THE RENAL VEIN WHAT WAS THE SIZE OF THE GONADAL VEIN PLEASE DESCRIBE THE BLEEDING THROUGH THE STAPLE LINE OOZING ETC HOW WAS THE BLEEDING ADDRESSED IT IS THE SURGEONS NORMAL PRACTICE TO USE A STAPLER ON THE GONADAL VEIN WERE ANY STAPLE FORMATION ISSUES NOTED IN THE PRIMARY PROCEDURE WAS IT CONFIRMED THAT THE ENTIRE VESSEL WAS PROXIMAL TO CUT LINE PRIOR TO FIRING WAS THERE ANY EXTRANEOUS TISSUE BEYOND THE CUT LINE WHEN DEVICE WAS FIRED WAS THE TIP OF DEVICE VISUALIZED PRIOR TO FIRING WERE ANY CLIPS USED IN THE AREA OF THE FIRING WHAT WERE THE POSTOPERATIVE COMPLICATIONS HOW LONG AFTER THE PRIMARY PROCEDURE WAS THE REOPERATION WHAT WAS DONE DURING THE REOPERATION WERE ANY STAPLE FORMATION ISSUES NOTED IN THE REOPERATION IF AVAILABLE PLEASE SEND THE VIDEO WHAT IS THE CURRENT PATIENT STATUS | Operational Problem Identified | Medical Evaluation result code
6771076 | 82037057 | D | MRI OF THE RIGHT SHOULDER PATIENT RETURNED TO THE FACILITY THE FOLLOWING DAY WITH SECOND DEGREE BURNS TO THE ANTERIOR AND POSTERIOR ASPECT OF THE RIGHT SHOULDER THE PATIENT WAS IMAGED WITH A PATIENT GOWN AND NO DIRECT CONTACT WITH THE SHOULDER COIL OCCURRED PATIENT WAS SEEN IN THE ER AND RELEASED WITH NO FURTHER ISSUES | Insufficient Information | Device problem code
6771076 | 82037057 | D | MRI OF THE RIGHT SHOULDER PATIENT RETURNED TO THE FACILITY THE FOLLOWING DAY WITH SECOND DEGREE BURNS TO THE ANTERIOR AND POSTERIOR ASPECT OF THE RIGHT SHOULDER THE PATIENT WAS IMAGED WITH A PATIENT GOWN AND NO DIRECT CONTACT WITH THE SHOULDER COIL OCCURRED PATIENT WAS SEEN IN THE ER AND RELEASED WITH NO FURTHER ISSUES | Second-Degree Burn | Patient problem code
6773496 | 82039877 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS LEFT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Insufficient Information | Device problem code
6773496 | 82039877 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS LEFT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Keratitis | Patient problem code
6773496 | 82039877 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS LEFT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Blurred Vision | Patient Problem Code
6773496 | 82039877 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS LEFT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Corneal Abrasion | Patient problem code
6773494 | 82039895 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS RIGHT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Insufficient Information | Device problem code
6773494 | 82039895 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS RIGHT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Keratitis | Patient problem code
6773494 | 82039895 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS RIGHT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Blurred Vision | Patient problem code
6773494 | 82039895 | D | AN OPTOMETRIST REPORTED A PATIENT PRESENTED WITH BLURRY VISION AND LIGHT SENSITIVITY IN BOTH EYES ONE DAY POST LASIK THE PATIENT WAS NOTED TO HAVE DIFFUSE LAMELLAR KERATITIS DLK WITH CORNEAL ABRASIONS THE PATIENT WAS GIVEN AN ORAL STEROID AND THE PREVIOUSLY PRESCRIBED TOPICAL STEROID WAS INCREASED ADDITIONAL INFORMATION RECEIVED DLK RESOLVED AND THE CORNEAL ABRASION HAS HEALED THERE ARE TWO RELATED REPORTS FOR THIS PATIENT THIS REPORT ADDRESSES THE PATIENTS RIGHT EYE AND ANOTHER MANUFACTURER REPORT WILL BE FILED FOR THE FELLOW EYE | Corneal Abrasion | Patient Problem Code
6771726 | 82052332 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS REPORTED THAT THE DISPLAY WAS FLASHING BUT THE IABP OPERATED NORMALLY THE FSE FOUND A LOOSE CONNECTION ON THE VIDEO DRIVER BOARD THE FSE THEN CLEANED ALL VIDEO DRIVER CONNECTIONS AND VERIFIED PROPER DISPLAY FUNCTION THE FSE PERFORMED ALL FUNCTIONAL OPERATIONAL AND ELECTRICAL SAFETY TESTS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Loose or Intermittent Connection | Device problem code
6771726 | 82052332 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS REPORTED THAT THE DISPLAY WAS FLASHING BUT THE IABP OPERATED NORMALLY THE FSE FOUND A LOOSE CONNECTION ON THE VIDEO DRIVER BOARD THE FSE THEN CLEANED ALL VIDEO DRIVER CONNECTIONS AND VERIFIED PROPER DISPLAY FUNCTION THE FSE PERFORMED ALL FUNCTIONAL OPERATIONAL AND ELECTRICAL SAFETY TESTS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Display or Visual Feedback Problem | Device Problem Code
6771726 | 82052332 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS REPORTED THAT THE DISPLAY WAS FLASHING BUT THE IABP OPERATED NORMALLY THE FSE FOUND A LOOSE CONNECTION ON THE VIDEO DRIVER BOARD THE FSE THEN CLEANED ALL VIDEO DRIVER CONNECTIONS AND VERIFIED PROPER DISPLAY FUNCTION THE FSE PERFORMED ALL FUNCTIONAL OPERATIONAL AND ELECTRICAL SAFETY TESTS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Detachment of Device or device Component | Device problem code
6771726 | 82052332 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS REPORTED THAT THE DISPLAY WAS FLASHING BUT THE IABP OPERATED NORMALLY THE FSE FOUND A LOOSE CONNECTION ON THE VIDEO DRIVER BOARD THE FSE THEN CLEANED ALL VIDEO DRIVER CONNECTIONS AND VERIFIED PROPER DISPLAY FUNCTION THE FSE PERFORMED ALL FUNCTIONAL OPERATIONAL AND ELECTRICAL SAFETY TESTS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Connection Problem | Device problem code
6771726 | 82052332 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS REPORTED THAT THE DISPLAY WAS FLASHING BUT THE IABP OPERATED NORMALLY THE FSE FOUND A LOOSE CONNECTION ON THE VIDEO DRIVER BOARD THE FSE THEN CLEANED ALL VIDEO DRIVER CONNECTIONS AND VERIFIED PROPER DISPLAY FUNCTION THE FSE PERFORMED ALL FUNCTIONAL OPERATIONAL AND ELECTRICAL SAFETY TESTS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Device Problem Found | Medical Evaluation result code
6774013 | 82053446 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD AN INTERMITTENT HEAT RASH CAUSED BY THE LIFEVEST THE PATIENT CONSULTED HIS PHYSICIAN WHO PRESCRIBED A LOTION AND A CREAM FOLLOWUP INDICATED THAT THE IRRITATION WAS ALMOST GONE WITH USE OF THE CREAM | Irritation | Patient problem code
6774013 | 82053446 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD AN INTERMITTENT HEAT RASH CAUSED BY THE LIFEVEST THE PATIENT CONSULTED HIS PHYSICIAN WHO PRESCRIBED A LOTION AND A CREAM FOLLOWUP INDICATED THAT THE IRRITATION WAS ALMOST GONE WITH USE OF THE CREAM | Insufficient Information | Device problem code
6774013 | 82053446 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD AN INTERMITTENT HEAT RASH CAUSED BY THE LIFEVEST THE PATIENT CONSULTED HIS PHYSICIAN WHO PRESCRIBED A LOTION AND A CREAM FOLLOWUP INDICATED THAT THE IRRITATION WAS ALMOST GONE WITH USE OF THE CREAM | Rash | Patient problem code
6768024 | 82056903 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE ANTIBODY SCREEN TEST WELL IMAGE IN QUESTION WHICH WAS VISUALLY POSITIVE THE FULL AND COMPLETE FACILITY TELEPHONE EXTENSION NUMBER IS B | No Device Problem Found | Medical Evaluation result code
6774309 | 82060967 | D | IT WAS REPORTED THE NURSE RECAPPED A ML BD TB SYRINGE WITH G X IN BD PRECISIONGLIDE DETACHABLE NEEDLE ONE HANDED THE NEEDLE WENT THROUGH THE SHIELD AND RESULTED IN A NEEDLE STICK INJURY AFTER USE MEDICAL INTERVENTIONS WERE PROVIDED A BLOOD SAMPLE FROM BOTH THE PATIENT AND CLINICIAN WAS COLLECTED TO ASCERTAIN THE POSSIBILITY OF HIV EXPOSURE AT THIS TIME | Injury | Patient problem code
6774309 | 82060967 | D | IT WAS REPORTED THE NURSE RECAPPED A ML BD TB SYRINGE WITH G X IN BD PRECISIONGLIDE DETACHABLE NEEDLE ONE HANDED THE NEEDLE WENT THROUGH THE SHIELD AND RESULTED IN A NEEDLE STICK INJURY AFTER USE MEDICAL INTERVENTIONS WERE PROVIDED A BLOOD SAMPLE FROM BOTH THE PATIENT AND CLINICIAN WAS COLLECTED TO ASCERTAIN THE POSSIBILITY OF HIV EXPOSURE AT THIS TIME | Insufficient Information | Device problem code
6768726 | 82063947 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMOTHERM WOULD NOT COOL THERE WAS NO PATIENT HARMED OR INJURED AND THE PROCEDURE CONTINUED SUCCESSFULLY CSZ ATTEMPTED TO TROUBLE SHOOT THE DEVICE OVER THE PHONE HAVING THE FACILITYS BIOMED CHECK THE COMPRESSOR THE COMPRESSOR WAS CONFIRMED ON AND WORKING THE DEVICE WAS RETURNED AND CSZ FOUND A SMALL REFRIGERANT LEAK BELOW THE FILTER DRYER THE FITTING BELOW THE FILTER DRYER NEEDED TO BE TIGHTENED THE FITTING WAS TIGHTENED AND THE DEVICE FUNCTIONS AS DESIGNED | Leak / Splash | Device problem code
6768726 | 82063947 | N | CSZ MEDICAL TECHNICAL SUPPORT RECEIVED A COMPLAINT STATING THE HEMOTHERM WOULD NOT COOL THERE WAS NO PATIENT HARMED OR INJURED AND THE PROCEDURE CONTINUED SUCCESSFULLY CSZ ATTEMPTED TO TROUBLE SHOOT THE DEVICE OVER THE PHONE HAVING THE FACILITYS BIOMED CHECK THE COMPRESSOR THE COMPRESSOR WAS CONFIRMED ON AND WORKING THE DEVICE WAS RETURNED AND CSZ FOUND A SMALL REFRIGERANT LEAK BELOW THE FILTER DRYER THE FITTING BELOW THE FILTER DRYER NEEDED TO BE TIGHTENED THE FITTING WAS TIGHTENED AND THE DEVICE FUNCTIONS AS DESIGNED | Leakage/Seal | Medical Evaluation result code
6767108 | 82064071 | N | INVESTIGATION THE RUN DATA FILE RDF WAS ANALYZED FOR THIS EVENT SIGNALS IN THE RDF SUGGEST AN OBSTRUCTION IN THE INLET LINE LED TO THE ALARMS GENERATED AND SUBSEQUENT REVIEW OF THE RDF ASSOCIATED WITH THIS COMPLAINT CONFIRMS THE OCCURRENCE OF EXCESSIVE INLET PRESSURE WAS TOO LOW ALARMS INLET PRESSURE ALARMS AND RETURN PRESSURE ALARMS STARTED SHORTLY AFTER ONSET OF THE RUN AND PERSISTED FOR THE DURATION LATER ON DURING THE PROCEDURE THERE WERE SEVERAL INLET PRESSURE SENSOR MALFUNCTIONED LOWLEVEL RESERVOIR SENSOR DID NOT DETECT FLUID AND PATIENTS FLUID BALANCE MAY BE LOWER THAN REPORTED ALARMS IF THE INLET LINE OR INLET LINE TRAP DEVELOPS AN OBSTRUCTION OR PARTIAL OCCLUSION FROM A SKIN PLUG OR CLOT THIS CAN CAUSE THE MALFUNCTION OF THE INLET PRESSURE SENSOR AND EXPLAIN THESE ALARMS IN ADDITION A DEVELOPMENT OF A CLOT IN THIS LOCATION MAY CAUSE THE INLET LINE TO STOP DRAWING ENTIRELY OR TO DRAW LESS FLUID INTO THE SYSTEM BECAUSE IT IS BLOCKING THE FLUID PATHWAY THIS IN TURN CAN CAUSE LESS FLUID TO EXIT THE CHANNEL VIA THE PASSIVE RBC LINE AND CAN RESULT IN THE RESERVOIR ALARMS PREVIOUSLY MENTIONED BECAUSE LESS FLUID IS BEING RETURNED TO THE RESERVOIR THAN THE SYSTEM EXPECTS SINCE THESE ALARMS OCCURRED LATER IN THE PROCEDURE IT IS SUSPECTED THAT CLUMPS FROM THE PATIENTS ACCESS SITE WERE PULLED INTO THE INLET LINE LATER ON DURING THE RUN INVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | Obstruction of Flow | Device problem code
6767108 | 82064071 | N | INVESTIGATION THE RUN DATA FILE RDF WAS ANALYZED FOR THIS EVENT SIGNALS IN THE RDF SUGGEST AN OBSTRUCTION IN THE INLET LINE LED TO THE ALARMS GENERATED AND SUBSEQUENT REVIEW OF THE RDF ASSOCIATED WITH THIS COMPLAINT CONFIRMS THE OCCURRENCE OF EXCESSIVE INLET PRESSURE WAS TOO LOW ALARMS INLET PRESSURE ALARMS AND RETURN PRESSURE ALARMS STARTED SHORTLY AFTER ONSET OF THE RUN AND PERSISTED FOR THE DURATION LATER ON DURING THE PROCEDURE THERE WERE SEVERAL INLET PRESSURE SENSOR MALFUNCTIONED LOWLEVEL RESERVOIR SENSOR DID NOT DETECT FLUID AND PATIENTS FLUID BALANCE MAY BE LOWER THAN REPORTED ALARMS IF THE INLET LINE OR INLET LINE TRAP DEVELOPS AN OBSTRUCTION OR PARTIAL OCCLUSION FROM A SKIN PLUG OR CLOT THIS CAN CAUSE THE MALFUNCTION OF THE INLET PRESSURE SENSOR AND EXPLAIN THESE ALARMS IN ADDITION A DEVELOPMENT OF A CLOT IN THIS LOCATION MAY CAUSE THE INLET LINE TO STOP DRAWING ENTIRELY OR TO DRAW LESS FLUID INTO THE SYSTEM BECAUSE IT IS BLOCKING THE FLUID PATHWAY THIS IN TURN CAN CAUSE LESS FLUID TO EXIT THE CHANNEL VIA THE PASSIVE RBC LINE AND CAN RESULT IN THE RESERVOIR ALARMS PREVIOUSLY MENTIONED BECAUSE LESS FLUID IS BEING RETURNED TO THE RESERVOIR THAN THE SYSTEM EXPECTS SINCE THESE ALARMS OCCURRED LATER IN THE PROCEDURE IT IS SUSPECTED THAT CLUMPS FROM THE PATIENTS ACCESS SITE WERE PULLED INTO THE INLET LINE LATER ON DURING THE RUN INVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | Operational Problem Identified | Medical Evaluation result code
6537659 | 82070102 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE POWER CONNECTOR NOT FULLY CONNECTED DEVICE WAS RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Connection Problem | Device Problem Code
6537659 | 82070102 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE POWER CONNECTOR NOT FULLY CONNECTED DEVICE WAS RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Scratched Material | Device problem code
6537659 | 82070102 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE POWER CONNECTOR NOT FULLY CONNECTED DEVICE WAS RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Short Circuit | Medical Evaluation result code
6770622 | 82147957 | N | REPORTABLE MALFUNCTION WITH INOMAX DSIR B WAS CREATED AS COMPLAINTB THE SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO OVERDELIVERY A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE FIELD THE ROOT CAUSE FOR THIS REPORT WAS PROPORTIONAL VALVE FAILURE THE DISTRIBUTOR WAS RECOMMONDED TO REPLACE THE PROPORTIONAL VALVE DUE TO THIS FINDING THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM THIS CASE DID NOT RESULT IN AN ADVERSE EVENTSERIOUS ADVERSE EVENT HOWEVER IT IS BEING SUBMITTED TO REGULATORY AUTHORITIES BECAUSE A SIMILAR FAILURE OCCURRED IN THE PAST WHICH RESULTED IN A SERIOUS ADVERSE EVENT MDR | Excess Flow or Over-Infusion | Device problem code
6770249 | 82149236 | D | IT WAS REPORTED THAT A F MYNXGRIP VASCULAR CLOSURE DEVICE VCD SEALANT WAS OOZING OUT OF GROIN THERE WAS NO REPORTED PATIENT INJURY THE PRODUCT IS AVAILABLE FOR RETURN THERE WERE NO DAMAGES OR ANOMALIES NOTED WHEN THE DEVICE WAS REMOVED FROM THE PACKAGE THE DID PREP NORMALLY THE DEPLOYER IS CERTIFED ON THE MYNX DEVICES | No Consequences Or Impact To Patient | Patient problem code
6770249 | 82149236 | D | IT WAS REPORTED THAT A F MYNXGRIP VASCULAR CLOSURE DEVICE VCD SEALANT WAS OOZING OUT OF GROIN THERE WAS NO REPORTED PATIENT INJURY THE PRODUCT IS AVAILABLE FOR RETURN THERE WERE NO DAMAGES OR ANOMALIES NOTED WHEN THE DEVICE WAS REMOVED FROM THE PACKAGE THE DID PREP NORMALLY THE DEPLOYER IS CERTIFED ON THE MYNX DEVICES | Unintended Ejection | Device problem code
6771012 | 82153180 | N | AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON B TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE FSE MADE AN ADJUSTMENT TO BOTH THE Z AND X CAMERA POSITIONS AND ALSO CLEANED THE READER MIRROR THE FSE ALSO ADJUISTED THE WASHER RESIDUAL VOLUME TO GRAM IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY POSITIVE | Use of Incorrect Control Settings | Device problem code
6771012 | 82153180 | N | AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON B TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE FSE MADE AN ADJUSTMENT TO BOTH THE Z AND X CAMERA POSITIONS AND ALSO CLEANED THE READER MIRROR THE FSE ALSO ADJUISTED THE WASHER RESIDUAL VOLUME TO GRAM IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY POSITIVE | Configuration Issue | Medical Evaluation result code
6777112 | 82154098 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SWITCH MODES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6777112 | 82154098 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SWITCH MODES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6777202 | 82158463 | N | THE HILLROM TECHNICIAN EVALUATED THE BED AND FOUND THE SNAPS ON THE CENTER CUSHIONS HAD BECOME UNSNAPPED THE TECHNICIAN RECONNECTED THE SNAPS ON THE CENTER CUSHIONS TO RESOLVE THE ISSUE THE PATIENT IS BEING TREATED WITH MUPIROCIN OINTMENT THE PRESSURE SORE IS NEARLY HEALED AND THE PATIENT WILL BE SEEING THE WOUND CARE DOCTOR AGAIN NEXT WEEK PER THE HILLROM SERVICE MANUAL DURING THE INITIAL INFLATION OF MATTRESS WITH BOLSTER THE INSTALLER SHOULD ADJUST THE BOLSTERS INTO THE PROPER LOCATION ON THE SURFACE A SEARCH OF THE HILLROM MAINTENANCE RECORDS DID NOT SHOW HILLROM PERFORMED ANY PREVENTATIVE MAINTENANCE ON THIS BED IT IS UNKNOWN IF THE FACILITY PERFORMS PREVENTATIVE MAINTENANCE ON THEIR BEDS BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | Inflation Problem | Device problem code
6777202 | 82158463 | N | THE HILLROM TECHNICIAN EVALUATED THE BED AND FOUND THE SNAPS ON THE CENTER CUSHIONS HAD BECOME UNSNAPPED THE TECHNICIAN RECONNECTED THE SNAPS ON THE CENTER CUSHIONS TO RESOLVE THE ISSUE THE PATIENT IS BEING TREATED WITH MUPIROCIN OINTMENT THE PRESSURE SORE IS NEARLY HEALED AND THE PATIENT WILL BE SEEING THE WOUND CARE DOCTOR AGAIN NEXT WEEK PER THE HILLROM SERVICE MANUAL DURING THE INITIAL INFLATION OF MATTRESS WITH BOLSTER THE INSTALLER SHOULD ADJUST THE BOLSTERS INTO THE PROPER LOCATION ON THE SURFACE A SEARCH OF THE HILLROM MAINTENANCE RECORDS DID NOT SHOW HILLROM PERFORMED ANY PREVENTATIVE MAINTENANCE ON THIS BED IT IS UNKNOWN IF THE FACILITY PERFORMS PREVENTATIVE MAINTENANCE ON THEIR BEDS BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | Maintenance Problem Identified | Medical Evaluation result code
6774536 | 82163626 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON B AND DETERMINED THAT THEY BOTH HAD A CLEAR HALO AROUND THE BUTTON WITH ALSO A PINKISH OUTER CIRCLE | Connection Problem | Device problem code
6774536 | 82163626 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON B AND DETERMINED THAT THEY BOTH HAD A CLEAR HALO AROUND THE BUTTON WITH ALSO A PINKISH OUTER CIRCLE | No Device Problem Found | Medical Evaluation result code
6777662 | 82169644 | D | CUSTOMER REPORTED VIA PHONE CALL HIGH BLOOD GLUCOSE LEVELS AND HOSPITALIZATION CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL CUSTOMER WENT TO THE HOSPITAL ON B CUSTOMER FELT SICK HAD HIGH BLOOD GLUCOSE KETONES AND THEY WERE DISCHARGED AFTER NINE DAYS AT THE HOSPITAL CUSTOMER DECLINED TO TROUBLESHOOT FOR HIGH BLOOD GLUCOSE LEVELS THE DEVICE WAS NOT RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6777662 | 82169644 | D | CUSTOMER REPORTED VIA PHONE CALL HIGH BLOOD GLUCOSE LEVELS AND HOSPITALIZATION CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL CUSTOMER WENT TO THE HOSPITAL ON B CUSTOMER FELT SICK HAD HIGH BLOOD GLUCOSE KETONES AND THEY WERE DISCHARGED AFTER NINE DAYS AT THE HOSPITAL CUSTOMER DECLINED TO TROUBLESHOOT FOR HIGH BLOOD GLUCOSE LEVELS THE DEVICE WAS NOT RETURNED FOR ANALYSIS | Fever | Patient problem code
6777662 | 82169644 | D | CUSTOMER REPORTED VIA PHONE CALL HIGH BLOOD GLUCOSE LEVELS AND HOSPITALIZATION CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL CUSTOMER WENT TO THE HOSPITAL ON B CUSTOMER FELT SICK HAD HIGH BLOOD GLUCOSE KETONES AND THEY WERE DISCHARGED AFTER NINE DAYS AT THE HOSPITAL CUSTOMER DECLINED TO TROUBLESHOOT FOR HIGH BLOOD GLUCOSE LEVELS THE DEVICE WAS NOT RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6779166 | 82331479 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY HAD A HIGH BLOOD GLUCOSE LEVEL OF MGDL THEY TREATED WITH THE INSULIN PUMP THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED THERE WAS NO INDICATION OF A MALFUNCTION FROM THE INSULIN PUMP IT WAS NOTED THAT THE CUSTOMER CALLED BACK WITHIN THE DAY TO PERFORM THE BASIC OCCLUSION TEST THE INSULIN PUMP PASSED THE BASIC OCCLUSION TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6779166 | 82331479 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY HAD A HIGH BLOOD GLUCOSE LEVEL OF MGDL THEY TREATED WITH THE INSULIN PUMP THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED THERE WAS NO INDICATION OF A MALFUNCTION FROM THE INSULIN PUMP IT WAS NOTED THAT THE CUSTOMER CALLED BACK WITHIN THE DAY TO PERFORM THE BASIC OCCLUSION TEST THE INSULIN PUMP PASSED THE BASIC OCCLUSION TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6778967 | 82333436 | D | IT WAS REPORTED BY THE CUSTOMERS SPOUSE THAT THE CUSTOMER RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON THE MORNING OF B WITH BLOOD GLUCOSE IN THE MGDL RANGE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE TABLETS AND FOOD TO TREAT THE CUSTOMERS LEVELS INCREASED AND DROPPED AGAIN WITHOUT THEM BEING CONNECTED TO THE PUMP THE CUSTOMER ROSE TO MGDL TO OVER MGDL ON THE MORNING OF THE CALL ON THE EVENING OF B THE CUSTOMER WAS HOSPITALIZED DUE TO A LOW OF MGDL THE CUSTOMER DID NOT RECEIVE ANY ALARMS ON THE PUMP THE PUMP OVER DELIVERED INSULIN INTO THE CUSTOMER THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE LATER INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED AND THE CUSTOMER FEELS THE PUMP OVER DELIVERED INSULIN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6778967 | 82333436 | D | IT WAS REPORTED BY THE CUSTOMERS SPOUSE THAT THE CUSTOMER RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON THE MORNING OF B WITH BLOOD GLUCOSE IN THE MGDL RANGE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE TABLETS AND FOOD TO TREAT THE CUSTOMERS LEVELS INCREASED AND DROPPED AGAIN WITHOUT THEM BEING CONNECTED TO THE PUMP THE CUSTOMER ROSE TO MGDL TO OVER MGDL ON THE MORNING OF THE CALL ON THE EVENING OF B THE CUSTOMER WAS HOSPITALIZED DUE TO A LOW OF MGDL THE CUSTOMER DID NOT RECEIVE ANY ALARMS ON THE PUMP THE PUMP OVER DELIVERED INSULIN INTO THE CUSTOMER THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE LATER INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED AND THE CUSTOMER FEELS THE PUMP OVER DELIVERED INSULIN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Excess Flow or Over-Infusion | Device problem code
6785019 | 82382577 | N | THE TROLLEY WITH THE PREPARED PATIENT WAS MOVED INTO THE MAGNET ROOM FOR THE EXAMINATION THE TROLLEY WAS MOVED TO THE DOCKING POSITION AT THE MR TABLE WHICH HAD BEEN MOVED DOWN THE MR TABLE WAS MOVED UPWARDS IT TOOK UP THE TROLLEY WITH THE PATIENT THE PATIENT HAD HIS HANDS AT THE SIDE OF THE TROLLEY TABLE TOP AND AT THE MOMENT THE MR TABLE TOOK UP THE TROLLEY TABLE THE PATIENTS FINGER GOT JAMMED BETWEEN TROLLEY TABLE AND MR TABLE THIS RESULTED IN A DISTAL PHALANX AMPUTATION ON THE RIGHT HAND THIS IS CONSIDERED AN UNFORTUNATE INCIDENT PLEASE NOTE THAT THERE ARE ARM RESTS AVAILABLE TO LIMIT THE CHANCE OF PATIENTS GRABBING THE TABLE TOP FURTHERMORE THERE ARE WARNINGS WITHIN THE INSTRUCTIONS FOR USE STATING THAT THE OPERATOR SHOULD ALWAYS CHECK THAT NO FINGERS CABLES OR OTHER ITEMS GET CAUGHT BETWEEN THE TABLE STRETCHER AND CARRIER | Adverse Event Without Identified Device or Use Problem | Device problem code
6783891 | 82388987 | D | THE FOLLOWING WAS REPORTED TO GORE ON AN UNKNOWN DATE IN THE PATIENT PRESENTED WITH AN ABDOMINAL AORTIC ANEURYSM THAT WAS TREATED WITH GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE IN A PROXIMAL TYPE I ENDOLEAK WAS IDENTIFIED AND TREATED ON B WITH A MEDTRONIC EXTENSION DEVICE AND A GORE AORTIC EXTENDER THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND THE ENDOLEAK WAS SOLVED | Hemorrhage | Patient problem code
6783891 | 82388987 | D | THE FOLLOWING WAS REPORTED TO GORE ON AN UNKNOWN DATE IN THE PATIENT PRESENTED WITH AN ABDOMINAL AORTIC ANEURYSM THAT WAS TREATED WITH GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE IN A PROXIMAL TYPE I ENDOLEAK WAS IDENTIFIED AND TREATED ON B WITH A MEDTRONIC EXTENSION DEVICE AND A GORE AORTIC EXTENDER THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND THE ENDOLEAK WAS SOLVED | Insufficient Information | Device problem code
6779156 | 82407858 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM WITH TWO GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE IMAGING IDENTIFIED A SUSPECTED TYPE II ENDOLEAK WITH ANEURYSM GROWTH OF MM ON B AN INTERVENTION WAS PERFORMED TO TREAT THE TYPE II ENDOLEAK WHEREBY A LUMBAR ARTERY WAS COILED IT WAS REPORTED AN AORTIC EXTENDER COMPONENT WAS IMPLANTED FOR PROXIMAL EXTENSION FROM THE TRUNKIPSILATERAL LEG COMPONENT A CONTRALATERAL LEG COMPONENT WAS ALSO IMPLANTED OFF OF THE TRUNK FOR DISTAL EXTENSION THE PHYSICIAN REPORTEDLY STATED HE WANTED TO HAVE MORE SEAL AS A PRECAUTION INCASE THE ANEURYSM CONTINUED TO GROW FINAL IMAGING REPORTEDLY IDENTIFIED THE ANEURYSM WAS EXCLUDED THE PATIENT TOLERATED THE PROCEDURE | Hemorrhage | Patient problem code
6779156 | 82407858 | D | ON B THE PATIENT UNDERWENT TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM WITH TWO GORE EXCLUDER AAA ENDOPROSTHESES ON AN UNKNOWN DATE IMAGING IDENTIFIED A SUSPECTED TYPE II ENDOLEAK WITH ANEURYSM GROWTH OF MM ON B AN INTERVENTION WAS PERFORMED TO TREAT THE TYPE II ENDOLEAK WHEREBY A LUMBAR ARTERY WAS COILED IT WAS REPORTED AN AORTIC EXTENDER COMPONENT WAS IMPLANTED FOR PROXIMAL EXTENSION FROM THE TRUNKIPSILATERAL LEG COMPONENT A CONTRALATERAL LEG COMPONENT WAS ALSO IMPLANTED OFF OF THE TRUNK FOR DISTAL EXTENSION THE PHYSICIAN REPORTEDLY STATED HE WANTED TO HAVE MORE SEAL AS A PRECAUTION INCASE THE ANEURYSM CONTINUED TO GROW FINAL IMAGING REPORTEDLY IDENTIFIED THE ANEURYSM WAS EXCLUDED THE PATIENT TOLERATED THE PROCEDURE | Insufficient Information | Device problem code
6786859 | 82446387 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6786859 | 82446387 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6786859 | 82446387 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Noise, Audible | Device problem code
6786560 | 82447115 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE RECENTLY HOSPITALIZED DUE TO A HIGH BLOOD GLUCOSE LEVEL OVER MGDL CUSTOMER ALSO REPORTED A BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF ADMISSION BUT HAD BEEN OFF FOR LESS THAN HOURS TROUBLESHOOTING WAS NOT PERFORMED THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6786560 | 82447115 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE RECENTLY HOSPITALIZED DUE TO A HIGH BLOOD GLUCOSE LEVEL OVER MGDL CUSTOMER ALSO REPORTED A BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF ADMISSION BUT HAD BEEN OFF FOR LESS THAN HOURS TROUBLESHOOTING WAS NOT PERFORMED THE INSULIN PUMP WAS NOT REPLACED OR RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6785923 | 82468903 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USERS PRODIGY DIABETES METER MALFUNCTIONED AND WOULD NOT ALLOW HIM TO TEST THE END USER WAS ATTEMPTING TO PERFORM A BLOOD GLUCOSE TEST AND THE METER SUBSEQUENTLY REVERTED TO THE SETTING MODE HE WAS NOT EXPERIENCING ANY SIGNIFICANT SYMPTOMS BUT OUT OF CONCERN FOR HIS BLOOD GLUCOSE LEVELS HE VISITED THE ER UPON ARRIVAL TO THE ER HIS BLOOD GLUCOSE READING WAS MGDL HE RECEIVED UNITS OF LANTUS AND AFTER MINUTES AT THE ER HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO FOLLOW HIS INSULIN REGIMEN NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | Defective Device | Device problem code
6785923 | 82468903 | D | IT WAS REPORTED THAT MEDICAL ATTENTION WAS SOUGHT ON B AT PM AFTER THE END USERS PRODIGY DIABETES METER MALFUNCTIONED AND WOULD NOT ALLOW HIM TO TEST THE END USER WAS ATTEMPTING TO PERFORM A BLOOD GLUCOSE TEST AND THE METER SUBSEQUENTLY REVERTED TO THE SETTING MODE HE WAS NOT EXPERIENCING ANY SIGNIFICANT SYMPTOMS BUT OUT OF CONCERN FOR HIS BLOOD GLUCOSE LEVELS HE VISITED THE ER UPON ARRIVAL TO THE ER HIS BLOOD GLUCOSE READING WAS MGDL HE RECEIVED UNITS OF LANTUS AND AFTER MINUTES AT THE ER HE WAS DISCHARGED WITH A BLOOD GLUCOSE READING OF MGDL AND INSTRUCTED TO FOLLOW HIS INSULIN REGIMEN NO ADDITIONAL DETAILS WERE PROVIDED IN REGARDS TO THIS MEDICAL EVENT | No Known Impact Or Consequence To Patient | Patient problem code
6777329 | 82472537 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND REPORTED NO PROBLEMS WERE FOUND DURING THE EVALUATION OF THE IABP FSE COMPLETED PM WITH FULL CALIBRATION FUNCTIONAL TESTING AND SAFETY CHECK TO FACTORY SPECIFICATIONS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6776989 | 82473170 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED FINDING MOISTER IN THE FILTER WHICH CAUSED THE ERROR THE FSE PERFORMED THE CONDENSATION REMOVAL PROCEDURE THE FSE ALSO PERFORMED A SYSTEM CALIBRATION AND LET THE IABP RUN AND PUMP FOR DAYS WITHOUT ANY ERRORS OCCURRING THE PROBLEM HAD BEEN RESOLVED THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6776555 | 82474509 | D | IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE AT THE SECOND FIRING IT WAS DIFFICULT TO STAPLE AND IT WAS NEEDED TO FORCE THE STAPLES WERE NOT PLACED PROPERLY ON THE STOMACH AND FRAGMENTS OF THE BLADE WERE FOUND IN PATIENTS ABDOMEN PER THE SURGEON TWO FACTORS COULD CONTRIBUTE TO THE INCIDENT THE CARTRIDGE PROTECTION WAS REMOVED BEFORE THE CARTRIDGE PLACEMENT AND DIFFICULTY HAPPENED WHEN CUTTING AGAIN THE PREVIOUS STAPLE LINE MECHANICAL SUTURE REINFORCEMENT WITH A THREAD MANUAL SUTURE NO DEHISCENCE WAS NOTED NO PATIENT CONSEQUENCE WAS REPORTED | Device Fragments In Patient | Patient problem code
6776555 | 82474509 | D | IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE AT THE SECOND FIRING IT WAS DIFFICULT TO STAPLE AND IT WAS NEEDED TO FORCE THE STAPLES WERE NOT PLACED PROPERLY ON THE STOMACH AND FRAGMENTS OF THE BLADE WERE FOUND IN PATIENTS ABDOMEN PER THE SURGEON TWO FACTORS COULD CONTRIBUTE TO THE INCIDENT THE CARTRIDGE PROTECTION WAS REMOVED BEFORE THE CARTRIDGE PLACEMENT AND DIFFICULTY HAPPENED WHEN CUTTING AGAIN THE PREVIOUS STAPLE LINE MECHANICAL SUTURE REINFORCEMENT WITH A THREAD MANUAL SUTURE NO DEHISCENCE WAS NOTED NO PATIENT CONSEQUENCE WAS REPORTED | Failure to Form Staple | Device problem code
6778963 | 82475973 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO INVESTIGATE THE FSE EVALUATED THE IABP UNIT AND RESEATED THE COLOR DRIVER CABLE WHICH CORRECTED THE ISSUE THE IABP UNIT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6789042 | 82509802 | D | ON B THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM MEASURING MM IN DIAMETER AND WAS IMPLANTED WITH GORE EXCLUDER ENDOPROSTHESES THE PATIENT TOLERATED THE PROCEDURE ON B A TYPE II ENDOLEAK FROM THE LEFT LUMBAR ARTERY WAS IDENTIFIED ON B THE PATIENT UNDERWENT EMBOLIZATION PROCEDURE USING ONYX TO TREAT THE TYPE II ENDOLEAK ON B THE PATIENT UNDERWENT REINTERVENTION FOR A TYPE II ENDOLEAK EMBOLIZATION WAS ATTEMPTED BUT WAS UNABLE TO BE PERFORMED DUE TO POOR ACCESS NO ANEURYSM ENLARGEMENT WAS REPORTED ON B THE PATIENT WAS AT THE HOSPITAL TO CHECK AN ENDOLEAK ON B THE CT A WAS PERFORMED ALSO WAS REPORTED THAT THE PATIENT HOSPITAL STAY WAS EXTENDED DUE TO A COURSE OF PRE AND POST PROCEDURAL IV RENAL PROTECTION THE PHYSICIAN WAS MONITORING THE PATIENT | Insufficient Information | Device problem code
6789042 | 82509802 | D | ON B THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM MEASURING MM IN DIAMETER AND WAS IMPLANTED WITH GORE EXCLUDER ENDOPROSTHESES THE PATIENT TOLERATED THE PROCEDURE ON B A TYPE II ENDOLEAK FROM THE LEFT LUMBAR ARTERY WAS IDENTIFIED ON B THE PATIENT UNDERWENT EMBOLIZATION PROCEDURE USING ONYX TO TREAT THE TYPE II ENDOLEAK ON B THE PATIENT UNDERWENT REINTERVENTION FOR A TYPE II ENDOLEAK EMBOLIZATION WAS ATTEMPTED BUT WAS UNABLE TO BE PERFORMED DUE TO POOR ACCESS NO ANEURYSM ENLARGEMENT WAS REPORTED ON B THE PATIENT WAS AT THE HOSPITAL TO CHECK AN ENDOLEAK ON B THE CT A WAS PERFORMED ALSO WAS REPORTED THAT THE PATIENT HOSPITAL STAY WAS EXTENDED DUE TO A COURSE OF PRE AND POST PROCEDURAL IV RENAL PROTECTION THE PHYSICIAN WAS MONITORING THE PATIENT | Hemorrhage | Patient problem code
6789380 | 82554378 | D | IT WAS REPORTED THAT A PATIENT UNDERWENT AN MR OF THE SPINE ON B THE PATIENT LATER CONTACTED THE MR DEPARTMENT ON B AND INFORMED THEM THAT THEY VISITED A LARYNGOLOGIST BECAUSE OF HEARING IMPAIRMENT IN THE LEFT EAR THE PATIENT SAID THAT THEIR HEARING WAS IMPROVED BUT STILL NOT THE SAME AS PRIOR TO THE MR EXAM THE PATIENT HAS NOT PRESENTED OR SENT ANY CONFIRMATION OF HEARING LOSS THAT HAS BEEN DIAGNOSED BY A PHYSICIAN THE PATIENT DID NOT PROVIDE INFORMATION ON WHETHER ANY TREATMENT WAS RECEIVED THE CUSTOMER CONFIRMED THAT HEARING PROTECTION HAD BEEN UTILIZED DURING THE EXAM | Hearing Impairment | Patient problem code
6789380 | 82554378 | D | IT WAS REPORTED THAT A PATIENT UNDERWENT AN MR OF THE SPINE ON B THE PATIENT LATER CONTACTED THE MR DEPARTMENT ON B AND INFORMED THEM THAT THEY VISITED A LARYNGOLOGIST BECAUSE OF HEARING IMPAIRMENT IN THE LEFT EAR THE PATIENT SAID THAT THEIR HEARING WAS IMPROVED BUT STILL NOT THE SAME AS PRIOR TO THE MR EXAM THE PATIENT HAS NOT PRESENTED OR SENT ANY CONFIRMATION OF HEARING LOSS THAT HAS BEEN DIAGNOSED BY A PHYSICIAN THE PATIENT DID NOT PROVIDE INFORMATION ON WHETHER ANY TREATMENT WAS RECEIVED THE CUSTOMER CONFIRMED THAT HEARING PROTECTION HAD BEEN UTILIZED DURING THE EXAM | Insufficient Information | Device problem code
6789732 | 82559162 | D | IT WAS REPORTED THAT AN OPERATING ROOM NURSE RECEIVED A NEEDLE STICK TO HER MIDDLE FINGER AFTER VENIPUNCTURE WHILE USING A G BD VENFLON PRO SAFETY PERIPHERAL SAFETY IV CATHETER THE SAFETY STOPPER WAS FIXED AT THE NEEDLE POINT LAB TESTS PERFORMED RESULTS WERE NEGATIVE NURSE REPORTS SHE WILL FOLLOWUP WITH ADDITIONAL TESTING IN WEEKS | Physical Resistance / Sticking | Device problem code
6789732 | 82559162 | D | IT WAS REPORTED THAT AN OPERATING ROOM NURSE RECEIVED A NEEDLE STICK TO HER MIDDLE FINGER AFTER VENIPUNCTURE WHILE USING A G BD VENFLON PRO SAFETY PERIPHERAL SAFETY IV CATHETER THE SAFETY STOPPER WAS FIXED AT THE NEEDLE POINT LAB TESTS PERFORMED RESULTS WERE NEGATIVE NURSE REPORTS SHE WILL FOLLOWUP WITH ADDITIONAL TESTING IN WEEKS | No Information | Patient problem code
6759010 | 82569534 | D | ON B ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER REPORTED THEY OBTAINED A TROPONIN I RESULT OF NGML ON A B FEMALE PATIENT WHO CAME TO ED COMPLAINING OF SHORTNESS OF BREATH RETURN PRODUCT IS AVAILABLE FOR INVESTIGATION THERE WAS NO ADDITIONAL PATIENT INFORMATION AVAILABLE AT THE TIME OF THIS REPORT PATIENT CARE WAS BASED ON THE NEGATIVE LAB RESULTS TEST DATE METHOD RESULT TIME RESULT B COBAS B ISTAT AT THE TIME OF THE EVENT THERE WAS NO INDICATION OF A PRODUCT MALFUNCTION BASED ON THE INFORMATION PROVIDED AND ASSESSED BY APOC HOWEVER THIS MDR IS A RETROSPECTIVE FILING IN RESPONSE TO AN OBSERVATION FROM AN FDA INSPECTION CONDUCTED MAY TH TO TH AT ABBOTT POINT OF CARE B | Adverse Event Without Identified Device or Use Problem | Device problem code
6759010 | 82569534 | D | ON B ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER REPORTED THEY OBTAINED A TROPONIN I RESULT OF NGML ON A B FEMALE PATIENT WHO CAME TO ED COMPLAINING OF SHORTNESS OF BREATH RETURN PRODUCT IS AVAILABLE FOR INVESTIGATION THERE WAS NO ADDITIONAL PATIENT INFORMATION AVAILABLE AT THE TIME OF THIS REPORT PATIENT CARE WAS BASED ON THE NEGATIVE LAB RESULTS TEST DATE METHOD RESULT TIME RESULT B COBAS B ISTAT AT THE TIME OF THE EVENT THERE WAS NO INDICATION OF A PRODUCT MALFUNCTION BASED ON THE INFORMATION PROVIDED AND ASSESSED BY APOC HOWEVER THIS MDR IS A RETROSPECTIVE FILING IN RESPONSE TO AN OBSERVATION FROM AN FDA INSPECTION CONDUCTED MAY TH TO TH AT ABBOTT POINT OF CARE B | Dyspnea | Patient problem code
6771008 | 82593627 | D | IT WAS REPORTED THAT THE PATIENTS TEMPERATURE DROPPED BELOW THE TARGET TEMPERATURE AT THE TIME OF THE CALL TO MSS THE PATIENT HAD APPROXIMATELY HOURS LEFT IN THE HYPOTHERMIA PHASE OF TREATMENT THE PATIENTS TEMPERATURE WAS C WITH A TARGET OF C AND A WATER TEMPERATURE OF C THE WATER TEMPERATURE HAD BEEN TRENDING IN THE MID TO HIGH TWENTIES THE FLOW RATE WAS OPTIMAL NO RECENT ALARMS WERE REPORTED BY THE NURSE THE NURSE TESTED THE HEATER IN MANUAL MODE AND THE WATER TEMPERATURE ONLY INCREASED TO C AFTER MINUTES MSS INSTRUCTED THE NURSE TO STOP THE DEVICE IT WAS LATER REPORTED THAT THE NURSE CONTINUED THERAPY ON THE DEVICE THE PATIENTS TEMPERATURE DECREASED TO C AT ONE POINT DURING THE HYPOTHERMIA PHASE OF THERAPY THE NURSE REPORTED THAT THEY HAD CHECKED THE DEVICE IN MANUAL MODE TO ENSURE THAT IT WAS ABLE TO HEAT AND COOL AND DETERMINED THAT THE DEVICE WAS ABLE TO HEAT AND COOL THE WATER THE NURSE STATED THAT HE ADJUSTED THE TARGET TEMPERATURE TO HELP THE PATIENT ACHIEVE THE TARGET TEMPERATURE THE PATIENT WAS ABLE TO COMPLETE THE REWARM PHASE OF THERAPY THE NURSE REPORTED THAT REWARM TOOK APPROXIMATELY HOURS WHEN IT WAS EXPECTED TO BE COMPLETED IN HOURS THE PATIENT WAS IN THE NORMOTHERMIA PHASE OF TREATMENT WHEN THE PATIENT EXPIRED PER THE NURSE THE PATIENT EXPIRED DUE TO NATURAL CAUSES THE NURSE REPORTED THAT THE DEATH WAS DUE TO PATIENT CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6771008 | 82593627 | D | IT WAS REPORTED THAT THE PATIENTS TEMPERATURE DROPPED BELOW THE TARGET TEMPERATURE AT THE TIME OF THE CALL TO MSS THE PATIENT HAD APPROXIMATELY HOURS LEFT IN THE HYPOTHERMIA PHASE OF TREATMENT THE PATIENTS TEMPERATURE WAS C WITH A TARGET OF C AND A WATER TEMPERATURE OF C THE WATER TEMPERATURE HAD BEEN TRENDING IN THE MID TO HIGH TWENTIES THE FLOW RATE WAS OPTIMAL NO RECENT ALARMS WERE REPORTED BY THE NURSE THE NURSE TESTED THE HEATER IN MANUAL MODE AND THE WATER TEMPERATURE ONLY INCREASED TO C AFTER MINUTES MSS INSTRUCTED THE NURSE TO STOP THE DEVICE IT WAS LATER REPORTED THAT THE NURSE CONTINUED THERAPY ON THE DEVICE THE PATIENTS TEMPERATURE DECREASED TO C AT ONE POINT DURING THE HYPOTHERMIA PHASE OF THERAPY THE NURSE REPORTED THAT THEY HAD CHECKED THE DEVICE IN MANUAL MODE TO ENSURE THAT IT WAS ABLE TO HEAT AND COOL AND DETERMINED THAT THE DEVICE WAS ABLE TO HEAT AND COOL THE WATER THE NURSE STATED THAT HE ADJUSTED THE TARGET TEMPERATURE TO HELP THE PATIENT ACHIEVE THE TARGET TEMPERATURE THE PATIENT WAS ABLE TO COMPLETE THE REWARM PHASE OF THERAPY THE NURSE REPORTED THAT REWARM TOOK APPROXIMATELY HOURS WHEN IT WAS EXPECTED TO BE COMPLETED IN HOURS THE PATIENT WAS IN THE NORMOTHERMIA PHASE OF TREATMENT WHEN THE PATIENT EXPIRED PER THE NURSE THE PATIENT EXPIRED DUE TO NATURAL CAUSES THE NURSE REPORTED THAT THE DEATH WAS DUE TO PATIENT CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Death | Patient problem code
6771008 | 82593627 | D | IT WAS REPORTED THAT THE PATIENTS TEMPERATURE DROPPED BELOW THE TARGET TEMPERATURE AT THE TIME OF THE CALL TO MSS THE PATIENT HAD APPROXIMATELY HOURS LEFT IN THE HYPOTHERMIA PHASE OF TREATMENT THE PATIENTS TEMPERATURE WAS C WITH A TARGET OF C AND A WATER TEMPERATURE OF C THE WATER TEMPERATURE HAD BEEN TRENDING IN THE MID TO HIGH TWENTIES THE FLOW RATE WAS OPTIMAL NO RECENT ALARMS WERE REPORTED BY THE NURSE THE NURSE TESTED THE HEATER IN MANUAL MODE AND THE WATER TEMPERATURE ONLY INCREASED TO C AFTER MINUTES MSS INSTRUCTED THE NURSE TO STOP THE DEVICE IT WAS LATER REPORTED THAT THE NURSE CONTINUED THERAPY ON THE DEVICE THE PATIENTS TEMPERATURE DECREASED TO C AT ONE POINT DURING THE HYPOTHERMIA PHASE OF THERAPY THE NURSE REPORTED THAT THEY HAD CHECKED THE DEVICE IN MANUAL MODE TO ENSURE THAT IT WAS ABLE TO HEAT AND COOL AND DETERMINED THAT THE DEVICE WAS ABLE TO HEAT AND COOL THE WATER THE NURSE STATED THAT HE ADJUSTED THE TARGET TEMPERATURE TO HELP THE PATIENT ACHIEVE THE TARGET TEMPERATURE THE PATIENT WAS ABLE TO COMPLETE THE REWARM PHASE OF THERAPY THE NURSE REPORTED THAT REWARM TOOK APPROXIMATELY HOURS WHEN IT WAS EXPECTED TO BE COMPLETED IN HOURS THE PATIENT WAS IN THE NORMOTHERMIA PHASE OF TREATMENT WHEN THE PATIENT EXPIRED PER THE NURSE THE PATIENT EXPIRED DUE TO NATURAL CAUSES THE NURSE REPORTED THAT THE DEATH WAS DUE TO PATIENT CONDITION AND DID NOT ALLEGE THE DEATH AGAINST THE DEVICE CLINICAL STATEMENT THE PATIENTS REPORTED DEATH WAS AN INCIDENTAL ELEMENT OF THE PATIENTS MEDICAL HISTORY AND IS UNRELATED TO THE REASON FOR THE COMPLAINT THERE IS NO INDICATION REPORT OR ALLEGATION THAT THE DEVICE MALFUNCTION WAS RELATED TO THE PATIENTS DEATH | Hypothermia | Patient problem code
6792267 | 82618494 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR OF THE INTRAAORTIC BALLOON PUMP IABP INVOLVED IN THE EVENT WAS REVIEWED THERE WERE NO NONCONFORMANCES IN THE PRODUCTION DHR RELATED TO THE REPORTED EVENT ON JULY THE CUSTOMER BIOMED REPORTED THAT THE IABP UNIT HAS BEEN QUARANTINED IN THE OFFICE SINCE THE EVENT THEY TRIED EXCHANGING THE SAFETY DISK AND BALLOON CATHETER WITH KNOWN GOOD ONES BUT THE ISSUE PERSISTED REPAIRS WILL BE MADE AT THE DISCRETION OF THE DEPARTMENT IF ADDITIONAL INFORMATION IS PROVIDED WE WILL REPORT THE INFORMATION ACCORDINGLY | No Findings Available | Medical Evaluation result code
6773300 | 82634326 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A FIELD SERVICE ENGINEER FSE STATED THAT THE CUSTOMER REPORTED THE IABP SHUT DOWN WHILE IN USE THE FSE INDICATED THAT HE FOUND ERROR CODE WHICH INDICATES CONDENSATION OR MOISTURE IN THE LINE AND REQUIRES PURGING THE LINE THE FSE PERFORMED THE CONDENSATION REMOVAL PROCEDURE AS WELL AS A COMPLETE CHECKOUT AND CALIBRATION THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS THE IABP WAS RETURNED TO THE CUSTOMER AND CLEARED FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6784309 | 82634415 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS UNABLE TO BE REVIEWED AS WE WERE NOT ABLE TO OBTAIN THE SERIAL NUMBER FOR THE IABP ASSOCIATED WITH THIS COMPLAINT THE COMPANY REPRESENTATIVE FROM THE BUSINESS UNIT BU STATED THE INTRAAORTIC BALLOON PUMP IABP WAS NOT THE ISSUE IT WAS DETERMINED THAT THE PROBLEM WAS THE TWO GAS BOTTLES WHICH WERE DELIVERED EMPTY | No Findings Available | Medical Evaluation result code
6793515 | 82654649 | D | IT WAS REPORTED PATIENT WAS DISCHARGED FROM THE HOSPITAL FOLLOWING A REVISION PROCEDURE ON ANTIBIOTICS FOR PRESUMED ASPIRATION PNEUMONIA ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE HOWEVER NO MORE IS AVAILABLE AT THIS TIME | Pneumonia | Patient problem code
6793515 | 82654649 | D | IT WAS REPORTED PATIENT WAS DISCHARGED FROM THE HOSPITAL FOLLOWING A REVISION PROCEDURE ON ANTIBIOTICS FOR PRESUMED ASPIRATION PNEUMONIA ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE HOWEVER NO MORE IS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
6793629 | 82654934 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH ULTRAMINI METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B AT PM SHE OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND ON B SHE ADMINISTERED UNITS OF HUMALOG INSULIN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF SHAKY AND SWEATY AND THAT SHE SELFTREATED WITH GLUCOSE TABLETS DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6793629 | 82654934 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH ULTRAMINI METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B AT PM SHE OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND ON B SHE ADMINISTERED UNITS OF HUMALOG INSULIN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF SHAKY AND SWEATY AND THAT SHE SELFTREATED WITH GLUCOSE TABLETS DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Shaking/Tremors | Patient problem code
6793629 | 82654934 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH ULTRAMINI METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B AT PM SHE OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND ON B SHE ADMINISTERED UNITS OF HUMALOG INSULIN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF SHAKY AND SWEATY AND THAT SHE SELFTREATED WITH GLUCOSE TABLETS DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
6793629 | 82654934 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH ULTRAMINI METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B AT PM SHE OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND ON B SHE ADMINISTERED UNITS OF HUMALOG INSULIN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF SHAKY AND SWEATY AND THAT SHE SELFTREATED WITH GLUCOSE TABLETS DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Sweating | Patient problem code
6793513 | 82655289 | D | IT WAS REPORTED PATIENT WAS DISCHARGED FROM THE HOSPITAL FOLLOWING A REVISION PROCEDURE ON ANTIBIOTICS FOR PRESUMED ASPIRATION PNEUMONIA ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE HOWEVER NO MORE IS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
6793513 | 82655289 | D | IT WAS REPORTED PATIENT WAS DISCHARGED FROM THE HOSPITAL FOLLOWING A REVISION PROCEDURE ON ANTIBIOTICS FOR PRESUMED ASPIRATION PNEUMONIA ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE HOWEVER NO MORE IS AVAILABLE AT THIS TIME | Aspiration | Patient Problem Code
6793611 | 82656392 | D | A PERITONEAL DIALYSIS PD PATIENTS CONTACT STATED THE PD PATIENT RECENTLY HAD GOTTEN OUT OF THE HOSPITAL AND HAD SWITCHED FROM USING TO DEXTROSE BAGS FOR PD TREATMENTS EVER SINCE SHE WAS IN THE HOSPITAL FOLLOWUP WITH THE PD PATIENT REVEALED SHE HAD DIABETES AND HAD BECOME HYPOGLYCEMIC AND WAS TAKEN TO THE HOSPITAL AND ADMITTED ON B THE PD PATIENT STATED SHE CONTINUED COMPLETING CONTINUOUS CYCLERASSISTED PERITONEAL DIALYSIS CCPD THERAPY WHILE HOSPITALIZED AND WAS DISCHARGED ON B ADDITIONAL FOLLOWUP WITH THE PD PATIENTS CLINIC REGISTERED NURSE RN REVEALED THE PATIENT WAS A KNOWN DIABETIC AND HAD TAKEN HER INSULIN ON B BUT HAD NOT EATEN WHICH CAUSED A DROP IN HER BLOOD SUGAR LEVELS THE PD RN STATED THE PATIENT ENDED UP BEING TAKEN TO THE HOSPITAL AND CONFIRMED THE PATIENT WAS HOSPITALIZED FOR HYPOGLYCEMIA FROM B PER PDRN THE PATIENTS HOSPITALIZED EVENT WAS UNRELATED TO THE PATIENTS PERITONEAL DIALYSIS REGIMEN OR THE PATIENTS CYCLER AND STATED AS OF B THE PATIENT WAS ABLE TO CONTINUE TO USE CONTINUOUS CYCLERASSISTED PERITONEAL DIALYSIS CCPD THERAPY WITH THE CYCLER WITHOUT FURTHER ISSUE MEDICAL RECORDS WERE REQUESTED | Insufficient Information | Device problem code
6793611 | 82656392 | D | A PERITONEAL DIALYSIS PD PATIENTS CONTACT STATED THE PD PATIENT RECENTLY HAD GOTTEN OUT OF THE HOSPITAL AND HAD SWITCHED FROM USING TO DEXTROSE BAGS FOR PD TREATMENTS EVER SINCE SHE WAS IN THE HOSPITAL FOLLOWUP WITH THE PD PATIENT REVEALED SHE HAD DIABETES AND HAD BECOME HYPOGLYCEMIC AND WAS TAKEN TO THE HOSPITAL AND ADMITTED ON B THE PD PATIENT STATED SHE CONTINUED COMPLETING CONTINUOUS CYCLERASSISTED PERITONEAL DIALYSIS CCPD THERAPY WHILE HOSPITALIZED AND WAS DISCHARGED ON B ADDITIONAL FOLLOWUP WITH THE PD PATIENTS CLINIC REGISTERED NURSE RN REVEALED THE PATIENT WAS A KNOWN DIABETIC AND HAD TAKEN HER INSULIN ON B BUT HAD NOT EATEN WHICH CAUSED A DROP IN HER BLOOD SUGAR LEVELS THE PD RN STATED THE PATIENT ENDED UP BEING TAKEN TO THE HOSPITAL AND CONFIRMED THE PATIENT WAS HOSPITALIZED FOR HYPOGLYCEMIA FROM B PER PDRN THE PATIENTS HOSPITALIZED EVENT WAS UNRELATED TO THE PATIENTS PERITONEAL DIALYSIS REGIMEN OR THE PATIENTS CYCLER AND STATED AS OF B THE PATIENT WAS ABLE TO CONTINUE TO USE CONTINUOUS CYCLERASSISTED PERITONEAL DIALYSIS CCPD THERAPY WITH THE CYCLER WITHOUT FURTHER ISSUE MEDICAL RECORDS WERE REQUESTED | Hypoglycemia | Patient problem code
6760984 | 82658743 | N | B THE ANALYSIS FOUND THAT ONE SR RELOAD WAS RETURNED WITH THE KNIFE NOT IN THE DOGHOUSE THIS IS CONSISTENT WITH AN IMPROPER LOADING TECHNIQUE THE CARTRIDGE RELOAD HAD THE SWING TAB IN THE LOCKED POSITION AND THE PROXIMAL DRIVERS UP AND THE REMAINING DRIVERS WERE DOWN WITH STAPLES PRESENT THIS CONDITION IS CONSISTENT WITH AN IMPROPER LOADING TECHNIQUE WHEN LOADING THE CARTRIDGE IN THE DEVICE MAKE SURE THE CARTRIDGE IS INSIDE THE CHANNEL TRACK AND THE PROPER LOADING TECHNIQUE IS BEING USED PLEASE REFER INSTRUCTIONS FOR USE THE CARTRIDGE RELOAD SWING TAB WAS RESET AND THE KNIFE WAS MANUALLY RETURNED IN ORDER TO FIRE THE CARTRIDGE THE CARTRIDGE WAS TESTED FOR FUNCTIONALITY WITH A TEST DEVICE AND FIRED WITHOUT ANY DIFFICULTIES THE STAPLE LINE AND CUT LINE WERE COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER SHAPE THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | No Device Problem Found | Medical Evaluation result code
6788204 | 82661636 | N | ALL INSTRUMENTS AFFECTED WERE REPROCESSED PRIOR TO USE IN PATIENT PROCEDURES THE USER FACILITY STATED AN EMPLOYEE STRUGGLED TO PUSH THE LOADING CART AFTER REMOVING AN INSTRUMENT RACK FROM THE STERILIZER THIS EMPLOYEE WAS UNABLE TO PUSH THE LOADING CART FROM THE HANDLE SIDE AND SUBSEQUENTLY TRIED TO PULL THE CART FROM THE FRONT WHICH RESULTED IN THE CART TIPPING AND AN INSTRUMENT RACK FALLING TO THE GROUND THE USER FACILITY STATED THE EMPLOYEE SUSTAINED MINOR INJURIES AS SHE MOVED TO AVOID THE FALLING INSTRUMENT RACK THE EMPLOYEE SELFTREATED AND RETURNED TO WORK THE EVOLUTION LOADING CART OPERATOR MANUAL STATES ON PAGE WARNING PREVENT PHYSICAL INJURY GET HELP IF LOADING CAR BECOMES STUCK OR DIFFICULT TO MOVE DO NOT ATTEMPT TO LIFT A LOADING CAR WITHOUT ASSISTANCE A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE LOADING CART SUBJECT OF THE REPORTED EVENT THE TECHNICIAN CONFIRMED THE LOADING CART WAS OPERATING PROPERLY THE TECHNICIAN FOUND THAT THE USER FACILITY HAS THREE EVOLUTION STERILIZERS THAT SIT AT DIFFERENT HEIGHTS DUE TO UNEVEN FLOORS AT THE USER FACILITY THE EVOLUTION LOADING CART OPERATOR MANUAL STATES ON PAGE PROPER TRANSFER CARRIAGE HEIGHT IS REQUIRED WHEN PLACING LOADING CAR IN STERILIZER CHAMBER THE TECHNICIAN FOUND THAT THE LOADING CART SUBJECT OF THE REPORTED EVENT WAS BEING USED WITH AN EVOLUTION STERILIZER FOR WHICH IT DID NOT PROPERLY ALIGN THE EVOLUTION STERILIZER OPERATOR MANUAL STATES ON PAGE VERIFY LOADING CAR IS PROPERLY ALIGNED UPON LOADING TRANSFER CARRIAGE THE TECHNICIAN FOUND THAT THE DIFFERENCES IN HEIGHT BETWEEN THE STERILIZER AND LOADING CART CAUSED THE INSTRUMENT RACK TO BE LOADED ONTO THE CART IMPROPERLY THIS CAUSED THE INSTRUMENT RACK TO LEAN WHILE SITTING ON THE LOADING CART MAKING IT DIFFICULT TO TRANSFER DUE TO THE UNEVEN WEIGHT DISTRIBUTION AND RESULTED IN THE REPORTED EVENT THE TECHNICIAN PERFORMED AN INSERVICE TRAINING WITH USER FACILITY PERSONNEL ON THE PROPER USE AND OPERATION OF THE EVOLUTION LOADING CART WITH THEIR EVOLUTION STERILIZERS THE TECHNICIAN SPECIFICALLY NOTED THE REQUIREMENT OF THE LOADING CART AND STERILIZER TO BE SET AT THE SAME HEIGHT IT WAS RECOMMENDED TO THE USER FACILITY TO LABEL EACH LOADING CART WITH A CORRESPONDING EVOLUTION STERILIZER TO PREVENT THIS ISSUE FROM RECURRING NO ADDITIONAL ISSUES HAVE BEEN REPORTED WITH THE EVOLUTION LOADING CARTS OR EVOLUTION STERILIZERS AT THE USER FACILITY | Problem with Sterilization | Device problem code
6785384 | 82665999 | N | ALL INSTRUMENTS AFFECTED WERE REPROCESSED PRIOR TO USE IN PATIENT PROCEDURES THE USER FACILITY STATED THE INSTRUMENT PACK WRAP WAS GETTING CAUGHT WITHIN THE STERILIZER DURING THE UNLOADING PROCESS CAUSING THEM TO RIP A STERIS SERVICE TECHNICIAN ARRIVED ONSITE INSPECTED THE UNIT AND CONFIRMED THE UNIT TO BE OPERATING ACCORDING TO SPECIFICATION THE TECHNICIAN IDENTIFIED USER FACILITY PERSONNEL WERE NOT PROPERLY UNLOADING THE STERILIZER RESULTING IN THE REPORTED EVENT THE VPRO MAX STERILIZER OPERATOR MANUAL STATES THE PROPER UNLOADING PRACTICES FOR THE STERILIZER IN SECTION ON PAGE THE TECHNICIAN NOTIFIED THE USER FACILITY OF THE PROPER UNLOADING PRACTICES WHEN USING THE VPRO MAX STERILIZER SPECIFICALLY HOW TO AVOID RIPPING THE WRAPPING ON THE INSTRUMENT PACKS NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Sterilization Problem Identified | Medical Evaluation result code
6785396 | 82666319 | N | NO REPORT OF DELAYS OR CANCELLATIONS DURING THE TIME OF THE REPORTED EVENT THE VPRO PLUS STERILIZER WAS RUNNING A CYCLE WHEN THE UNIT BEGAN TO PRODUCE AN ODOR THROUGHOUT THE ROOM THE UNIT DID NOT ALARM AND THE CYCLE COMPLETED SUCCESSFULLY IT IS UNKNOWN IF THE EMPLOYEE SUBJECT OF THE REPORTED SOUGHT OR RECEIVED MEDICAL TREATMENT A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THE UNIT WAS OPERATING ACCORDING TO SPECIFICATIONS THE TECHNICIAN DID NOT DETECT AN ODOR DURING A TEST CYCLE THE TECHNICIAN TESTED THE STERILIZER AND RETURNED THE UNIT TO SERVICE THE VPRO PLUS STERILIZER WAS MANUFACTURED IN AND IS UNDER STERIS SERVICE AGREEMENT NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Device Emits Odor | Device problem code
6785396 | 82666319 | N | NO REPORT OF DELAYS OR CANCELLATIONS DURING THE TIME OF THE REPORTED EVENT THE VPRO PLUS STERILIZER WAS RUNNING A CYCLE WHEN THE UNIT BEGAN TO PRODUCE AN ODOR THROUGHOUT THE ROOM THE UNIT DID NOT ALARM AND THE CYCLE COMPLETED SUCCESSFULLY IT IS UNKNOWN IF THE EMPLOYEE SUBJECT OF THE REPORTED SOUGHT OR RECEIVED MEDICAL TREATMENT A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THE UNIT WAS OPERATING ACCORDING TO SPECIFICATIONS THE TECHNICIAN DID NOT DETECT AN ODOR DURING A TEST CYCLE THE TECHNICIAN TESTED THE STERILIZER AND RETURNED THE UNIT TO SERVICE THE VPRO PLUS STERILIZER WAS MANUFACTURED IN AND IS UNDER STERIS SERVICE AGREEMENT NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Device Alarm System | Device problem code
6785396 | 82666319 | N | NO REPORT OF DELAYS OR CANCELLATIONS DURING THE TIME OF THE REPORTED EVENT THE VPRO PLUS STERILIZER WAS RUNNING A CYCLE WHEN THE UNIT BEGAN TO PRODUCE AN ODOR THROUGHOUT THE ROOM THE UNIT DID NOT ALARM AND THE CYCLE COMPLETED SUCCESSFULLY IT IS UNKNOWN IF THE EMPLOYEE SUBJECT OF THE REPORTED SOUGHT OR RECEIVED MEDICAL TREATMENT A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THE UNIT WAS OPERATING ACCORDING TO SPECIFICATIONS THE TECHNICIAN DID NOT DETECT AN ODOR DURING A TEST CYCLE THE TECHNICIAN TESTED THE STERILIZER AND RETURNED THE UNIT TO SERVICE THE VPRO PLUS STERILIZER WAS MANUFACTURED IN AND IS UNDER STERIS SERVICE AGREEMENT NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Operational Problem Identified | Medical Evaluation result code
6789227 | 82667121 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6793771 | 82670028 | N | UDI UNAVAILABLE | Insufficient Information | Device Problem Code
6793751 | 82672735 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF HARDENING LUMPY AND NODULAR ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS WARNINGS PRODUCT USE AT SPECIFIC SITES IN WHICH AN ACTIVE INFLAMMATORY PROCESS SKIN ERUPTIONS SUCH AS CYSTS PIMPLES RASHES OR HIVES OR INFECTION IS PRESENT SHOULD BE DEFERRED UNTIL THE UNDERLYING PROCESS HAS BEEN CONTROLLED INJECTION SITE RESPONSES CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND LASTING DAYS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLES AND INJECTION SITE RESPONSES BY SEVERITY AND DURATION AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS HEALTH CARE PROFESSIONAL INSTRUCTIONS THE INJECTION TECHNIQUE MAY VARY WITH REGARD TO THE ANGLE AND ORIENTATION OF THE BEVEL THE DEPTH OF INJECTION AND THE QUANTITY ADMINISTERED A TUNNELING TECHNIQUE SERIAL PUNCTURE TECHNIQUE FANNING TECHNIQUE OR A COMBINATION HAS BEEN USED TO ACHIEVE OPTIMAL RESULTS INJECTING THE PRODUCT TOO SUPERFICIALLY MAY RESULT IN VISIBLE LUMPS ANDOR DISCOLORATION | No Findings Available | Medical Evaluation result code
6793885 | 82677682 | N | NO PATIENT INFORMATION PROVIDED AS NO PATIENT WAS INVOLVED IN THIS CONCERN A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT THE NAVIGATION SYSTEM WOULD BECOME UNRESPONSIVE WHEN LOADING A BLANK CD HOWEVER WHEN THE CD WAS NOT BLANK IT WOULD LOAD WITHOUT ISSUE THE HARDWARE SOFTWARE AND INSTRUMENTS PASSED THE SYSTEM CHECKOUT THE SYSTEM WAS FOUND TO BE FULLY FUNCTIONAL | Component or Accessory Incompatibility | Device problem code
6793885 | 82677682 | N | NO PATIENT INFORMATION PROVIDED AS NO PATIENT WAS INVOLVED IN THIS CONCERN A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT THE NAVIGATION SYSTEM WOULD BECOME UNRESPONSIVE WHEN LOADING A BLANK CD HOWEVER WHEN THE CD WAS NOT BLANK IT WOULD LOAD WITHOUT ISSUE THE HARDWARE SOFTWARE AND INSTRUMENTS PASSED THE SYSTEM CHECKOUT THE SYSTEM WAS FOUND TO BE FULLY FUNCTIONAL | No Device Problem Found | Medical Evaluation result code
6786730 | 82679239 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE COMPANY REPRESENTATIVE FROM THE BUSINESS UNIT BU REPORTED THAT THEY TRIED TO FIX THE CONNECTION PROBLEM BY ADJUSTING THE POSITION OF THE AC AND EXTERNAL CONNECTORS BUT STILL THE PROBLEM PERSISTED THEY TRIED TO ADJUST LOCKING METAL BY ADJUSTING THE PLATE THE COMPANY REPRESENTATIVE REPORTED THAT THE MENTIONED COMPLAINT HAS BEEN RECTIFIED BY ADJUSTING THE CONSOLE CHASSIS AND LOCK IN THE CART ALL PREVENTATIVE MAINTENANCE PM CHECKS WERE PERFORMED AND THE IABP UNIT HAS PASSED ALL TESTS | Expiration Date Error | Device problem code
6786730 | 82679239 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE COMPANY REPRESENTATIVE FROM THE BUSINESS UNIT BU REPORTED THAT THEY TRIED TO FIX THE CONNECTION PROBLEM BY ADJUSTING THE POSITION OF THE AC AND EXTERNAL CONNECTORS BUT STILL THE PROBLEM PERSISTED THEY TRIED TO ADJUST LOCKING METAL BY ADJUSTING THE PLATE THE COMPANY REPRESENTATIVE REPORTED THAT THE MENTIONED COMPLAINT HAS BEEN RECTIFIED BY ADJUSTING THE CONSOLE CHASSIS AND LOCK IN THE CART ALL PREVENTATIVE MAINTENANCE PM CHECKS WERE PERFORMED AND THE IABP UNIT HAS PASSED ALL TESTS | No Device Problem Found | Medical Evaluation result code
6794052 | 82680258 | D | IT WAS REPORTED THAT A PATIENT WAS NOTED TO BE ANEMIC WHEN SEEN BY HOME HEALTH THE PATIENT WAS CALLED AND DIRECTED TO COME IN AS A DIRECT ADMIT WHEN SEEN EARLY ON THE MORNING OF B THAT PATIENT WAS ASYMPTOMATIC WITH NO SHORTNESS OF BREATH AND NO ORTHOSTATIC SYMPTOMS A COMPUTERIZED TOMOGRAPHY CT OF THE PATIENTS CHEST ABDOMEN AND PELVIS WAS NOT ABLE TO INDICATE THE CAUSE OF THE PATIENTS ANEMIA HE WAS TREATED WITH FOUR UNITS OF PACKED CELLS WITH IMPROVEMENT IN HIS HEMOGLOBIN HGB AND HEMATOCRIT HCT OF NOTE THE PATIENT IS REPORTED TO BE CHRONICALLY ILL WITH ONGOING DRIVELINE DL INFECTION RECURRENT SEPSIS CHRONIC RENAL INSUFFICIENCY AND CHRONIC PAIN HE REQUESTED TO BE MADE A DO NOT RESUSCITATEINTERVENE AND HIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS TURNED OFF HE WAS DISCHARGED HOME ON B WITH NO EVIDENCE OF BLEEDING THE PATIENTS MEDICAL TEAM FEELS THAT THE PATIENTS ANEMIA IS RELATED TO HIS CHRONIC ILLNESS ANDOR A SLOW GASTROINTESTINAL GI BLEED THE EVENT WAS REPORTED TO HAVE BEEN RESOLVED ON B THE PRIMARY INVESTIGATOR REPORTED THAT THE EVENT WAS RELATED TO THE PATIENTS CONDITION AND UNRELATED TO THE STUDY DEVICE OR TO THE IMPLANT PROCEDURE NO FURTHER INFORMATION HAS BEEN PROVIDED | Adverse Event Without Identified Device or Use Problem | Device problem code
6794052 | 82680258 | D | IT WAS REPORTED THAT A PATIENT WAS NOTED TO BE ANEMIC WHEN SEEN BY HOME HEALTH THE PATIENT WAS CALLED AND DIRECTED TO COME IN AS A DIRECT ADMIT WHEN SEEN EARLY ON THE MORNING OF B THAT PATIENT WAS ASYMPTOMATIC WITH NO SHORTNESS OF BREATH AND NO ORTHOSTATIC SYMPTOMS A COMPUTERIZED TOMOGRAPHY CT OF THE PATIENTS CHEST ABDOMEN AND PELVIS WAS NOT ABLE TO INDICATE THE CAUSE OF THE PATIENTS ANEMIA HE WAS TREATED WITH FOUR UNITS OF PACKED CELLS WITH IMPROVEMENT IN HIS HEMOGLOBIN HGB AND HEMATOCRIT HCT OF NOTE THE PATIENT IS REPORTED TO BE CHRONICALLY ILL WITH ONGOING DRIVELINE DL INFECTION RECURRENT SEPSIS CHRONIC RENAL INSUFFICIENCY AND CHRONIC PAIN HE REQUESTED TO BE MADE A DO NOT RESUSCITATEINTERVENE AND HIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS TURNED OFF HE WAS DISCHARGED HOME ON B WITH NO EVIDENCE OF BLEEDING THE PATIENTS MEDICAL TEAM FEELS THAT THE PATIENTS ANEMIA IS RELATED TO HIS CHRONIC ILLNESS ANDOR A SLOW GASTROINTESTINAL GI BLEED THE EVENT WAS REPORTED TO HAVE BEEN RESOLVED ON B THE PRIMARY INVESTIGATOR REPORTED THAT THE EVENT WAS RELATED TO THE PATIENTS CONDITION AND UNRELATED TO THE STUDY DEVICE OR TO THE IMPLANT PROCEDURE NO FURTHER INFORMATION HAS BEEN PROVIDED | Infection | Patient problem code
6794052 | 82680258 | D | IT WAS REPORTED THAT A PATIENT WAS NOTED TO BE ANEMIC WHEN SEEN BY HOME HEALTH THE PATIENT WAS CALLED AND DIRECTED TO COME IN AS A DIRECT ADMIT WHEN SEEN EARLY ON THE MORNING OF B THAT PATIENT WAS ASYMPTOMATIC WITH NO SHORTNESS OF BREATH AND NO ORTHOSTATIC SYMPTOMS A COMPUTERIZED TOMOGRAPHY CT OF THE PATIENTS CHEST ABDOMEN AND PELVIS WAS NOT ABLE TO INDICATE THE CAUSE OF THE PATIENTS ANEMIA HE WAS TREATED WITH FOUR UNITS OF PACKED CELLS WITH IMPROVEMENT IN HIS HEMOGLOBIN HGB AND HEMATOCRIT HCT OF NOTE THE PATIENT IS REPORTED TO BE CHRONICALLY ILL WITH ONGOING DRIVELINE DL INFECTION RECURRENT SEPSIS CHRONIC RENAL INSUFFICIENCY AND CHRONIC PAIN HE REQUESTED TO BE MADE A DO NOT RESUSCITATEINTERVENE AND HIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS TURNED OFF HE WAS DISCHARGED HOME ON B WITH NO EVIDENCE OF BLEEDING THE PATIENTS MEDICAL TEAM FEELS THAT THE PATIENTS ANEMIA IS RELATED TO HIS CHRONIC ILLNESS ANDOR A SLOW GASTROINTESTINAL GI BLEED THE EVENT WAS REPORTED TO HAVE BEEN RESOLVED ON B THE PRIMARY INVESTIGATOR REPORTED THAT THE EVENT WAS RELATED TO THE PATIENTS CONDITION AND UNRELATED TO THE STUDY DEVICE OR TO THE IMPLANT PROCEDURE NO FURTHER INFORMATION HAS BEEN PROVIDED | Sepsis | Patient problem code
6794052 | 82680258 | D | IT WAS REPORTED THAT A PATIENT WAS NOTED TO BE ANEMIC WHEN SEEN BY HOME HEALTH THE PATIENT WAS CALLED AND DIRECTED TO COME IN AS A DIRECT ADMIT WHEN SEEN EARLY ON THE MORNING OF B THAT PATIENT WAS ASYMPTOMATIC WITH NO SHORTNESS OF BREATH AND NO ORTHOSTATIC SYMPTOMS A COMPUTERIZED TOMOGRAPHY CT OF THE PATIENTS CHEST ABDOMEN AND PELVIS WAS NOT ABLE TO INDICATE THE CAUSE OF THE PATIENTS ANEMIA HE WAS TREATED WITH FOUR UNITS OF PACKED CELLS WITH IMPROVEMENT IN HIS HEMOGLOBIN HGB AND HEMATOCRIT HCT OF NOTE THE PATIENT IS REPORTED TO BE CHRONICALLY ILL WITH ONGOING DRIVELINE DL INFECTION RECURRENT SEPSIS CHRONIC RENAL INSUFFICIENCY AND CHRONIC PAIN HE REQUESTED TO BE MADE A DO NOT RESUSCITATEINTERVENE AND HIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS TURNED OFF HE WAS DISCHARGED HOME ON B WITH NO EVIDENCE OF BLEEDING THE PATIENTS MEDICAL TEAM FEELS THAT THE PATIENTS ANEMIA IS RELATED TO HIS CHRONIC ILLNESS ANDOR A SLOW GASTROINTESTINAL GI BLEED THE EVENT WAS REPORTED TO HAVE BEEN RESOLVED ON B THE PRIMARY INVESTIGATOR REPORTED THAT THE EVENT WAS RELATED TO THE PATIENTS CONDITION AND UNRELATED TO THE STUDY DEVICE OR TO THE IMPLANT PROCEDURE NO FURTHER INFORMATION HAS BEEN PROVIDED | Renal Failure | Patient problem code
6794052 | 82680258 | D | IT WAS REPORTED THAT A PATIENT WAS NOTED TO BE ANEMIC WHEN SEEN BY HOME HEALTH THE PATIENT WAS CALLED AND DIRECTED TO COME IN AS A DIRECT ADMIT WHEN SEEN EARLY ON THE MORNING OF B THAT PATIENT WAS ASYMPTOMATIC WITH NO SHORTNESS OF BREATH AND NO ORTHOSTATIC SYMPTOMS A COMPUTERIZED TOMOGRAPHY CT OF THE PATIENTS CHEST ABDOMEN AND PELVIS WAS NOT ABLE TO INDICATE THE CAUSE OF THE PATIENTS ANEMIA HE WAS TREATED WITH FOUR UNITS OF PACKED CELLS WITH IMPROVEMENT IN HIS HEMOGLOBIN HGB AND HEMATOCRIT HCT OF NOTE THE PATIENT IS REPORTED TO BE CHRONICALLY ILL WITH ONGOING DRIVELINE DL INFECTION RECURRENT SEPSIS CHRONIC RENAL INSUFFICIENCY AND CHRONIC PAIN HE REQUESTED TO BE MADE A DO NOT RESUSCITATEINTERVENE AND HIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS TURNED OFF HE WAS DISCHARGED HOME ON B WITH NO EVIDENCE OF BLEEDING THE PATIENTS MEDICAL TEAM FEELS THAT THE PATIENTS ANEMIA IS RELATED TO HIS CHRONIC ILLNESS ANDOR A SLOW GASTROINTESTINAL GI BLEED THE EVENT WAS REPORTED TO HAVE BEEN RESOLVED ON B THE PRIMARY INVESTIGATOR REPORTED THAT THE EVENT WAS RELATED TO THE PATIENTS CONDITION AND UNRELATED TO THE STUDY DEVICE OR TO THE IMPLANT PROCEDURE NO FURTHER INFORMATION HAS BEEN PROVIDED | Pain | Patient problem code
6794080 | 82681336 | D | THE REPORTED VALVE WAS IMPLANTED TO A PATIENT VIA VPSHUNT WITH THE SETTING OF B AFTER DISCHARGED FROM THE HOSPITAL THE PATIENT CLAIMED VOMITING AFTER THE PATIENT TOOK A BATH AT HOME SO THAT THE PATIENT VISITED TO THE HOSPITAL ON B THE SURGEON CHECKED THE PRESSURE SETTING THEN THE SURGEON FOUND THAT THE SETTING WAS UNINTENTIONALLY CHANGED FROM TO AFTER CHANGED THE PRESSURE SETTING BACK TO THE PATIENTS CONDITION WAS IMPROVED AND NOW IT IS FINE THE SURGEON REQUESTED THE INVESTIGATION THAT THE REASON WHY THE PRESSURE SETTING UNINTENTIONALLY CHANGED ALTHOUGH THERE WAS AN MRI LOCKING MECHANISM WITH THE VALVE NO FURTHER INFORMATION WAS PROVIDED BY THE HOSPITAL | Inappropriate or Unexpected Reset | Device problem code
6794080 | 82681336 | D | THE REPORTED VALVE WAS IMPLANTED TO A PATIENT VIA VPSHUNT WITH THE SETTING OF B AFTER DISCHARGED FROM THE HOSPITAL THE PATIENT CLAIMED VOMITING AFTER THE PATIENT TOOK A BATH AT HOME SO THAT THE PATIENT VISITED TO THE HOSPITAL ON B THE SURGEON CHECKED THE PRESSURE SETTING THEN THE SURGEON FOUND THAT THE SETTING WAS UNINTENTIONALLY CHANGED FROM TO AFTER CHANGED THE PRESSURE SETTING BACK TO THE PATIENTS CONDITION WAS IMPROVED AND NOW IT IS FINE THE SURGEON REQUESTED THE INVESTIGATION THAT THE REASON WHY THE PRESSURE SETTING UNINTENTIONALLY CHANGED ALTHOUGH THERE WAS AN MRI LOCKING MECHANISM WITH THE VALVE NO FURTHER INFORMATION WAS PROVIDED BY THE HOSPITAL | Vomiting | Patient problem code
6794224 | 82685398 | N | A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT THE SCREWDRIVER USED IN THE PROCEDURE WAS NEWER AND NOT COMPATIBLE WITH THE APPLICATION SOFTWARE INSTALLED ONTO THE NAVIGATION SYSTEM THE REPRESENTATIVE REPORTED THAT THEY WERE ABLE TO VERIFY INSTRUMENTS IN A TRACKER WITHOUT FAULT THE HARDWARE SOFTWARE AND INSTRUMENTS PASSED THE SYSTEM CHECKOUT THE SYSTEM WAS FOUND TO BE FULLY FUNCTIONAL | No Device Problem Found | Medical Evaluation result code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Fever | Patient problem code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Vomiting | Patient problem code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Insufficient Information | Device problem code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6794247 | 82686460 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B AT AFTERNOON WITH BLOOD GLUCOSE OF MGDL CUSTOMER WAS HOSPITALIZED TWICE BECAUSE OF DIABETIC KETOACIDOSIS THE TUBING AND CANNULA WAS SENT BACK THE HOSPITAL NAME WAS B CUSTOMER TURNED WHITE AND HIS EYE ROLLED IN THE BACK OF HIS HEAD AND HE WAS DEHYDRATED CUSTOMER WAS VERY SKINNY CUSTOMER WAS VOMITING AND HAD FEVER CUSTOMER WAS PUT IN IV AND ALMOST WENT DIABETIC COMMA CUSTOMER WAS RELEASED AFTER TWO WEEKS CUSTOMER WAS NOT WEARING THE INSULIN PUMP WHILE IN THE HOSPITAL HOWEVER CUSTOMER WAS NOT OFF THE INSULIN PUMP FOR MORE THAN HOURS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Dehydrated | Patient problem code
6794298 | 82686969 | D | MEDTRONIC RECEIVED INFORMATION THAT EIGHT DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6794298 | 82686969 | D | MEDTRONIC RECEIVED INFORMATION THAT EIGHT DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6794485 | 82693521 | D | ON B A PATIENT IN THE USA UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY PEG TUBE WITH JEJUNAL PEGJ TUBE THE PROCEDURE WENT WELL AND THE PATIENT WAS SENT HOME LATER THAT DAY SHE WAS HOSPITALIZED DUE TO INTOLERABLE PAIN A CHEST XRAY WAS PERFORMED WHICH DIAGNOSED PNEUMONIA THE DOCTOR WAS UNSURE IF THE PNEUMONIA WAS ASPIRATION PNEUMONIA AS A RESULT OF THE PROCEDURE OR IF IT WAS A PREEXISTING BACTERIAL PNEUMONIA THE PATIENT WAS TREATED WITH UNKNOWN ANTIBIOTICS FOR BOTH TYPES THE PATIENT HAS NOT STARTED DUOPA MEDICATION YET ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL | Pneumonia | Patient problem code
6794485 | 82693521 | D | ON B A PATIENT IN THE USA UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY PEG TUBE WITH JEJUNAL PEGJ TUBE THE PROCEDURE WENT WELL AND THE PATIENT WAS SENT HOME LATER THAT DAY SHE WAS HOSPITALIZED DUE TO INTOLERABLE PAIN A CHEST XRAY WAS PERFORMED WHICH DIAGNOSED PNEUMONIA THE DOCTOR WAS UNSURE IF THE PNEUMONIA WAS ASPIRATION PNEUMONIA AS A RESULT OF THE PROCEDURE OR IF IT WAS A PREEXISTING BACTERIAL PNEUMONIA THE PATIENT WAS TREATED WITH UNKNOWN ANTIBIOTICS FOR BOTH TYPES THE PATIENT HAS NOT STARTED DUOPA MEDICATION YET ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL | Insufficient Information | Device problem code
6794485 | 82693521 | D | ON B A PATIENT IN THE USA UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY PEG TUBE WITH JEJUNAL PEGJ TUBE THE PROCEDURE WENT WELL AND THE PATIENT WAS SENT HOME LATER THAT DAY SHE WAS HOSPITALIZED DUE TO INTOLERABLE PAIN A CHEST XRAY WAS PERFORMED WHICH DIAGNOSED PNEUMONIA THE DOCTOR WAS UNSURE IF THE PNEUMONIA WAS ASPIRATION PNEUMONIA AS A RESULT OF THE PROCEDURE OR IF IT WAS A PREEXISTING BACTERIAL PNEUMONIA THE PATIENT WAS TREATED WITH UNKNOWN ANTIBIOTICS FOR BOTH TYPES THE PATIENT HAS NOT STARTED DUOPA MEDICATION YET ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL | Aspiration | Patient problem code
6794485 | 82693521 | D | ON B A PATIENT IN THE USA UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY PEG TUBE WITH JEJUNAL PEGJ TUBE THE PROCEDURE WENT WELL AND THE PATIENT WAS SENT HOME LATER THAT DAY SHE WAS HOSPITALIZED DUE TO INTOLERABLE PAIN A CHEST XRAY WAS PERFORMED WHICH DIAGNOSED PNEUMONIA THE DOCTOR WAS UNSURE IF THE PNEUMONIA WAS ASPIRATION PNEUMONIA AS A RESULT OF THE PROCEDURE OR IF IT WAS A PREEXISTING BACTERIAL PNEUMONIA THE PATIENT WAS TREATED WITH UNKNOWN ANTIBIOTICS FOR BOTH TYPES THE PATIENT HAS NOT STARTED DUOPA MEDICATION YET ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL | Pain | Patient problem code
6794689 | 82699420 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device Problem Code
6794705 | 82699572 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device Problem Code
6794509 | 82699595 | D | IT WAS REPORTED FROM THE CLINICAL STUDY SITE THAT THE PATIENT PRESENTED ON B WITH WEEK HISTORY OF MELENA AND PROGRESSIVE SHORTNESS OF BREATH IN THE SETTING OF VOLUME OVERLOAD FOUND TO HAVE PARAINFLUENZA VIRUS AND DECOMPENSATED HEART FAILURE RESOLVED WITH CONSERVATIVE TREATMENT ESOPHAGOGASTRODUODENOSCOPY EGD WAS OFFERED BUT PATIENT DECLINED INSTEAD OPTING FOR OUTPATIENT EGD EGD TO BE SCHEDULED AS OUTPATIENT IF NEEDED AS PER GASTROINTESTINAL GI RECOMMENDATIONS SUBJECT DISCHARGED ON B NO ADDITIONAL INFORMATION AVAILABLE | Virus | Patient problem code
6794509 | 82699595 | D | IT WAS REPORTED FROM THE CLINICAL STUDY SITE THAT THE PATIENT PRESENTED ON B WITH WEEK HISTORY OF MELENA AND PROGRESSIVE SHORTNESS OF BREATH IN THE SETTING OF VOLUME OVERLOAD FOUND TO HAVE PARAINFLUENZA VIRUS AND DECOMPENSATED HEART FAILURE RESOLVED WITH CONSERVATIVE TREATMENT ESOPHAGOGASTRODUODENOSCOPY EGD WAS OFFERED BUT PATIENT DECLINED INSTEAD OPTING FOR OUTPATIENT EGD EGD TO BE SCHEDULED AS OUTPATIENT IF NEEDED AS PER GASTROINTESTINAL GI RECOMMENDATIONS SUBJECT DISCHARGED ON B NO ADDITIONAL INFORMATION AVAILABLE | Insufficient Information | Device problem code
6794509 | 82699595 | D | IT WAS REPORTED FROM THE CLINICAL STUDY SITE THAT THE PATIENT PRESENTED ON B WITH WEEK HISTORY OF MELENA AND PROGRESSIVE SHORTNESS OF BREATH IN THE SETTING OF VOLUME OVERLOAD FOUND TO HAVE PARAINFLUENZA VIRUS AND DECOMPENSATED HEART FAILURE RESOLVED WITH CONSERVATIVE TREATMENT ESOPHAGOGASTRODUODENOSCOPY EGD WAS OFFERED BUT PATIENT DECLINED INSTEAD OPTING FOR OUTPATIENT EGD EGD TO BE SCHEDULED AS OUTPATIENT IF NEEDED AS PER GASTROINTESTINAL GI RECOMMENDATIONS SUBJECT DISCHARGED ON B NO ADDITIONAL INFORMATION AVAILABLE | Heart Failure | Patient problem code
6794509 | 82699595 | D | IT WAS REPORTED FROM THE CLINICAL STUDY SITE THAT THE PATIENT PRESENTED ON B WITH WEEK HISTORY OF MELENA AND PROGRESSIVE SHORTNESS OF BREATH IN THE SETTING OF VOLUME OVERLOAD FOUND TO HAVE PARAINFLUENZA VIRUS AND DECOMPENSATED HEART FAILURE RESOLVED WITH CONSERVATIVE TREATMENT ESOPHAGOGASTRODUODENOSCOPY EGD WAS OFFERED BUT PATIENT DECLINED INSTEAD OPTING FOR OUTPATIENT EGD EGD TO BE SCHEDULED AS OUTPATIENT IF NEEDED AS PER GASTROINTESTINAL GI RECOMMENDATIONS SUBJECT DISCHARGED ON B NO ADDITIONAL INFORMATION AVAILABLE | Dyspnea | Patient problem code
6794723 | 82700610 | D | B CLINICAL STUDY IT WAS REPORTED THERE WAS RESTENOSIS IN B A MM JETSTREAM XC ATHERECTOMY CATHETER WAS SELECTED FOR AN ATHERECTOMY PROCEDURE IN THE RIGHT PROXIMAL SUPERFICIAL FEMORAL ARTERY THE TARGET VESSEL HAD STENOSIS REFERENCE VESSEL DIAMETER OF MM LENGTH OF MM AND CLASSIFIED AS A TASC II C LESION POST TREATMENT PERCUTANEOUS TRANSLUMINAL BALLOON ANGIOPLASTY PTA WAS PERFORMED WITH FINAL RESIDUAL STENOSIS TWO DAYS LATER THE SUBJECT WAS DISCHARGED ON PRASUGREL AND ASPIRIN IN B DAYS POST INDEX PROCEDURE SUBJECT WAS NOTED WITH RIGHT SUPERFICIAL FEMORAL ARTERY RESTENOSIS NO ACTION WAS TAKEN TO TREAT THE EVENT AT THE TIME OF REPORTING THE EVENT WAS CONSIDERED TO BE NOT RECOVEREDNOT RESOLVED | Insufficient Information | Device problem code
6794723 | 82700610 | D | B CLINICAL STUDY IT WAS REPORTED THERE WAS RESTENOSIS IN B A MM JETSTREAM XC ATHERECTOMY CATHETER WAS SELECTED FOR AN ATHERECTOMY PROCEDURE IN THE RIGHT PROXIMAL SUPERFICIAL FEMORAL ARTERY THE TARGET VESSEL HAD STENOSIS REFERENCE VESSEL DIAMETER OF MM LENGTH OF MM AND CLASSIFIED AS A TASC II C LESION POST TREATMENT PERCUTANEOUS TRANSLUMINAL BALLOON ANGIOPLASTY PTA WAS PERFORMED WITH FINAL RESIDUAL STENOSIS TWO DAYS LATER THE SUBJECT WAS DISCHARGED ON PRASUGREL AND ASPIRIN IN B DAYS POST INDEX PROCEDURE SUBJECT WAS NOTED WITH RIGHT SUPERFICIAL FEMORAL ARTERY RESTENOSIS NO ACTION WAS TAKEN TO TREAT THE EVENT AT THE TIME OF REPORTING THE EVENT WAS CONSIDERED TO BE NOT RECOVEREDNOT RESOLVED | Stenosis | Patient problem code
6794794 | 82706126 | D | IT WAS REPORTED THAT A PERICARDIAL EFFUSION AND PLEURAL EFFUSION OCCURRED POST PROCEDURE A MM WATCHMAN LAA CLOSURE DEVICE DELIVERY SYSTEM WAS SUCCESSFULLY IMPLANTED DURING A LAA CLOSURE PROCEDURE THE PATIENT WAS DISCHARGED ONE DAY POST PROCEDURE AND THEIR STATUS WAS FINE AT THEIR ONE WEEK FOLLOW UP CHECK GROIN CHECK THEY WERE FINE EIGHTEEN DAYS POST PROCEDURE THE PATIENT PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH THE PATIENT WAS ADMINISTERED LASIX BECAUSE THEY THOUGHT IT MAY HAVE BEEN DUE TO HEART FAILURE EXACERBATION THEY ORDERED AN ECHOCARDIOGRAM WHEN THEY NOTICED A SUBSTANTIAL PERICARDIAL EFFUSION IN THE PERICARDIAL SPACE THE PATIENT ALSO HAD A PLEURAL EFFUSION THAT MAY HAVE BEEN CAUSED BY THE SHORTNESS OF BREATH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THEY WERE ADMITTED FOR MONITORING AND THEIR STATUS WAS FINE THEIR HEMATOCRIT WAS AND WAS NOTED TO BE THE DAY BEFORE THE PROCEDURE THERE WAS NO CLOT OR TAMPONADE NOTICED THEY DECIDED TO STOP THE WARFARIN AND PERFORM A REPEAT ECHOCARDIOGRAM THE NEXT DAY THE NEXT DAY A TRANSTHORACIC ECHOCARDIOGRAM TTE WAS PERFORMED AND THERE WAS NO CHANGE TO THE PE OR PLEURAL EFFUSION AND THE PATIENT WAS DISCHARGED | Heart Failure | Patient problem code
6794794 | 82706126 | D | IT WAS REPORTED THAT A PERICARDIAL EFFUSION AND PLEURAL EFFUSION OCCURRED POST PROCEDURE A MM WATCHMAN LAA CLOSURE DEVICE DELIVERY SYSTEM WAS SUCCESSFULLY IMPLANTED DURING A LAA CLOSURE PROCEDURE THE PATIENT WAS DISCHARGED ONE DAY POST PROCEDURE AND THEIR STATUS WAS FINE AT THEIR ONE WEEK FOLLOW UP CHECK GROIN CHECK THEY WERE FINE EIGHTEEN DAYS POST PROCEDURE THE PATIENT PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH THE PATIENT WAS ADMINISTERED LASIX BECAUSE THEY THOUGHT IT MAY HAVE BEEN DUE TO HEART FAILURE EXACERBATION THEY ORDERED AN ECHOCARDIOGRAM WHEN THEY NOTICED A SUBSTANTIAL PERICARDIAL EFFUSION IN THE PERICARDIAL SPACE THE PATIENT ALSO HAD A PLEURAL EFFUSION THAT MAY HAVE BEEN CAUSED BY THE SHORTNESS OF BREATH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THEY WERE ADMITTED FOR MONITORING AND THEIR STATUS WAS FINE THEIR HEMATOCRIT WAS AND WAS NOTED TO BE THE DAY BEFORE THE PROCEDURE THERE WAS NO CLOT OR TAMPONADE NOTICED THEY DECIDED TO STOP THE WARFARIN AND PERFORM A REPEAT ECHOCARDIOGRAM THE NEXT DAY THE NEXT DAY A TRANSTHORACIC ECHOCARDIOGRAM TTE WAS PERFORMED AND THERE WAS NO CHANGE TO THE PE OR PLEURAL EFFUSION AND THE PATIENT WAS DISCHARGED | Dyspnea | Patient problem code
6794794 | 82706126 | D | IT WAS REPORTED THAT A PERICARDIAL EFFUSION AND PLEURAL EFFUSION OCCURRED POST PROCEDURE A MM WATCHMAN LAA CLOSURE DEVICE DELIVERY SYSTEM WAS SUCCESSFULLY IMPLANTED DURING A LAA CLOSURE PROCEDURE THE PATIENT WAS DISCHARGED ONE DAY POST PROCEDURE AND THEIR STATUS WAS FINE AT THEIR ONE WEEK FOLLOW UP CHECK GROIN CHECK THEY WERE FINE EIGHTEEN DAYS POST PROCEDURE THE PATIENT PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH THE PATIENT WAS ADMINISTERED LASIX BECAUSE THEY THOUGHT IT MAY HAVE BEEN DUE TO HEART FAILURE EXACERBATION THEY ORDERED AN ECHOCARDIOGRAM WHEN THEY NOTICED A SUBSTANTIAL PERICARDIAL EFFUSION IN THE PERICARDIAL SPACE THE PATIENT ALSO HAD A PLEURAL EFFUSION THAT MAY HAVE BEEN CAUSED BY THE SHORTNESS OF BREATH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THEY WERE ADMITTED FOR MONITORING AND THEIR STATUS WAS FINE THEIR HEMATOCRIT WAS AND WAS NOTED TO BE THE DAY BEFORE THE PROCEDURE THERE WAS NO CLOT OR TAMPONADE NOTICED THEY DECIDED TO STOP THE WARFARIN AND PERFORM A REPEAT ECHOCARDIOGRAM THE NEXT DAY THE NEXT DAY A TRANSTHORACIC ECHOCARDIOGRAM TTE WAS PERFORMED AND THERE WAS NO CHANGE TO THE PE OR PLEURAL EFFUSION AND THE PATIENT WAS DISCHARGED | Insufficient Information | Device problem code
6794794 | 82706126 | D | IT WAS REPORTED THAT A PERICARDIAL EFFUSION AND PLEURAL EFFUSION OCCURRED POST PROCEDURE A MM WATCHMAN LAA CLOSURE DEVICE DELIVERY SYSTEM WAS SUCCESSFULLY IMPLANTED DURING A LAA CLOSURE PROCEDURE THE PATIENT WAS DISCHARGED ONE DAY POST PROCEDURE AND THEIR STATUS WAS FINE AT THEIR ONE WEEK FOLLOW UP CHECK GROIN CHECK THEY WERE FINE EIGHTEEN DAYS POST PROCEDURE THE PATIENT PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH THE PATIENT WAS ADMINISTERED LASIX BECAUSE THEY THOUGHT IT MAY HAVE BEEN DUE TO HEART FAILURE EXACERBATION THEY ORDERED AN ECHOCARDIOGRAM WHEN THEY NOTICED A SUBSTANTIAL PERICARDIAL EFFUSION IN THE PERICARDIAL SPACE THE PATIENT ALSO HAD A PLEURAL EFFUSION THAT MAY HAVE BEEN CAUSED BY THE SHORTNESS OF BREATH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THEY WERE ADMITTED FOR MONITORING AND THEIR STATUS WAS FINE THEIR HEMATOCRIT WAS AND WAS NOTED TO BE THE DAY BEFORE THE PROCEDURE THERE WAS NO CLOT OR TAMPONADE NOTICED THEY DECIDED TO STOP THE WARFARIN AND PERFORM A REPEAT ECHOCARDIOGRAM THE NEXT DAY THE NEXT DAY A TRANSTHORACIC ECHOCARDIOGRAM TTE WAS PERFORMED AND THERE WAS NO CHANGE TO THE PE OR PLEURAL EFFUSION AND THE PATIENT WAS DISCHARGED | Pleural Effusion | Patient problem code
6794794 | 82706126 | D | IT WAS REPORTED THAT A PERICARDIAL EFFUSION AND PLEURAL EFFUSION OCCURRED POST PROCEDURE A MM WATCHMAN LAA CLOSURE DEVICE DELIVERY SYSTEM WAS SUCCESSFULLY IMPLANTED DURING A LAA CLOSURE PROCEDURE THE PATIENT WAS DISCHARGED ONE DAY POST PROCEDURE AND THEIR STATUS WAS FINE AT THEIR ONE WEEK FOLLOW UP CHECK GROIN CHECK THEY WERE FINE EIGHTEEN DAYS POST PROCEDURE THE PATIENT PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH THE PATIENT WAS ADMINISTERED LASIX BECAUSE THEY THOUGHT IT MAY HAVE BEEN DUE TO HEART FAILURE EXACERBATION THEY ORDERED AN ECHOCARDIOGRAM WHEN THEY NOTICED A SUBSTANTIAL PERICARDIAL EFFUSION IN THE PERICARDIAL SPACE THE PATIENT ALSO HAD A PLEURAL EFFUSION THAT MAY HAVE BEEN CAUSED BY THE SHORTNESS OF BREATH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THEY WERE ADMITTED FOR MONITORING AND THEIR STATUS WAS FINE THEIR HEMATOCRIT WAS AND WAS NOTED TO BE THE DAY BEFORE THE PROCEDURE THERE WAS NO CLOT OR TAMPONADE NOTICED THEY DECIDED TO STOP THE WARFARIN AND PERFORM A REPEAT ECHOCARDIOGRAM THE NEXT DAY THE NEXT DAY A TRANSTHORACIC ECHOCARDIOGRAM TTE WAS PERFORMED AND THERE WAS NO CHANGE TO THE PE OR PLEURAL EFFUSION AND THE PATIENT WAS DISCHARGED | Pericardial Effusion | Patient problem code
6794950 | 82708162 | N | THE PRODUCT EVALUATION WAS PERFORMED BY THE DISTRIBUTOR THE ISSUE OF BED EXIT NOT ALARMING WAS RESOLVED BY PLUGGING THE UNIT BACK INTO THE WALL OUTLET | Device Alarm System | Device problem code
6794983 | 82709030 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing Information | Device problem code
6794983 | 82709030 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | No Findings Available | Medical Evaluation result code
6794786 | 82709114 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT ON B THE PATIENT NOTED A RISE IN THEIR WHITE BLOOD CELLS COUNT WBC TO SECONDARY TO A DRIVELINE INFECTION ADDRESSED IN B AND THE RELATED BACTEREMIA THE PATIENT WAS ASYMPTOMATIC AND DID NOT PRESENT WITH ANY OTHER CLINICAL SYMPTOMS THE PATIENT WAS NOT TREATED FOR THE MENTIONED EVENT AND THEIR WBC FELL BACK INTO NORMAL RANGE BY B IT WAS NOTED THAT UPON ADMISSION THE PATIENT WAS PUT ON CHRONIC OUTPATIENT IV ANTIBIOTICS VIA PERIPHERALLY INSERTED CENTRAL CATHETER PICC AND WAS TRANSITIONED TO NEW IV THERAPIES DURING ADMISSION AND DISCHARGED HOME WITH RESOLVED BACTEREMIA THE PATIENT IS NOTED TO BE STABLE WITH CLEAR CULTURES ON NEW SUPPRESSIVE ANTIBIOTIC THERAPY FOR THE ADDRESSED DRIVELINE INFECTION THE PATIENT WAS DISCHARGED HOME FROM THE HOSPITAL ON B NO ADDITIONAL INFORMATION PROVIDED | Insufficient Information | Device problem code
6794786 | 82709114 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT ON B THE PATIENT NOTED A RISE IN THEIR WHITE BLOOD CELLS COUNT WBC TO SECONDARY TO A DRIVELINE INFECTION ADDRESSED IN B AND THE RELATED BACTEREMIA THE PATIENT WAS ASYMPTOMATIC AND DID NOT PRESENT WITH ANY OTHER CLINICAL SYMPTOMS THE PATIENT WAS NOT TREATED FOR THE MENTIONED EVENT AND THEIR WBC FELL BACK INTO NORMAL RANGE BY B IT WAS NOTED THAT UPON ADMISSION THE PATIENT WAS PUT ON CHRONIC OUTPATIENT IV ANTIBIOTICS VIA PERIPHERALLY INSERTED CENTRAL CATHETER PICC AND WAS TRANSITIONED TO NEW IV THERAPIES DURING ADMISSION AND DISCHARGED HOME WITH RESOLVED BACTEREMIA THE PATIENT IS NOTED TO BE STABLE WITH CLEAR CULTURES ON NEW SUPPRESSIVE ANTIBIOTIC THERAPY FOR THE ADDRESSED DRIVELINE INFECTION THE PATIENT WAS DISCHARGED HOME FROM THE HOSPITAL ON B NO ADDITIONAL INFORMATION PROVIDED | Infection | Patient problem code
6795222 | 82715999 | D | MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING A B FEMALE PATIENT WITH A HISTORY OF PERCUTANEOUS MITRAL COMMISSUROTOMY YEARS PREVIOUSLY FOR RHEUMATIC MITRAL STENOSIS IN MORE RECENT YEARS THE PATIENT HAS BEEN ADMITTED FOR CONGESTIVE HEART FAILURE AND PULMONARY DYSFUNCTION ELECTROCARDIOGRAPHY REVEALED CHRONIC ATRIAL FIBRILLATION ECHOCARDIOGRAPHY SHOWED SEVERE MITRAL STENOSIS AND SEVERE TRICUSPID REGURGITATION CHEST COMPUTED TOMOGRAPHY SHOWED THE RIGHT ATRIUM WAS MORE SEVERELY ENLARGED THAN THE LEFT AND COMPRESSED THE RIGHT LOWER LUNG SURGERY WAS PERFORMED THROUGH A MEDIAN STERNOTOMY THE PERICARDIUM WAS EXTENSIVELY ADHERED TO THE HEART AND WAS DISSECTED AWAY TRICUSPID ANNULOPLASTY WAS PERFORMED WITH A MM NONMEDTRONIC ANNULOPLASTY RING AFTER AORTIC CROSSCLAMPING MITRAL VALVE REPLACEMENT WAS PERFORMED WITH A MM MEDTRONIC MOSAIC PORCINE BIOPROSTHESIS SERIAL NUMBERS NOT PROVIDED ALTHOUGH THE PATIENT REQUIRED REEXPLORATION FOR MASSIVE HEMOTHORAX ON POSTOPERATIVE DAY SHE WAS EXTUBATED ON POSTOPERATIVE DAY AND DISCHARGED WITH NO EVIDENCE OF HEART OR RESPIRATORY FAILURE AFTER REHABILITATION ON POSTOPERATIVE DAY NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED | Insufficient Information | Device problem code
6795222 | 82715999 | D | MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING A B FEMALE PATIENT WITH A HISTORY OF PERCUTANEOUS MITRAL COMMISSUROTOMY YEARS PREVIOUSLY FOR RHEUMATIC MITRAL STENOSIS IN MORE RECENT YEARS THE PATIENT HAS BEEN ADMITTED FOR CONGESTIVE HEART FAILURE AND PULMONARY DYSFUNCTION ELECTROCARDIOGRAPHY REVEALED CHRONIC ATRIAL FIBRILLATION ECHOCARDIOGRAPHY SHOWED SEVERE MITRAL STENOSIS AND SEVERE TRICUSPID REGURGITATION CHEST COMPUTED TOMOGRAPHY SHOWED THE RIGHT ATRIUM WAS MORE SEVERELY ENLARGED THAN THE LEFT AND COMPRESSED THE RIGHT LOWER LUNG SURGERY WAS PERFORMED THROUGH A MEDIAN STERNOTOMY THE PERICARDIUM WAS EXTENSIVELY ADHERED TO THE HEART AND WAS DISSECTED AWAY TRICUSPID ANNULOPLASTY WAS PERFORMED WITH A MM NONMEDTRONIC ANNULOPLASTY RING AFTER AORTIC CROSSCLAMPING MITRAL VALVE REPLACEMENT WAS PERFORMED WITH A MM MEDTRONIC MOSAIC PORCINE BIOPROSTHESIS SERIAL NUMBERS NOT PROVIDED ALTHOUGH THE PATIENT REQUIRED REEXPLORATION FOR MASSIVE HEMOTHORAX ON POSTOPERATIVE DAY SHE WAS EXTUBATED ON POSTOPERATIVE DAY AND DISCHARGED WITH NO EVIDENCE OF HEART OR RESPIRATORY FAILURE AFTER REHABILITATION ON POSTOPERATIVE DAY NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
6795368 | 82750175 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED LESS INVASIVE IMPLANTATION OF HEARTWARE LEFT VENTRICULAR ASSIST DEVICE THE ARTICLE INCLUDED A CASE STUDY DESCRIBING THE PROCEDURE INVOLVING A PATIENT IN CARDIOGENIC SHOCK WHO WAS SUCCESSFULLY IMPLANTED WITH A VENTRICULAR ASSIST DEVICE VAD VIA A LEFT ANTERIOR SMALL THORACOTOMY WITH MINIUPPER STERNOTOMY ECMO SUPPORT AND LOWER HEPARIN DOSE ADMINISTRATION THIS PATIENT WAS DISCHARGED FROM HOSPITAL AFTER ALMOST A MONTH AND IS NOW ELIGIBLE FOR CARDIAC TRANSPLANTATION THE ARTICLE ALSO DISCUSSES THE COMPLICATIONS OF TWENTYONE PATIENTS IMPLANTED WITH THE SAME SURGICAL APPROACH BETWEEN JANUARY AND DECEMBER ONE OF THESE PATIENTS WAS REPORTED TO HAVE EXPIRED THE AUTHORS CONCLUDED THAT MINIINVASIVE CARDIAC SURGERY COMPARED WITH CONVENTIONAL FULL STERNOTOMY CARDIAC SURGERY IS ASSOCIATED WITH DECREASED BLEEDING BLOOD PRODUCT TRANSFUSION STERNAL WOUND INFECTION SCAR DISSATISFACTION VENTILATION TIME INTENSIVE CARE UNIT STAY HOSPITAL LENGTH OF STAY AND REDUCED TIME TO RETURN TO NORMAL ACTIVITY THEREBY RESULTING IN BETTER OUTCOMES NO FURTHER INFORMATION HAS BEEN PROVIDED | Shock, Cardiogenic | Patient problem code
6795368 | 82750175 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED LESS INVASIVE IMPLANTATION OF HEARTWARE LEFT VENTRICULAR ASSIST DEVICE THE ARTICLE INCLUDED A CASE STUDY DESCRIBING THE PROCEDURE INVOLVING A PATIENT IN CARDIOGENIC SHOCK WHO WAS SUCCESSFULLY IMPLANTED WITH A VENTRICULAR ASSIST DEVICE VAD VIA A LEFT ANTERIOR SMALL THORACOTOMY WITH MINIUPPER STERNOTOMY ECMO SUPPORT AND LOWER HEPARIN DOSE ADMINISTRATION THIS PATIENT WAS DISCHARGED FROM HOSPITAL AFTER ALMOST A MONTH AND IS NOW ELIGIBLE FOR CARDIAC TRANSPLANTATION THE ARTICLE ALSO DISCUSSES THE COMPLICATIONS OF TWENTYONE PATIENTS IMPLANTED WITH THE SAME SURGICAL APPROACH BETWEEN JANUARY AND DECEMBER ONE OF THESE PATIENTS WAS REPORTED TO HAVE EXPIRED THE AUTHORS CONCLUDED THAT MINIINVASIVE CARDIAC SURGERY COMPARED WITH CONVENTIONAL FULL STERNOTOMY CARDIAC SURGERY IS ASSOCIATED WITH DECREASED BLEEDING BLOOD PRODUCT TRANSFUSION STERNAL WOUND INFECTION SCAR DISSATISFACTION VENTILATION TIME INTENSIVE CARE UNIT STAY HOSPITAL LENGTH OF STAY AND REDUCED TIME TO RETURN TO NORMAL ACTIVITY THEREBY RESULTING IN BETTER OUTCOMES NO FURTHER INFORMATION HAS BEEN PROVIDED | Scar Tissue | Patient problem code
6795368 | 82750175 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED LESS INVASIVE IMPLANTATION OF HEARTWARE LEFT VENTRICULAR ASSIST DEVICE THE ARTICLE INCLUDED A CASE STUDY DESCRIBING THE PROCEDURE INVOLVING A PATIENT IN CARDIOGENIC SHOCK WHO WAS SUCCESSFULLY IMPLANTED WITH A VENTRICULAR ASSIST DEVICE VAD VIA A LEFT ANTERIOR SMALL THORACOTOMY WITH MINIUPPER STERNOTOMY ECMO SUPPORT AND LOWER HEPARIN DOSE ADMINISTRATION THIS PATIENT WAS DISCHARGED FROM HOSPITAL AFTER ALMOST A MONTH AND IS NOW ELIGIBLE FOR CARDIAC TRANSPLANTATION THE ARTICLE ALSO DISCUSSES THE COMPLICATIONS OF TWENTYONE PATIENTS IMPLANTED WITH THE SAME SURGICAL APPROACH BETWEEN JANUARY AND DECEMBER ONE OF THESE PATIENTS WAS REPORTED TO HAVE EXPIRED THE AUTHORS CONCLUDED THAT MINIINVASIVE CARDIAC SURGERY COMPARED WITH CONVENTIONAL FULL STERNOTOMY CARDIAC SURGERY IS ASSOCIATED WITH DECREASED BLEEDING BLOOD PRODUCT TRANSFUSION STERNAL WOUND INFECTION SCAR DISSATISFACTION VENTILATION TIME INTENSIVE CARE UNIT STAY HOSPITAL LENGTH OF STAY AND REDUCED TIME TO RETURN TO NORMAL ACTIVITY THEREBY RESULTING IN BETTER OUTCOMES NO FURTHER INFORMATION HAS BEEN PROVIDED | Hemorrhage | Patient problem code
6795368 | 82750175 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED LESS INVASIVE IMPLANTATION OF HEARTWARE LEFT VENTRICULAR ASSIST DEVICE THE ARTICLE INCLUDED A CASE STUDY DESCRIBING THE PROCEDURE INVOLVING A PATIENT IN CARDIOGENIC SHOCK WHO WAS SUCCESSFULLY IMPLANTED WITH A VENTRICULAR ASSIST DEVICE VAD VIA A LEFT ANTERIOR SMALL THORACOTOMY WITH MINIUPPER STERNOTOMY ECMO SUPPORT AND LOWER HEPARIN DOSE ADMINISTRATION THIS PATIENT WAS DISCHARGED FROM HOSPITAL AFTER ALMOST A MONTH AND IS NOW ELIGIBLE FOR CARDIAC TRANSPLANTATION THE ARTICLE ALSO DISCUSSES THE COMPLICATIONS OF TWENTYONE PATIENTS IMPLANTED WITH THE SAME SURGICAL APPROACH BETWEEN JANUARY AND DECEMBER ONE OF THESE PATIENTS WAS REPORTED TO HAVE EXPIRED THE AUTHORS CONCLUDED THAT MINIINVASIVE CARDIAC SURGERY COMPARED WITH CONVENTIONAL FULL STERNOTOMY CARDIAC SURGERY IS ASSOCIATED WITH DECREASED BLEEDING BLOOD PRODUCT TRANSFUSION STERNAL WOUND INFECTION SCAR DISSATISFACTION VENTILATION TIME INTENSIVE CARE UNIT STAY HOSPITAL LENGTH OF STAY AND REDUCED TIME TO RETURN TO NORMAL ACTIVITY THEREBY RESULTING IN BETTER OUTCOMES NO FURTHER INFORMATION HAS BEEN PROVIDED | Insufficient Information | Device problem code
6795368 | 82750175 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED LESS INVASIVE IMPLANTATION OF HEARTWARE LEFT VENTRICULAR ASSIST DEVICE THE ARTICLE INCLUDED A CASE STUDY DESCRIBING THE PROCEDURE INVOLVING A PATIENT IN CARDIOGENIC SHOCK WHO WAS SUCCESSFULLY IMPLANTED WITH A VENTRICULAR ASSIST DEVICE VAD VIA A LEFT ANTERIOR SMALL THORACOTOMY WITH MINIUPPER STERNOTOMY ECMO SUPPORT AND LOWER HEPARIN DOSE ADMINISTRATION THIS PATIENT WAS DISCHARGED FROM HOSPITAL AFTER ALMOST A MONTH AND IS NOW ELIGIBLE FOR CARDIAC TRANSPLANTATION THE ARTICLE ALSO DISCUSSES THE COMPLICATIONS OF TWENTYONE PATIENTS IMPLANTED WITH THE SAME SURGICAL APPROACH BETWEEN JANUARY AND DECEMBER ONE OF THESE PATIENTS WAS REPORTED TO HAVE EXPIRED THE AUTHORS CONCLUDED THAT MINIINVASIVE CARDIAC SURGERY COMPARED WITH CONVENTIONAL FULL STERNOTOMY CARDIAC SURGERY IS ASSOCIATED WITH DECREASED BLEEDING BLOOD PRODUCT TRANSFUSION STERNAL WOUND INFECTION SCAR DISSATISFACTION VENTILATION TIME INTENSIVE CARE UNIT STAY HOSPITAL LENGTH OF STAY AND REDUCED TIME TO RETURN TO NORMAL ACTIVITY THEREBY RESULTING IN BETTER OUTCOMES NO FURTHER INFORMATION HAS BEEN PROVIDED | Infection | Patient problem code
6796147 | 82751056 | D | THE PATIENT WAS INITIALLY IMPLANTED WITH A BIRUCATED AND SUPRARENAL EXTENSION ON B DURING A ROUTINE FOLLOW UP IT WAS DISCOVERED THAT THE PATIENT HAD AN ENDOLEAK TYPE IIIB OF THE PROXIMAL EXTENSION THE PHYSICIAN COMPLETED A SECONDARY ON B TO RELINE WITH TWO VELA SUPRARENAL EXTENSION AND A BIFURCATED STENT PATIENT IS REPORTED AS DOING GOOD POST PROCEDURE | Insufficient Information | Device problem code
6796147 | 82751056 | D | THE PATIENT WAS INITIALLY IMPLANTED WITH A BIRUCATED AND SUPRARENAL EXTENSION ON B DURING A ROUTINE FOLLOW UP IT WAS DISCOVERED THAT THE PATIENT HAD AN ENDOLEAK TYPE IIIB OF THE PROXIMAL EXTENSION THE PHYSICIAN COMPLETED A SECONDARY ON B TO RELINE WITH TWO VELA SUPRARENAL EXTENSION AND A BIFURCATED STENT PATIENT IS REPORTED AS DOING GOOD POST PROCEDURE | Hemorrhage | Patient problem code
6796316 | 82752706 | D | IT WAS REPORTED THAT ON B A XMM XIENCE XPEDITION STENT WAS IMPLANTED IN THE MID LEFT ANTERIOR DESCENDING LAD CORONARY ARTERY STENOSED LESION MEDICATION WAS PROVIDED AND THE PATIENT RECOVERED WELL AND WAS DISCHARGED ON B ON B THE PATIENT HAD BEEN HOSPITALIZED FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE PROXIMAL LAD WAS STENOSED AND MID LAD STENOSED STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT AND MEDICATION WAS PROVIDED ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL ON B THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE MID TO DISTAL LAD REMAINED STENOSED THE STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT DRUG COATED BALLOON DILATATION WAS PERFORMED AS TREATMENT ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL PER PHYSICIAN THE EVENTS WERE POSSIBLY RELATED TO THE DEVICE AND INDEX PROCEDURE THERE WAS NO ADDITIONAL INFORMATION PROVIDED | Insufficient Information | Device problem code
6796316 | 82752706 | D | IT WAS REPORTED THAT ON B A XMM XIENCE XPEDITION STENT WAS IMPLANTED IN THE MID LEFT ANTERIOR DESCENDING LAD CORONARY ARTERY STENOSED LESION MEDICATION WAS PROVIDED AND THE PATIENT RECOVERED WELL AND WAS DISCHARGED ON B ON B THE PATIENT HAD BEEN HOSPITALIZED FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE PROXIMAL LAD WAS STENOSED AND MID LAD STENOSED STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT AND MEDICATION WAS PROVIDED ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL ON B THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE MID TO DISTAL LAD REMAINED STENOSED THE STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT DRUG COATED BALLOON DILATATION WAS PERFORMED AS TREATMENT ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL PER PHYSICIAN THE EVENTS WERE POSSIBLY RELATED TO THE DEVICE AND INDEX PROCEDURE THERE WAS NO ADDITIONAL INFORMATION PROVIDED | Stenosis | Patient problem code
6796316 | 82752706 | D | IT WAS REPORTED THAT ON B A XMM XIENCE XPEDITION STENT WAS IMPLANTED IN THE MID LEFT ANTERIOR DESCENDING LAD CORONARY ARTERY STENOSED LESION MEDICATION WAS PROVIDED AND THE PATIENT RECOVERED WELL AND WAS DISCHARGED ON B ON B THE PATIENT HAD BEEN HOSPITALIZED FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE PROXIMAL LAD WAS STENOSED AND MID LAD STENOSED STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT AND MEDICATION WAS PROVIDED ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL ON B THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL FOR CHEST PAIN CORONARY ANGIOGRAPHY WAS PERFORMED AND THE MID TO DISTAL LAD REMAINED STENOSED THE STENOSIS WAS WITHIN MM OF THE XMM XIENCE XPEDITION STENT DRUG COATED BALLOON DILATATION WAS PERFORMED AS TREATMENT ON B THE PATIENT WAS DISCHARGED FROM THE HOSPITAL PER PHYSICIAN THE EVENTS WERE POSSIBLY RELATED TO THE DEVICE AND INDEX PROCEDURE THERE WAS NO ADDITIONAL INFORMATION PROVIDED | Chest Pain | Patient problem code
6796310 | 82753231 | D | AS REPORTED APPROXIMATELY YEARS POST TRANSAPICAL TAVR PROCEDURE IN THE AORTIC POSITION THE PATIENT BECAME SYMPTOMATIC AND UNDERWENT RIGHT SUBCLAVIAN ARTERY VALVE IN VALVE SAPIEN IN SAPIEN PROCEDURE DUE TO VALVE DEGENERATION RIGHT SUBCLAVIAN ARTERY ACCESS WAS CHOSEN AS THE PATIENT DID NOT HAVE GOOD FEMORAL ACCESS THE SAPIEN WAS SUCCESSFULLY DEPLOYED IN A AORTICVENTRICULAR POSITION THE PATIENT LEFT IN STABLE CONDITION THE PATIENTS VALVE DEGENERATION WAS ATTRIBUTED TO THE PATIENTS DIABETES | Adverse Event Without Identified Device or Use Problem | Device problem code
6796310 | 82753231 | D | AS REPORTED APPROXIMATELY YEARS POST TRANSAPICAL TAVR PROCEDURE IN THE AORTIC POSITION THE PATIENT BECAME SYMPTOMATIC AND UNDERWENT RIGHT SUBCLAVIAN ARTERY VALVE IN VALVE SAPIEN IN SAPIEN PROCEDURE DUE TO VALVE DEGENERATION RIGHT SUBCLAVIAN ARTERY ACCESS WAS CHOSEN AS THE PATIENT DID NOT HAVE GOOD FEMORAL ACCESS THE SAPIEN WAS SUCCESSFULLY DEPLOYED IN A AORTICVENTRICULAR POSITION THE PATIENT LEFT IN STABLE CONDITION THE PATIENTS VALVE DEGENERATION WAS ATTRIBUTED TO THE PATIENTS DIABETES | No Consequences Or Impact To Patient | Patient Problem Code
6796584 | 82754821 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS INSERTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM DIAMETER ABDOMINAL AORTIC ANEURYSM CURRENTLY THE RIGHT ILIAC IS MM MM FROM PROXIMAL TO DISTAL THE LEFT ILIAC IS MM MM FROM PROXIMAL TO DISTAL IT WAS REPORTED THAT A FOLLOWUP CT WAS DONE AND SHOWED THAT THE ANEURYSM HAS GROWN FROM CM TO CM IN MONTHS NEITHER LIMB HAD WALL APPOSITION IN THE ILIAC ARTERIES AND THE ILIACS HAD DILATED LARGER THAN THE MM DIAMETER LIMBS THERE WAS A MM GAP FROM EDGE OF THE LIMBS TO THE HYPOGASTRIC ARTERY ON BOTH SIDES THE PHYSICIAN ELECTED TO IMPLANT ENDURANT STENT GRAFT LIMBS BILATERALLY XX THE ENDOLEAK APPEARED TO BE RESOLVED PER THE PHYSICIAN THE CAUSE OF THE EVENT WAS ANATOMY RELATED THE LIMBS PULLED UP AS THE ANEURYSM SAC GREW AND THE ILIACS DILATED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Aneurysm | Patient problem code
6796584 | 82754821 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS INSERTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM DIAMETER ABDOMINAL AORTIC ANEURYSM CURRENTLY THE RIGHT ILIAC IS MM MM FROM PROXIMAL TO DISTAL THE LEFT ILIAC IS MM MM FROM PROXIMAL TO DISTAL IT WAS REPORTED THAT A FOLLOWUP CT WAS DONE AND SHOWED THAT THE ANEURYSM HAS GROWN FROM CM TO CM IN MONTHS NEITHER LIMB HAD WALL APPOSITION IN THE ILIAC ARTERIES AND THE ILIACS HAD DILATED LARGER THAN THE MM DIAMETER LIMBS THERE WAS A MM GAP FROM EDGE OF THE LIMBS TO THE HYPOGASTRIC ARTERY ON BOTH SIDES THE PHYSICIAN ELECTED TO IMPLANT ENDURANT STENT GRAFT LIMBS BILATERALLY XX THE ENDOLEAK APPEARED TO BE RESOLVED PER THE PHYSICIAN THE CAUSE OF THE EVENT WAS ANATOMY RELATED THE LIMBS PULLED UP AS THE ANEURYSM SAC GREW AND THE ILIACS DILATED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Hemorrhage | Patient problem code
6796584 | 82754821 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS INSERTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM DIAMETER ABDOMINAL AORTIC ANEURYSM CURRENTLY THE RIGHT ILIAC IS MM MM FROM PROXIMAL TO DISTAL THE LEFT ILIAC IS MM MM FROM PROXIMAL TO DISTAL IT WAS REPORTED THAT A FOLLOWUP CT WAS DONE AND SHOWED THAT THE ANEURYSM HAS GROWN FROM CM TO CM IN MONTHS NEITHER LIMB HAD WALL APPOSITION IN THE ILIAC ARTERIES AND THE ILIACS HAD DILATED LARGER THAN THE MM DIAMETER LIMBS THERE WAS A MM GAP FROM EDGE OF THE LIMBS TO THE HYPOGASTRIC ARTERY ON BOTH SIDES THE PHYSICIAN ELECTED TO IMPLANT ENDURANT STENT GRAFT LIMBS BILATERALLY XX THE ENDOLEAK APPEARED TO BE RESOLVED PER THE PHYSICIAN THE CAUSE OF THE EVENT WAS ANATOMY RELATED THE LIMBS PULLED UP AS THE ANEURYSM SAC GREW AND THE ILIACS DILATED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Insufficient Information | Device problem code
6796575 | 82757901 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM IT WAS REPORTED THAT HELIFX ENDOANCHORS WERE PLACED TO TREAT A TYPE IA ENDOLEAK IT WAS NOTED THAT OF THE ENDOANCHORS DID NOT ADEQUATELY PENETRATE THE AORTIC WALL BECAUSE THE DISTANCE BETWEEN THE STENT GRAFT AND THE AORTIC WALL WAS TOO GREAT ANOTHER MANUFACTURERS BARE METAL STENT WAS ALSO PLACED IN THE PROXIMAL NECK THE TYPE IA ENDOLEAK WAS REDUCED BUT STILL PRESENT A FOLLOW UP CT DONE ON SHOWED THAT THERE WERE NO ENDOLEAKS PRESENT THE PHYSICIAN INVESTIGATOR ASSESSED THE EVENT TO BE RELATED TO THE INDEX PROCEDURE A TYPE IA ENDOLEAK WAS AGAIN SEEN IT WAS NOTED THAT THE TYPE IA ENDOLEAK REMAINED FROM THE FAILURE OF THE ENDOANCHORS TO SEAL THE ANEURYSM SAC AS PREVIOUSLY NOTED THE PATIENT WAS HOSPITALIZED AND THE ENDOLEAK REPORTEDLY RESOLVED WITH TREATMENT A FIXATION POINT TYPE IA ENDOLEAK WAS THEN OBSERVED THE PATIENT WAS READMITTED TO THE HOSPITAL AND ADDITIONAL TREATMENT WAS PROVIDED THE EVENT REPORTEDLY RESOLVED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Aneurysm | Patient problem code
6796575 | 82757901 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM IT WAS REPORTED THAT HELIFX ENDOANCHORS WERE PLACED TO TREAT A TYPE IA ENDOLEAK IT WAS NOTED THAT OF THE ENDOANCHORS DID NOT ADEQUATELY PENETRATE THE AORTIC WALL BECAUSE THE DISTANCE BETWEEN THE STENT GRAFT AND THE AORTIC WALL WAS TOO GREAT ANOTHER MANUFACTURERS BARE METAL STENT WAS ALSO PLACED IN THE PROXIMAL NECK THE TYPE IA ENDOLEAK WAS REDUCED BUT STILL PRESENT A FOLLOW UP CT DONE ON SHOWED THAT THERE WERE NO ENDOLEAKS PRESENT THE PHYSICIAN INVESTIGATOR ASSESSED THE EVENT TO BE RELATED TO THE INDEX PROCEDURE A TYPE IA ENDOLEAK WAS AGAIN SEEN IT WAS NOTED THAT THE TYPE IA ENDOLEAK REMAINED FROM THE FAILURE OF THE ENDOANCHORS TO SEAL THE ANEURYSM SAC AS PREVIOUSLY NOTED THE PATIENT WAS HOSPITALIZED AND THE ENDOLEAK REPORTEDLY RESOLVED WITH TREATMENT A FIXATION POINT TYPE IA ENDOLEAK WAS THEN OBSERVED THE PATIENT WAS READMITTED TO THE HOSPITAL AND ADDITIONAL TREATMENT WAS PROVIDED THE EVENT REPORTEDLY RESOLVED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Hemorrhage | Patient problem code
6796575 | 82757901 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM IT WAS REPORTED THAT HELIFX ENDOANCHORS WERE PLACED TO TREAT A TYPE IA ENDOLEAK IT WAS NOTED THAT OF THE ENDOANCHORS DID NOT ADEQUATELY PENETRATE THE AORTIC WALL BECAUSE THE DISTANCE BETWEEN THE STENT GRAFT AND THE AORTIC WALL WAS TOO GREAT ANOTHER MANUFACTURERS BARE METAL STENT WAS ALSO PLACED IN THE PROXIMAL NECK THE TYPE IA ENDOLEAK WAS REDUCED BUT STILL PRESENT A FOLLOW UP CT DONE ON SHOWED THAT THERE WERE NO ENDOLEAKS PRESENT THE PHYSICIAN INVESTIGATOR ASSESSED THE EVENT TO BE RELATED TO THE INDEX PROCEDURE A TYPE IA ENDOLEAK WAS AGAIN SEEN IT WAS NOTED THAT THE TYPE IA ENDOLEAK REMAINED FROM THE FAILURE OF THE ENDOANCHORS TO SEAL THE ANEURYSM SAC AS PREVIOUSLY NOTED THE PATIENT WAS HOSPITALIZED AND THE ENDOLEAK REPORTEDLY RESOLVED WITH TREATMENT A FIXATION POINT TYPE IA ENDOLEAK WAS THEN OBSERVED THE PATIENT WAS READMITTED TO THE HOSPITAL AND ADDITIONAL TREATMENT WAS PROVIDED THE EVENT REPORTEDLY RESOLVED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Insufficient Information | Device problem code
6791052 | 82760609 | D | A HEALTHCARE PROFESSIONAL REPORTED THAT DURING A VITRECTOMY PROCEDURE THERE WAS A LOSS OF INTRAOCULAR PRESSURE IOP AND THE PATIENTS EYE COLLAPSED PATIENT IMPACT IS UNKNOWN AT THIS TIME ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT RECEIVED | Collapse | Patient problem code
6791052 | 82760609 | D | A HEALTHCARE PROFESSIONAL REPORTED THAT DURING A VITRECTOMY PROCEDURE THERE WAS A LOSS OF INTRAOCULAR PRESSURE IOP AND THE PATIENTS EYE COLLAPSED PATIENT IMPACT IS UNKNOWN AT THIS TIME ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT RECEIVED | Intraocular Pressure, Delayed, Uncontrolled | Patient problem code
6791052 | 82760609 | D | A HEALTHCARE PROFESSIONAL REPORTED THAT DURING A VITRECTOMY PROCEDURE THERE WAS A LOSS OF INTRAOCULAR PRESSURE IOP AND THE PATIENTS EYE COLLAPSED PATIENT IMPACT IS UNKNOWN AT THIS TIME ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT RECEIVED | Insufficient Information | Device problem code
6770275 | 82766731 | N | B BATCH NFK ADDITIONAL INFORMATION RECEIVED THE ECHELON SEEMED TO INITIALLY FIRE BUT THEN STOP AND AUTOMATICALLY REVERSE EVEN WITH THE TRIGGER DEPRESSED NOTHING UNUSUAL OF THE TISSUE WAS FOUND IN THE SPECIMEN WHEN IT WAS REMOVED FROM THE PATIENT OR THE DEVICE WHEN INSPECTED | Fire | Device problem code
6770275 | 82766731 | N | B BATCH NFK ADDITIONAL INFORMATION RECEIVED THE ECHELON SEEMED TO INITIALLY FIRE BUT THEN STOP AND AUTOMATICALLY REVERSE EVEN WITH THE TRIGGER DEPRESSED NOTHING UNUSUAL OF THE TISSUE WAS FOUND IN THE SPECIMEN WHEN IT WAS REMOVED FROM THE PATIENT OR THE DEVICE WHEN INSPECTED | Operational Problem Identified | Medical Evaluation result code
6796803 | 82772534 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO PACE A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS SET TO PACE AT BPM AND OUTPUTTED PACER CURRENT AT BPM COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Component or Accessory Incompatibility | Device Problem Code
6796803 | 82772534 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO PACE A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS SET TO PACE AT BPM AND OUTPUTTED PACER CURRENT AT BPM COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6796780 | 82774468 | D | INFORMATION WAS RECEIVED FROM A CONSUMER VIA A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR FOR CHRONIC LOW BACK PAIN AND SPINAL PAIN IT WAS REPORTED BY THE CONSUMER THAT THEY WERE HAVING BURNING SHARP PAIN AT THE INS SITE THE IMPEDANCES WERE ALL ON THE LEAD THE MANUFACTURER REPRESENTATIVE TURNED OFF THAT LEAD AND WAS ONLY USING THE LEAD THE PATIENT WAS GETTING BILATERAL COVERAGE THE PATIENT DENIED FALLS OR ACCIDENTS THAT COULD HAVE CONTRIBUTED TO THE EVENT AFTER THE REPROGRAMMING THE PATIENT CONFIRMED THE BURNING PAIN WENT AWAY ONCE ONLY THE LEAD WAS USED THE ISSUE WAS NOT YET FULLY RESOLVED IT WAS UNKNOWN IF A SURGICAL INTERVENTION WAS PLANNED PATIENT WEIGHT AND MEDICAL HISTORY WERE UNKNOWN THE PATIENT STATUS WAS CONFIRMED TO BE ALIVE WITH NO INJURY IT WAS UNKNOWN WHEN THIS EVENT STARTED NO FURTHER COMPLICATIONS WERE REPORTED | Pain | Patient problem code
6796780 | 82774468 | D | INFORMATION WAS RECEIVED FROM A CONSUMER VIA A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR FOR CHRONIC LOW BACK PAIN AND SPINAL PAIN IT WAS REPORTED BY THE CONSUMER THAT THEY WERE HAVING BURNING SHARP PAIN AT THE INS SITE THE IMPEDANCES WERE ALL ON THE LEAD THE MANUFACTURER REPRESENTATIVE TURNED OFF THAT LEAD AND WAS ONLY USING THE LEAD THE PATIENT WAS GETTING BILATERAL COVERAGE THE PATIENT DENIED FALLS OR ACCIDENTS THAT COULD HAVE CONTRIBUTED TO THE EVENT AFTER THE REPROGRAMMING THE PATIENT CONFIRMED THE BURNING PAIN WENT AWAY ONCE ONLY THE LEAD WAS USED THE ISSUE WAS NOT YET FULLY RESOLVED IT WAS UNKNOWN IF A SURGICAL INTERVENTION WAS PLANNED PATIENT WEIGHT AND MEDICAL HISTORY WERE UNKNOWN THE PATIENT STATUS WAS CONFIRMED TO BE ALIVE WITH NO INJURY IT WAS UNKNOWN WHEN THIS EVENT STARTED NO FURTHER COMPLICATIONS WERE REPORTED | Burning Sensation | Patient problem code
6796780 | 82774468 | D | INFORMATION WAS RECEIVED FROM A CONSUMER VIA A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR FOR CHRONIC LOW BACK PAIN AND SPINAL PAIN IT WAS REPORTED BY THE CONSUMER THAT THEY WERE HAVING BURNING SHARP PAIN AT THE INS SITE THE IMPEDANCES WERE ALL ON THE LEAD THE MANUFACTURER REPRESENTATIVE TURNED OFF THAT LEAD AND WAS ONLY USING THE LEAD THE PATIENT WAS GETTING BILATERAL COVERAGE THE PATIENT DENIED FALLS OR ACCIDENTS THAT COULD HAVE CONTRIBUTED TO THE EVENT AFTER THE REPROGRAMMING THE PATIENT CONFIRMED THE BURNING PAIN WENT AWAY ONCE ONLY THE LEAD WAS USED THE ISSUE WAS NOT YET FULLY RESOLVED IT WAS UNKNOWN IF A SURGICAL INTERVENTION WAS PLANNED PATIENT WEIGHT AND MEDICAL HISTORY WERE UNKNOWN THE PATIENT STATUS WAS CONFIRMED TO BE ALIVE WITH NO INJURY IT WAS UNKNOWN WHEN THIS EVENT STARTED NO FURTHER COMPLICATIONS WERE REPORTED | Insufficient Information | Device problem code
6794300 | 82780414 | D | AS REPORTED THE PATIENT COMPLAINED STRONGLY ABOUT THE INTRODUCTION AND WITHDRAWAL OF THE AVANTI TRANSRADIAL SHEATH SI AVANTI TRANSRAD KIT CM A DEVICE THAT IS NOT HYDROPHILIC THERE WAS NO REPORTED PATIENT INJURY THE INDEX PROCEDURE WAS A CORONARY ANGIOPLASTY PROCEDURE THERE WERE NO KINKS OR OTHER DAMAGES NOTED PRIOR TO INSERTING THE PRODUCT INTO THE PATIENT THE DEVICE WAS PREPPED ACCORDING TO INSTRUCTIONS FOR USE IFU THE DEVICE WAS PREP NORMALLY THE PATIENT COMPLAINED OF PAIN AT THE TIME OF INSERTING THE INTRODUCER THE PROCEDURE WAS CONTINUED AND TO TREAT THE REACTION AN ANALGESIC WAS PLACED FOR TREATMENT AFTER THE TREATMENT THE REACTION RESOLVED IMMEDIATELY THERE WAS A LITTLE RESISTANCE ADVANCING THE DEVICE THE PRODUCT WAS REMOVED INTACT IN ONE PIECE FROM THE PATIENT THERE WAS NO INJURY TO THE PATIENT AND THE PATIENT IS IN PERFECT CONDITION | No Consequences Or Impact To Patient | Patient problem code
6794300 | 82780414 | D | AS REPORTED THE PATIENT COMPLAINED STRONGLY ABOUT THE INTRODUCTION AND WITHDRAWAL OF THE AVANTI TRANSRADIAL SHEATH SI AVANTI TRANSRAD KIT CM A DEVICE THAT IS NOT HYDROPHILIC THERE WAS NO REPORTED PATIENT INJURY THE INDEX PROCEDURE WAS A CORONARY ANGIOPLASTY PROCEDURE THERE WERE NO KINKS OR OTHER DAMAGES NOTED PRIOR TO INSERTING THE PRODUCT INTO THE PATIENT THE DEVICE WAS PREPPED ACCORDING TO INSTRUCTIONS FOR USE IFU THE DEVICE WAS PREP NORMALLY THE PATIENT COMPLAINED OF PAIN AT THE TIME OF INSERTING THE INTRODUCER THE PROCEDURE WAS CONTINUED AND TO TREAT THE REACTION AN ANALGESIC WAS PLACED FOR TREATMENT AFTER THE TREATMENT THE REACTION RESOLVED IMMEDIATELY THERE WAS A LITTLE RESISTANCE ADVANCING THE DEVICE THE PRODUCT WAS REMOVED INTACT IN ONE PIECE FROM THE PATIENT THERE WAS NO INJURY TO THE PATIENT AND THE PATIENT IS IN PERFECT CONDITION | Adverse Event Without Identified Device or Use Problem | Device problem code
6794283 | 82781367 | D | AS REPORTED THE PATIENT COMPLAINED STRONGLY ABOUT THE INTRODUCTION AND WITHDRAWAL OF THE AVANTI TRANSRADIAL SHEATH SI AVANTI TRANSRAD KIT CM A DEVICE THAT IS NOT HYDROPHILIC THERE WAS NO REPORTED PATIENT INJURY THE INDEX PROCEDURE WAS A CORONARY ANGIOPLASTY PROCEDURE THERE WERE NO KINKS OR OTHER DAMAGES NOTED PRIOR TO INSERTING THE PRODUCT INTO THE PATIENT THE DEVICE WAS PREPPED ACCORDING TO INSTRUCTIONS FOR USE IFU THE DEVICE WAS PREP NORMALLY THE PATIENT COMPLAINED OF PAIN AT THE TIME OF INSERTING THE INTRODUCER THE PROCEDURE WAS CONTINUED AND TO TREAT THE REACTION AN ANALGESIC WAS PLACED FOR TREATMENT AFTER THE TREATMENT THE REACTION RESOLVED IMMEDIATELY THERE WAS A LITTLE RESISTANCE ADVANCING THE DEVICE THE PRODUCT WAS REMOVED INTACT IN ONE PIECE FROM THE PATIENT THERE WAS NO INJURY TO THE PATIENT AND THE PATIENT IS IN PERFECT CONDITION | No Consequences Or Impact To Patient | Patient problem code
6794283 | 82781367 | D | AS REPORTED THE PATIENT COMPLAINED STRONGLY ABOUT THE INTRODUCTION AND WITHDRAWAL OF THE AVANTI TRANSRADIAL SHEATH SI AVANTI TRANSRAD KIT CM A DEVICE THAT IS NOT HYDROPHILIC THERE WAS NO REPORTED PATIENT INJURY THE INDEX PROCEDURE WAS A CORONARY ANGIOPLASTY PROCEDURE THERE WERE NO KINKS OR OTHER DAMAGES NOTED PRIOR TO INSERTING THE PRODUCT INTO THE PATIENT THE DEVICE WAS PREPPED ACCORDING TO INSTRUCTIONS FOR USE IFU THE DEVICE WAS PREP NORMALLY THE PATIENT COMPLAINED OF PAIN AT THE TIME OF INSERTING THE INTRODUCER THE PROCEDURE WAS CONTINUED AND TO TREAT THE REACTION AN ANALGESIC WAS PLACED FOR TREATMENT AFTER THE TREATMENT THE REACTION RESOLVED IMMEDIATELY THERE WAS A LITTLE RESISTANCE ADVANCING THE DEVICE THE PRODUCT WAS REMOVED INTACT IN ONE PIECE FROM THE PATIENT THERE WAS NO INJURY TO THE PATIENT AND THE PATIENT IS IN PERFECT CONDITION | Adverse Event Without Identified Device or Use Problem | Device problem code
6775361 | 82781435 | D | THE HOSPITAL REPORTED THAT DURING AN ENDOSCOPIC VEIN HARVESTING PROCEDURE VASOVIEW HEMOPRO CAUTERY FORCEPS WERE NOT CUTTING BURNING THROUGH TISSUE EFFECTIVELY TAKING EXTENDED PERIOD TO CUT THROUGH TISSUE OR NOT CUTTING AT ALL ALL THE CABLE CONNECTIONS WERE CHECKED AND REATTACHED THE SETTING ON THE POWER SUPPLY WAS INCREASED THE JAWS WERE CLEANED BUT THERE WAS NO CHANGE IN EFFECTIVENESS A REPLACEMENT DEVICE WAS USED TO COMPLETE THE PROCEDURE THE HOSPITAL DID NOT REPORT ANY PATIENT EFFECTS | Failure to Cut | Device Problem Code
6775361 | 82781435 | D | THE HOSPITAL REPORTED THAT DURING AN ENDOSCOPIC VEIN HARVESTING PROCEDURE VASOVIEW HEMOPRO CAUTERY FORCEPS WERE NOT CUTTING BURNING THROUGH TISSUE EFFECTIVELY TAKING EXTENDED PERIOD TO CUT THROUGH TISSUE OR NOT CUTTING AT ALL ALL THE CABLE CONNECTIONS WERE CHECKED AND REATTACHED THE SETTING ON THE POWER SUPPLY WAS INCREASED THE JAWS WERE CLEANED BUT THERE WAS NO CHANGE IN EFFECTIVENESS A REPLACEMENT DEVICE WAS USED TO COMPLETE THE PROCEDURE THE HOSPITAL DID NOT REPORT ANY PATIENT EFFECTS | No Known Impact Or Consequence To Patient | Patient problem code
6783867 | 82783805 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6783867 | 82783805 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6797322 | 82786335 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED FOR DAY AS A RESULT OF BEING DIAGNOSED WITH DIABETIC KETOACIDOSIS AFTER WEARING THE POD ARM A HAZARD ALARM WAS EMITTED HOWEVER NO FURTHER DETAILS COULD BE OBTAINED FOR THIS EVENT DESPITE GOOD FAITH EFFORT TO DO SO | Insufficient Information | Device problem code
6797322 | 82786335 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED FOR DAY AS A RESULT OF BEING DIAGNOSED WITH DIABETIC KETOACIDOSIS AFTER WEARING THE POD ARM A HAZARD ALARM WAS EMITTED HOWEVER NO FURTHER DETAILS COULD BE OBTAINED FOR THIS EVENT DESPITE GOOD FAITH EFFORT TO DO SO | Diabetic Ketoacidosis | Patient problem code
6783872 | 82788596 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6783872 | 82788596 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6797388 | 82788732 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE ASSOCIATED DEVICE INAPPROPRIATELY RECOGNIZED THESE ADULT ELECTRODE PADS AS PEDIATRIC ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Failure to Sense | Device Problem Code
6797388 | 82788732 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE ASSOCIATED DEVICE INAPPROPRIATELY RECOGNIZED THESE ADULT ELECTRODE PADS AS PEDIATRIC ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Known Impact Or Consequence To Patient | Patient problem code
6797425 | 82788736 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS DKA THE CAUSE OF THE HIGH BG WAS NOT KNOWN THERE WAS NO ISSUES OBSERVED WITH THE CARTRIDGE INFUSION TUBING OR CANNULA THE CUSTOMER WAS TREATED WITH INTRAVENOUS INSULIN AND FLUIDS THE CUSTOMER WAS DISCHARGED ON B AS THE CONTACT DID NOT HAVE ACCESS TO THE CUSTOMERS PUMP AT THE TIME OF THE REPORT TANDEM TECHNICAL SUPPORT WAS UNABLE TO PERFORM A PUMP SYSTEM CHECK THE PUMP DATA SHOWED THAT THE CARTRIDGE HAD BEEN USED FOR DAYS CTS INFORMED THE CUSTOMER THAT HUMALOG WAS LABELED FOR HOUR USE WITH THE CARTRIDGE THE CONTACT ACKNOWLEDGED THE INFORMATION | Hyperglycemia | Patient Problem Code
6797425 | 82788736 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS DKA THE CAUSE OF THE HIGH BG WAS NOT KNOWN THERE WAS NO ISSUES OBSERVED WITH THE CARTRIDGE INFUSION TUBING OR CANNULA THE CUSTOMER WAS TREATED WITH INTRAVENOUS INSULIN AND FLUIDS THE CUSTOMER WAS DISCHARGED ON B AS THE CONTACT DID NOT HAVE ACCESS TO THE CUSTOMERS PUMP AT THE TIME OF THE REPORT TANDEM TECHNICAL SUPPORT WAS UNABLE TO PERFORM A PUMP SYSTEM CHECK THE PUMP DATA SHOWED THAT THE CARTRIDGE HAD BEEN USED FOR DAYS CTS INFORMED THE CUSTOMER THAT HUMALOG WAS LABELED FOR HOUR USE WITH THE CARTRIDGE THE CONTACT ACKNOWLEDGED THE INFORMATION | Diabetic Ketoacidosis | Patient problem code
6797425 | 82788736 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS OVER MGDL THE CUSTOMER WAS HOSPITALIZED FOR DIABETIC KETOACIDOSIS DKA THE CAUSE OF THE HIGH BG WAS NOT KNOWN THERE WAS NO ISSUES OBSERVED WITH THE CARTRIDGE INFUSION TUBING OR CANNULA THE CUSTOMER WAS TREATED WITH INTRAVENOUS INSULIN AND FLUIDS THE CUSTOMER WAS DISCHARGED ON B AS THE CONTACT DID NOT HAVE ACCESS TO THE CUSTOMERS PUMP AT THE TIME OF THE REPORT TANDEM TECHNICAL SUPPORT WAS UNABLE TO PERFORM A PUMP SYSTEM CHECK THE PUMP DATA SHOWED THAT THE CARTRIDGE HAD BEEN USED FOR DAYS CTS INFORMED THE CUSTOMER THAT HUMALOG WAS LABELED FOR HOUR USE WITH THE CARTRIDGE THE CONTACT ACKNOWLEDGED THE INFORMATION | Adverse Event Without Identified Device or Use Problem | Device problem code
6797392 | 82788866 | D | DURING A DENTAL PROCEDURE A PATIENT HAD CONTACT WITH A LIP RETRACTOR THAT HAD BEEN PREVIOUSLY CLEANED WITH SPOROX II STERILIZATION AND DISINFECTANT SOLUTION THE PATIENT FELT A BURNING SENSATION FIRST IN HER HAND AND IT TURNED WHITE THEN AFTER TEN MINUTES SHE FELT A BURNING SENSATION THAT WENT FROM HER ESOPHAGUS TO HER STOMACH AN HOUR LATER SHE WAS TAKEN TO THE ER THE REACTION WAS IMMEDIATE AFTER CONTACT WITH LIP RETRACTOR SHE HAD NO PREVIOUS SYMPTOMS BEFORE THE CONTACT THE EVENT OUTCOME IS UNKNOWN AS OF THIS MDR EVALUATION | Insufficient Information | Device problem code
6797392 | 82788866 | D | DURING A DENTAL PROCEDURE A PATIENT HAD CONTACT WITH A LIP RETRACTOR THAT HAD BEEN PREVIOUSLY CLEANED WITH SPOROX II STERILIZATION AND DISINFECTANT SOLUTION THE PATIENT FELT A BURNING SENSATION FIRST IN HER HAND AND IT TURNED WHITE THEN AFTER TEN MINUTES SHE FELT A BURNING SENSATION THAT WENT FROM HER ESOPHAGUS TO HER STOMACH AN HOUR LATER SHE WAS TAKEN TO THE ER THE REACTION WAS IMMEDIATE AFTER CONTACT WITH LIP RETRACTOR SHE HAD NO PREVIOUS SYMPTOMS BEFORE THE CONTACT THE EVENT OUTCOME IS UNKNOWN AS OF THIS MDR EVALUATION | Burning Sensation | Patient problem code
6783881 | 82789122 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6783881 | 82789122 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6797437 | 82791439 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV FOR HYDRATION AND INSULIN WAS DELIVERED THROUGH IV | Diabetic Ketoacidosis | Patient problem code
6797437 | 82791439 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV FOR HYDRATION AND INSULIN WAS DELIVERED THROUGH IV | Insufficient Information | Device problem code
6779144 | 82793314 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP WAS NOT ABLE TO REWIND CUSTOMER BLOOD GLUCOSE READING WAS MGDL CUSTOMER STATES PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD CUSTOMER WAS NOT ABLE TO CHECK HISTORY FOR ANY ALARMS CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE PUMP AND REVERT TO BACKUP PLAN PER HEALTHCARE PROVIDER INSTRUCTIONS THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6779144 | 82793314 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THE INSULIN PUMP WAS NOT ABLE TO REWIND CUSTOMER BLOOD GLUCOSE READING WAS MGDL CUSTOMER STATES PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD CUSTOMER WAS NOT ABLE TO CHECK HISTORY FOR ANY ALARMS CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE PUMP AND REVERT TO BACKUP PLAN PER HEALTHCARE PROVIDER INSTRUCTIONS THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Pumping Problem | Device problem code
6777316 | 82808650 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING B | No Findings Available | Medical Evaluation result code
6797458 | 82808889 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6797458 | 82808889 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6797449 | 82809454 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Hyperglycemia | Patient Problem Code
6797449 | 82809454 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6797260 | 82809661 | D | INFORMATION WAS RECEIVED VIA A MANUFACTURER REPRESENTATIVE FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING BURNING WHEN STIMULATION IS ON STARTING ON B A MANUFACTURER REPRESENTATIVE IS MEETING WITH THE PATIENT ON B TO EVALUATE THE SYSTEM THE ISSUE WAS NOT RESOLVED AT THE TIME OF THE REPORT THERE WERE NO FURTHER COMPLICATIONS REPORTED | Insufficient Information | Device problem code
6797260 | 82809661 | D | INFORMATION WAS RECEIVED VIA A MANUFACTURER REPRESENTATIVE FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING BURNING WHEN STIMULATION IS ON STARTING ON B A MANUFACTURER REPRESENTATIVE IS MEETING WITH THE PATIENT ON B TO EVALUATE THE SYSTEM THE ISSUE WAS NOT RESOLVED AT THE TIME OF THE REPORT THERE WERE NO FURTHER COMPLICATIONS REPORTED | Burning Sensation | Patient problem code
6794975 | 82816118 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AT MONTHS POSTOPERATION AND THERE WAS A RIDGE WIDTH PROBLEM NO WEIGHT OF PATIENT WAS GIVEN THE PATIENT HAD A HISTORY OF TRAUMA AND HAD A SOCKET GRAFT PRIOR TO PLACING THE IMPLANT ALSO THE PATIENT HAD BONE TYPE III AND NO PREEXISTING CONDITIONS BUT EXPERIENCED GENERAL BONE LOSS THERE WAS NO SERIOUS INJURY OR HARM TO THE PATIENT THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES OSSEOINTEGRATION HAS BEEN SHOWN TO DEPEND ON THE TYPE AND VOLUME OF BONE AVAILABLE AT THE SITE OF IMPLANTATION FAILURE TO OSSEOINTEGRATION IS AN INHERENT RISK IN IMPLANT SURGERY AND OCCURS OF ALL IMPLANT CASES ACCORDING TO PUBLISHED LITERATURE ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | No Findings Available | Medical Evaluation result code
6794785 | 82816297 | N | THE DENTIST REPORTED AN B PATIENT WITH BONE TYPE III EXPERIENCED IMPLANT FAILURE TO OSSEOINTEGRATE AFTER MONTHS AND THE PATIENT HAD RIDGE WIDTH PROBLEM NO WEIGHT OF THE PATIENT WAS GIVEN THE PATIENT HAD NO HEALTH ISSUES OR PREEXISTING CONDITIONS BUT EXPERIENCE GENERAL BONE LOSS THERE WAS NO SERIOUS INJURY OR HARM TO THE PATIENT THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES PUBLISHED STUDIES SHOW THAT OF IMPLANTS FAIL TO OSSEOINTEGRATE AND THIS IS AN INHERENT RISK OF IMPLANT SURGERY POOR BONE QUALITY AND AGE GREATER THAN YEARS ARE KNOWN TO CONTRIBUTE TO AN INCREASED POSSIBILITY OF DENTAL IMPLANTS FAILING TO OSSEOINTEGRATE ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | No Device Problem Found | Medical Evaluation result code
6794918 | 82816807 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER OR WEEKS PERIOD AND THERE WAS A RIDGE WIDTH ISSUE AND THAT THE DENTIST BLEW OUT THE RIDGE NO WEIGHT OF THE PATIENT WAS GIVEN THE PATIENT HAD BONE TYPE III AND NO PREEXISTING CONDITIONS BUT EXPERIENCE GENERAL BONE LOSS THERE WAS NO SERIOUS INJURY OR HARM TO THE PATIENT THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES OSSEOINTEGRATION HAS BEEN SHOWN TO DEPEND ON THE TYPE AND VOLUME OF BONE AVAILABLE AT THE SITE OF IMPLANTATION FAILURE TO OSSEOINTEGRATE IS AN INHERENT RISK IN IMPLANT SURGERY AND OCCURS IN OF ALL IMPLANT CASES ACCORDING TO PUBLISHED LITERATURE ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | No Device Problem Found | Medical Evaluation result code
6798380 | 82824642 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Diabetic Ketoacidosis | Patient problem code
6798380 | 82824642 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6798380 | 82824642 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B WITH HIGH BLOOD GLUCOSE THEY WERE TREATED AND RELEASED ON THE SAME NIGHT THE CUSTOMER WAS ALSO HOSPITALIZED ON B DUE TO A HIGH BLOOD GLUCOSE LEVEL OF MGDL AND DIABETIC KETOACIDOSIS THE CUSTOMER STATED THEY KEPT TREATING THEIR BLOOD GLUCOSE BUT IT WOULD NOT COME DOWN THEY WERE CURRENTLY AT THE HOSPITAL WITH AN INSULIN DRIP THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED ON THE INSULIN PUMP AND INSULIN EXITED WITHOUT INCIDENT THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE BEEPED DURING THE TONE TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Noise, Audible | Device problem code
6798570 | 82845029 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device Problem Code
6799326 | 82859585 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE LEVEL OF MGDL CUSTOMER DECLINED TROUBLESHOOTING FOR LOW BLOOD GLUCOSE READING CUSTOMER STATED NOT WEARING THE INSULIN PUMP CAUSED LOW BLOOD GLUCOSE READING INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Insufficient Information | Device problem code
6799326 | 82859585 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE LEVEL OF MGDL CUSTOMER DECLINED TROUBLESHOOTING FOR LOW BLOOD GLUCOSE READING CUSTOMER STATED NOT WEARING THE INSULIN PUMP CAUSED LOW BLOOD GLUCOSE READING INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Hypoglycemia | Patient problem code
6798849 | 82862370 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN ANTIBACTERIAL MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Infection | Patient problem code
6798849 | 82862370 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN ANTIBACTERIAL MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Insufficient Information | Device problem code
6798849 | 82862370 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN ANTIBACTERIAL MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Abscess | Patient problem code
6798849 | 82862370 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN ANTIBACTERIAL MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Bruise | Patient problem code
6799290 | 82864268 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON B WITH UNSPECIFIED SIGNSSYMPTOMS OF A STROKE AT THE TIME OF ADMISSION THE PATIENT HAD A INR OF A CT SCAN REVEALED A CLOT IN THE MIDDLE CEREBRAL ARTERY A THROMBECTOMY WAS ATTEMPTED BUT WAS UNSUCCESSFUL AS OF B THE PATIENT HAD RESIDUAL LEFT VISUAL IMPAIRMENT IT WAS REPORTED THAT THE PATIENT WOULD BE MONITORED AND MEDICALLY MANAGED ADDITIONAL INFORMATION WAS REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device problem code
6799290 | 82864268 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON B WITH UNSPECIFIED SIGNSSYMPTOMS OF A STROKE AT THE TIME OF ADMISSION THE PATIENT HAD A INR OF A CT SCAN REVEALED A CLOT IN THE MIDDLE CEREBRAL ARTERY A THROMBECTOMY WAS ATTEMPTED BUT WAS UNSUCCESSFUL AS OF B THE PATIENT HAD RESIDUAL LEFT VISUAL IMPAIRMENT IT WAS REPORTED THAT THE PATIENT WOULD BE MONITORED AND MEDICALLY MANAGED ADDITIONAL INFORMATION WAS REQUESTED BUT NOT YET PROVIDED | Cerebrovascular Accident | Patient problem code
6799290 | 82864268 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON B WITH UNSPECIFIED SIGNSSYMPTOMS OF A STROKE AT THE TIME OF ADMISSION THE PATIENT HAD A INR OF A CT SCAN REVEALED A CLOT IN THE MIDDLE CEREBRAL ARTERY A THROMBECTOMY WAS ATTEMPTED BUT WAS UNSUCCESSFUL AS OF B THE PATIENT HAD RESIDUAL LEFT VISUAL IMPAIRMENT IT WAS REPORTED THAT THE PATIENT WOULD BE MONITORED AND MEDICALLY MANAGED ADDITIONAL INFORMATION WAS REQUESTED BUT NOT YET PROVIDED | Vision, Impaired | Patient problem code
6798872 | 82871047 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Insufficient Information | Device problem code
6798872 | 82871047 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Bruise | Patient problem code
6798872 | 82871047 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Infection | Patient problem code
6799842 | 82873230 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT FELT SICK WAS SUBSEQUENTLY TAKEN TO THE EMERGENCY ROOM AND TREATED WITH ADMINISTRATION OF FAST ACTING INSULIN IT SHOULD BE NOTED THAT PATIENT HAD DIFFICULTY DURING THE ACTIVATION PROCESS AND NOTED IT WAS HARD TO FILL THE POD THE POD HAD EMITTED AN OCCLUSION ALARM AT SOME POINT DURING THIS EVENT HOWEVER ITS UNCLEAR AS TO THE POINT IN TIME NO FURTHER INFORMATION IS AVAILABLE | Filling Problem | Device problem code
6799842 | 82873230 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT FELT SICK WAS SUBSEQUENTLY TAKEN TO THE EMERGENCY ROOM AND TREATED WITH ADMINISTRATION OF FAST ACTING INSULIN IT SHOULD BE NOTED THAT PATIENT HAD DIFFICULTY DURING THE ACTIVATION PROCESS AND NOTED IT WAS HARD TO FILL THE POD THE POD HAD EMITTED AN OCCLUSION ALARM AT SOME POINT DURING THIS EVENT HOWEVER ITS UNCLEAR AS TO THE POINT IN TIME NO FURTHER INFORMATION IS AVAILABLE | No Known Impact Or Consequence To Patient | Patient Problem Code
6799849 | 82873502 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Infection | Patient problem code
6799849 | 82873502 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Bruise | Patient problem code
6799849 | 82873502 | D | IT WAS REPORTED THAT AN INFECTION CAUSING BRUISING AND PUSS ALONG WITH A BLOOD GLUCOSE OF MGDL HAD OCCURRED WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM PATIENT VISITED PHYSICIAN AND WAS GIVEN AN UNKNOWN MEDICINE AS TREATMENT IT WAS NOT CLEARLY KNOWN IF THE SAID MEDICINE WAS PRESCRIPTION OR NOT FURTHERMORE IT WAS NOT CLEARLY KNOWN IF THE INFECTION WAS CLINICALLY OR SELF DIAGNOSED | Insufficient Information | Device problem code
6778920 | 82875652 | N | B PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER FILE NUMBER DUE TO SOFTWARE ERROR PUMP RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6778920 | 82875652 | N | B PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM LINE NUMBER FILE NUMBER DUE TO SOFTWARE ERROR PUMP RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | No Device Problem Found | Medical Evaluation result code
6799501 | 82878882 | D | THIS COMPLAINT IS FROM A LITERATURE SOURCE IT WAS REPORTED THAT PATIENTS WITH ATRIAL FIBRILLATION DISEASE UNDERWENT CATHETER ABLATION FROM TO THE PATIENTS WERE DIVIDED INTO GROUPS AN RMN WITH REMOTE CONTROLLED MAGNETIC NAVIGATION N GROUP AND A MAN MANUAL ABLATIONN AMONG THEM ONE PATIENT WITH A BRANCH RETINAL VEIN OCCLUSION IN THE RMN GROUP BASED ON THE FACTS OF THE CASE AND THE AUTHORS ASSESSMENT THERE WERE NO REPORTED MALFUNCTION WITH ANY OF THE BWI CATHETERS AND SYSTEMS USED DURING THE CASE THUS THIS EVENT IS UNRELATED TO THE DEVICE AND MOST LIKELY RELATED TO THE PROCEDURE TITLE REMOTE MAGNETIC CATHETER NAVIGATION VERSUS CONVENTIONAL ABLATION IN ATRIAL FIBRILLATION ABLATIONFLUOROSCOPY REDUCTION THE PRIMARY PURPOSE OF THIS STUDY WAS TO FLUOROSCOPY AND TOTAL PROCEDURAL TIMES BETWEEN RMN AND CONVENTIONAL MANUAL MAN ABLATION OF ATRIAL FIBRILLATION SECONDARY OBJECTIVES WERE TO COMPARE THE EFFICACY AND SAFETY OF BOTH ABLATION TECHNIQUES SUSPECT DEVICE IS A NAVISTAR RMT THERMOCOOL HOWEVER CATALOG AND LOT NUMBER IS UNKNOWN | Atrial Fibrillation | Patient problem code
6799501 | 82878882 | D | THIS COMPLAINT IS FROM A LITERATURE SOURCE IT WAS REPORTED THAT PATIENTS WITH ATRIAL FIBRILLATION DISEASE UNDERWENT CATHETER ABLATION FROM TO THE PATIENTS WERE DIVIDED INTO GROUPS AN RMN WITH REMOTE CONTROLLED MAGNETIC NAVIGATION N GROUP AND A MAN MANUAL ABLATIONN AMONG THEM ONE PATIENT WITH A BRANCH RETINAL VEIN OCCLUSION IN THE RMN GROUP BASED ON THE FACTS OF THE CASE AND THE AUTHORS ASSESSMENT THERE WERE NO REPORTED MALFUNCTION WITH ANY OF THE BWI CATHETERS AND SYSTEMS USED DURING THE CASE THUS THIS EVENT IS UNRELATED TO THE DEVICE AND MOST LIKELY RELATED TO THE PROCEDURE TITLE REMOTE MAGNETIC CATHETER NAVIGATION VERSUS CONVENTIONAL ABLATION IN ATRIAL FIBRILLATION ABLATIONFLUOROSCOPY REDUCTION THE PRIMARY PURPOSE OF THIS STUDY WAS TO FLUOROSCOPY AND TOTAL PROCEDURAL TIMES BETWEEN RMN AND CONVENTIONAL MANUAL MAN ABLATION OF ATRIAL FIBRILLATION SECONDARY OBJECTIVES WERE TO COMPARE THE EFFICACY AND SAFETY OF BOTH ABLATION TECHNIQUES SUSPECT DEVICE IS A NAVISTAR RMT THERMOCOOL HOWEVER CATALOG AND LOT NUMBER IS UNKNOWN | Adverse Event Without Identified Device or Use Problem | Device problem code
6800067 | 82891420 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED ON B THAT THE PUMP POWER WAS SPIKING HIGHER THAN NORMAL PUMP POWER AND THEN RESULTANT FLOWS PERIODICALLY FOR THE PREVIOUS FEW DAYS THE PATIENTS LACTATE DEHYDROGENASE LDH WAS JUST OVER UL AND THEIR URINE WAS SLIGHTLY TEA COLORED THE PATIENT WAS ADMITTED INTO THE HOSPITAL THE SUBMITTED LOG FILE WAS REVIEWED BY THE MANUFACTURERS TECHNICAL SERVICES REPRESENTATIVE AND THE DATA SHOWED A TYPE OF TRAJECTORY THAT HAS BEEN LINKED TO THE POSSIBLE PRESENCE OF THROMBUS IT WAS REPORTED ON B THAT THE ECHOCARDIOGRAM WAS NEGATIVE THE PATIENT WAS GIVEN HEPARIN AND THE ELEVATED LDH RESOLVED THE PATIENT WAS DISCHARGED TO HOME ON AN UNSPECIFIED DATE AND IS BEING FOLLOWED UP IN THE CLINIC | Thrombus | Patient problem code
6800067 | 82891420 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED ON B THAT THE PUMP POWER WAS SPIKING HIGHER THAN NORMAL PUMP POWER AND THEN RESULTANT FLOWS PERIODICALLY FOR THE PREVIOUS FEW DAYS THE PATIENTS LACTATE DEHYDROGENASE LDH WAS JUST OVER UL AND THEIR URINE WAS SLIGHTLY TEA COLORED THE PATIENT WAS ADMITTED INTO THE HOSPITAL THE SUBMITTED LOG FILE WAS REVIEWED BY THE MANUFACTURERS TECHNICAL SERVICES REPRESENTATIVE AND THE DATA SHOWED A TYPE OF TRAJECTORY THAT HAS BEEN LINKED TO THE POSSIBLE PRESENCE OF THROMBUS IT WAS REPORTED ON B THAT THE ECHOCARDIOGRAM WAS NEGATIVE THE PATIENT WAS GIVEN HEPARIN AND THE ELEVATED LDH RESOLVED THE PATIENT WAS DISCHARGED TO HOME ON AN UNSPECIFIED DATE AND IS BEING FOLLOWED UP IN THE CLINIC | Power Problem | Device problem code
6794706 | 82893250 | N | THE HILLROM TECHNICIAN WAS UNABLE TO LOCATE THE BED PER THE HILLROM USER MANUAL WARNING THE BED EXIT ALARM SYSTEM IS NOT INTENDED AS A SUBSTITUTE FOR GOOD NURSING PRACTICES THE BED EXIT ALARM SYSTEM MUST BE USED IN CONJUNCTION WITH A SOUND RISK ASSESSMENT AND PROTOCOL PER THE HILLROM USER MANUAL IF THE BED EXIT ALARM DOES NOT ARM AND ALL THREE MODE INDICATORS ARE FLASHING REMOVE THE PATIENT AND ZERO THE BED EXIT SYSTEM A SEARCH OF THE HILLROM MAINTENANCE RECORDS DID NOT SHOW HILLROM PERFORMED ANY PREVENTATIVE MAINTENANCE ON THIS BED IT IS UNKNOWN IF THE FACILITY PERFORMS PREVENTATIVE MAINTENANCE ON THEIR BEDS HILLROM HAS CONTACTED THE ACCOUNT THREE TIMES ASKING IF THEY ARE ABLE TO LOCATE THE BED WITH THIS ALLEGED ISSUE THE ACCOUNT IS UNABLE TO LOCATE THE BED AND IS NO LONGER LOOKING FOR IT BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | No Findings Available | Medical Evaluation result code
6799130 | 82894504 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE ILLUMINATOR WAS INSTALLED INTO A TEST SYSTEM AND POWERED ON LAMP MODULE MESSAGE AND ERROR APPEARED THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Component Incompatible | Device problem code
6799130 | 82894504 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE ILLUMINATOR WAS INSTALLED INTO A TEST SYSTEM AND POWERED ON LAMP MODULE MESSAGE AND ERROR APPEARED THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Incompatible Component/ Accessory | Medical Evaluation result code
6800181 | 82896335 | D | THE CUSTOMER REPORTED THE FOLLOWING A PATIENT UNDERGOING A CT SCAN TO RULE OUT A PULMONARY EMBOLISM SUFFERED AN EXTRAVASATION OF AN UNKNOWN AMOUNT OF CONTRAST WHILE CONNECTED TO THE ENVISION CT INJECTOR THE SITE WAS UNABLE TO CONFIRM WHICH ARM THE EXTRAVASATION OCCURRED WHEN THE PATIENT COMPLAINED OF PAIN DURING THE CONTRAST INJECTION THE TECHNOLOGIST ABORTED THE INJECTION THE STUDY WAS INCOMPLETE AND THE PATIENT REFUSED A REPEAT EXAM WARM AND COLD COMPRESSES WERE APPLIED TO THE AFFECTED LIMB AND THE PATIENT WAS DISCHARGED TO HOME THE PATIENT RETURNED TO THE EMERGENCY DEPARTMENT A SHORT TIME LATER AFTER WHICH AN UNDISCLOSED MEDICAL TREATMENT WAS PROVIDED | Insufficient Information | Device problem code
6800181 | 82896335 | D | THE CUSTOMER REPORTED THE FOLLOWING A PATIENT UNDERGOING A CT SCAN TO RULE OUT A PULMONARY EMBOLISM SUFFERED AN EXTRAVASATION OF AN UNKNOWN AMOUNT OF CONTRAST WHILE CONNECTED TO THE ENVISION CT INJECTOR THE SITE WAS UNABLE TO CONFIRM WHICH ARM THE EXTRAVASATION OCCURRED WHEN THE PATIENT COMPLAINED OF PAIN DURING THE CONTRAST INJECTION THE TECHNOLOGIST ABORTED THE INJECTION THE STUDY WAS INCOMPLETE AND THE PATIENT REFUSED A REPEAT EXAM WARM AND COLD COMPRESSES WERE APPLIED TO THE AFFECTED LIMB AND THE PATIENT WAS DISCHARGED TO HOME THE PATIENT RETURNED TO THE EMERGENCY DEPARTMENT A SHORT TIME LATER AFTER WHICH AN UNDISCLOSED MEDICAL TREATMENT WAS PROVIDED | Extravasation | Patient problem code
6800181 | 82896335 | D | THE CUSTOMER REPORTED THE FOLLOWING A PATIENT UNDERGOING A CT SCAN TO RULE OUT A PULMONARY EMBOLISM SUFFERED AN EXTRAVASATION OF AN UNKNOWN AMOUNT OF CONTRAST WHILE CONNECTED TO THE ENVISION CT INJECTOR THE SITE WAS UNABLE TO CONFIRM WHICH ARM THE EXTRAVASATION OCCURRED WHEN THE PATIENT COMPLAINED OF PAIN DURING THE CONTRAST INJECTION THE TECHNOLOGIST ABORTED THE INJECTION THE STUDY WAS INCOMPLETE AND THE PATIENT REFUSED A REPEAT EXAM WARM AND COLD COMPRESSES WERE APPLIED TO THE AFFECTED LIMB AND THE PATIENT WAS DISCHARGED TO HOME THE PATIENT RETURNED TO THE EMERGENCY DEPARTMENT A SHORT TIME LATER AFTER WHICH AN UNDISCLOSED MEDICAL TREATMENT WAS PROVIDED | Pain | Patient problem code
6800216 | 82896897 | D | IT WAS REPORTED THROMBUS ON THE DEVICE OCCURRED IN B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS SUCCESSFULLY PERFORMED A MM WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PATIENT WAS DISCHARGED ON WARFARIN AND MG ASA THE PATIENT RETURNED FOR A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE IN B WHICH REVEALED THROMBUS ON THE DEVICE THE CLOSURE DEVICE WAS NOTED TO HAVE AN EXCELLENT SEAL AND ITS POSITION WAS UNCHANGED THE PATIENT HAD FOLLOWED THE MEDICATION REGIME SHE WAS SUPRATHERAPEUTIC THE WEEK BEFORE AND WAS REDOSED THE PATIENT WAS NOTED TO HAVE A LABILE INTERNATIONAL NORMALIZED RATIO INR THROUGHOUT THIS HER INR AT THE TIME OF THE TEE WAS SHE WAS PLACED BACK ON ELIQUIS AND WILL HAVE ANOTHER TEE IN B NO FURTHER PATIENT COMPLICATIONS WERE REPORTED | Thrombus | Patient problem code
6800216 | 82896897 | D | IT WAS REPORTED THROMBUS ON THE DEVICE OCCURRED IN B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS SUCCESSFULLY PERFORMED A MM WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PATIENT WAS DISCHARGED ON WARFARIN AND MG ASA THE PATIENT RETURNED FOR A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE IN B WHICH REVEALED THROMBUS ON THE DEVICE THE CLOSURE DEVICE WAS NOTED TO HAVE AN EXCELLENT SEAL AND ITS POSITION WAS UNCHANGED THE PATIENT HAD FOLLOWED THE MEDICATION REGIME SHE WAS SUPRATHERAPEUTIC THE WEEK BEFORE AND WAS REDOSED THE PATIENT WAS NOTED TO HAVE A LABILE INTERNATIONAL NORMALIZED RATIO INR THROUGHOUT THIS HER INR AT THE TIME OF THE TEE WAS SHE WAS PLACED BACK ON ELIQUIS AND WILL HAVE ANOTHER TEE IN B NO FURTHER PATIENT COMPLICATIONS WERE REPORTED | Insufficient Information | Device problem code
6800326 | 82901763 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SORE AND RED EXTREMELY HARD AND LUMPY AND INFLAMMATORY PROCESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN THE INJECTION TECHNIQUE FOR JUVÉDERM VOLUMA XC WITH REGARD TO THE ANGLE AND ORIENTATION OF THE BEVEL THE DEPTH SUBCUTANEOUS ANDOR SUBMUSCULARSUPRAPERIOSTEAL OF INJECTION AND THE QUANTITY ADMINISTERED MAY VARY DEPENDING ON THE AREA BEING TREATED INJECTION OF JUVÉDERM VOLUMA XC TOO SUPERFICIALLY INTRADERMALLY OR IN LARGE VOLUMES OVER A SMALL AREA MAY RESULT IN VISIBLE AND PERSISTENT LUMPS ANDOR DISCOLORATION JUVÉDERM VOLUMA XC SHOULD BE DISTRIBUTED IN SMALL ALIQUOTS SMALL BOLUSES OF ML TO ML OVER A LARGE AREA TO REDUCE THE RISK OF PERSISTENT LUMPINESS | No Findings Available | Medical Evaluation result code
6797105 | 82902234 | D | PATIENT HOSPITALIZED DUE TO HICKMAN MALFUNCTION ON B NEW HICKMAN CENTRAL CATHETER PLACED IN ON B AND PT DISCHARGED SAME DAY NO OTHER INFORMATION PROVIDED MANUFACTURER OF HICKMAN IS UNKNOWN REPORTED TO B PATIENTCAREGIVER DOSE OR AMOUNT NKM FREQUENCY CONTINUOUS ROUTE IV DATES OF USE B TO CURRENT DIAGNOSIS OR REASON FOR USE PAH | Defective Device | Device problem code
6797105 | 82902234 | D | PATIENT HOSPITALIZED DUE TO HICKMAN MALFUNCTION ON B NEW HICKMAN CENTRAL CATHETER PLACED IN ON B AND PT DISCHARGED SAME DAY NO OTHER INFORMATION PROVIDED MANUFACTURER OF HICKMAN IS UNKNOWN REPORTED TO B PATIENTCAREGIVER DOSE OR AMOUNT NKM FREQUENCY CONTINUOUS ROUTE IV DATES OF USE B TO CURRENT DIAGNOSIS OR REASON FOR USE PAH | No Known Impact Or Consequence To Patient | Patient problem code
6800404 | 82903044 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS WORKING WHEN THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6800404 | 82903044 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS WORKING WHEN THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6800419 | 82903382 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER WAS SHOWERING WHEN THE ADHESIVE PAD FAILED CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT AND A BOLUS OF UNITS WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6800419 | 82903382 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER WAS SHOWERING WHEN THE ADHESIVE PAD FAILED CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT AND A BOLUS OF UNITS WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6794040 | 82903814 | N | FINDINGS PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MINOR SCRATCHED LCD WINDOW B | Component Missing | Device problem code
6794040 | 82903814 | N | FINDINGS PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MINOR SCRATCHED LCD WINDOW B | Scratched Material | Device problem code
6794040 | 82903814 | N | FINDINGS PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MINOR SCRATCHED LCD WINDOW B | Packaging Problem Identified | medical Evaluation result code
6800430 | 82904050 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED A PERMANENT PACEMAKER WAS IMPLANTED THE FOLLOWING DAY NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6800430 | 82904050 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED A PERMANENT PACEMAKER WAS IMPLANTED THE FOLLOWING DAY NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6800382 | 82904318 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800382 | 82904318 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6800047 | 82905799 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6800047 | 82905799 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6800467 | 82909162 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6800467 | 82909162 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800520 | 82909375 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6800520 | 82909375 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800441 | 82909465 | D | IT WAS REPORTED BY THE PATIENT THAT THEY EXPERIENCED AN INFECTION AT INFUSION SITE LEG WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS PRESCRIBED ANTIBIOTIC CIPRO DOSAGE UNKNOWN | Insufficient Information | Device problem code
6800441 | 82909465 | D | IT WAS REPORTED BY THE PATIENT THAT THEY EXPERIENCED AN INFECTION AT INFUSION SITE LEG WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS PRESCRIBED ANTIBIOTIC CIPRO DOSAGE UNKNOWN | Infection | Patient problem code
6800501 | 82909518 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6800501 | 82909518 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | Device Dislodged or Dislocated | Device problem code
6800508 | 82909557 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE LEG | Device Dislodged or Dislocated | Device problem code
6800508 | 82909557 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6800533 | 82909620 | N | EVALUATION SUMMARY POST MARKET VIGILANCE PMV LED AN EVALUATION OF ONE DEVICE THE EVENT REPORT ALLEGES THE PRODUCT WAS USED IN A SURGICAL PROCEDURE THE VISUAL INSPECTION OF THE RETURNED PRODUCT NOTED THE OBTURATOR DISSECTOR BALLOON TROCAR BALLOON LOCK COLLAR AND VALVE SEAL APPEARED INTACT THE INSUFFLATION BULB WAS RECEIVED THE INFLATION SYRINGE WAS NOT RECEIVED PMV PERFORMED FUNCTIONAL TESTING THE DISSECTOR AND TROCAR BALLOONS WERE INFLATED NO LEAKS WERE DETECTED ALL OTHER COMPONENTS FUNCTIONED PROPERLY RECORDS FROM EACH MANUFACTURING LOT ARE THOROUGHLY REVIEWED TO ENSURE THAT PRODUCTS ARE RELEASED MEETING ALL QUALITY RELEASE SPECIFICATIONS AT THE TIME OF MANUFACTURE ANALYSIS CONCLUDED THERE WERE NO ASSEMBLY COMPONENT RELATED FAILURES THEREFORE A DEVICE HISTORY RECORD WILL NOT BE PERFORMED OUR INVESTIGATION WAS UNABLE TO DETERMINE A ROOT CAUSE OR ESTABLISH A RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED INCIDENT SHOULD NEW INFORMATION BECOME AVAILABLE THE FILE WILL BE REOPENED AND THE INVESTIGATION SUMMARY WILL BE AMENDED AS APPROPRIATE | No Findings Available | Medical Evaluation result code
6800530 | 82909640 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6800530 | 82909640 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6800496 | 82910318 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800496 | 82910318 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6800544 | 82910791 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENTS BLOOD GLUCOSE CARBOHYDRATE AND INSULIN HISTORY IS AS FOLLOWS B AT PM THE PATIENT WENT TO THE EMERGENCY ROOM DUE TO LOW BG LEVELS WHERE THE PATIENT WAS TREATED WITH IV FLUIDS | Insufficient Information | Device problem code
6800544 | 82910791 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENTS BLOOD GLUCOSE CARBOHYDRATE AND INSULIN HISTORY IS AS FOLLOWS B AT PM THE PATIENT WENT TO THE EMERGENCY ROOM DUE TO LOW BG LEVELS WHERE THE PATIENT WAS TREATED WITH IV FLUIDS | Hypoglycemia | Patient Problem Code
6800039 | 82910829 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6800039 | 82910829 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6800041 | 82910852 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800041 | 82910852 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6799854 | 82911179 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | Device Dislodged or Dislocated | Device problem code
6799854 | 82911179 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6800607 | 82913866 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE BACK | No Known Impact Or Consequence To Patient | Patient Problem Code
6800607 | 82913866 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE BACK | Device Dislodged or Dislocated | Device problem code
6800483 | 82915573 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800483 | 82915573 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6799850 | 82918861 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6799850 | 82918861 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE LEG | Device Dislodged or Dislocated | Device problem code
6799882 | 82919030 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT DETECTED INSULIN VISUALLY AND COULD SMELL INSULIN ON THE ADHESIVE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6799882 | 82919030 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT DETECTED INSULIN VISUALLY AND COULD SMELL INSULIN ON THE ADHESIVE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6799890 | 82919306 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON ADHESIVE AND COULD SMELL ON FINGERS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN THE POD HAS BEEN DISCARDED THUS UNAVAILABLE FOR RETURN | Positioning Problem | Device problem code
6799890 | 82919306 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON ADHESIVE AND COULD SMELL ON FINGERS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN THE POD HAS BEEN DISCARDED THUS UNAVAILABLE FOR RETURN | No Known Impact Or Consequence To Patient | Patient Problem Code
6796146 | 82919776 | N | INFORMATION RECEIVED FROM THE MANUFACTURERS SERVICE TECHNICIAN ON B STATED THAT THE HOSPITALS BIOMED WAS ABLE TO CLEAR THE MID ERROR MESSAGE AND THE THERMOGARD CONSOLE B WAS PLACED BACK IN SERVICE HOWEVER FURTHER INFORMATION WAS NOT PROVIDED AS THE CUSTOMER RETAINED THE SERVICE RECORD HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINT REPORTED FOR THE THERMOGARD CONSOLE WITH SERIAL NUMBER B | Device Displays Incorrect Message | Device problem code
6796146 | 82919776 | N | INFORMATION RECEIVED FROM THE MANUFACTURERS SERVICE TECHNICIAN ON B STATED THAT THE HOSPITALS BIOMED WAS ABLE TO CLEAR THE MID ERROR MESSAGE AND THE THERMOGARD CONSOLE B WAS PLACED BACK IN SERVICE HOWEVER FURTHER INFORMATION WAS NOT PROVIDED AS THE CUSTOMER RETAINED THE SERVICE RECORD HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINT REPORTED FOR THE THERMOGARD CONSOLE WITH SERIAL NUMBER B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6800788 | 82920651 | N | B INITIAL IMPLANT FOR THIS PATIENT SYSTEM CONSISTS OF RNS LOCATED IN THE RIGHT PARIETAL REGION AND STRIP LEADS PORT CL SN B RIGHT FRONT INTERHEMISPHERIC PORT CL SN B RIGHT PARIETAL SENSOR NOT CONNECTED CL SN B RIGHT FRONTAL MOTOR | Connection Problem | Device problem code
6800788 | 82920651 | N | B INITIAL IMPLANT FOR THIS PATIENT SYSTEM CONSISTS OF RNS LOCATED IN THE RIGHT PARIETAL REGION AND STRIP LEADS PORT CL SN B RIGHT FRONT INTERHEMISPHERIC PORT CL SN B RIGHT PARIETAL SENSOR NOT CONNECTED CL SN B RIGHT FRONTAL MOTOR | Wired Communication Problem | Medical Evaluation result code
6801453 | 82940277 | N | RESULTS OF INVESTIGATION VYAIRE MEDICAL WAS ABLE TO CONFIRM THE CUSTOMERS REPORTED ISSUE EXTERNAL INSPECTION OF THE PIGTAIL REVEALED THE PIGTAIL WAS COMPLETELY BURNT ALL TESTING WAS PERFORMED USING A GOOD KNOWN AC ADAPTER AS WELL AS A GOOD KNOWN TEST PATIENT CIRCUIT VYAIRE MEDICAL INSTALLED A GOOD KNOWN TEST PIGTAIL AND THE VENTILATOR PASSED AN EXTENDED HOUR RUN IN TEST AT THE CUSTOMERS SETTINGS WITHOUT ANY UNUSUAL ALARMS OR CONDITIONS INTERNAL INSPECTION DID NOT REVEAL ANY ABNORMALITIES WITH THE VENTILATOR | No Device Problem Found | Medical Evaluation result code
6801313 | 82942480 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS THE DISTRIBUTOR PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED MOISTURE INT HE BREATHING SYSTEM THE POPOFF AND EXHALATION VALVES WERE CLEANED AND DRIED THE UNIT WAS RETURNED TO SERVICE | Moisture Damage | Device problem code
6801313 | 82942480 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS THE DISTRIBUTOR PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED MOISTURE INT HE BREATHING SYSTEM THE POPOFF AND EXHALATION VALVES WERE CLEANED AND DRIED THE UNIT WAS RETURNED TO SERVICE | Storage Problem Identified | Medical Evaluation result code
6786232 | 82944348 | N | THE SYSTEM WAS USED FOR TREATMENT THIS CASE IS REPORTABLE AS A MDR DUE TO THE REPORTABLE MALFUNCTION PRESSURE DOME MEMBRANE LEAK SINCE THIS REPORTABLE MALFUNCTION IS ONLY ASSOCIATED WITH THE KIT THIS MDR WILL ONLY BE AGAINST THE KIT A BATCH RECORD REVIEW FOR KIT LOT E WAS CONDUCTED THERE WERE NO NONCONFORMANCES RELATED TO THE COMPLAINT THIS LOT MET ALL RELEASE REQUIREMENTS A REVIEW OF E FOR THE REPORTED ISSUE SHOWS NO TRENDS TRENDS WERE REVIEWED FOR COMPLAINT CATEGORIES PRESSURE DOME MEMBRANE LEAK AND ALARM COLLECT LINE AIR DETECTED NO TRENDS WERE DETECTED FOR EACH COMPLAINT CATEGORY THE KIT AND SMARTCARD WERE RETURNED FOR ANALYSIS REVIEW OF THE SMARTCARD DATA CONFIRMED THE OCCURRENCE OF THE COLLECT LINE AIR DETECTED ALARM EXAMINATION OF THE RETURNED KIT FOUND THE COLLECT PRESSURE DOME DIAPHRAGM WAS DISPLACED AND NOT PROPERLY SECURED TO THE PRESSURE DOME HOUSING THE COLLECT PRESSURE DOME AND COMPONENTS WERE PRESSURE TESTED AND NO LEAKS WERE FOUND THE INVESTIGATION DETERMINED THE CAUSE FOR THE COLLECT LINE AIR DETECTED ALARMS WAS THE LEAK AT THE COLLECT PRESSURE DOME DIAPHRAGM NOT BEING FULLY SEATED ONTO THE HOUSING ALL PRESSURE DOMES ARE TESTED FOR LEAKS PRIOR TO PRODUCT RELEASE THEREFORE IT IS UNLIKELY THE NONINTEGRAL SEAL OBSERVED ON THE PRODUCT RETURN WAS A RESULT OF MANUFACTURING THE CAUSE OF THE DISPLACED COLLECT PRESSURE DOME DIAPHRAGM WAS MOST LIKELY IMPROPER INSTALLATION OF THE COLLECT PRESSURE DOME ASSEMBLY BY THE END USER THIS INVESTIGATION IS NOW COMPLETE B | Leak / Splash | Device problem code
6786232 | 82944348 | N | THE SYSTEM WAS USED FOR TREATMENT THIS CASE IS REPORTABLE AS A MDR DUE TO THE REPORTABLE MALFUNCTION PRESSURE DOME MEMBRANE LEAK SINCE THIS REPORTABLE MALFUNCTION IS ONLY ASSOCIATED WITH THE KIT THIS MDR WILL ONLY BE AGAINST THE KIT A BATCH RECORD REVIEW FOR KIT LOT E WAS CONDUCTED THERE WERE NO NONCONFORMANCES RELATED TO THE COMPLAINT THIS LOT MET ALL RELEASE REQUIREMENTS A REVIEW OF E FOR THE REPORTED ISSUE SHOWS NO TRENDS TRENDS WERE REVIEWED FOR COMPLAINT CATEGORIES PRESSURE DOME MEMBRANE LEAK AND ALARM COLLECT LINE AIR DETECTED NO TRENDS WERE DETECTED FOR EACH COMPLAINT CATEGORY THE KIT AND SMARTCARD WERE RETURNED FOR ANALYSIS REVIEW OF THE SMARTCARD DATA CONFIRMED THE OCCURRENCE OF THE COLLECT LINE AIR DETECTED ALARM EXAMINATION OF THE RETURNED KIT FOUND THE COLLECT PRESSURE DOME DIAPHRAGM WAS DISPLACED AND NOT PROPERLY SECURED TO THE PRESSURE DOME HOUSING THE COLLECT PRESSURE DOME AND COMPONENTS WERE PRESSURE TESTED AND NO LEAKS WERE FOUND THE INVESTIGATION DETERMINED THE CAUSE FOR THE COLLECT LINE AIR DETECTED ALARMS WAS THE LEAK AT THE COLLECT PRESSURE DOME DIAPHRAGM NOT BEING FULLY SEATED ONTO THE HOUSING ALL PRESSURE DOMES ARE TESTED FOR LEAKS PRIOR TO PRODUCT RELEASE THEREFORE IT IS UNLIKELY THE NONINTEGRAL SEAL OBSERVED ON THE PRODUCT RETURN WAS A RESULT OF MANUFACTURING THE CAUSE OF THE DISPLACED COLLECT PRESSURE DOME DIAPHRAGM WAS MOST LIKELY IMPROPER INSTALLATION OF THE COLLECT PRESSURE DOME ASSEMBLY BY THE END USER THIS INVESTIGATION IS NOW COMPLETE B | Manufacturing Process Problem Identified | Medical Evaluation result code
6801732 | 82948230 | D | IT WAS REPORTED THAT THE CUSTOMER WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL FOR DIABETIC KETOACIDOSIS DKA THE BLOOD GLUCOSE LEVEL BG WAS MGDL WITH A LARGE LEVEL OF KETONES THE KETONE LEVEL WAS CONSIDERED TO BE DANGEROUS BY A HEALTHCARE PROVIDER THE CAUSE OF THE DKABG WAS NOT KNOWN THE CUSTOMER WAS TREATED WITH INTRAVENOUS INSULIN AND FLUIDS AND WAS DISCHARGED ON THE SAME DAY WITH THE BG AND KETONES RESOLVED | Insufficient Information | Device problem code
6801732 | 82948230 | D | IT WAS REPORTED THAT THE CUSTOMER WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL FOR DIABETIC KETOACIDOSIS DKA THE BLOOD GLUCOSE LEVEL BG WAS MGDL WITH A LARGE LEVEL OF KETONES THE KETONE LEVEL WAS CONSIDERED TO BE DANGEROUS BY A HEALTHCARE PROVIDER THE CAUSE OF THE DKABG WAS NOT KNOWN THE CUSTOMER WAS TREATED WITH INTRAVENOUS INSULIN AND FLUIDS AND WAS DISCHARGED ON THE SAME DAY WITH THE BG AND KETONES RESOLVED | No Known Impact Or Consequence To Patient | Patient problem code
6801764 | 82948538 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6801764 | 82948538 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | Device Dislodged or Dislocated | Device problem code
6797911 | 82948688 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6797911 | 82948688 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6799826 | 82949440 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6799826 | 82949440 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6798327 | 82950224 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE PATIENT WAS UNSURE IF THE CANNULA HAD DEPLOYED INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6798327 | 82950224 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE PATIENT WAS UNSURE IF THE CANNULA HAD DEPLOYED INDICATING A NEEDLE MECHANISM FAILURE | Activation, Positioning or Separation Problem | Device problem code
6798786 | 82950336 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6798786 | 82950336 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6798366 | 82950338 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE HIPBUTTOCKS | Device Dislodged or Dislocated | Device problem code
6798366 | 82950338 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6798854 | 82950979 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE CUSTOMER WAS SITTING WHEN THE ADHESIVE FAILED NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A COUPLE HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6798854 | 82950979 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE CUSTOMER WAS SITTING WHEN THE ADHESIVE FAILED NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A COUPLE HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6798382 | 82951412 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6798382 | 82951412 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6797850 | 82953733 | D | THE CUSTOMER REPORTED THAT SHE DID NOT BELIEVE THE CANNULA WAS PROPERLY INSERTED DUE TO FEELING AND SMELLING INSULIN ON HER SKIN HER BLOOD GLUCOSE REACHED IN THE MIDS MGDL WHILE WEARING WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS | Activation, Positioning or Separation Problem | Device problem code
6797850 | 82953733 | D | THE CUSTOMER REPORTED THAT SHE DID NOT BELIEVE THE CANNULA WAS PROPERLY INSERTED DUE TO FEELING AND SMELLING INSULIN ON HER SKIN HER BLOOD GLUCOSE REACHED IN THE MIDS MGDL WHILE WEARING WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS | No Known Impact Or Consequence To Patient | Patient Problem Code
6802016 | 82955718 | D | THE FOLLOWING WAS PUBLISHED IN THE JOURNAL OF CARDIOLOGY CASES DURING A PULMONARY VEIN ISOLATION PVI PROCEDURE FOR PAROXYSMAL ATRIAL FIBRILLATION A PERICARDIAL EFFUSION OCCURRED EIGHT MONTHS FOLLOWING THE PVI EXERTIONAL DYSPNEA APPEARED TRANSTHORACIC ECHOGRAPHY CONFIRMED A PERICARDIAL EFFUSION FOR WHICH A PERICARDIOCENTESIS WAS PERFORMED TO STABILIZED THE PATIENT THE PATIENT WAS FOLLOWED AND DISCHARGED THE CAUSE OF THE PERICARDIAL EFFUSION REMAINED UNKNOWN | Pericardial Effusion | Patient problem code
6802016 | 82955718 | D | THE FOLLOWING WAS PUBLISHED IN THE JOURNAL OF CARDIOLOGY CASES DURING A PULMONARY VEIN ISOLATION PVI PROCEDURE FOR PAROXYSMAL ATRIAL FIBRILLATION A PERICARDIAL EFFUSION OCCURRED EIGHT MONTHS FOLLOWING THE PVI EXERTIONAL DYSPNEA APPEARED TRANSTHORACIC ECHOGRAPHY CONFIRMED A PERICARDIAL EFFUSION FOR WHICH A PERICARDIOCENTESIS WAS PERFORMED TO STABILIZED THE PATIENT THE PATIENT WAS FOLLOWED AND DISCHARGED THE CAUSE OF THE PERICARDIAL EFFUSION REMAINED UNKNOWN | Insufficient Information | Device problem code
6802016 | 82955718 | D | THE FOLLOWING WAS PUBLISHED IN THE JOURNAL OF CARDIOLOGY CASES DURING A PULMONARY VEIN ISOLATION PVI PROCEDURE FOR PAROXYSMAL ATRIAL FIBRILLATION A PERICARDIAL EFFUSION OCCURRED EIGHT MONTHS FOLLOWING THE PVI EXERTIONAL DYSPNEA APPEARED TRANSTHORACIC ECHOGRAPHY CONFIRMED A PERICARDIAL EFFUSION FOR WHICH A PERICARDIOCENTESIS WAS PERFORMED TO STABILIZED THE PATIENT THE PATIENT WAS FOLLOWED AND DISCHARGED THE CAUSE OF THE PERICARDIAL EFFUSION REMAINED UNKNOWN | Dyspnea | Patient problem code
6797203 | 82959347 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6797203 | 82959347 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6802762 | 82976520 | D | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED ELEVATED BLOOD GLUCOSE LEVELS RANGING FROM MGDL TO ADDRESS THE BG LEVEL THE CUSTOMER CHANGED THE INFUSION SITE ADMINISTERED MANUAL INJECTIONS AND DELIVERED A CORRECTION BOLUS THE CUSTOMER DECLINED TO PERFORM TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT AND DECLINED FURTHER FOLLOWUP TO ENSURE THAT THE HIGH BGS WERE RESOLVED ADDITIONALLY THE CUSTOMER REPORTED RECEIVING A RESUME PUMP ALARM WHICH OCCURRED WHILE THE CUSTOMER WAS ASLEEP AROUND AM TO AM REVIEW OF THE PUMP DATA SHOWED THAT THE CUSTOMER HAD UNLOCKED THE PUMP SCREEN AND SUSPENDED INSULIN DELIVERY TWICE DURING AM THEN AFTER THE RESUME PUMP ALARMS HAD DECLARED THEY HAD BEEN CLEARED BY THE USER AND INSULIN DELIVERY HAD BEEN RESUMED THE CUSTOMER ACKNOWLEDGED THE INFORMATION PROVIDED BY TANDEM TECHNICAL SUPPORT | Hyperglycemia | Patient problem code
6802762 | 82976520 | D | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED ELEVATED BLOOD GLUCOSE LEVELS RANGING FROM MGDL TO ADDRESS THE BG LEVEL THE CUSTOMER CHANGED THE INFUSION SITE ADMINISTERED MANUAL INJECTIONS AND DELIVERED A CORRECTION BOLUS THE CUSTOMER DECLINED TO PERFORM TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT AND DECLINED FURTHER FOLLOWUP TO ENSURE THAT THE HIGH BGS WERE RESOLVED ADDITIONALLY THE CUSTOMER REPORTED RECEIVING A RESUME PUMP ALARM WHICH OCCURRED WHILE THE CUSTOMER WAS ASLEEP AROUND AM TO AM REVIEW OF THE PUMP DATA SHOWED THAT THE CUSTOMER HAD UNLOCKED THE PUMP SCREEN AND SUSPENDED INSULIN DELIVERY TWICE DURING AM THEN AFTER THE RESUME PUMP ALARMS HAD DECLARED THEY HAD BEEN CLEARED BY THE USER AND INSULIN DELIVERY HAD BEEN RESUMED THE CUSTOMER ACKNOWLEDGED THE INFORMATION PROVIDED BY TANDEM TECHNICAL SUPPORT | Insufficient Information | Device problem code
6802662 | 82988460 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE MULTIGAS UNIT WAS GIVING AN INCORRECT O READING NO ERROR MESSAGE WAS GIVEN BUT THE DOCTOR NOTICED THE DISCREPANCY THE UNIT WAS IN USE ON A PATIENT HOWEVER NO PATIENT HARM WAS REPORTED THE MULTIGAS UNIT WAS SENT IN FOR EVALUATION AND REPAIR THE UNIT WAS CLEANED AND EVALUATED AND THE REPORTED PROBLEM OF NOT READING O EXCHANGE CORRECTLY COULD NOT BE DUPLICATED THROUGH TESTING TROUBLESHOOTING AND EXTENDED OPERATION OF THE DEVICE THE UNIT WAS TESTED PER OPERATORSSERVICE MANUAL AND OPERATES TO MANUFACTURERS SPECIFICATIONS THE UNIT WAS SHIPPED BACK TO THE CUSTOMER NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH CFR SECTION IF ADDITIONAL INFORMATION BECOMES AVAILABLE | Device Displays Incorrect Message | Device problem code
6802662 | 82988460 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE MULTIGAS UNIT WAS GIVING AN INCORRECT O READING NO ERROR MESSAGE WAS GIVEN BUT THE DOCTOR NOTICED THE DISCREPANCY THE UNIT WAS IN USE ON A PATIENT HOWEVER NO PATIENT HARM WAS REPORTED THE MULTIGAS UNIT WAS SENT IN FOR EVALUATION AND REPAIR THE UNIT WAS CLEANED AND EVALUATED AND THE REPORTED PROBLEM OF NOT READING O EXCHANGE CORRECTLY COULD NOT BE DUPLICATED THROUGH TESTING TROUBLESHOOTING AND EXTENDED OPERATION OF THE DEVICE THE UNIT WAS TESTED PER OPERATORSSERVICE MANUAL AND OPERATES TO MANUFACTURERS SPECIFICATIONS THE UNIT WAS SHIPPED BACK TO THE CUSTOMER NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH CFR SECTION IF ADDITIONAL INFORMATION BECOMES AVAILABLE | No Findings Available | Medical Evaluation result code
6797431 | 82988836 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Insufficient Information | Device problem code
6792733 | 82993490 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THE TEST WELL IMAGES WERE CONSISTENT WITH THE ERRONEOUS OUTCOME AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND IT TO BE OPERATING AS EXPECTED | Connection Problem | Device problem code
6792733 | 82993490 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THE TEST WELL IMAGES WERE CONSISTENT WITH THE ERRONEOUS OUTCOME AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND IT TO BE OPERATING AS EXPECTED | No Device Problem Found | Medical Evaluation result code
6800059 | 82997821 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6800059 | 82997821 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800050 | 82997937 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6800050 | 82997937 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6799868 | 83007586 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED A RANGE OF TO MGDL THROUGHOUT THE DAY WHILE WEARING THE POD BETWEEN AND HOURS ADHESIVE WAS WET AND PATIENT COULD SMELL INSULIN ON THE SKIN THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM POD HAS BEEN DISCARDED THEREFORE NOT AVAILABLE FOR RETURN | Positioning Problem | Device problem code
6799868 | 83007586 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED A RANGE OF TO MGDL THROUGHOUT THE DAY WHILE WEARING THE POD BETWEEN AND HOURS ADHESIVE WAS WET AND PATIENT COULD SMELL INSULIN ON THE SKIN THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM POD HAS BEEN DISCARDED THEREFORE NOT AVAILABLE FOR RETURN | No Known Impact Or Consequence To Patient | Patient Problem Code
6792611 | 83008458 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE B THE GETINGE STM SERVICE TERRITORY MANAGER PERFORMED THE SOLENOID DRIVER BOARD FIELD SAFETY CORRECTIVE ACTION AS PER SERVICE BULLETIN THE STM ALSO PERFORMED A FULL CALIBRATION AND FUNCTIONAL TESTS IN ACCORDANCE WITH THE SYSTEMS SERVICE MANUAL THE IABP FAILED THE BATTERY RUN TIME TEST WITHIN MINUTES THE STM ADVISED THE CUSTOMER TO REPLACE THE BATTERIES AND ALSO INDICATED THAT THE CUSTOMER WOULD CLEAR THE IABP FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6800266 | 83011622 | N | BACKUP BATTERY IS AN ALTERNATE POWER SOURCE THE VENTILATOR WILL FUNCTION WITH THE PRIMARY AC POWER SOURCE AND CONTINUES TO VENTILATE THE PATIENT IF THE BACKUP BATTERY IS NOT FUNCTIONING HOWEVER IN THE EVENT OF AN AC POWER LOSS THE VENTILATOR WILL BE UNABLE TO SUSTAIN OPERATION IF ATTEMPTED TO POWER ON USING A DISCHARGED BATTERY THIS EVENT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO PATIENT HARM THE PATIENT MUST BE PLACED ON ANOTHER MEANS OF VENTILATOR SUPPORT | Premature Discharge of Battery | Device problem code
6800266 | 83011622 | N | BACKUP BATTERY IS AN ALTERNATE POWER SOURCE THE VENTILATOR WILL FUNCTION WITH THE PRIMARY AC POWER SOURCE AND CONTINUES TO VENTILATE THE PATIENT IF THE BACKUP BATTERY IS NOT FUNCTIONING HOWEVER IN THE EVENT OF AN AC POWER LOSS THE VENTILATOR WILL BE UNABLE TO SUSTAIN OPERATION IF ATTEMPTED TO POWER ON USING A DISCHARGED BATTERY THIS EVENT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO PATIENT HARM THE PATIENT MUST BE PLACED ON ANOTHER MEANS OF VENTILATOR SUPPORT | Loss of Power | Device problem code
6800266 | 83011622 | N | BACKUP BATTERY IS AN ALTERNATE POWER SOURCE THE VENTILATOR WILL FUNCTION WITH THE PRIMARY AC POWER SOURCE AND CONTINUES TO VENTILATE THE PATIENT IF THE BACKUP BATTERY IS NOT FUNCTIONING HOWEVER IN THE EVENT OF AN AC POWER LOSS THE VENTILATOR WILL BE UNABLE TO SUSTAIN OPERATION IF ATTEMPTED TO POWER ON USING A DISCHARGED BATTERY THIS EVENT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO PATIENT HARM THE PATIENT MUST BE PLACED ON ANOTHER MEANS OF VENTILATOR SUPPORT | Power Source Problem Identified | Medical Evaluation result code
6794272 | 83014518 | D | DURING A PATIENT PROCEDURE THE BRIDGE BALLOON WOULD NOT TRACK OVER THE WIRE KEPT GETTING HUNG UP AND WOULDNT PASS THROUGH THERE WAS NO HARM TO THE PATIENT AND THE PATIENT DISCHARGED PER OPERATIVE PLAN | No Consequences Or Impact To Patient | Patient problem code
6794272 | 83014518 | D | DURING A PATIENT PROCEDURE THE BRIDGE BALLOON WOULD NOT TRACK OVER THE WIRE KEPT GETTING HUNG UP AND WOULDNT PASS THROUGH THERE WAS NO HARM TO THE PATIENT AND THE PATIENT DISCHARGED PER OPERATIVE PLAN | Failure to Analyze Signal | Device problem code
6796510 | 83014521 | N | AFTER FURTHER INVESTIGATION DEVELOPMENT PERSONNEL DETERMINED THIS WAS A DEFECT IN THE CODE THAT SKIPS CORRUPTED IMAGES THE CODE WAS CHANGED IN MERGE PACS VERSION THE CONDITION OF SKIPPING CORRUPTED IMAGES WAS REMOVED AND A LOAD ERROR IS DISPLAYED TO A USER WHILE SCROLLING THROUGH A STUDY WITH CORRUPTED IMAGES THE USER MAY BE ABLE TO RESEND THE STUDY IMAGES FROM THE MODALITY WITHOUT ANY CORRUPTION WHICH WOULD THEN ALLOW FOR SCANNING THROUGH IMAGES WITHOUT ISSUE | Defective Device | Device problem code
6796510 | 83014521 | N | AFTER FURTHER INVESTIGATION DEVELOPMENT PERSONNEL DETERMINED THIS WAS A DEFECT IN THE CODE THAT SKIPS CORRUPTED IMAGES THE CODE WAS CHANGED IN MERGE PACS VERSION THE CONDITION OF SKIPPING CORRUPTED IMAGES WAS REMOVED AND A LOAD ERROR IS DISPLAYED TO A USER WHILE SCROLLING THROUGH A STUDY WITH CORRUPTED IMAGES THE USER MAY BE ABLE TO RESEND THE STUDY IMAGES FROM THE MODALITY WITHOUT ANY CORRUPTION WHICH WOULD THEN ALLOW FOR SCANNING THROUGH IMAGES WITHOUT ISSUE | Device Displays Incorrect Message | Device Problem Code
6796510 | 83014521 | N | AFTER FURTHER INVESTIGATION DEVELOPMENT PERSONNEL DETERMINED THIS WAS A DEFECT IN THE CODE THAT SKIPS CORRUPTED IMAGES THE CODE WAS CHANGED IN MERGE PACS VERSION THE CONDITION OF SKIPPING CORRUPTED IMAGES WAS REMOVED AND A LOAD ERROR IS DISPLAYED TO A USER WHILE SCROLLING THROUGH A STUDY WITH CORRUPTED IMAGES THE USER MAY BE ABLE TO RESEND THE STUDY IMAGES FROM THE MODALITY WITHOUT ANY CORRUPTION WHICH WOULD THEN ALLOW FOR SCANNING THROUGH IMAGES WITHOUT ISSUE | No Device Problem Found | Medical Evaluation result code
6801126 | 83014549 | D | PRIOR TO A PATIENT PROCEDURE AND DURING PREPARATION FOR USE THE PHYSICIAN NOTICED A KINK OR CRIMP CLOSE TO THE BALLOON ANOTHER DEVICE WAS OPENED THERE WERE NO REPORTED PATIENT INJURIES AND THE PATIENT WAS DISCHARGED PER OPERATIVE PLAN | No Known Impact Or Consequence To Patient | Patient problem code
6801126 | 83014549 | D | PRIOR TO A PATIENT PROCEDURE AND DURING PREPARATION FOR USE THE PHYSICIAN NOTICED A KINK OR CRIMP CLOSE TO THE BALLOON ANOTHER DEVICE WAS OPENED THERE WERE NO REPORTED PATIENT INJURIES AND THE PATIENT WAS DISCHARGED PER OPERATIVE PLAN | Insufficient Information | Device problem code
6803622 | 83015570 | N | B ON AUGUST THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS REOPERATION FEW DAYS AFTER PRIMARY CEMENTLESS THA DUE TO LEG LENGTH DISCREPANCY THE OPERATED LIMB WAS FOUND TO BE MM LONGER THAN CONTRALATERAL THIS IS NOT A PROBLEM ORIGINATED BY A FAULTY OR MALFUNCTIONING DEVICE BATCH REVIEW PERFORMED ON AUGUST B | Inadequacy of Device Shape and/or Size | Device problem code
6803622 | 83015570 | N | B ON AUGUST THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS REOPERATION FEW DAYS AFTER PRIMARY CEMENTLESS THA DUE TO LEG LENGTH DISCREPANCY THE OPERATED LIMB WAS FOUND TO BE MM LONGER THAN CONTRALATERAL THIS IS NOT A PROBLEM ORIGINATED BY A FAULTY OR MALFUNCTIONING DEVICE BATCH REVIEW PERFORMED ON AUGUST B | Deformation Problem | Medical Evaluation result code
6803736 | 83026062 | N | B THE PATIENT WAS INITIALLY IMPLANTED ON B THIS IS THE PATIENTS FOURTH NEUROSTIMULATOR CL SN B RIGHT POSTERIOR FRONTAL INFERIOR CONNECTED TO PORT CL SN B RIGHT POSTERIOR FRONTAL INFERIOR SUPERIOR CONNECTED TO PORT NOT CONNECTED CL SN B RIGHT LATERAL FRONTAL AND CL SN B RIGHT MEDIAL FRONTAL | Connection Problem | Device problem code
6803736 | 83026062 | N | B THE PATIENT WAS INITIALLY IMPLANTED ON B THIS IS THE PATIENTS FOURTH NEUROSTIMULATOR CL SN B RIGHT POSTERIOR FRONTAL INFERIOR CONNECTED TO PORT CL SN B RIGHT POSTERIOR FRONTAL INFERIOR SUPERIOR CONNECTED TO PORT NOT CONNECTED CL SN B RIGHT LATERAL FRONTAL AND CL SN B RIGHT MEDIAL FRONTAL | No Device Problem Found | Medical Evaluation result code
6804184 | 83030006 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device problem code
6803999 | 83031008 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED SINUS ISSUES THE PAST COUPLE OF DAYS AND ALSO NOTICED WITHIN THE LAST DAY OR SO THAT THEY HAVE BEEN UNABLE TO FULLY VOID COMPLETELY THEY ARE NOT FEELING TOO GOD AND CAN HARDLY GO WHEN VOIDING THE PATIENT WAS TOLD TO INCREASE STIMULATION FOR POSSIBLE SYMPTOM RELIEF IMPLANT IS B NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6803999 | 83031008 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED SINUS ISSUES THE PAST COUPLE OF DAYS AND ALSO NOTICED WITHIN THE LAST DAY OR SO THAT THEY HAVE BEEN UNABLE TO FULLY VOID COMPLETELY THEY ARE NOT FEELING TOO GOD AND CAN HARDLY GO WHEN VOIDING THE PATIENT WAS TOLD TO INCREASE STIMULATION FOR POSSIBLE SYMPTOM RELIEF IMPLANT IS B NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Known Impact Or Consequence To Patient | Patient problem code
6792750 | 83047565 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THE TEST WELL IMAGES WERE CONSISTENT WITH THE ERRONEOUS OUTCOME AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND IT TO BE OPERATING AS EXPECTED | No Device Problem Found | Medical Evaluation result code
6793017 | 83047674 | N | NO BLOOD SAMPLE OR ELECTRONIC CONNECTION WAS AVAILABLE OR POSSIBLE | No Device Problem Found | Medical Evaluation result code
6805017 | 83053185 | N | DEVICE EVALUATION IS NOT NECESSARY BECAUSE THE REPORTED MIGRATION HAS BEEN DETERMINED AS NOT RELATED TO VNS THERAPY | No Device Problem Found | Medical Evaluation result code
6806138 | 83093649 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT HE CHANGED HIS INFUSION SET AND THEIR BLOOD GLUCOSE LEVEL WENT HIGH THE CUSTOMERS BLOOD GLUCOSE LEVEL AT THE TIME OF THE INCIDENT WAS MGDL THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER TREATED THE HIGH BLOOD GLUCOSE WITH THE INSULIN PUMP TROUBLESHOOTING WAS PERFORMED AND INSULIN EXITED WITHOUT INCIDENT THE INSULIN PASSED THE BASIC OCCLUSION TEST WHILE TROUBLESHOOTING THE CUSTOMERS BLOOD GLUCOSE LEVEL WENT DOWN TO MGDL THE CUSTOMER WAS ADVISED TO MONITOR HIS BLOOD GLUCOSE THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6806138 | 83093649 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT HE CHANGED HIS INFUSION SET AND THEIR BLOOD GLUCOSE LEVEL WENT HIGH THE CUSTOMERS BLOOD GLUCOSE LEVEL AT THE TIME OF THE INCIDENT WAS MGDL THE CUSTOMERS CURRENT BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER TREATED THE HIGH BLOOD GLUCOSE WITH THE INSULIN PUMP TROUBLESHOOTING WAS PERFORMED AND INSULIN EXITED WITHOUT INCIDENT THE INSULIN PASSED THE BASIC OCCLUSION TEST WHILE TROUBLESHOOTING THE CUSTOMERS BLOOD GLUCOSE LEVEL WENT DOWN TO MGDL THE CUSTOMER WAS ADVISED TO MONITOR HIS BLOOD GLUCOSE THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6806712 | 83095966 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED A DOWNSTREAM OCCLUSION ALARM WHEN NO OCCLUSION WAS PRESENT THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Known Impact Or Consequence To Patient | Patient problem code
6806712 | 83095966 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED A DOWNSTREAM OCCLUSION ALARM WHEN NO OCCLUSION WAS PRESENT THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | Pumping Problem | Device problem code
6805264 | 83096767 | N | THIS REPORT HAS BEEN IDENTIFIED AS B BRAUN INTERNAL REPORT B THE INSTRUCTIONS FOR USE OF THE PRODUCT HAVE BEEN CHECKED AND THE FOLLOWING SIDE EFFECTS ARE LISTED IN VERY RARE CASES THERE MAY BE A MILD BURNING SENSATION AFTER APPLICATION OF PRONTOSAN BUT THIS USUALLY DISSIPATES AFTER A FEW MINUTES PRONTOSAN CAN CAUSE ALLERGIC REACTIONS SUCH AS ITCHING URTICARIA AND RASHES EXANTHEMA IN RARE CASES LESS THAN OUT OF ANAPHYLACTIC SHOCK HAS BEEN REPORTED THE PRODUCT QUANTITIES IN FOR PRONTOSAN WOUND IRRIGATION SOLUTION WERE OVER MILLION UNITS NO NEGATIVE TREND CAN BE OBSERVED WE CONTINUOUSLY MONITOR PRODUCT INCIDENT REPORTS TO IDENTIFY TRENDS WHICH MIGHT AFFECT THE QUALITY OF OUR PRODUCTS WE WILL MAINTAIN THIS REPORT FOR FUTURE REFERENCE AND CONTINUE TO MONITOR OTHER REPORTS FOR SIMILAR NATURE IF ADDITIONAL PERTINENT INFORMATION BECOMES AVAILABLE A FOLLOW UP REPORT WILL BE SUBMITTED NO SAMPLES WERE SENT FOR ANALYTICAL INVESTIGATION NO BATCH NUMBER IS AVAILABLE WITHOUT THE SAMPLE ANDOR LOT NUMBER FURTHER EVALUATION IS NOT POSSIBLE IF A SAMPLE ANDOR ADDITIONAL PERTINENT INFORMATION BECOMES AVAILABLE A FOLLOW UP REPORT WILL BE SUBMITTED | Results Pending Completion of Investigation | Medical Evaluation result code
6795106 | 83100607 | N | REPORTABLE MALFUNCTION WITH INOMAX DSIR B WAS CREATED AS COMPLAINTB THE SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO OVERDELIVERY A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE FIELD THE ROOT CAUSE FOR THIS REPORT WAS PROPORTIONAL VALVE FAILURE THE DISTRIBUTOR WAS RECOMMENDED TO REPLACE THE PROPORTIONAL VALVE DUE TO THIS FINDING THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM THIS CASE DID NOT RESULT IN AN ADVERSE EVENTSERIOUS ADVERSE EVENT HOWEVER IT IS BEING SUBMITTED TO REGULATORY AUTHORITIES BECAUSE A SIMILAR FAILURE OCCURRED IN THE PAST WHICH RESULTED IN A SERIOUS ADVERSE EVENT MDR | Failure to Deliver | Device problem code
6795106 | 83100607 | N | REPORTABLE MALFUNCTION WITH INOMAX DSIR B WAS CREATED AS COMPLAINTB THE SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO OVERDELIVERY A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE FIELD THE ROOT CAUSE FOR THIS REPORT WAS PROPORTIONAL VALVE FAILURE THE DISTRIBUTOR WAS RECOMMENDED TO REPLACE THE PROPORTIONAL VALVE DUE TO THIS FINDING THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM THIS CASE DID NOT RESULT IN AN ADVERSE EVENTSERIOUS ADVERSE EVENT HOWEVER IT IS BEING SUBMITTED TO REGULATORY AUTHORITIES BECAUSE A SIMILAR FAILURE OCCURRED IN THE PAST WHICH RESULTED IN A SERIOUS ADVERSE EVENT MDR | Operational Problem Identified | Medical Evaluation result code
6801753 | 83103732 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6801753 | 83103732 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6801516 | 83104040 | N | ANALYSIS WAS UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT TEST DUE TO BLANK DISPLAY BLANK DISPLAY DUE TO BOARD CONNECTOR NOT PLUGGED IN PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | No Display / Image | Device Problem Code
6801516 | 83104040 | N | ANALYSIS WAS UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT TEST DUE TO BLANK DISPLAY BLANK DISPLAY DUE TO BOARD CONNECTOR NOT PLUGGED IN PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
6801516 | 83104040 | N | ANALYSIS WAS UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT TEST DUE TO BLANK DISPLAY BLANK DISPLAY DUE TO BOARD CONNECTOR NOT PLUGGED IN PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Connection Problem | Device problem code
6801516 | 83104040 | N | ANALYSIS WAS UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT TEST DUE TO BLANK DISPLAY BLANK DISPLAY DUE TO BOARD CONNECTOR NOT PLUGGED IN PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Non-Functional Defect | Medical Evaluation result code
6795165 | 83108550 | N | THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON JUL WHICH PERFORMED AS EXPECTED IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE TEST WELLS IN QUESTION WHICH WERE VISUALLY POSITIVE | Connection Problem | Device problem code
6795165 | 83108550 | N | THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON JUL WHICH PERFORMED AS EXPECTED IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE TEST WELLS IN QUESTION WHICH WERE VISUALLY POSITIVE | No Device Problem Found | Medical Evaluation result code
6802702 | 83110161 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Incompatible Component/ Accessory | Medical Evaluation result code
6806804 | 83116208 | N | THE ORING WHICH SEALS THE CO CANISTER TO THE BREATHING SYSTEM HAD DUST FROM THE SODASORB CAUSING AN INCOMPLETE SEAL THE INCOMPLETE SEAL CAUSED THE PATIENT TO REBREATH CO B | Environmental Particulates | Device problem code
6806804 | 83116208 | N | THE ORING WHICH SEALS THE CO CANISTER TO THE BREATHING SYSTEM HAD DUST FROM THE SODASORB CAUSING AN INCOMPLETE SEAL THE INCOMPLETE SEAL CAUSED THE PATIENT TO REBREATH CO B | Leakage/Seal | Medical Evaluation result code
6801755 | 83116300 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE B THE GETINGE STM SERVICE TERRITORY MANAGER PERFORMED THE SOLENOID DRIVER BOARD FIELD SAFETY CORRECTIVE ACTION AS PER SERVICE BULLETIN THE STM ALSO PERFORMED A FULL CALIBRATION AND FUNCTIONAL TESTS IN ACCORDANCE WITH THE SYSTEMS SERVICE MANUAL THE IABP FAILED THE BATTERY RUN TIME TEST WITHIN MINUTES THE STM ADVISED THE CUSTOMER TO REPLACE THE BATTERIES AND ALSO INDICATED THAT THE CUSTOMER WOULD CLEAR THE IABP FOR CLINICAL USE | Calibration Problem | Device problem code
6801755 | 83116300 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE B THE GETINGE STM SERVICE TERRITORY MANAGER PERFORMED THE SOLENOID DRIVER BOARD FIELD SAFETY CORRECTIVE ACTION AS PER SERVICE BULLETIN THE STM ALSO PERFORMED A FULL CALIBRATION AND FUNCTIONAL TESTS IN ACCORDANCE WITH THE SYSTEMS SERVICE MANUAL THE IABP FAILED THE BATTERY RUN TIME TEST WITHIN MINUTES THE STM ADVISED THE CUSTOMER TO REPLACE THE BATTERIES AND ALSO INDICATED THAT THE CUSTOMER WOULD CLEAR THE IABP FOR CLINICAL USE | Failure to Calibrate | Device Problem Code
6803115 | 83116700 | N | EVALUATION SUMMARY THE STENT HAD MOVED SLIGHTLY DISTALLY ON THE BALLOON AND WAS NOT POSITIONED BETWEEN THE MARKER BANDS AS PER SPECIFICATIONS DEFORMATION WAS EVIDENT TO A NUMBER OF THE MID STENT SEGMENTS WITH STRUTS RAISED BUNCHED AND OVERLAPPING DEFORMATION WAS EVIDENT TO A NUMBER OF THE MOST PROXIMAL STENT SEGMENTS WITH STRUTS RAISED CRIMP IMPRESSIONS WERE VISIBLE ON THE EXPOSED PROXIMAL AND MID BALLOON SURFACE DEFORMATION WAS EVIDENT TO THE DISTAL TIP B | Material Deformation | Device problem code
6803115 | 83116700 | N | EVALUATION SUMMARY THE STENT HAD MOVED SLIGHTLY DISTALLY ON THE BALLOON AND WAS NOT POSITIONED BETWEEN THE MARKER BANDS AS PER SPECIFICATIONS DEFORMATION WAS EVIDENT TO A NUMBER OF THE MID STENT SEGMENTS WITH STRUTS RAISED BUNCHED AND OVERLAPPING DEFORMATION WAS EVIDENT TO A NUMBER OF THE MOST PROXIMAL STENT SEGMENTS WITH STRUTS RAISED CRIMP IMPRESSIONS WERE VISIBLE ON THE EXPOSED PROXIMAL AND MID BALLOON SURFACE DEFORMATION WAS EVIDENT TO THE DISTAL TIP B | Deformation Problem | Medical Evaluation result code
6798442 | 83125917 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER THE COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE SERVICE TERRITORY MANAGER STM FOUND A HELIUM LEAK AT THE LOWER FITTING OF THE LOW PRESSURE HELIUM REGULATOR THE STM TIGHTENED THE FITTING THE IABP THEN PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS CLEARED FOR CLINICAL USE AND RETURNED TO THE CUSTOMER | No Device Problem Found | Medical Evaluation result code
6803186 | 83132506 | N | THE AFFECTED SAVINA AND ITS LOGBOOK WERE ANALYZED BY THE MANUFACTURER ON THE DAYS BEFORE THE EVENT THE USER LOG SHOWS THAT AN INTENSIVE USE OF THE DEVICE HAS TAKEN PLACE BOTH ON INTERNAL AND EXTERNAL BATTERY WITHOUT ALLOWING A SUFFICIENT CHARGE OF THE DEVICE ON B THE DEVICE HAS BEEN USED ALTERNATELY ON BATTERY AND MAINS POWER SEVERAL TIMES BETWEEN AM AND PM AT PM ACCORDING TO THE LOGBOOK A COMPLETE POWER FAILURE OCCURRED DURING OPERATION PRECEDING THAT THE INTERNAL BATTERY WAS IN USE SINCE PM WHICH WAS ACKNOWLEDGED BY THE USER ACCORDING TO LOG ENTRIES THE VISUAL AND ACOUSTIC ALARMS INTERNAL BATTERY UNDER AND INTERNAL BATTERY UNDER OCCURRED DURING THIS PERIOD A TEST OF THE DEVICE SHOWED NO ABNORMALITIES WITH REGARD TO THE POWER SUPPLY THE POWER FAIL ALARM AND THE OTHER ALARM FUNCTIONS THE BATTERIES WERE FULLY CHARGED IN THE LABORATORY AND TESTED UNDER THE STANDARD FUNCTION OF THE DEVICE THIS RESULTED IN A BATTERY RUN TIME OF MINUTES ON EXTERNAL BATTERY AND MINUTES ON INTERNAL BATTERY IN THE TEST IT WAS ALSO FOUND THAT WHEN THE INTERNAL BATTERY IS FULLY CHARGED THE LAST HIGHPRIORITY INDICATION OF THE DISCHARGED BATTERY COMES SIX MINUTES BEFORE THE DEVICE IS SWITCHED OFF THE RUN TIME OF THE INTERNAL BATTERY IS WITHIN THE SPECIFICATION BUT THE EXTERNAL BATTERY IS OUT OF SPECIFICATION FROM THE ANALYSIS OF THE AVAILABLE INFORMATION WE CONCLUDE THAT THE DEVICE FAILED DUE TO A DEPLETED BATTERY WHICH WAS INADEQUATELY CHARGED DURING INTENSIVE USE WE COULD NOT CONFIRM THE REPORTED MISSING ALARMING ACCORDING TO THE LOGBOOK THE SPECIFIED ALARMS WERE GENERATED BY THE DEVICE BUT DUE TO THE LOW INITIAL CHARGE OF THE INTERNAL BATTERY IT WAS TECHNICALLY IMPOSSIBLE TO ENSURE A USE PERIOD OF MINUTES AFTER THE LAST BATTERY ALARM | Component Missing | Device problem code
6803186 | 83132506 | N | THE AFFECTED SAVINA AND ITS LOGBOOK WERE ANALYZED BY THE MANUFACTURER ON THE DAYS BEFORE THE EVENT THE USER LOG SHOWS THAT AN INTENSIVE USE OF THE DEVICE HAS TAKEN PLACE BOTH ON INTERNAL AND EXTERNAL BATTERY WITHOUT ALLOWING A SUFFICIENT CHARGE OF THE DEVICE ON B THE DEVICE HAS BEEN USED ALTERNATELY ON BATTERY AND MAINS POWER SEVERAL TIMES BETWEEN AM AND PM AT PM ACCORDING TO THE LOGBOOK A COMPLETE POWER FAILURE OCCURRED DURING OPERATION PRECEDING THAT THE INTERNAL BATTERY WAS IN USE SINCE PM WHICH WAS ACKNOWLEDGED BY THE USER ACCORDING TO LOG ENTRIES THE VISUAL AND ACOUSTIC ALARMS INTERNAL BATTERY UNDER AND INTERNAL BATTERY UNDER OCCURRED DURING THIS PERIOD A TEST OF THE DEVICE SHOWED NO ABNORMALITIES WITH REGARD TO THE POWER SUPPLY THE POWER FAIL ALARM AND THE OTHER ALARM FUNCTIONS THE BATTERIES WERE FULLY CHARGED IN THE LABORATORY AND TESTED UNDER THE STANDARD FUNCTION OF THE DEVICE THIS RESULTED IN A BATTERY RUN TIME OF MINUTES ON EXTERNAL BATTERY AND MINUTES ON INTERNAL BATTERY IN THE TEST IT WAS ALSO FOUND THAT WHEN THE INTERNAL BATTERY IS FULLY CHARGED THE LAST HIGHPRIORITY INDICATION OF THE DISCHARGED BATTERY COMES SIX MINUTES BEFORE THE DEVICE IS SWITCHED OFF THE RUN TIME OF THE INTERNAL BATTERY IS WITHIN THE SPECIFICATION BUT THE EXTERNAL BATTERY IS OUT OF SPECIFICATION FROM THE ANALYSIS OF THE AVAILABLE INFORMATION WE CONCLUDE THAT THE DEVICE FAILED DUE TO A DEPLETED BATTERY WHICH WAS INADEQUATELY CHARGED DURING INTENSIVE USE WE COULD NOT CONFIRM THE REPORTED MISSING ALARMING ACCORDING TO THE LOGBOOK THE SPECIFIED ALARMS WERE GENERATED BY THE DEVICE BUT DUE TO THE LOW INITIAL CHARGE OF THE INTERNAL BATTERY IT WAS TECHNICALLY IMPOSSIBLE TO ENSURE A USE PERIOD OF MINUTES AFTER THE LAST BATTERY ALARM | Premature Discharge of Battery | Device problem code
6803186 | 83132506 | N | THE AFFECTED SAVINA AND ITS LOGBOOK WERE ANALYZED BY THE MANUFACTURER ON THE DAYS BEFORE THE EVENT THE USER LOG SHOWS THAT AN INTENSIVE USE OF THE DEVICE HAS TAKEN PLACE BOTH ON INTERNAL AND EXTERNAL BATTERY WITHOUT ALLOWING A SUFFICIENT CHARGE OF THE DEVICE ON B THE DEVICE HAS BEEN USED ALTERNATELY ON BATTERY AND MAINS POWER SEVERAL TIMES BETWEEN AM AND PM AT PM ACCORDING TO THE LOGBOOK A COMPLETE POWER FAILURE OCCURRED DURING OPERATION PRECEDING THAT THE INTERNAL BATTERY WAS IN USE SINCE PM WHICH WAS ACKNOWLEDGED BY THE USER ACCORDING TO LOG ENTRIES THE VISUAL AND ACOUSTIC ALARMS INTERNAL BATTERY UNDER AND INTERNAL BATTERY UNDER OCCURRED DURING THIS PERIOD A TEST OF THE DEVICE SHOWED NO ABNORMALITIES WITH REGARD TO THE POWER SUPPLY THE POWER FAIL ALARM AND THE OTHER ALARM FUNCTIONS THE BATTERIES WERE FULLY CHARGED IN THE LABORATORY AND TESTED UNDER THE STANDARD FUNCTION OF THE DEVICE THIS RESULTED IN A BATTERY RUN TIME OF MINUTES ON EXTERNAL BATTERY AND MINUTES ON INTERNAL BATTERY IN THE TEST IT WAS ALSO FOUND THAT WHEN THE INTERNAL BATTERY IS FULLY CHARGED THE LAST HIGHPRIORITY INDICATION OF THE DISCHARGED BATTERY COMES SIX MINUTES BEFORE THE DEVICE IS SWITCHED OFF THE RUN TIME OF THE INTERNAL BATTERY IS WITHIN THE SPECIFICATION BUT THE EXTERNAL BATTERY IS OUT OF SPECIFICATION FROM THE ANALYSIS OF THE AVAILABLE INFORMATION WE CONCLUDE THAT THE DEVICE FAILED DUE TO A DEPLETED BATTERY WHICH WAS INADEQUATELY CHARGED DURING INTENSIVE USE WE COULD NOT CONFIRM THE REPORTED MISSING ALARMING ACCORDING TO THE LOGBOOK THE SPECIFIED ALARMS WERE GENERATED BY THE DEVICE BUT DUE TO THE LOW INITIAL CHARGE OF THE INTERNAL BATTERY IT WAS TECHNICALLY IMPOSSIBLE TO ENSURE A USE PERIOD OF MINUTES AFTER THE LAST BATTERY ALARM | Electrical Power Problem | Device Problem Code
6803186 | 83132506 | N | THE AFFECTED SAVINA AND ITS LOGBOOK WERE ANALYZED BY THE MANUFACTURER ON THE DAYS BEFORE THE EVENT THE USER LOG SHOWS THAT AN INTENSIVE USE OF THE DEVICE HAS TAKEN PLACE BOTH ON INTERNAL AND EXTERNAL BATTERY WITHOUT ALLOWING A SUFFICIENT CHARGE OF THE DEVICE ON B THE DEVICE HAS BEEN USED ALTERNATELY ON BATTERY AND MAINS POWER SEVERAL TIMES BETWEEN AM AND PM AT PM ACCORDING TO THE LOGBOOK A COMPLETE POWER FAILURE OCCURRED DURING OPERATION PRECEDING THAT THE INTERNAL BATTERY WAS IN USE SINCE PM WHICH WAS ACKNOWLEDGED BY THE USER ACCORDING TO LOG ENTRIES THE VISUAL AND ACOUSTIC ALARMS INTERNAL BATTERY UNDER AND INTERNAL BATTERY UNDER OCCURRED DURING THIS PERIOD A TEST OF THE DEVICE SHOWED NO ABNORMALITIES WITH REGARD TO THE POWER SUPPLY THE POWER FAIL ALARM AND THE OTHER ALARM FUNCTIONS THE BATTERIES WERE FULLY CHARGED IN THE LABORATORY AND TESTED UNDER THE STANDARD FUNCTION OF THE DEVICE THIS RESULTED IN A BATTERY RUN TIME OF MINUTES ON EXTERNAL BATTERY AND MINUTES ON INTERNAL BATTERY IN THE TEST IT WAS ALSO FOUND THAT WHEN THE INTERNAL BATTERY IS FULLY CHARGED THE LAST HIGHPRIORITY INDICATION OF THE DISCHARGED BATTERY COMES SIX MINUTES BEFORE THE DEVICE IS SWITCHED OFF THE RUN TIME OF THE INTERNAL BATTERY IS WITHIN THE SPECIFICATION BUT THE EXTERNAL BATTERY IS OUT OF SPECIFICATION FROM THE ANALYSIS OF THE AVAILABLE INFORMATION WE CONCLUDE THAT THE DEVICE FAILED DUE TO A DEPLETED BATTERY WHICH WAS INADEQUATELY CHARGED DURING INTENSIVE USE WE COULD NOT CONFIRM THE REPORTED MISSING ALARMING ACCORDING TO THE LOGBOOK THE SPECIFIED ALARMS WERE GENERATED BY THE DEVICE BUT DUE TO THE LOW INITIAL CHARGE OF THE INTERNAL BATTERY IT WAS TECHNICALLY IMPOSSIBLE TO ENSURE A USE PERIOD OF MINUTES AFTER THE LAST BATTERY ALARM | Battery Problem | Device Problem Code
6803186 | 83132506 | N | THE AFFECTED SAVINA AND ITS LOGBOOK WERE ANALYZED BY THE MANUFACTURER ON THE DAYS BEFORE THE EVENT THE USER LOG SHOWS THAT AN INTENSIVE USE OF THE DEVICE HAS TAKEN PLACE BOTH ON INTERNAL AND EXTERNAL BATTERY WITHOUT ALLOWING A SUFFICIENT CHARGE OF THE DEVICE ON B THE DEVICE HAS BEEN USED ALTERNATELY ON BATTERY AND MAINS POWER SEVERAL TIMES BETWEEN AM AND PM AT PM ACCORDING TO THE LOGBOOK A COMPLETE POWER FAILURE OCCURRED DURING OPERATION PRECEDING THAT THE INTERNAL BATTERY WAS IN USE SINCE PM WHICH WAS ACKNOWLEDGED BY THE USER ACCORDING TO LOG ENTRIES THE VISUAL AND ACOUSTIC ALARMS INTERNAL BATTERY UNDER AND INTERNAL BATTERY UNDER OCCURRED DURING THIS PERIOD A TEST OF THE DEVICE SHOWED NO ABNORMALITIES WITH REGARD TO THE POWER SUPPLY THE POWER FAIL ALARM AND THE OTHER ALARM FUNCTIONS THE BATTERIES WERE FULLY CHARGED IN THE LABORATORY AND TESTED UNDER THE STANDARD FUNCTION OF THE DEVICE THIS RESULTED IN A BATTERY RUN TIME OF MINUTES ON EXTERNAL BATTERY AND MINUTES ON INTERNAL BATTERY IN THE TEST IT WAS ALSO FOUND THAT WHEN THE INTERNAL BATTERY IS FULLY CHARGED THE LAST HIGHPRIORITY INDICATION OF THE DISCHARGED BATTERY COMES SIX MINUTES BEFORE THE DEVICE IS SWITCHED OFF THE RUN TIME OF THE INTERNAL BATTERY IS WITHIN THE SPECIFICATION BUT THE EXTERNAL BATTERY IS OUT OF SPECIFICATION FROM THE ANALYSIS OF THE AVAILABLE INFORMATION WE CONCLUDE THAT THE DEVICE FAILED DUE TO A DEPLETED BATTERY WHICH WAS INADEQUATELY CHARGED DURING INTENSIVE USE WE COULD NOT CONFIRM THE REPORTED MISSING ALARMING ACCORDING TO THE LOGBOOK THE SPECIFIED ALARMS WERE GENERATED BY THE DEVICE BUT DUE TO THE LOW INITIAL CHARGE OF THE INTERNAL BATTERY IT WAS TECHNICALLY IMPOSSIBLE TO ENSURE A USE PERIOD OF MINUTES AFTER THE LAST BATTERY ALARM | Power Source Problem Identified | Medical Evaluation result code
6798926 | 83133203 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | Connection Problem | Device problem code
6798926 | 83133203 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Device Problem Found | Medical Evaluation result code
6807020 | 83133819 | D | THE CUSTOMER REPORTED THAT HIS BLOOD GLUCOSE READ HIGH MGDL WHILE WEARING THE POD ON HIS HIPBUTTOCKS FOR LESS THAN AN HOUR HE STATED THAT THE CANNULA APPEARED TO HAVE A SLIGHT BEND TO IT WHEN THE POD WAS REMOVED THE ADHESIVE WAS NOT SECURE TREATMENT INCLUDED PUTTING ON A NEW POD AND DOING A CORRECTION | Hyperglycemia | Patient Problem Code
6807020 | 83133819 | D | THE CUSTOMER REPORTED THAT HIS BLOOD GLUCOSE READ HIGH MGDL WHILE WEARING THE POD ON HIS HIPBUTTOCKS FOR LESS THAN AN HOUR HE STATED THAT THE CANNULA APPEARED TO HAVE A SLIGHT BEND TO IT WHEN THE POD WAS REMOVED THE ADHESIVE WAS NOT SECURE TREATMENT INCLUDED PUTTING ON A NEW POD AND DOING A CORRECTION | Material Twisted / Bent | Device problem code
6785170 | 83138491 | N | RELIABILITY ANALYSIS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | No Findings Available | Medical Evaluation result code
6794113 | 83140620 | N | THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON JUL WHICH PERFORMED AS EXPECTED IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELLS IN QUESTION ON JUL WHICH APPEARED AS VISUALLY POSITIVE | Connection Problem | Device problem code
6794113 | 83140620 | N | THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON JUL WHICH PERFORMED AS EXPECTED IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELLS IN QUESTION ON JUL WHICH APPEARED AS VISUALLY POSITIVE | No Device Problem Found | Medical Evaluation result code
6807458 | 83143358 | N | B DATE SENT BATCH UNK ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DO THE SURGEONS BELIEVE THE ADVERSE EVENT WAS RELATED TO THE ECHELON STAPLER WHAT COLOR CATEGORY BLACK OR GREEN WERE THE ADVERSE EVENT OF BLEEDING EVENT DESCRIPTION CONTINUED THE REMAINING PATIENTS REQUIRED NO ADDITIONAL INTERVENTION THREE PATIENTS REQUIRED READMISSION FOR DEHYDRATION WHICH RESOLVED IN DAYS OF REHYDRATION WITH INTRAVENOUS FLUIDS ONE PATIENT PRESENTED WITH NEARCOMPLETE GASTRIC OBSTRUCTION SHORTLY FOLLOWING OPERATION UPPER GI SERIES FLUOROSCOPY DEMONSTRATED A KINK AT THE UPPER ASPECT OF THE SLEEVE THE PATIENT WENT THROUGH UNEVENTFUL LAPAROSCOPIC CONVERSION OPERATION TO ROUXENY GASTRIC BYPASS ON THE TH WEEK POST LSG ONE PATIENT WAS ADMITTED FOR WOUND INFECTION THAT REQUIRED INCISION AND DRAINAGE AT THE SITE WHERE THE GASTRIC SPECIMEN WAS EXTRACTED DURING LSG | Obstruction of Flow | Device problem code
6807458 | 83143358 | N | B DATE SENT BATCH UNK ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DO THE SURGEONS BELIEVE THE ADVERSE EVENT WAS RELATED TO THE ECHELON STAPLER WHAT COLOR CATEGORY BLACK OR GREEN WERE THE ADVERSE EVENT OF BLEEDING EVENT DESCRIPTION CONTINUED THE REMAINING PATIENTS REQUIRED NO ADDITIONAL INTERVENTION THREE PATIENTS REQUIRED READMISSION FOR DEHYDRATION WHICH RESOLVED IN DAYS OF REHYDRATION WITH INTRAVENOUS FLUIDS ONE PATIENT PRESENTED WITH NEARCOMPLETE GASTRIC OBSTRUCTION SHORTLY FOLLOWING OPERATION UPPER GI SERIES FLUOROSCOPY DEMONSTRATED A KINK AT THE UPPER ASPECT OF THE SLEEVE THE PATIENT WENT THROUGH UNEVENTFUL LAPAROSCOPIC CONVERSION OPERATION TO ROUXENY GASTRIC BYPASS ON THE TH WEEK POST LSG ONE PATIENT WAS ADMITTED FOR WOUND INFECTION THAT REQUIRED INCISION AND DRAINAGE AT THE SITE WHERE THE GASTRIC SPECIMEN WAS EXTRACTED DURING LSG | Operational Problem Identified | Medical Evaluation result code
6807458 | 83143358 | N | B DATE SENT BATCH UNK ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DO THE SURGEONS BELIEVE THE ADVERSE EVENT WAS RELATED TO THE ECHELON STAPLER WHAT COLOR CATEGORY BLACK OR GREEN WERE THE ADVERSE EVENT OF BLEEDING EVENT DESCRIPTION CONTINUED THE REMAINING PATIENTS REQUIRED NO ADDITIONAL INTERVENTION THREE PATIENTS REQUIRED READMISSION FOR DEHYDRATION WHICH RESOLVED IN DAYS OF REHYDRATION WITH INTRAVENOUS FLUIDS ONE PATIENT PRESENTED WITH NEARCOMPLETE GASTRIC OBSTRUCTION SHORTLY FOLLOWING OPERATION UPPER GI SERIES FLUOROSCOPY DEMONSTRATED A KINK AT THE UPPER ASPECT OF THE SLEEVE THE PATIENT WENT THROUGH UNEVENTFUL LAPAROSCOPIC CONVERSION OPERATION TO ROUXENY GASTRIC BYPASS ON THE TH WEEK POST LSG ONE PATIENT WAS ADMITTED FOR WOUND INFECTION THAT REQUIRED INCISION AND DRAINAGE AT THE SITE WHERE THE GASTRIC SPECIMEN WAS EXTRACTED DURING LSG | Leak / Splash | Device problem code
6807458 | 83143358 | N | B DATE SENT BATCH UNK ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DO THE SURGEONS BELIEVE THE ADVERSE EVENT WAS RELATED TO THE ECHELON STAPLER WHAT COLOR CATEGORY BLACK OR GREEN WERE THE ADVERSE EVENT OF BLEEDING EVENT DESCRIPTION CONTINUED THE REMAINING PATIENTS REQUIRED NO ADDITIONAL INTERVENTION THREE PATIENTS REQUIRED READMISSION FOR DEHYDRATION WHICH RESOLVED IN DAYS OF REHYDRATION WITH INTRAVENOUS FLUIDS ONE PATIENT PRESENTED WITH NEARCOMPLETE GASTRIC OBSTRUCTION SHORTLY FOLLOWING OPERATION UPPER GI SERIES FLUOROSCOPY DEMONSTRATED A KINK AT THE UPPER ASPECT OF THE SLEEVE THE PATIENT WENT THROUGH UNEVENTFUL LAPAROSCOPIC CONVERSION OPERATION TO ROUXENY GASTRIC BYPASS ON THE TH WEEK POST LSG ONE PATIENT WAS ADMITTED FOR WOUND INFECTION THAT REQUIRED INCISION AND DRAINAGE AT THE SITE WHERE THE GASTRIC SPECIMEN WAS EXTRACTED DURING LSG | Leakage/Seal | Medical Evaluation result code
6807751 | 83151019 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6807763 | 83151023 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Material Discolored | Device problem code
6807763 | 83151023 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807772 | 83151274 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807753 | 83151844 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807756 | 83151882 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807764 | 83152367 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807832 | 83153199 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6807613 | 83159756 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED INSULIN ON SKIN AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6807613 | 83159756 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED INSULIN ON SKIN AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6808284 | 83166197 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ONE HOUR AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6808284 | 83166197 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ONE HOUR AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6798815 | 83167073 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6798815 | 83167073 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6808401 | 83175236 | D | IT WAS REPORTED THAT A PATIENT IS EXPERIENCING OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT IS REPORTED TO HAVE COMPLETE HIP PERFORMANCE ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Insufficient Information | Device problem code
6808401 | 83175236 | D | IT WAS REPORTED THAT A PATIENT IS EXPERIENCING OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT IS REPORTED TO HAVE COMPLETE HIP PERFORMANCE ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Osteolysis | Patient problem code
6808401 | 83175236 | D | IT WAS REPORTED THAT A PATIENT IS EXPERIENCING OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT IS REPORTED TO HAVE COMPLETE HIP PERFORMANCE ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Calcification | Patient problem code
6808466 | 83194310 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THEY WERE IN PAIN BUT DIDNT SAY FROM WHAT PATIENT ALSO SAID THEIR HEALTHCARE PROVIDERS HCP OFFICE SAID TO CHANGE THEIR BANDAGES BUT NOW THE HCP SAID NOT TO AND THEY ARENT SURE WHAT TO DO PATIENT WAS ADVISED TO CONNECT WITH THEIR HCP THE NEXT DAY PATIENT SAID THEY WERENT ABLE TO VOID IN THE MIDDLE OF THE NIGHT THEY USUALLY HAVE TO INCREASE STIMULATION IN THE MORNING AND HAD QUESTIONS ABOUT STIMULATION PATIENT WAS CONCERNED ON B PATIENT SAID THE RIGHT SIDE OF THEIR STOMACH WAS HURTING REALLY BAD SO THEY WERE ADVISED TO DECREASE STIMULATION FOUR DAYS LATER PATIENT HAS PAIN IS STICK TO THEIR STOMACH AND DIDNT EAT FOR DAYS THIS TOOK PLACE ON B NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Known Impact Or Consequence To Patient | Patient problem code
6808466 | 83194310 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THEY WERE IN PAIN BUT DIDNT SAY FROM WHAT PATIENT ALSO SAID THEIR HEALTHCARE PROVIDERS HCP OFFICE SAID TO CHANGE THEIR BANDAGES BUT NOW THE HCP SAID NOT TO AND THEY ARENT SURE WHAT TO DO PATIENT WAS ADVISED TO CONNECT WITH THEIR HCP THE NEXT DAY PATIENT SAID THEY WERENT ABLE TO VOID IN THE MIDDLE OF THE NIGHT THEY USUALLY HAVE TO INCREASE STIMULATION IN THE MORNING AND HAD QUESTIONS ABOUT STIMULATION PATIENT WAS CONCERNED ON B PATIENT SAID THE RIGHT SIDE OF THEIR STOMACH WAS HURTING REALLY BAD SO THEY WERE ADVISED TO DECREASE STIMULATION FOUR DAYS LATER PATIENT HAS PAIN IS STICK TO THEIR STOMACH AND DIDNT EAT FOR DAYS THIS TOOK PLACE ON B NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6808466 | 83194310 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THEY WERE IN PAIN BUT DIDNT SAY FROM WHAT PATIENT ALSO SAID THEIR HEALTHCARE PROVIDERS HCP OFFICE SAID TO CHANGE THEIR BANDAGES BUT NOW THE HCP SAID NOT TO AND THEY ARENT SURE WHAT TO DO PATIENT WAS ADVISED TO CONNECT WITH THEIR HCP THE NEXT DAY PATIENT SAID THEY WERENT ABLE TO VOID IN THE MIDDLE OF THE NIGHT THEY USUALLY HAVE TO INCREASE STIMULATION IN THE MORNING AND HAD QUESTIONS ABOUT STIMULATION PATIENT WAS CONCERNED ON B PATIENT SAID THE RIGHT SIDE OF THEIR STOMACH WAS HURTING REALLY BAD SO THEY WERE ADVISED TO DECREASE STIMULATION FOUR DAYS LATER PATIENT HAS PAIN IS STICK TO THEIR STOMACH AND DIDNT EAT FOR DAYS THIS TOOK PLACE ON B NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6808623 | 83194659 | D | REPORT OF VARILIFTLX REVISION SURGERY TO REPOSITION IMPLANT DUE TO ONGOING BACK PAIN AFTER PATIENT FALL | Fall | Patient problem code
6808623 | 83194659 | D | REPORT OF VARILIFTLX REVISION SURGERY TO REPOSITION IMPLANT DUE TO ONGOING BACK PAIN AFTER PATIENT FALL | Insufficient Information | Device problem code
6808623 | 83194659 | D | REPORT OF VARILIFTLX REVISION SURGERY TO REPOSITION IMPLANT DUE TO ONGOING BACK PAIN AFTER PATIENT FALL | Pain | Patient problem code
6809439 | 83195151 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HIS ONETOUCH VERIOIQ READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION AS THE PATIENT WAS UNAVAILABLE WHEN FOLLOWUP WAS ATTEMPTED THE PATIENT REPORTED THAT THE DEVICE ISSUE BEGAN ON B AND REPORTED OBTAINING RESULTS OF AND MGDL ON THE SUBJECT METER WITHIN MINUTES OF EACH OTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE IN RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES HIS DIABETES WITH ORAL MEDICATION DIET AND EXERCISE AND DENIED MAKING ANY CHANGES TO HIS USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT AT AN UNSPECIFIED TIME AFTER THE ALLEGED ISSUE OCCURRED AT PM TO PM HE DEVELOPED A SYMPTOM OF SHAKY HANDS HOWEVER DENIED RECEIVING ANY TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE PATIENT HAD FOLLOWED THE CORRECT TESTING PROCEDURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY WHEN A CONTROL SOLUTION TEST WAS PERFORMED THE RESULTS WERE WITHIN RANGE PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
6809439 | 83195151 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HIS ONETOUCH VERIOIQ READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION AS THE PATIENT WAS UNAVAILABLE WHEN FOLLOWUP WAS ATTEMPTED THE PATIENT REPORTED THAT THE DEVICE ISSUE BEGAN ON B AND REPORTED OBTAINING RESULTS OF AND MGDL ON THE SUBJECT METER WITHIN MINUTES OF EACH OTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE IN RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES HIS DIABETES WITH ORAL MEDICATION DIET AND EXERCISE AND DENIED MAKING ANY CHANGES TO HIS USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT AT AN UNSPECIFIED TIME AFTER THE ALLEGED ISSUE OCCURRED AT PM TO PM HE DEVELOPED A SYMPTOM OF SHAKY HANDS HOWEVER DENIED RECEIVING ANY TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE PATIENT HAD FOLLOWED THE CORRECT TESTING PROCEDURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY WHEN A CONTROL SOLUTION TEST WAS PERFORMED THE RESULTS WERE WITHIN RANGE PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Insufficient Information | Device problem code
6809439 | 83195151 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HIS ONETOUCH VERIOIQ READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION AS THE PATIENT WAS UNAVAILABLE WHEN FOLLOWUP WAS ATTEMPTED THE PATIENT REPORTED THAT THE DEVICE ISSUE BEGAN ON B AND REPORTED OBTAINING RESULTS OF AND MGDL ON THE SUBJECT METER WITHIN MINUTES OF EACH OTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE IN RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES HIS DIABETES WITH ORAL MEDICATION DIET AND EXERCISE AND DENIED MAKING ANY CHANGES TO HIS USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT AT AN UNSPECIFIED TIME AFTER THE ALLEGED ISSUE OCCURRED AT PM TO PM HE DEVELOPED A SYMPTOM OF SHAKY HANDS HOWEVER DENIED RECEIVING ANY TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE PATIENT HAD FOLLOWED THE CORRECT TESTING PROCEDURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY WHEN A CONTROL SOLUTION TEST WAS PERFORMED THE RESULTS WERE WITHIN RANGE PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | No Code Available | Patient problem code
6809439 | 83195151 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HIS ONETOUCH VERIOIQ READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION AS THE PATIENT WAS UNAVAILABLE WHEN FOLLOWUP WAS ATTEMPTED THE PATIENT REPORTED THAT THE DEVICE ISSUE BEGAN ON B AND REPORTED OBTAINING RESULTS OF AND MGDL ON THE SUBJECT METER WITHIN MINUTES OF EACH OTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE IN RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES HIS DIABETES WITH ORAL MEDICATION DIET AND EXERCISE AND DENIED MAKING ANY CHANGES TO HIS USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT AT AN UNSPECIFIED TIME AFTER THE ALLEGED ISSUE OCCURRED AT PM TO PM HE DEVELOPED A SYMPTOM OF SHAKY HANDS HOWEVER DENIED RECEIVING ANY TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE PATIENT HAD FOLLOWED THE CORRECT TESTING PROCEDURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY WHEN A CONTROL SOLUTION TEST WAS PERFORMED THE RESULTS WERE WITHIN RANGE PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Shaking/Tremors | Patient problem code
6808827 | 83197226 | D | THE CUSTOMER REPORTED THAT DURING A PROCEDURE THE PATIENT FELT A BURNING SENSATION FROM THE DEVICE THE PATIENT WAS NOTED TO HAVE AN INJURY | Insufficient Information | Device problem code
6808827 | 83197226 | D | THE CUSTOMER REPORTED THAT DURING A PROCEDURE THE PATIENT FELT A BURNING SENSATION FROM THE DEVICE THE PATIENT WAS NOTED TO HAVE AN INJURY | Injury | Patient problem code
6808827 | 83197226 | D | THE CUSTOMER REPORTED THAT DURING A PROCEDURE THE PATIENT FELT A BURNING SENSATION FROM THE DEVICE THE PATIENT WAS NOTED TO HAVE AN INJURY | Burning Sensation | Patient problem code
6809095 | 83198244 | N | FINAL ANALYSIS FOUND BURN MARKS ON THE MAIN CIRCUIT BOARD OF THE AC POWER ADAPTER WHICH RESULTED IN A POWER UP ANOMALY | Power Problem | Device problem code
6809095 | 83198244 | N | FINAL ANALYSIS FOUND BURN MARKS ON THE MAIN CIRCUIT BOARD OF THE AC POWER ADAPTER WHICH RESULTED IN A POWER UP ANOMALY | Power Source Problem Identified | medical Evaluation result code
6796523 | 83206916 | N | DURING TROUBLESHOOTING EFFORTS WITH THE CUSTOMER AND MERGE TECHNICAL SUPPORT THE PATIENT DATA MODULE PDM WAS TESTED WITH THE TESTPLUG AND THE PDM WAS ABLE TO CAPTURE PRESSURES ON ALL FOUR PORTS THE CUSTOMER REPORTED THAT THEY WOULD REPLACE THEIR CABLES MERGE HEALTHCARE IS CONTINUING TO FOLLOW UP WITH THE CUSTOMER FOR ADDITIONAL PATIENT INFORMATION AND RESOLUTION OF EVENT IF FURTHER INFORMATION IS RECEIVED FROM THE CUSTOMER A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDING TO HEMOB MERGE HEMO USER MANUAL P CHANNEL THROUGH ALLOWS THE USER TO SELECT A LOCATION LABEL FOR RECORDING INVASIVE PRESSURES THIS INCLUDES SETTING UP AND SELECTING PRESSURE LABELS TO ALLOW FOR COMMON RECORDINGS TO BE CAPTURED IF A TRANSDUCER OR CABLE IS NOT CONNECTED THE AFFECTED CHANNELS WILL BE GRAYED OUT P PROBLEM INVASIVE PRESSURE READINGS AND WAVEFORMS ARE NOT DISPLAYED ON THE HEMO MONITOR SOLUTION POSSIBLE SOLUTIONS INCLUDE PROPER CONNECTION OF THE TRANSDUCER CABLE INTO THE PDM FOR THE NECESSARY PRESSURE CHANNEL THE TRANSDUCER CABLE CONNECTION INTO THE CUSTOMERSUPPLIED TRANSDUCER AND THE INTEGRITY OF THE AIRFREE LINE TO ENABLE A MEASUREMENT FROM THE MOST DISTAL CATHETER LOCATION TO THE HEMO SCREEN ENSURE THE CORRECT PRESSURE CHANNEL IS SELECTED FOR DISPLAY AND ON THE CORRECT SCALE ADDITIONALLY FLUID SUCH AS SALINE OR CONTRAST AT A CONNECTION POINT MAY COMPROMISE A SECURE FIT OR PROVIDE ERRONEOUS VALUES FOR DISPLAY CONSIDER REPLACING TRANSDUCER CABLE FROM PDM CUSTOMERSUPPLIED TRANSDUCER KIT OR CHECKING PLACEMENT OF STOPCOCKS TO OPEN OR CLOSED AS APPROPRIATE | Connection Problem | Device problem code
6796523 | 83206916 | N | DURING TROUBLESHOOTING EFFORTS WITH THE CUSTOMER AND MERGE TECHNICAL SUPPORT THE PATIENT DATA MODULE PDM WAS TESTED WITH THE TESTPLUG AND THE PDM WAS ABLE TO CAPTURE PRESSURES ON ALL FOUR PORTS THE CUSTOMER REPORTED THAT THEY WOULD REPLACE THEIR CABLES MERGE HEALTHCARE IS CONTINUING TO FOLLOW UP WITH THE CUSTOMER FOR ADDITIONAL PATIENT INFORMATION AND RESOLUTION OF EVENT IF FURTHER INFORMATION IS RECEIVED FROM THE CUSTOMER A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDING TO HEMOB MERGE HEMO USER MANUAL P CHANNEL THROUGH ALLOWS THE USER TO SELECT A LOCATION LABEL FOR RECORDING INVASIVE PRESSURES THIS INCLUDES SETTING UP AND SELECTING PRESSURE LABELS TO ALLOW FOR COMMON RECORDINGS TO BE CAPTURED IF A TRANSDUCER OR CABLE IS NOT CONNECTED THE AFFECTED CHANNELS WILL BE GRAYED OUT P PROBLEM INVASIVE PRESSURE READINGS AND WAVEFORMS ARE NOT DISPLAYED ON THE HEMO MONITOR SOLUTION POSSIBLE SOLUTIONS INCLUDE PROPER CONNECTION OF THE TRANSDUCER CABLE INTO THE PDM FOR THE NECESSARY PRESSURE CHANNEL THE TRANSDUCER CABLE CONNECTION INTO THE CUSTOMERSUPPLIED TRANSDUCER AND THE INTEGRITY OF THE AIRFREE LINE TO ENABLE A MEASUREMENT FROM THE MOST DISTAL CATHETER LOCATION TO THE HEMO SCREEN ENSURE THE CORRECT PRESSURE CHANNEL IS SELECTED FOR DISPLAY AND ON THE CORRECT SCALE ADDITIONALLY FLUID SUCH AS SALINE OR CONTRAST AT A CONNECTION POINT MAY COMPROMISE A SECURE FIT OR PROVIDE ERRONEOUS VALUES FOR DISPLAY CONSIDER REPLACING TRANSDUCER CABLE FROM PDM CUSTOMERSUPPLIED TRANSDUCER KIT OR CHECKING PLACEMENT OF STOPCOCKS TO OPEN OR CLOSED AS APPROPRIATE | No Device Problem Found | Medical Evaluation result code
6809682 | 83214311 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAS A UTI AND STARTED TAKING MEDICATION FOR IT YESTERDAY ON B PATIENT STILL HAD A UTI NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Urinary Tract | Patient problem code
6809682 | 83214311 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAS A UTI AND STARTED TAKING MEDICATION FOR IT YESTERDAY ON B PATIENT STILL HAD A UTI NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6806026 | 83215421 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6806026 | 83215421 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6809908 | 83221995 | N | PATIENT INFORMATION COULD NOT BE OBTAINED AFTER MULTIPLE ATTEMPTS ATTEMPTS WERE MADE AS FOLLOWS PHONE PHONE VOICEMAIL B A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE REPORTED COMPLAINT INSPECTION OF THE EQUIPMENT NOTED THAT THE FLUSH BUTTON WAS CAUGHT ON THE TABLETOP THE TABLETOP WAS REMOVED INSPECTED AND REINSTALLED | Misassembled During Installation | Device problem code
6809908 | 83221995 | N | PATIENT INFORMATION COULD NOT BE OBTAINED AFTER MULTIPLE ATTEMPTS ATTEMPTS WERE MADE AS FOLLOWS PHONE PHONE VOICEMAIL B A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE REPORTED COMPLAINT INSPECTION OF THE EQUIPMENT NOTED THAT THE FLUSH BUTTON WAS CAUGHT ON THE TABLETOP THE TABLETOP WAS REMOVED INSPECTED AND REINSTALLED | Assembly Problem Identified | Medical Evaluation result code
6809906 | 83222554 | D | IT WAS REPORTED THAT THE END USER STARTED USING AQUACEL AG EXTRA TO THE RIGHT AND LEFT ISCHIUM PRESSURE ULCERS SOMETIME IN B THE DRESSING WAS CHANGED DAILY BY HER HUSBAND OR HOME HEALTH NURSE THE PERIWOUND SKIN WAS CLEANED WITH MEDLINE SKINTEGRITY WOUND CLEANSER AQUACEL AG EXTRA WAS APPLIED THEN COVERED WITH GAUZE AND SECURED WITH SELFADHESIVE ROLL DRESSING END USER DEVELOPED CELLULITIS IE RED WEEPY EDEMATOUS SKIN SURROUNDING THE WOUNDS WHICH THEN SPREAD TO ENTIRE BODY UP TO HER MOUTH IT WAS REPORTED THAT THE END USER EXPERIENCED FROTHING AT THE MOUTH AT THIS TIME AND THE WOUND BED DEVELOPED ESCHAR A WOUND CULTURE WAS TAKEN AND END USER WAS STARTED ON PRESCRIPTION BACTRIM THE END USER FEELS THE CELLULITIS AND OTHER SYMPTOMS ARE A RESULT OF AN ALLERGIC REACTION TO THE AQUACEL AG EXTRA HOWEVER IT IS NOT KNOWN IF THE DOCTOR PROVIDED THE SAME DIAGNOSIS IT WAS ALSO REPORTED THAT THE END USER AT FIRST DID NOT THINK THE PRODUCT WAS THE CAUSE OF THE ALLERGIC REACTION BUT AT THE BEGINNING OF B THE AQUACEL AG EXTRA WAS DISCONTINUED IT WAS NOTED THAT THE END USER RETURNED TO USING THE PREVIOUS DRESSING DAKINS SOLUTION AND THE ISSUE WAS RESOLVED COMPLETELY WITHIN DAYS THE SKIN IS NOW FLESHTONE AND NOT EDEMATOUS OR WEEPY AND THE WOUND BED IS PINK THE END USER WAS ADVISED TO DISCONTINUE THE USE OF AQUACEL AG EXTRA NO FURTHER INFORMATION HAS BEEN PROVIDED | Pressure Sores/Ulcers | Patient problem code
6809906 | 83222554 | D | IT WAS REPORTED THAT THE END USER STARTED USING AQUACEL AG EXTRA TO THE RIGHT AND LEFT ISCHIUM PRESSURE ULCERS SOMETIME IN B THE DRESSING WAS CHANGED DAILY BY HER HUSBAND OR HOME HEALTH NURSE THE PERIWOUND SKIN WAS CLEANED WITH MEDLINE SKINTEGRITY WOUND CLEANSER AQUACEL AG EXTRA WAS APPLIED THEN COVERED WITH GAUZE AND SECURED WITH SELFADHESIVE ROLL DRESSING END USER DEVELOPED CELLULITIS IE RED WEEPY EDEMATOUS SKIN SURROUNDING THE WOUNDS WHICH THEN SPREAD TO ENTIRE BODY UP TO HER MOUTH IT WAS REPORTED THAT THE END USER EXPERIENCED FROTHING AT THE MOUTH AT THIS TIME AND THE WOUND BED DEVELOPED ESCHAR A WOUND CULTURE WAS TAKEN AND END USER WAS STARTED ON PRESCRIPTION BACTRIM THE END USER FEELS THE CELLULITIS AND OTHER SYMPTOMS ARE A RESULT OF AN ALLERGIC REACTION TO THE AQUACEL AG EXTRA HOWEVER IT IS NOT KNOWN IF THE DOCTOR PROVIDED THE SAME DIAGNOSIS IT WAS ALSO REPORTED THAT THE END USER AT FIRST DID NOT THINK THE PRODUCT WAS THE CAUSE OF THE ALLERGIC REACTION BUT AT THE BEGINNING OF B THE AQUACEL AG EXTRA WAS DISCONTINUED IT WAS NOTED THAT THE END USER RETURNED TO USING THE PREVIOUS DRESSING DAKINS SOLUTION AND THE ISSUE WAS RESOLVED COMPLETELY WITHIN DAYS THE SKIN IS NOW FLESHTONE AND NOT EDEMATOUS OR WEEPY AND THE WOUND BED IS PINK THE END USER WAS ADVISED TO DISCONTINUE THE USE OF AQUACEL AG EXTRA NO FURTHER INFORMATION HAS BEEN PROVIDED | Hypersensitivity | Patient problem code
6809906 | 83222554 | D | IT WAS REPORTED THAT THE END USER STARTED USING AQUACEL AG EXTRA TO THE RIGHT AND LEFT ISCHIUM PRESSURE ULCERS SOMETIME IN B THE DRESSING WAS CHANGED DAILY BY HER HUSBAND OR HOME HEALTH NURSE THE PERIWOUND SKIN WAS CLEANED WITH MEDLINE SKINTEGRITY WOUND CLEANSER AQUACEL AG EXTRA WAS APPLIED THEN COVERED WITH GAUZE AND SECURED WITH SELFADHESIVE ROLL DRESSING END USER DEVELOPED CELLULITIS IE RED WEEPY EDEMATOUS SKIN SURROUNDING THE WOUNDS WHICH THEN SPREAD TO ENTIRE BODY UP TO HER MOUTH IT WAS REPORTED THAT THE END USER EXPERIENCED FROTHING AT THE MOUTH AT THIS TIME AND THE WOUND BED DEVELOPED ESCHAR A WOUND CULTURE WAS TAKEN AND END USER WAS STARTED ON PRESCRIPTION BACTRIM THE END USER FEELS THE CELLULITIS AND OTHER SYMPTOMS ARE A RESULT OF AN ALLERGIC REACTION TO THE AQUACEL AG EXTRA HOWEVER IT IS NOT KNOWN IF THE DOCTOR PROVIDED THE SAME DIAGNOSIS IT WAS ALSO REPORTED THAT THE END USER AT FIRST DID NOT THINK THE PRODUCT WAS THE CAUSE OF THE ALLERGIC REACTION BUT AT THE BEGINNING OF B THE AQUACEL AG EXTRA WAS DISCONTINUED IT WAS NOTED THAT THE END USER RETURNED TO USING THE PREVIOUS DRESSING DAKINS SOLUTION AND THE ISSUE WAS RESOLVED COMPLETELY WITHIN DAYS THE SKIN IS NOW FLESHTONE AND NOT EDEMATOUS OR WEEPY AND THE WOUND BED IS PINK THE END USER WAS ADVISED TO DISCONTINUE THE USE OF AQUACEL AG EXTRA NO FURTHER INFORMATION HAS BEEN PROVIDED | Cellulitis | Patient problem code
6809906 | 83222554 | D | IT WAS REPORTED THAT THE END USER STARTED USING AQUACEL AG EXTRA TO THE RIGHT AND LEFT ISCHIUM PRESSURE ULCERS SOMETIME IN B THE DRESSING WAS CHANGED DAILY BY HER HUSBAND OR HOME HEALTH NURSE THE PERIWOUND SKIN WAS CLEANED WITH MEDLINE SKINTEGRITY WOUND CLEANSER AQUACEL AG EXTRA WAS APPLIED THEN COVERED WITH GAUZE AND SECURED WITH SELFADHESIVE ROLL DRESSING END USER DEVELOPED CELLULITIS IE RED WEEPY EDEMATOUS SKIN SURROUNDING THE WOUNDS WHICH THEN SPREAD TO ENTIRE BODY UP TO HER MOUTH IT WAS REPORTED THAT THE END USER EXPERIENCED FROTHING AT THE MOUTH AT THIS TIME AND THE WOUND BED DEVELOPED ESCHAR A WOUND CULTURE WAS TAKEN AND END USER WAS STARTED ON PRESCRIPTION BACTRIM THE END USER FEELS THE CELLULITIS AND OTHER SYMPTOMS ARE A RESULT OF AN ALLERGIC REACTION TO THE AQUACEL AG EXTRA HOWEVER IT IS NOT KNOWN IF THE DOCTOR PROVIDED THE SAME DIAGNOSIS IT WAS ALSO REPORTED THAT THE END USER AT FIRST DID NOT THINK THE PRODUCT WAS THE CAUSE OF THE ALLERGIC REACTION BUT AT THE BEGINNING OF B THE AQUACEL AG EXTRA WAS DISCONTINUED IT WAS NOTED THAT THE END USER RETURNED TO USING THE PREVIOUS DRESSING DAKINS SOLUTION AND THE ISSUE WAS RESOLVED COMPLETELY WITHIN DAYS THE SKIN IS NOW FLESHTONE AND NOT EDEMATOUS OR WEEPY AND THE WOUND BED IS PINK THE END USER WAS ADVISED TO DISCONTINUE THE USE OF AQUACEL AG EXTRA NO FURTHER INFORMATION HAS BEEN PROVIDED | Insufficient Information | Device problem code
6809906 | 83222554 | D | IT WAS REPORTED THAT THE END USER STARTED USING AQUACEL AG EXTRA TO THE RIGHT AND LEFT ISCHIUM PRESSURE ULCERS SOMETIME IN B THE DRESSING WAS CHANGED DAILY BY HER HUSBAND OR HOME HEALTH NURSE THE PERIWOUND SKIN WAS CLEANED WITH MEDLINE SKINTEGRITY WOUND CLEANSER AQUACEL AG EXTRA WAS APPLIED THEN COVERED WITH GAUZE AND SECURED WITH SELFADHESIVE ROLL DRESSING END USER DEVELOPED CELLULITIS IE RED WEEPY EDEMATOUS SKIN SURROUNDING THE WOUNDS WHICH THEN SPREAD TO ENTIRE BODY UP TO HER MOUTH IT WAS REPORTED THAT THE END USER EXPERIENCED FROTHING AT THE MOUTH AT THIS TIME AND THE WOUND BED DEVELOPED ESCHAR A WOUND CULTURE WAS TAKEN AND END USER WAS STARTED ON PRESCRIPTION BACTRIM THE END USER FEELS THE CELLULITIS AND OTHER SYMPTOMS ARE A RESULT OF AN ALLERGIC REACTION TO THE AQUACEL AG EXTRA HOWEVER IT IS NOT KNOWN IF THE DOCTOR PROVIDED THE SAME DIAGNOSIS IT WAS ALSO REPORTED THAT THE END USER AT FIRST DID NOT THINK THE PRODUCT WAS THE CAUSE OF THE ALLERGIC REACTION BUT AT THE BEGINNING OF B THE AQUACEL AG EXTRA WAS DISCONTINUED IT WAS NOTED THAT THE END USER RETURNED TO USING THE PREVIOUS DRESSING DAKINS SOLUTION AND THE ISSUE WAS RESOLVED COMPLETELY WITHIN DAYS THE SKIN IS NOW FLESHTONE AND NOT EDEMATOUS OR WEEPY AND THE WOUND BED IS PINK THE END USER WAS ADVISED TO DISCONTINUE THE USE OF AQUACEL AG EXTRA NO FURTHER INFORMATION HAS BEEN PROVIDED | Ulcer | Patient problem code
6804162 | 83223666 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO INVESTIGATE AFTER INSPECTING THE IABP UNITS LOGS THE FSE FOUND ELECTRICAL CODE TO CORRECT THE ISSUE THE FSE PERFORMED SOLENOID DRIVER BOARD FIELD SAFETY CORRECTIVE ACTION THE FSE PERFORMED A FULL CALIBRATION AND FUNCTIONAL TEST IN ACCORDANCE WITH THE SYSTEMS RESPECTIVE SERVICE MANUAL THE IABP WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Findings Available | Medical Evaluation result code
6810047 | 83226044 | N | UPON ANALYSIS A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX RETRACTED AND CLOGGED WITH BLOODTISSUE THE STYLET WAS INSERTED INSIDE THE LEAD AND BLOODBODY FLUIDS WERE NOTED INSIDE THE INNER COIL THIS WAS MOST LIKELY THE CAUSE OF THE STYLET BEING IMMOBILE INSIDE THE LEAD THE LEAD WAS CUT IN ORDER TO EVALUATE THE HELIX FUNCTION BY EXCLUDING THE OVER TORQUED AND OVERLAPPING REGION OF THE INNER COIL AFTER CUTTING THE LEAD AND CLEANING THE HELIX COULD BE EXTENDED AND RETRACTED FROM THE INNER COIL FULL HELIX EXTENSION LENGTH WAS MEASURED AND FOUND TO MEET SPECIFICATION | No Device Problem Found | Medical Evaluation result code
6810164 | 83228972 | D | IT WAS REPORTED THAT WHILE USING SPECTRUM PUMPS A NURSE HAD FREQUENT ISSUES WITH UPSTREAM OCCLUSION ALARMS WHEN NO OCCLUSION WAS PRESENT THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Known Impact Or Consequence To Patient | Patient problem code
6810164 | 83228972 | D | IT WAS REPORTED THAT WHILE USING SPECTRUM PUMPS A NURSE HAD FREQUENT ISSUES WITH UPSTREAM OCCLUSION ALARMS WHEN NO OCCLUSION WAS PRESENT THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | Insufficient Information | Device problem code
6805143 | 83230773 | N | THERE HAVE BEEN A SIGNIFICANT NUMBER OF TUBE REPLACEMENTS OR REMOVALS AND THIS IS THE FIRST SUCH EVENT TO BE REPORTED MATERIAL SEPARATION HAS BEEN ANTICIPATED IN THE ATUBE RISK ASSESSMENT AND MITIGATED AS FAR AS POSSIBLE BY DESIGN AND TESTING OF THE ATUBE AND ASSESSMENT OF THE RISKS ASSOCIATED WITH REMOVAL OF SEPARATED PIECES OF THE ATUBE BY ENDOSCOPIC MEANS WHAT HAS BEEN LEARNED AS A RESULT OF THIS EVENT IS THAT PLACEMENT OF THE SNARE MAY BE IMPORTANT PARTICULARLY IN OLDER IMPLANTS AS A RESULT THE INSTRUCTIONS FOR ATUBE REMOVAL ARE BEING REVISED TO SPECIFY PLACEMENT OF THE SNARE AS AN ADDITIONAL MEANS OF RISK REDUCTION THE CLINICIANS ARE NOW REMINDED NOT TO ATTACH THE SNARE AT THE ASPIRATION HOLES BUT RATHER AT THE FULL CIRCUMFERENCE OF THE ATUBE FOR MORE SECURE PURCHASE DURING REMOVAL | Material Separation | Device problem code
6805143 | 83230773 | N | THERE HAVE BEEN A SIGNIFICANT NUMBER OF TUBE REPLACEMENTS OR REMOVALS AND THIS IS THE FIRST SUCH EVENT TO BE REPORTED MATERIAL SEPARATION HAS BEEN ANTICIPATED IN THE ATUBE RISK ASSESSMENT AND MITIGATED AS FAR AS POSSIBLE BY DESIGN AND TESTING OF THE ATUBE AND ASSESSMENT OF THE RISKS ASSOCIATED WITH REMOVAL OF SEPARATED PIECES OF THE ATUBE BY ENDOSCOPIC MEANS WHAT HAS BEEN LEARNED AS A RESULT OF THIS EVENT IS THAT PLACEMENT OF THE SNARE MAY BE IMPORTANT PARTICULARLY IN OLDER IMPLANTS AS A RESULT THE INSTRUCTIONS FOR ATUBE REMOVAL ARE BEING REVISED TO SPECIFY PLACEMENT OF THE SNARE AS AN ADDITIONAL MEANS OF RISK REDUCTION THE CLINICIANS ARE NOW REMINDED NOT TO ATTACH THE SNARE AT THE ASPIRATION HOLES BUT RATHER AT THE FULL CIRCUMFERENCE OF THE ATUBE FOR MORE SECURE PURCHASE DURING REMOVAL | Fracture Problem | Medical Evaluation result code
6809523 | 83232930 | N | AGE AT TIME OF EVENT B B DEVICE EVALUATED BY MFRRETURNED PRODUCT CONSISTED OF A SOLENT PROXI THROMBECTOMY SYSTEM THE PUMP EFFLUENTSUPPLY LINE SHAFT AND TIP WERE MICROSCOPICALLY TACTILE AND VISUALLY INSPECTED THERE WAS BLOOD OUTSIDE AND INSIDE THE DEVICE THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS NOT CONNECTED WHEN THE DEVICE WAS RECEIVED THE BARB FITTING NUT WAS CONNECTED TO THE SNAP FIT ADAPTER AND FUNCTIONAL TESTING WAS PERFORMED FUNCTIONAL TESTING WAS PERFORMED BY PLACING THE DEVICE IN THE ANGIOJET CONSOLE FUNCTIONAL TESTING CONSISTED OF RUNNING THE COMPLAINT DEVICE THROUGH THE FULL PRIMING CYCLE AND SECONDS IN THROMBECTOMY MODE THE DEVICE RAN WITHIN NORMAL RANGE KPSI WITHOUT ANY LEAKS ISSUE OR ALARMS THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS ANALYZED WHEN THE DEVICE WAS MICROSCOPICALLY EXAMINED IT WAS FOUND THAT THERE WAS NO ADHESIVE ON THE BARB FITTING NUT TO THE SNAP FIT ADAPTER THE MANUFACTURING BATCH RECORD REVIEW CONFIRMED THAT THE DEVICE MET ALL MATERIAL ASSEMBLY AND PERFORMANCE SPECIFICATIONS THE ROOT CAUSE HAS BEEN DETERMINED TO BE MANUFACTURING RELATED B | Connection Problem | Device problem code
6809523 | 83232930 | N | AGE AT TIME OF EVENT B B DEVICE EVALUATED BY MFRRETURNED PRODUCT CONSISTED OF A SOLENT PROXI THROMBECTOMY SYSTEM THE PUMP EFFLUENTSUPPLY LINE SHAFT AND TIP WERE MICROSCOPICALLY TACTILE AND VISUALLY INSPECTED THERE WAS BLOOD OUTSIDE AND INSIDE THE DEVICE THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS NOT CONNECTED WHEN THE DEVICE WAS RECEIVED THE BARB FITTING NUT WAS CONNECTED TO THE SNAP FIT ADAPTER AND FUNCTIONAL TESTING WAS PERFORMED FUNCTIONAL TESTING WAS PERFORMED BY PLACING THE DEVICE IN THE ANGIOJET CONSOLE FUNCTIONAL TESTING CONSISTED OF RUNNING THE COMPLAINT DEVICE THROUGH THE FULL PRIMING CYCLE AND SECONDS IN THROMBECTOMY MODE THE DEVICE RAN WITHIN NORMAL RANGE KPSI WITHOUT ANY LEAKS ISSUE OR ALARMS THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS ANALYZED WHEN THE DEVICE WAS MICROSCOPICALLY EXAMINED IT WAS FOUND THAT THERE WAS NO ADHESIVE ON THE BARB FITTING NUT TO THE SNAP FIT ADAPTER THE MANUFACTURING BATCH RECORD REVIEW CONFIRMED THAT THE DEVICE MET ALL MATERIAL ASSEMBLY AND PERFORMANCE SPECIFICATIONS THE ROOT CAUSE HAS BEEN DETERMINED TO BE MANUFACTURING RELATED B | Component or Accessory Incompatibility | Device Problem Code
6809523 | 83232930 | N | AGE AT TIME OF EVENT B B DEVICE EVALUATED BY MFRRETURNED PRODUCT CONSISTED OF A SOLENT PROXI THROMBECTOMY SYSTEM THE PUMP EFFLUENTSUPPLY LINE SHAFT AND TIP WERE MICROSCOPICALLY TACTILE AND VISUALLY INSPECTED THERE WAS BLOOD OUTSIDE AND INSIDE THE DEVICE THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS NOT CONNECTED WHEN THE DEVICE WAS RECEIVED THE BARB FITTING NUT WAS CONNECTED TO THE SNAP FIT ADAPTER AND FUNCTIONAL TESTING WAS PERFORMED FUNCTIONAL TESTING WAS PERFORMED BY PLACING THE DEVICE IN THE ANGIOJET CONSOLE FUNCTIONAL TESTING CONSISTED OF RUNNING THE COMPLAINT DEVICE THROUGH THE FULL PRIMING CYCLE AND SECONDS IN THROMBECTOMY MODE THE DEVICE RAN WITHIN NORMAL RANGE KPSI WITHOUT ANY LEAKS ISSUE OR ALARMS THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS ANALYZED WHEN THE DEVICE WAS MICROSCOPICALLY EXAMINED IT WAS FOUND THAT THERE WAS NO ADHESIVE ON THE BARB FITTING NUT TO THE SNAP FIT ADAPTER THE MANUFACTURING BATCH RECORD REVIEW CONFIRMED THAT THE DEVICE MET ALL MATERIAL ASSEMBLY AND PERFORMANCE SPECIFICATIONS THE ROOT CAUSE HAS BEEN DETERMINED TO BE MANUFACTURING RELATED B | Leak / Splash | Device problem code
6809523 | 83232930 | N | AGE AT TIME OF EVENT B B DEVICE EVALUATED BY MFRRETURNED PRODUCT CONSISTED OF A SOLENT PROXI THROMBECTOMY SYSTEM THE PUMP EFFLUENTSUPPLY LINE SHAFT AND TIP WERE MICROSCOPICALLY TACTILE AND VISUALLY INSPECTED THERE WAS BLOOD OUTSIDE AND INSIDE THE DEVICE THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS NOT CONNECTED WHEN THE DEVICE WAS RECEIVED THE BARB FITTING NUT WAS CONNECTED TO THE SNAP FIT ADAPTER AND FUNCTIONAL TESTING WAS PERFORMED FUNCTIONAL TESTING WAS PERFORMED BY PLACING THE DEVICE IN THE ANGIOJET CONSOLE FUNCTIONAL TESTING CONSISTED OF RUNNING THE COMPLAINT DEVICE THROUGH THE FULL PRIMING CYCLE AND SECONDS IN THROMBECTOMY MODE THE DEVICE RAN WITHIN NORMAL RANGE KPSI WITHOUT ANY LEAKS ISSUE OR ALARMS THE BARB FITTING NUT TO THE SNAP FIT ADAPTER WAS ANALYZED WHEN THE DEVICE WAS MICROSCOPICALLY EXAMINED IT WAS FOUND THAT THERE WAS NO ADHESIVE ON THE BARB FITTING NUT TO THE SNAP FIT ADAPTER THE MANUFACTURING BATCH RECORD REVIEW CONFIRMED THAT THE DEVICE MET ALL MATERIAL ASSEMBLY AND PERFORMANCE SPECIFICATIONS THE ROOT CAUSE HAS BEEN DETERMINED TO BE MANUFACTURING RELATED B | No Findings Available | Medical Evaluation result code
6809821 | 83235816 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMERS BLOOD GLUCOSE LEVELS STARTED TO RISE ON B HOWEVER THE CUSTOMER WAS UNAWARE OF THE CAUSE ON B THE CUSTOMER CONTINUED TO EXPERIENCE ELEVATED BLOOD GLUCOSE LEVELS WHEN THEY NOTICED THE ADHESIVE HAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT HOURS AFTER THE EVENT AND MANUAL INJECTIONS WERE ADMINISTERED TO TREAT THE HYPERGLYCEMIA B | Device Dislodged or Dislocated | Device problem code
6809821 | 83235816 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMERS BLOOD GLUCOSE LEVELS STARTED TO RISE ON B HOWEVER THE CUSTOMER WAS UNAWARE OF THE CAUSE ON B THE CUSTOMER CONTINUED TO EXPERIENCE ELEVATED BLOOD GLUCOSE LEVELS WHEN THEY NOTICED THE ADHESIVE HAD FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT HOURS AFTER THE EVENT AND MANUAL INJECTIONS WERE ADMINISTERED TO TREAT THE HYPERGLYCEMIA B | Hyperglycemia | Patient problem code
6798921 | 83243544 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Device Problem Found | Medical Evaluation result code
6802176 | 83245534 | N | THE RETURNED CAC ADAPTERS CAC AND CAC WERE FOUND TO BE FULLY FUNCTIONAL FOLLOWING A LIGHTNING STRIKE VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICES PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING PER THE INSTRUCTIONS FOR USE IFU PATIENTS ARE INSTRUCTED TO ALWAYS KEEP A SPARE CONTROLLER AVAILABLE AT ALL TIMES PER THE INSTRUCTIONS FOR USE IFU PATIENTS ARE INSTRUCTED TO ALWAYS INVESTIGATE AND IF POSSIBLE CORRECT THE CAUSE OF ANY ALARM SILENCING AN ALARM DOES NOT RESOLVE THE ALARM CONDITION THE INSTRUCTIONS FOR USE IFU INCLUDES THE FOLLOWING WARNING NEVER DISCONNECT BOTH POWER SOURCES BATTERIES AND AC OR DC ADAPTER AT THE SAME TIME SINCE THIS WILL STOP THE PUMP AT LEAST ONE POWER SOURCE MUST BE CONNECTED AT ALL TIMES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA CAC ADAPTER CAC US MANUFACTURING DATE EXPIRATION DATE B | Disconnection | Device problem code
6802176 | 83245534 | N | THE RETURNED CAC ADAPTERS CAC AND CAC WERE FOUND TO BE FULLY FUNCTIONAL FOLLOWING A LIGHTNING STRIKE VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICES PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING PER THE INSTRUCTIONS FOR USE IFU PATIENTS ARE INSTRUCTED TO ALWAYS KEEP A SPARE CONTROLLER AVAILABLE AT ALL TIMES PER THE INSTRUCTIONS FOR USE IFU PATIENTS ARE INSTRUCTED TO ALWAYS INVESTIGATE AND IF POSSIBLE CORRECT THE CAUSE OF ANY ALARM SILENCING AN ALARM DOES NOT RESOLVE THE ALARM CONDITION THE INSTRUCTIONS FOR USE IFU INCLUDES THE FOLLOWING WARNING NEVER DISCONNECT BOTH POWER SOURCES BATTERIES AND AC OR DC ADAPTER AT THE SAME TIME SINCE THIS WILL STOP THE PUMP AT LEAST ONE POWER SOURCE MUST BE CONNECTED AT ALL TIMES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA CAC ADAPTER CAC US MANUFACTURING DATE EXPIRATION DATE B | No Findings Available | Medical Evaluation result code
6810765 | 83245684 | N | B A THIRD PARTY SERVICE AGENT EVALUATED THE CUSTOMERS DEVICE AND VERIFIED THE REPORTED ISSUE IT WAS DETERMINED THAT A BAD CONNECTION IN THE INTERNAL PADDLES BEING USED IN CONJUNCTION WITH THE DEVICE HAD CAUSED THE REPORTED ISSUE NO FURTHER DETAILS WERE PROVIDED NO IDENTIFYING INFORMATION ABOUT THE INTERNAL PADDLES WAS PROVIDED BY THE THIRD PARTY SERVICE AGENT AFTER OBSERVING PROPER DEVICE OPERATION THROUGH FUNCTIONAL AND PERFORMANCE TESTING THE UNIT WAS RETURNED TO THE CUSTOMER FOR USE NEITHER THE DEVICE NOR THE INTERNAL PADDLES HAVE BEEN RETURNED TO PHYSIOCONTROL FOR EVALUATION | No Device Problem Found | Medical Evaluation result code
6810868 | 83249397 | N | REVIEW OF THE UNITS SOFTWARE LOG SHOW THAT THE DEVICE OPERATED FOR A TOTAL OF COMPRESSIONS THE LOG SHOWS THERE WERE THREE RESCUE EVENTS AND A NUMBER OF SMALL RUNS WITH EITHER ERRATIC LOADS ADHOC TESTS OR VERY STABLE LOADS SPRINGS OR MANIKINS B IN EVERY CASE THE UNIT GAVE COMPRESSIONS WELL WITHIN TOLERANCES BOTH IN TIMING AND POSITION DEPTHTHERE WAS NO INDICATION OF ANY MALFUNCTION TESTING OF THE DEVICE BY DEFIBTECH WHICH INCLUDED RUNNING THE DEVICE ON A TEST FIXTURE FOR MIN COMPRESSIONS AND A VISUAL INSPECTION OF THE DEVICES INTERNALS CONNECTIONS CABLING SOLDERING ETC SHOWED THAT THE DEVICE IS FUNCTIONING PROPERLY AS DESIGNED | No Device Problem Found | Medical Evaluation result code
6798939 | 83254728 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Device Problem Found | Medical Evaluation result code
6801530 | 83255346 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE WITH SLIGHT RED BLOOD CELL ADHERENCE | No Device Problem Found | Medical Evaluation result code
6803458 | 83255672 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY POSITIVE | No Device Problem Found | Medical Evaluation result code
6811471 | 83271871 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device problem code
6799236 | 83273425 | N | ONE CAC ADAPTER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA FUNCTIONAL TESTING ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN OPEN POSITIVE WIRE AS A RESULT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN CONNECTION WITHIN THE CONNECTOR PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE THE CONTROLLER REQUIRES TWO POWER SOURCES FOR SAFETY EITHER TWO BATTERIES OR ONE BATTERY AND AN AC ADAPTER OR DC ADAPTER PROPER CONNECTION IS VERIFIED WHEN THE ACDC INDICATOR ON THE CONTROLLER TURNS GREEN AND THE CORRESPONDING BATTERY INDICATOR TURNS OFF IF THE ACDC INDICATOR DOESNT TURN GREEN THE CONTROLLER IS USING BATTERY POWER AND THE POWER DISCONNECT ALARM WILL SOUND THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6799236 | 83273425 | N | ONE CAC ADAPTER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA FUNCTIONAL TESTING ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN OPEN POSITIVE WIRE AS A RESULT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN CONNECTION WITHIN THE CONNECTOR PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE THE CONTROLLER REQUIRES TWO POWER SOURCES FOR SAFETY EITHER TWO BATTERIES OR ONE BATTERY AND AN AC ADAPTER OR DC ADAPTER PROPER CONNECTION IS VERIFIED WHEN THE ACDC INDICATOR ON THE CONTROLLER TURNS GREEN AND THE CORRESPONDING BATTERY INDICATOR TURNS OFF IF THE ACDC INDICATOR DOESNT TURN GREEN THE CONTROLLER IS USING BATTERY POWER AND THE POWER DISCONNECT ALARM WILL SOUND THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6799236 | 83273425 | N | ONE CAC ADAPTER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA FUNCTIONAL TESTING ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN OPEN POSITIVE WIRE AS A RESULT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN CONNECTION WITHIN THE CONNECTOR PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE THE CONTROLLER REQUIRES TWO POWER SOURCES FOR SAFETY EITHER TWO BATTERIES OR ONE BATTERY AND AN AC ADAPTER OR DC ADAPTER PROPER CONNECTION IS VERIFIED WHEN THE ACDC INDICATOR ON THE CONTROLLER TURNS GREEN AND THE CORRESPONDING BATTERY INDICATOR TURNS OFF IF THE ACDC INDICATOR DOESNT TURN GREEN THE CONTROLLER IS USING BATTERY POWER AND THE POWER DISCONNECT ALARM WILL SOUND THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Noise, Audible | Device problem code
6799236 | 83273425 | N | ONE CAC ADAPTER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA FUNCTIONAL TESTING ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN OPEN POSITIVE WIRE AS A RESULT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN CONNECTION WITHIN THE CONNECTOR PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE THE CONTROLLER REQUIRES TWO POWER SOURCES FOR SAFETY EITHER TWO BATTERIES OR ONE BATTERY AND AN AC ADAPTER OR DC ADAPTER PROPER CONNECTION IS VERIFIED WHEN THE ACDC INDICATOR ON THE CONTROLLER TURNS GREEN AND THE CORRESPONDING BATTERY INDICATOR TURNS OFF IF THE ACDC INDICATOR DOESNT TURN GREEN THE CONTROLLER IS USING BATTERY POWER AND THE POWER DISCONNECT ALARM WILL SOUND THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Non-Functional Defect | Medical Evaluation result code
6811618 | 83279219 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER CLINICAL PROCEDURE NO PATIENTS ETHNICITY AND RACE WERE GIVEN PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Failure to Osseointegrate | Device problem code
6811618 | 83279219 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER CLINICAL PROCEDURE NO PATIENTS ETHNICITY AND RACE WERE GIVEN PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Results Pending Completion of Investigation | Medical Evaluation result code
6811776 | 83293257 | N | REPORT SOURCE B | Insufficient Information | Device Problem Code
6811472 | 83293408 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6811455 | 83293583 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6811467 | 83293763 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6811470 | 83293900 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6811580 | 83296111 | N | NO REPORT OF PATIENT INVOLVEMENT GE HEALTHCARE RECOMMENDED TO THE HOSPITAL TO REPLACE THE COMPACT FLASH CARD | Insufficient Information | Device problem code
6811917 | 83310589 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Scratched Material | Device problem code
6811917 | 83310589 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
6811917 | 83310589 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
6811917 | 83310589 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Fracture Problem | Medical Evaluation result code
6794442 | 83313614 | N | LOT REVIEW A REVIEW OF LOT SHOWED B UNITS MANUFACTURED TESTED INSPECTED AND RELEASED IN MAY CITING NO EXCEPTION DOCUMENTS RECEIVING INSPECTION RECEIVED ONE USED CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW SMARTSITE INFUSION SET CAREFUSION THE AS RECEIVED ONE CH SAMPLE WAS PRIMED AND SPINNING FEATURE WASNT ACTIVATED FUNCTIONAL TESTING EACH OF THE RETURNED CH SPINNING SPIROS WERE CONNECTED TO THE NEW SMARTSITE INFUSION PUMP SET THE CONNECTION RESIDUAL AND DISCONNECT TORQUE VALUES WERE MEASURED AND MET PRODUCT SPECIFICATION QE ANALYSIS SUMMARY THE CH SPINNING SPIROS WERE ALL RETURNED WITHOUT THE SPIN FEATURES ACTIVATED EACH OF THE CH SPINNING SPIROS WERE SECURELY CONNECTED TO THE SMARTSITE INFUSION PUMP SET B MALE LUER WHEN THE SPIN FEATURE WAS ACTIVATED DURING CONNECTION ALL SPIROS MET PRODUCT SPECIFICATION AND DID NOT DISCONNECT WHEN ATTACHED ICU MEDICAL WILL MONITOR AND TREND | Disconnection | Device problem code
6794442 | 83313614 | N | LOT REVIEW A REVIEW OF LOT SHOWED B UNITS MANUFACTURED TESTED INSPECTED AND RELEASED IN MAY CITING NO EXCEPTION DOCUMENTS RECEIVING INSPECTION RECEIVED ONE USED CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW SMARTSITE INFUSION SET CAREFUSION THE AS RECEIVED ONE CH SAMPLE WAS PRIMED AND SPINNING FEATURE WASNT ACTIVATED FUNCTIONAL TESTING EACH OF THE RETURNED CH SPINNING SPIROS WERE CONNECTED TO THE NEW SMARTSITE INFUSION PUMP SET THE CONNECTION RESIDUAL AND DISCONNECT TORQUE VALUES WERE MEASURED AND MET PRODUCT SPECIFICATION QE ANALYSIS SUMMARY THE CH SPINNING SPIROS WERE ALL RETURNED WITHOUT THE SPIN FEATURES ACTIVATED EACH OF THE CH SPINNING SPIROS WERE SECURELY CONNECTED TO THE SMARTSITE INFUSION PUMP SET B MALE LUER WHEN THE SPIN FEATURE WAS ACTIVATED DURING CONNECTION ALL SPIROS MET PRODUCT SPECIFICATION AND DID NOT DISCONNECT WHEN ATTACHED ICU MEDICAL WILL MONITOR AND TREND | Failure to Disconnect | Device problem code
6794442 | 83313614 | N | LOT REVIEW A REVIEW OF LOT SHOWED B UNITS MANUFACTURED TESTED INSPECTED AND RELEASED IN MAY CITING NO EXCEPTION DOCUMENTS RECEIVING INSPECTION RECEIVED ONE USED CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW SMARTSITE INFUSION SET CAREFUSION THE AS RECEIVED ONE CH SAMPLE WAS PRIMED AND SPINNING FEATURE WASNT ACTIVATED FUNCTIONAL TESTING EACH OF THE RETURNED CH SPINNING SPIROS WERE CONNECTED TO THE NEW SMARTSITE INFUSION PUMP SET THE CONNECTION RESIDUAL AND DISCONNECT TORQUE VALUES WERE MEASURED AND MET PRODUCT SPECIFICATION QE ANALYSIS SUMMARY THE CH SPINNING SPIROS WERE ALL RETURNED WITHOUT THE SPIN FEATURES ACTIVATED EACH OF THE CH SPINNING SPIROS WERE SECURELY CONNECTED TO THE SMARTSITE INFUSION PUMP SET B MALE LUER WHEN THE SPIN FEATURE WAS ACTIVATED DURING CONNECTION ALL SPIROS MET PRODUCT SPECIFICATION AND DID NOT DISCONNECT WHEN ATTACHED ICU MEDICAL WILL MONITOR AND TREND | Connection Problem | Device problem code
6794442 | 83313614 | N | LOT REVIEW A REVIEW OF LOT SHOWED B UNITS MANUFACTURED TESTED INSPECTED AND RELEASED IN MAY CITING NO EXCEPTION DOCUMENTS RECEIVING INSPECTION RECEIVED ONE USED CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW CH ONCOLOGY KIT WSPINNING SPIROS CCLIP RED CAP LOT ONE NEW SMARTSITE INFUSION SET CAREFUSION THE AS RECEIVED ONE CH SAMPLE WAS PRIMED AND SPINNING FEATURE WASNT ACTIVATED FUNCTIONAL TESTING EACH OF THE RETURNED CH SPINNING SPIROS WERE CONNECTED TO THE NEW SMARTSITE INFUSION PUMP SET THE CONNECTION RESIDUAL AND DISCONNECT TORQUE VALUES WERE MEASURED AND MET PRODUCT SPECIFICATION QE ANALYSIS SUMMARY THE CH SPINNING SPIROS WERE ALL RETURNED WITHOUT THE SPIN FEATURES ACTIVATED EACH OF THE CH SPINNING SPIROS WERE SECURELY CONNECTED TO THE SMARTSITE INFUSION PUMP SET B MALE LUER WHEN THE SPIN FEATURE WAS ACTIVATED DURING CONNECTION ALL SPIROS MET PRODUCT SPECIFICATION AND DID NOT DISCONNECT WHEN ATTACHED ICU MEDICAL WILL MONITOR AND TREND | Non-Functional Defect | Medical Evaluation result code
6813010 | 83367135 | D | IT WAS REPORTED THAT A PATIENT SHOWS CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT IS REPORTED TO HAVE COMPLETE HIP PERFORMANCE ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Insufficient Information | Device problem code
6813010 | 83367135 | D | IT WAS REPORTED THAT A PATIENT SHOWS CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT IS REPORTED TO HAVE COMPLETE HIP PERFORMANCE ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Calcification | Patient problem code
6813209 | 83376257 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER CLINICAL PROCEDURE AND THE IMPLANT WAS REMOVED DUE TO PAIN ANDOR NUMBNESS THE PATIENT HAD PREEXISTING CONDITIONS WHICH INCLUDE HEART DISEASE HIGH BLOOD PRESSURE AND DIABETES TYPE II ALSO HE WAS REPORTED TO HAVE BONE TYPE II QUALITY AND EXPERIENCED GENERAL BONE LOSS INFECTION AND GRANULATED TISSUE AT OR AROUND THE IMPLANT SITE HOWEVER THERE WAS NO HARM CAUSE TO THE PATIENT THE DENTIST DID NOT KNOW THE PATIENTS ETHNICITY AND RACE PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Failure to Osseointegrate | Device problem code
6813209 | 83376257 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER CLINICAL PROCEDURE AND THE IMPLANT WAS REMOVED DUE TO PAIN ANDOR NUMBNESS THE PATIENT HAD PREEXISTING CONDITIONS WHICH INCLUDE HEART DISEASE HIGH BLOOD PRESSURE AND DIABETES TYPE II ALSO HE WAS REPORTED TO HAVE BONE TYPE II QUALITY AND EXPERIENCED GENERAL BONE LOSS INFECTION AND GRANULATED TISSUE AT OR AROUND THE IMPLANT SITE HOWEVER THERE WAS NO HARM CAUSE TO THE PATIENT THE DENTIST DID NOT KNOW THE PATIENTS ETHNICITY AND RACE PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Component or Accessory Incompatibility | Device problem code
6813209 | 83376257 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AFTER CLINICAL PROCEDURE AND THE IMPLANT WAS REMOVED DUE TO PAIN ANDOR NUMBNESS THE PATIENT HAD PREEXISTING CONDITIONS WHICH INCLUDE HEART DISEASE HIGH BLOOD PRESSURE AND DIABETES TYPE II ALSO HE WAS REPORTED TO HAVE BONE TYPE II QUALITY AND EXPERIENCED GENERAL BONE LOSS INFECTION AND GRANULATED TISSUE AT OR AROUND THE IMPLANT SITE HOWEVER THERE WAS NO HARM CAUSE TO THE PATIENT THE DENTIST DID NOT KNOW THE PATIENTS ETHNICITY AND RACE PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Interoperability Problem Identified | Medical Evaluation result code
6814506 | 83420733 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE LEVEL OVER MGDL CUSTOMER FEELS OK TO TROUBLESHOOT FOR LOW BLOOD GLUCOSE READING CUSTOMER CURRENT BLOOD GLUCOSE READING WAS MGDL CUSTOMER BLOOD GLUCOSE READING AT THE TIME OF THE INCIDENT WAS MGDL CUSTOMER TREATED THEIR LOW BLOOD GLUCOSE READING WITH FOOD CUSTOMER INFUSION SET WAS CHANGED ON B CUSTOMER WAS USING NOVALOG INSULIN THE RESERVOIR MIRRORS SAME AMOUNT OF INSULIN WITH INSULIN PUMP SCREEN CUSTOMER RECENTLY STARTED USING THE INSULIN PUMP CUSTOMER STATED THE INSULIN PUMP WAS NOT EXPOSED TO MAGNETIC FIELDS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | Insufficient Information | Device problem code
6814506 | 83420733 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED LOW BLOOD GLUCOSE LEVEL OVER MGDL CUSTOMER FEELS OK TO TROUBLESHOOT FOR LOW BLOOD GLUCOSE READING CUSTOMER CURRENT BLOOD GLUCOSE READING WAS MGDL CUSTOMER BLOOD GLUCOSE READING AT THE TIME OF THE INCIDENT WAS MGDL CUSTOMER TREATED THEIR LOW BLOOD GLUCOSE READING WITH FOOD CUSTOMER INFUSION SET WAS CHANGED ON B CUSTOMER WAS USING NOVALOG INSULIN THE RESERVOIR MIRRORS SAME AMOUNT OF INSULIN WITH INSULIN PUMP SCREEN CUSTOMER RECENTLY STARTED USING THE INSULIN PUMP CUSTOMER STATED THE INSULIN PUMP WAS NOT EXPOSED TO MAGNETIC FIELDS INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6813406 | 83422880 | D | IT WAS REPORTED A PATIENT HAS OSTEOLYSIS AND CALCIFICATIONS ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Insufficient Information | Device problem code
6813406 | 83422880 | D | IT WAS REPORTED A PATIENT HAS OSTEOLYSIS AND CALCIFICATIONS ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Osteolysis | Patient problem code
6813406 | 83422880 | D | IT WAS REPORTED A PATIENT HAS OSTEOLYSIS AND CALCIFICATIONS ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Calcification | Patient problem code
6814633 | 83424853 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE TABLETS TO TREAT THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS THE CUSTOMER DID NOT REWIND THE PUMP WHEN CHANGING OUT THE SET AND THE PUMP GAVE HIM UNITS OF INSULIN TROUBLESHOOTING WAS COMPLETED AND THE CUSTOMER WILL MONITOR THE PUMP CUSTOMER ALSO REPORTED THAT THEY WERE EXPERIENCING HIGHS OF TO MGDL AFTER LEAVING THE HOSPITAL THE CUSTOMER EXPERIENCED HIGH SYMPTOMS SUCH AS FEELING TIRED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6814633 | 83424853 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE TABLETS TO TREAT THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS THE CUSTOMER DID NOT REWIND THE PUMP WHEN CHANGING OUT THE SET AND THE PUMP GAVE HIM UNITS OF INSULIN TROUBLESHOOTING WAS COMPLETED AND THE CUSTOMER WILL MONITOR THE PUMP CUSTOMER ALSO REPORTED THAT THEY WERE EXPERIENCING HIGHS OF TO MGDL AFTER LEAVING THE HOSPITAL THE CUSTOMER EXPERIENCED HIGH SYMPTOMS SUCH AS FEELING TIRED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6814611 | 83425044 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6814658 | 83454813 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6815831 | 83457045 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HAVING HIGH BLOOD GLUCOSE WITHIN THE RANGE OF MGDL TO MGDL THEY COULD ONLY BRING THEIR BLOOD GLUCOSE DOWN WITH MANUAL INJECTIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MGDL THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED THE INSULIN PUMPS SETTINGS WERE PROGRAMMED ACCURATELY THEY DID NOT RECEIVE ANY ALARM SINCE THE HIGH BLOOD GLUCOSE BEGAN THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Device Alarm System | Device problem code
6815831 | 83457045 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HAVING HIGH BLOOD GLUCOSE WITHIN THE RANGE OF MGDL TO MGDL THEY COULD ONLY BRING THEIR BLOOD GLUCOSE DOWN WITH MANUAL INJECTIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL DURING THE INCIDENT WAS MGDL THEIR CURRENT BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED THE INSULIN PUMPS SETTINGS WERE PROGRAMMED ACCURATELY THEY DID NOT RECEIVE ANY ALARM SINCE THE HIGH BLOOD GLUCOSE BEGAN THE DEVICE PASSED THE BASIC OCCLUSION TEST THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6815384 | 83458757 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL THEN WENT TO MGDL AND THEN TO MGDL THE CUSTOMER WAS GIVEN FOOD AND INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO REPORTED ISSUES WITH SENSOR GLUCOSE VERSUS BLOOD GLUCOSE DIFFERENCES THE CUSTOMER ALSO REPORTED THAT THEY HAD A LOW INCIDENT THE FRIDAY BEFORE THE LATEST INCIDENT AS WELL AS TWO WEEK PRIOR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6815384 | 83458757 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL THEN WENT TO MGDL AND THEN TO MGDL THE CUSTOMER WAS GIVEN FOOD AND INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED CUSTOMER ALSO REPORTED ISSUES WITH SENSOR GLUCOSE VERSUS BLOOD GLUCOSE DIFFERENCES THE CUSTOMER ALSO REPORTED THAT THEY HAD A LOW INCIDENT THE FRIDAY BEFORE THE LATEST INCIDENT AS WELL AS TWO WEEK PRIOR THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6813325 | 83459572 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD SKIN IRRITATION UNDER HER LEFT BREAST DESCRIBED AS A HEAT RASH THE PATIENT CONSULTED HER PHYSICIAN WHO PRESCRIBED A MEDICATION FOLLOWUP INDICATED THAT THE IRRITATION IMPROVED WITH USE OF THE MEDICATION | Rash | Patient problem code
6813325 | 83459572 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD SKIN IRRITATION UNDER HER LEFT BREAST DESCRIBED AS A HEAT RASH THE PATIENT CONSULTED HER PHYSICIAN WHO PRESCRIBED A MEDICATION FOLLOWUP INDICATED THAT THE IRRITATION IMPROVED WITH USE OF THE MEDICATION | Skin Irritation | Patient problem code
6813325 | 83459572 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD SKIN IRRITATION UNDER HER LEFT BREAST DESCRIBED AS A HEAT RASH THE PATIENT CONSULTED HER PHYSICIAN WHO PRESCRIBED A MEDICATION FOLLOWUP INDICATED THAT THE IRRITATION IMPROVED WITH USE OF THE MEDICATION | Insufficient Information | Device problem code
6813325 | 83459572 | D | A DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD SKIN IRRITATION UNDER HER LEFT BREAST DESCRIBED AS A HEAT RASH THE PATIENT CONSULTED HER PHYSICIAN WHO PRESCRIBED A MEDICATION FOLLOWUP INDICATED THAT THE IRRITATION IMPROVED WITH USE OF THE MEDICATION | Irritation | Patient problem code
6814663 | 83465683 | N | B DATE SENT BATCH UNK ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY RECTAL CARCINOMA IN WHAT TYPE OF PROCEDURE WAS THE DEVICE USED ANTERIOR RESECTION WHO FIRED THE DEVICE AND WHAT IS HISHER EXPERIENCE COLORECTAL FELLOW WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING UNKNOWN TO ME DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE NO AUDIBLE CRUNCH AS USUALLY OCCURS WHEN THE STAPLER FIRES WHAT IS MEANT BY THE DEVICE MISFIRED THE STAPLES DID NOT CLOSE BUT THE BOWEL WAS CUT WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE I BELIEVE THEY WERE FULL DONUTS AS THE STAPLER CUT THE BOWEL BUT DID NOT FIRE THE STAPLES WAS THE WASHER INSPECTED IF SO PLEASE DESCRIBE UNKNOWN DONT REALLY UNDERSTAND THE QUESTION WERE THERE ANY STAPLE FORMATION ISSUES IF SO WHAT WAS THEIR SHAPE THE STAPLES WERE NOT VISUALIZED COULD YOU PLEASE PROVIDE A COPY OF DATEX REPORT THAT IS A PRIVATE AND CONFIDENTIAL DOCUMENT I AM UNABLE TO RELEASE DUE TO OUR CONFIDENTIALITY POLICY WHAT IS THE CURRENT PATIENT STATUS PATIENT HAS PERMANENT COLOSTOMY AS A RESULT OF THE FAILED ANASTAMOSIS | Misfire | Device problem code
6814663 | 83465683 | N | B DATE SENT BATCH UNK ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY RECTAL CARCINOMA IN WHAT TYPE OF PROCEDURE WAS THE DEVICE USED ANTERIOR RESECTION WHO FIRED THE DEVICE AND WHAT IS HISHER EXPERIENCE COLORECTAL FELLOW WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING UNKNOWN TO ME DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE NO AUDIBLE CRUNCH AS USUALLY OCCURS WHEN THE STAPLER FIRES WHAT IS MEANT BY THE DEVICE MISFIRED THE STAPLES DID NOT CLOSE BUT THE BOWEL WAS CUT WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE I BELIEVE THEY WERE FULL DONUTS AS THE STAPLER CUT THE BOWEL BUT DID NOT FIRE THE STAPLES WAS THE WASHER INSPECTED IF SO PLEASE DESCRIBE UNKNOWN DONT REALLY UNDERSTAND THE QUESTION WERE THERE ANY STAPLE FORMATION ISSUES IF SO WHAT WAS THEIR SHAPE THE STAPLES WERE NOT VISUALIZED COULD YOU PLEASE PROVIDE A COPY OF DATEX REPORT THAT IS A PRIVATE AND CONFIDENTIAL DOCUMENT I AM UNABLE TO RELEASE DUE TO OUR CONFIDENTIALITY POLICY WHAT IS THE CURRENT PATIENT STATUS PATIENT HAS PERMANENT COLOSTOMY AS A RESULT OF THE FAILED ANASTAMOSIS | Failure to Form Staple | Device problem code
6814663 | 83465683 | N | B DATE SENT BATCH UNK ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY RECTAL CARCINOMA IN WHAT TYPE OF PROCEDURE WAS THE DEVICE USED ANTERIOR RESECTION WHO FIRED THE DEVICE AND WHAT IS HISHER EXPERIENCE COLORECTAL FELLOW WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING UNKNOWN TO ME DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE NO AUDIBLE CRUNCH AS USUALLY OCCURS WHEN THE STAPLER FIRES WHAT IS MEANT BY THE DEVICE MISFIRED THE STAPLES DID NOT CLOSE BUT THE BOWEL WAS CUT WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE I BELIEVE THEY WERE FULL DONUTS AS THE STAPLER CUT THE BOWEL BUT DID NOT FIRE THE STAPLES WAS THE WASHER INSPECTED IF SO PLEASE DESCRIBE UNKNOWN DONT REALLY UNDERSTAND THE QUESTION WERE THERE ANY STAPLE FORMATION ISSUES IF SO WHAT WAS THEIR SHAPE THE STAPLES WERE NOT VISUALIZED COULD YOU PLEASE PROVIDE A COPY OF DATEX REPORT THAT IS A PRIVATE AND CONFIDENTIAL DOCUMENT I AM UNABLE TO RELEASE DUE TO OUR CONFIDENTIALITY POLICY WHAT IS THE CURRENT PATIENT STATUS PATIENT HAS PERMANENT COLOSTOMY AS A RESULT OF THE FAILED ANASTAMOSIS | Operational Problem Identified | Medical Evaluation result code
6814292 | 83466308 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION COMPLAINANT INDICATED THAT DURING SUBSEQUENT TESTING THE MALFUNCTION WAS NOT DUPLICATED | No Patient Involvement | Patient problem code
6814292 | 83466308 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION COMPLAINANT INDICATED THAT DURING SUBSEQUENT TESTING THE MALFUNCTION WAS NOT DUPLICATED | Defective Device | Device Problem Code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Pain, Abdominal | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Constipation | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Flatus | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Burning Sensation | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Insufficient Information | Device problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Diarrhea | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Pain | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Discomfort | Patient problem code
6816256 | 83468862 | D | THIS CASE WAS INITIALLY RECEIVED VIA REGULATORY AUTHORITY ANSM REFERENCE NUMBER B ON AUG THIS SPONTANEOUS CASE WAS REPORTED BY A CONSUMER AND DESCRIBES THE OCCURRENCE OF ABDOMINAL PAIN ABDOMINAL PAIN IN A B FEMALE PATIENT WHO HAD ESSURE INSERTED THE OCCURRENCE OF ADDITIONAL NONSERIOUS EVENTS IS DETAILED BELOW ON B THE PATIENT HAD ESSURE INSERTED ON B YEAR AFTER INSERTION OF ESSURE THE PATIENT EXPERIENCED ABDOMINAL PAIN SERIOUSNESS CRITERIA MEDICALLY SIGNIFICANT AND INTERVENTION REQUIRED FLATULENCE A LOT OF FLATULENCE DYSMENORRHEA PAIN WORSENED IN RELATION TO MY MENSTRUAL CYCLE AND GASTROINTESTINAL MOTILITY DISORDER INTESTINAL TRANSIT BECAME TOTALLY DISRUPTED WITH ALTERNATING DIARRHEA AND CONSTIPATION IN THE PATIENT EXPERIENCED ABDOMINAL DISCOMFORT BURNING SENSATIONS IN THE LOWER ABDOMEN THE PATIENT WAS TREATED WITH SURGERY SALPINGECTOMY ON B AND ENDOMETRIAL RESECTION IN B AND SURGERY COLONOSCOPY ESSURE WAS REMOVED ON B AT THE TIME OF THE REPORT THE ABDOMINAL PAIN FLATULENCE DYSMENORRHOEA GASTROINTESTINAL MOTILITY DISORDER AND ABDOMINAL DISCOMFORT OUTCOME WAS UNKNOWN THE REPORTER PROVIDED NO CAUSALITY ASSESSMENT FOR ABDOMINAL DISCOMFORT ABDOMINAL PAIN DYSMENORRHOEA FLATULENCE AND GASTROINTESTINAL MOTILITY DISORDER WITH ESSURE DIAGNOSTIC RESULTS NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE BODY WEIGHT WAS REPORTED TO BE B | Menstrual Irregularities | Patient problem code
6814958 | 83476219 | N | A COMPLETE INVESTIGATION WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT CODE LOT NUMBER OR SAMPLE WAS PROVIDED QUESTIONS WERE SENT TO THE AUTHOR OF THE ARTICLE NO RESPONSE WAS RECEIVED THE STUDY CONCLUDED THAT CEVAR IS FEASIBLE AND EFFICIENT TO TREAT JAA IN PATIENTS UNSUITABLE FOR FENESTRATED BRANCHED ENDOGRAFTS FBE HOWEVER IN THIS PRELIMINARY EXPERIENCE COMPLICATIONS OF DEVICES INSERTION AND TYPE I ENDOLEAKS WERE NOT RARE ASSOCIATED FILES | No Device Problem Found | Medical Evaluation result code
6797808 | 83477764 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHILE THE CUSTOMER WAS AT WORK CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6797808 | 83477764 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHILE THE CUSTOMER WAS AT WORK CAUSING THE POD TO FALL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6813083 | 83477967 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6813083 | 83477967 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6812959 | 83482864 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6812959 | 83482864 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6809959 | 83486961 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6808974 | 83502159 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE THE IMPLANT WAS EASILY TURNED OUT WHILE REMOVING THE COVER SCREW AT THE TIME OF UNCOVERING THERE WAS NO HARM CAUSED TO THE PATIENT NO PATIENTS ETHNICITY AND RACE WAS GIVEN THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES PUBLISHED STUDIES SHOW THAT FAILURE TO OSSEOINTEGRATE OCCURS IN TO OF ALL CASES AND IS CONSIDERED AN INHERENT RISK IN IMPLANT SURGERY ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | Loss of Osseointegration | Device problem code
6808974 | 83502159 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE THE IMPLANT WAS EASILY TURNED OUT WHILE REMOVING THE COVER SCREW AT THE TIME OF UNCOVERING THERE WAS NO HARM CAUSED TO THE PATIENT NO PATIENTS ETHNICITY AND RACE WAS GIVEN THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES PUBLISHED STUDIES SHOW THAT FAILURE TO OSSEOINTEGRATE OCCURS IN TO OF ALL CASES AND IS CONSIDERED AN INHERENT RISK IN IMPLANT SURGERY ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | Operational Problem Identified | Medical Evaluation result code
6816957 | 83504869 | N | A COMPLETE LEAD WAS RETURNED IN ONE PIECE VISUAL INSPECTION REVEALED THE GREEN PTFE COATING OF THE STYLET TO BE CLOGGED INSIDE THE INNER COIL IN THE CONNECTOR PIN REGION WITH THE INNER COIL OFFSET THE PTFE COATING OF THE STYLET INSIDE THE INNER COIL IN THE CONNECTOR PIN REGION LIKELY CONTRIBUTED TO THE REPORTED INCIDENT ELECTRICAL TESTING FOUND NO INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS ALL DAMAGES NOTED ARE CONSISTENT WITH THOSE OCCURRING AT THE TIME OF PROCEDURE | Complete Blockage | Device problem code
6816957 | 83504869 | N | A COMPLETE LEAD WAS RETURNED IN ONE PIECE VISUAL INSPECTION REVEALED THE GREEN PTFE COATING OF THE STYLET TO BE CLOGGED INSIDE THE INNER COIL IN THE CONNECTOR PIN REGION WITH THE INNER COIL OFFSET THE PTFE COATING OF THE STYLET INSIDE THE INNER COIL IN THE CONNECTOR PIN REGION LIKELY CONTRIBUTED TO THE REPORTED INCIDENT ELECTRICAL TESTING FOUND NO INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS ALL DAMAGES NOTED ARE CONSISTENT WITH THOSE OCCURRING AT THE TIME OF PROCEDURE | No Device Problem Found | Medical Evaluation result code
6811933 | 83510835 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Scratched Material | Device problem code
6811933 | 83510835 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
6811933 | 83510835 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MISSING RESERVOIR TUBE ORING | No Device Problem Found | Medical Evaluation result code
6817348 | 83516882 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Osteolysis | Patient problem code
6817348 | 83516882 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
6817348 | 83516882 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Toxicity | Patient problem code
6817348 | 83516882 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Calcification | Patient problem code
6817336 | 83517332 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Toxicity | Patient problem code
6817336 | 83517332 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Calcification | Patient problem code
6817336 | 83517332 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
6817336 | 83517332 | D | IT WAS REPORTED THE PATIENT IS EXPERIENCING ELEVATED METAL IONS OSTEOLYSIS AND CALCIFICATIONS FOLLOWING A TOTAL HIP ARTHROPLASTY PATIENT WAS REPORTED TO HAVE COMPLETE HIP PERFORMANCE NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Osteolysis | Patient problem code
6817579 | 83519131 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS REVIEWED AND NO NONCONFORMANCES RELATED TO THE REPORTED EVENT WERE NOTED IT WAS DETERMINED THAT THE IABP DID NOT MALFUNCTION | No Device Problem Found | Medical Evaluation result code
6817598 | 83520145 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED THE FOLLOWING DAY NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6817598 | 83520145 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED THE FOLLOWING DAY NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6817656 | 83529172 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND WAS NOT ABLE TO CONFIRM THE REPORTED COMPLAINT HOWEVER A REVIEW OF THE LOGS INDICATED AN ALARM OCCURRED THAT MAY INDICATE AN LOSS OF MECHANICAL VENTILATION IT WAS RECOMMENDED TO REPLACE THE CENTRAL PROCESSING UNIT IF THE ISSUE REOCCURS | Ventilation Problem in Device Environment | Device Problem Code
6817656 | 83529172 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND WAS NOT ABLE TO CONFIRM THE REPORTED COMPLAINT HOWEVER A REVIEW OF THE LOGS INDICATED AN ALARM OCCURRED THAT MAY INDICATE AN LOSS OF MECHANICAL VENTILATION IT WAS RECOMMENDED TO REPLACE THE CENTRAL PROCESSING UNIT IF THE ISSUE REOCCURS | Non-Functional Defect | Medical Evaluation result code
6812635 | 83529670 | N | ADDITIONAL MANUFACTURING NARRATIVE THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Component or Accessory Incompatibility | Device problem code
6812635 | 83529670 | N | ADDITIONAL MANUFACTURING NARRATIVE THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Device Displays Incorrect Message | Device problem code
6812635 | 83529670 | N | ADDITIONAL MANUFACTURING NARRATIVE THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6817239 | 83529724 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER FOR THE DISCORDANT ELEVATED BHCG RESULT OBTAINED ON THE DIMENSION VISTA INSTRUMENT NABT NONSPECIFIC ANTIBODY BLOCKING TUBE AND HBT HETEROPHILIC ANTIBODY BLOCKING TUBE TUBES WERE SENT TO CUSTOMER BY SIEMENS HEALTHCARE DIAGNOSTICS FOR INTERFERENCE INVESTIGATION THE RESULTS OF A SAMPLE FROM THE PATIENT AFTER TREATMENT BY THE CUSTOMER OF THE SAMPLE WERE AS FOLLOWS BHCG ON NABT MIUML BHCG ON HBT MIUML THE SIEMENS HEALTHCARE DIAGNOSTICS HEADQUARTERS SUPPORT CENTER HSC REPRESENTATIVE EVALUATED THE INFORMATION PROVIDED ON THE INDIVIDUAL PATIENT OVER MULTIPLE TEST DATES AND THE NABT AND HBT TUBE TREATMENT RESULTS BASED UPON THE NABT AND HBT RESULTS AND THE MIUML RESULTS WITH ALTERNATE METHODOLOGIES HSC CONCLUDED THAT THE PATIENT HAS A HETEROPHILIC ANTIBODY THAT IS INTERFERING WITH THE VISTA BHCG TEST THE BHCG FLEX REAGENT CARTRIDGE INSTRUCTIONS FOR USE STATES PATIENT SAMPLES MAY CONTAIN HETEROPHILIC ANTIBODIES THAT COULD REACT IN IMMUNOASSAYS TO GIVE FALSELY ELEVATED OR DEPRESSED RESULTS THIS ASSAY HAS BEEN DESIGNED TO MINIMIZE INTERFERENCE FROM HETEROPHILIC ANTIBODIES NEVERTHELESS COMPLETE ELIMINATION OF THIS INTERFERENCE FROM ALL PATIENT SPECIMENS CANNOT BE GUARANTEED A TEST RESULT THAT IS INCONSISTENT WITH THE CLINICAL PICTURE AND PATIENT HISTORY SHOULD BE INTERPRETED WITH CAUTION THE DEVICE IS PERFORMING WITHIN SPECIFICATIONS NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED MDRS AND WERE FILED FOR THE SAME EVENT | Partial Blockage | Device Problem Code
6817239 | 83529724 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER FOR THE DISCORDANT ELEVATED BHCG RESULT OBTAINED ON THE DIMENSION VISTA INSTRUMENT NABT NONSPECIFIC ANTIBODY BLOCKING TUBE AND HBT HETEROPHILIC ANTIBODY BLOCKING TUBE TUBES WERE SENT TO CUSTOMER BY SIEMENS HEALTHCARE DIAGNOSTICS FOR INTERFERENCE INVESTIGATION THE RESULTS OF A SAMPLE FROM THE PATIENT AFTER TREATMENT BY THE CUSTOMER OF THE SAMPLE WERE AS FOLLOWS BHCG ON NABT MIUML BHCG ON HBT MIUML THE SIEMENS HEALTHCARE DIAGNOSTICS HEADQUARTERS SUPPORT CENTER HSC REPRESENTATIVE EVALUATED THE INFORMATION PROVIDED ON THE INDIVIDUAL PATIENT OVER MULTIPLE TEST DATES AND THE NABT AND HBT TUBE TREATMENT RESULTS BASED UPON THE NABT AND HBT RESULTS AND THE MIUML RESULTS WITH ALTERNATE METHODOLOGIES HSC CONCLUDED THAT THE PATIENT HAS A HETEROPHILIC ANTIBODY THAT IS INTERFERING WITH THE VISTA BHCG TEST THE BHCG FLEX REAGENT CARTRIDGE INSTRUCTIONS FOR USE STATES PATIENT SAMPLES MAY CONTAIN HETEROPHILIC ANTIBODIES THAT COULD REACT IN IMMUNOASSAYS TO GIVE FALSELY ELEVATED OR DEPRESSED RESULTS THIS ASSAY HAS BEEN DESIGNED TO MINIMIZE INTERFERENCE FROM HETEROPHILIC ANTIBODIES NEVERTHELESS COMPLETE ELIMINATION OF THIS INTERFERENCE FROM ALL PATIENT SPECIMENS CANNOT BE GUARANTEED A TEST RESULT THAT IS INCONSISTENT WITH THE CLINICAL PICTURE AND PATIENT HISTORY SHOULD BE INTERPRETED WITH CAUTION THE DEVICE IS PERFORMING WITHIN SPECIFICATIONS NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED MDRS AND WERE FILED FOR THE SAME EVENT | Results Pending Completion of Investigation | Medical Evaluation result code
6817404 | 83530661 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER FOR THE DISCORDANT ELEVATED BHCG RESULTS OBTAINED ON THE DIMENSION VISTA INSTRUMENT NABT NONSPECIFIC ANTIBODY BLOCKING TUBE AND HBT HETEROPHILIC ANTIBODY BLOCKING TUBE TUBES WERE SENT TO CUSTOMER BY SIEMENS HEALTHCARE DIAGNOSTICS FOR INTERFERENCE INVESTIGATION THE RESULTS OF A SAMPLE FROM THE PATIENT AFTER TREATMENT BY THE ACCOUNT OF THE SAMPLE WERE AS FOLLOWS BHCG ON NABT MIUML BHCG ON HBT MIUML THE SIEMENS HEALTHCARE DIAGNOSTICS HEADQUARTERS SUPPORT CENTER HSC REPRESENTATIVE EVALUATED THE INFORMATION PROVIDED ON THE INDIVIDUAL PATIENT OVER MULTIPLE TEST DATES AND THE NABT AND HBT TUBE TREATMENT RESULTS BASED UPON THE NABT AND HBT RESULTS AND THE MIUML RESULTS WITH ALTERNATE METHODOLOGIES HSC CONCLUDED THAT THE PATIENT HAS A HETEROPHILIC ANTIBODY THAT IS INTERFERING WITH THE VISTA BHCG TEST THE BHCG FLEX REAGENT CARTRIDGE INSTRUCTIONS FOR USE STATES PATIENT SAMPLES MAY CONTAIN HETEROPHILIC ANTIBODIES THAT COULD REACT IN IMMUNOASSAYS TO GIVE FALSELY ELEVATED OR DEPRESSED RESULTS THIS ASSAY HAS BEEN DESIGNED TO MINIMIZE INTERFERENCE FROM HETEROPHILIC ANTIBODIES NEVERTHELESS COMPLETE ELIMINATION OF THIS INTERFERENCE FROM ALL PATIENT SPECIMENS CANNOT BE GUARANTEED A TEST RESULT THAT IS INCONSISTENT WITH THE CLINICAL PICTURE AND PATIENT HISTORY SHOULD BE INTERPRETED WITH CAUTION THE DEVICE IS PERFORMING WITHIN SPECIFICATIONS NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED MDRS AND WERE FILED FOR THE SAME EVENT | Partial Blockage | Device Problem Code
6817404 | 83530661 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER FOR THE DISCORDANT ELEVATED BHCG RESULTS OBTAINED ON THE DIMENSION VISTA INSTRUMENT NABT NONSPECIFIC ANTIBODY BLOCKING TUBE AND HBT HETEROPHILIC ANTIBODY BLOCKING TUBE TUBES WERE SENT TO CUSTOMER BY SIEMENS HEALTHCARE DIAGNOSTICS FOR INTERFERENCE INVESTIGATION THE RESULTS OF A SAMPLE FROM THE PATIENT AFTER TREATMENT BY THE ACCOUNT OF THE SAMPLE WERE AS FOLLOWS BHCG ON NABT MIUML BHCG ON HBT MIUML THE SIEMENS HEALTHCARE DIAGNOSTICS HEADQUARTERS SUPPORT CENTER HSC REPRESENTATIVE EVALUATED THE INFORMATION PROVIDED ON THE INDIVIDUAL PATIENT OVER MULTIPLE TEST DATES AND THE NABT AND HBT TUBE TREATMENT RESULTS BASED UPON THE NABT AND HBT RESULTS AND THE MIUML RESULTS WITH ALTERNATE METHODOLOGIES HSC CONCLUDED THAT THE PATIENT HAS A HETEROPHILIC ANTIBODY THAT IS INTERFERING WITH THE VISTA BHCG TEST THE BHCG FLEX REAGENT CARTRIDGE INSTRUCTIONS FOR USE STATES PATIENT SAMPLES MAY CONTAIN HETEROPHILIC ANTIBODIES THAT COULD REACT IN IMMUNOASSAYS TO GIVE FALSELY ELEVATED OR DEPRESSED RESULTS THIS ASSAY HAS BEEN DESIGNED TO MINIMIZE INTERFERENCE FROM HETEROPHILIC ANTIBODIES NEVERTHELESS COMPLETE ELIMINATION OF THIS INTERFERENCE FROM ALL PATIENT SPECIMENS CANNOT BE GUARANTEED A TEST RESULT THAT IS INCONSISTENT WITH THE CLINICAL PICTURE AND PATIENT HISTORY SHOULD BE INTERPRETED WITH CAUTION THE DEVICE IS PERFORMING WITHIN SPECIFICATIONS NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED MDRS AND WERE FILED FOR THE SAME EVENT | Results Pending Completion of Investigation | Medical Evaluation result code
6811269 | 83530974 | N | B AS PER THE BIOMEDICAL ENGINEER MULTIPLE ATTEMPTS WITH DIFFERENT OXYGEN O SENSORS WERE MADE AND THE EPGS COULD NOT BE CALIBRATED A MECHANICAL FLOW METER WAS IN USE AND IT DID NOT MATCH THE CENTRAL CONTROL MONITOR CCM THE SUSPECT DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION | Failure to Calibrate | Device problem code
6811269 | 83530974 | N | B AS PER THE BIOMEDICAL ENGINEER MULTIPLE ATTEMPTS WITH DIFFERENT OXYGEN O SENSORS WERE MADE AND THE EPGS COULD NOT BE CALIBRATED A MECHANICAL FLOW METER WAS IN USE AND IT DID NOT MATCH THE CENTRAL CONTROL MONITOR CCM THE SUSPECT DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION | Failure to Calibrate | Medical Evaluation result code
6817429 | 83530977 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS OUTSIDE WORKING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT FIVE MINUTES AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6817429 | 83530977 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS OUTSIDE WORKING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT FIVE MINUTES AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6817718 | 83531065 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED A RANGE OF TO MGDL WHILE WEARING THE POD PATIENTS BG WENT UP SO HIGH EVEN AFTER BOLUS HAD TO INJECT WITH INSULIN THUS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | Hyperglycemia | Patient Problem Code
6817718 | 83531065 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED A RANGE OF TO MGDL WHILE WEARING THE POD PATIENTS BG WENT UP SO HIGH EVEN AFTER BOLUS HAD TO INJECT WITH INSULIN THUS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | Positioning Problem | Device problem code
6808034 | 83531145 | N | THE IMRIS FIELD SERVICE ENGINEER MET WITH PERSONNEL FROM CLINICAL ENGINEERING TO INSPECT THE OR TABLE THEY CONCLUDED THAT THE BACK SECTION WAS NOT FULLY SEATED AND LOCKED INTO THE TABLE AT THE TIME WHEN THE INCIDENT OCCURRED THEY FULLY TESTED TABLE MOVEMENT AND BACK SECTION LOCKING MECHANISM TO ENSURE PROPER OPERATION ALL TESTING PASSED IMRIS CLINICAL EDUCATION TEAM SHALL COMMUNICATE WITH B OPERATING ROOM OR PERSONNEL TO ENSURE THAT THEY ARE FOLLOWING THE INSTRUCTIONS IN THE IFU AND TO REEDUCATE CLEANING CREW AND CLINICAL STAFF THAT PREPARES THE OPERATING ROOM OR FOR PROPER USAGE OF THE OR TABLE BACK SECTION B | Mechanical Jam | Device problem code
6808034 | 83531145 | N | THE IMRIS FIELD SERVICE ENGINEER MET WITH PERSONNEL FROM CLINICAL ENGINEERING TO INSPECT THE OR TABLE THEY CONCLUDED THAT THE BACK SECTION WAS NOT FULLY SEATED AND LOCKED INTO THE TABLE AT THE TIME WHEN THE INCIDENT OCCURRED THEY FULLY TESTED TABLE MOVEMENT AND BACK SECTION LOCKING MECHANISM TO ENSURE PROPER OPERATION ALL TESTING PASSED IMRIS CLINICAL EDUCATION TEAM SHALL COMMUNICATE WITH B OPERATING ROOM OR PERSONNEL TO ENSURE THAT THEY ARE FOLLOWING THE INSTRUCTIONS IN THE IFU AND TO REEDUCATE CLEANING CREW AND CLINICAL STAFF THAT PREPARES THE OPERATING ROOM OR FOR PROPER USAGE OF THE OR TABLE BACK SECTION B | Activation, Positioning or Separation Problem | Device problem code
6808034 | 83531145 | N | THE IMRIS FIELD SERVICE ENGINEER MET WITH PERSONNEL FROM CLINICAL ENGINEERING TO INSPECT THE OR TABLE THEY CONCLUDED THAT THE BACK SECTION WAS NOT FULLY SEATED AND LOCKED INTO THE TABLE AT THE TIME WHEN THE INCIDENT OCCURRED THEY FULLY TESTED TABLE MOVEMENT AND BACK SECTION LOCKING MECHANISM TO ENSURE PROPER OPERATION ALL TESTING PASSED IMRIS CLINICAL EDUCATION TEAM SHALL COMMUNICATE WITH B OPERATING ROOM OR PERSONNEL TO ENSURE THAT THEY ARE FOLLOWING THE INSTRUCTIONS IN THE IFU AND TO REEDUCATE CLEANING CREW AND CLINICAL STAFF THAT PREPARES THE OPERATING ROOM OR FOR PROPER USAGE OF THE OR TABLE BACK SECTION B | Mechanical Problem Identified | Medical Evaluation result code
6817738 | 83532223 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED A RANGE OF TO MGDL WHILE WEARING THE POD DUE TO BG LEVELS FALLING DOWN FROM AN ELEVATED LEVEL AFTER A POD CHANGE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | Positioning Problem | Device problem code
6817738 | 83532223 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED A RANGE OF TO MGDL WHILE WEARING THE POD DUE TO BG LEVELS FALLING DOWN FROM AN ELEVATED LEVEL AFTER A POD CHANGE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6817729 | 83532379 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD DUE TO BG LEVELS FALLING DOWN FROM AN ELEVATED LEVEL AFTER A POD CHANGE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | Positioning Problem | Device problem code
6817729 | 83532379 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD DUE TO BG LEVELS FALLING DOWN FROM AN ELEVATED LEVEL AFTER A POD CHANGE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK POD IS NO LONGER AVAILABLE AS IT HAS BEEN DISCARDED BY THE PATIENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6812644 | 83534884 | N | THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Device Displays Incorrect Message | Device problem code
6812644 | 83534884 | N | THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Component or Accessory Incompatibility | Device problem code
6812644 | 83534884 | N | THE RETURNED SENSOR WAS EVALUATED THE UNIT PASSED VISUAL INSPECTION AND SOFTWARE TESTING HOWEVER THE DEVICE DID NOT PASS FUNCTIONAL TESTING THE SENSOR WAS FAILED CONTINUITY DUE TO AN OPEN EMITTER ASSEMBLY DURING TESTING THE SENSOR WAS ABLE TO GENERATE A REPLACE SENSOR ERROR MESSAGE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6809642 | 83536266 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT MAUDE REPORT MW ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED IT WAS REPORTED THAT THE SECOND DEVICE UTILIZED WOULD NOT SEAL PROPERLY PLEASE CLARIFY THIS INFORMATION THE INTEGRITY OF THE STAPLE LINE WAS NOT ACHIEVED THE STAPLE LINE WAS NOT FULLY SECURED BECAUSE SEVERAL OF THE STAPLES DID NOT ACTUALLY STAPLE WERE THERE STAPLES MISSING FROM THE STAPLE LINE YES WHAT WAS THE SHAPE OF THE DEPLOYED STAPLES BFORMATION IRREGULAR LEGS STRAIGHT THE SECURE STAPLES WHICH WERE VISUALIZED WERE B SHAPED THE OTHERS WERE NOT FULLY CLOSED DID THE DEVICE STAPLE AND CUT AS INTENDED HOWEVER THE PATIENT TISSUE DID NOT HOLD TOGETHER NOT QUITE THE INTEGRITY OF THE STAPLE LINE WAS NOT MAINTAINED | No Findings Available | Medical Evaluation result code
6817962 | 83538291 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THE EXPERIENCED URGENCY AND LACK OF VOIDING MIGHT BE DUE TO A BLADDER INFECTION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Urinary Tract | Patient problem code
6817962 | 83538291 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THE EXPERIENCED URGENCY AND LACK OF VOIDING MIGHT BE DUE TO A BLADDER INFECTION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6817962 | 83538291 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THE EXPERIENCED URGENCY AND LACK OF VOIDING MIGHT BE DUE TO A BLADDER INFECTION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient problem code
6817962 | 83538291 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THE EXPERIENCED URGENCY AND LACK OF VOIDING MIGHT BE DUE TO A BLADDER INFECTION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Micturation Urgency | Patient problem code
6818713 | 83603812 | N | DEVICE EVALUATION PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SPECIMEN CUP VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND CLEAR RESIDUE ON LENS WORK ORDER SEARCH NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT B | Component Missing | Device problem code
6818713 | 83603812 | N | DEVICE EVALUATION PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SPECIMEN CUP VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND CLEAR RESIDUE ON LENS WORK ORDER SEARCH NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6818675 | 83603864 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6818087 | 83604422 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6814444 | 83615052 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Insufficient Information | Device problem code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Hemorrhage | Patient problem code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Cardiomyopathy | Patient problem code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Ischemia | Patient problem code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Discharge | Patient problem code
6819890 | 83615275 | D | A REPORT WAS RECEIVED VIA A LITERATURE ARTICLE ENTITLED MINIMALLY INVASIVE THORACIC LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION CASE SERIES DEMONSTRATING AN INTEGRATED MULTIDISCIPLINARY STRATEGY THIS ARTICLE DESCRIBES THE AUTHORS INITIAL EXPERIENCE WITH LVAD IMPLANTATION VIA A MINIMALLY INVASIVE SURGICAL APPROACH WITHOUT CARDIOPULMONARY BYPASS CPB THIRTEEN PATIENTS WITH ADVANCED HEART FAILURE IMPLANTED DEEMED TO BE GOOD CANDIDATES FOR THIS APPROACH OVER A MONTH PERIOD WERE INCLUDED IN THIS RETROSPECTIVE OVERVIEW THESE PATIENTS INCLUDED SEVEN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY AND SIX PATIENTS WITH NONISCHEMIC CARDIOMYOPATHY THE GROUP CONSISTED OF TEN MALE PATIENTS AND THREE FEMALE PATIENTS WITH A MEAN AGE OF YEARS THREE OF THE PATIENTS ARE REPORTED TO HAVE HAD PREVIOUS STERNOTOMY ALL PATIENTS RECOVERED AND WERE DISCHARGED FROM THE HOSPITAL THE ARTICLE FURTHER REPORTS THAT INTRAOPERATIVELY ONE PATIENT REQUIRED UNPLANNED CPB TO CONTROL BLEEDING IN ADDITION TWO PATIENTS REQUIRED REINTUBATION WITHIN HOURS AND THREE PATIENTS UNDERWENT CARDIAC TRANSPLANTATION WITHIN SIX MONTHS OF THEIR IMPLANT THE AUTHORS CONCLUDED THAT MINIMALLY INVASIVE OFFPUMP LVAD IMPLANTATION IS AN EMERGING ALTERNATIVE TO PLACEMENT BY MIDLINE STERNOTOMY BASED ON THEIR INITIAL EXPERIENCE AN OFFPUMP THORACIC STRATEGY MAY BE A DESIRABLE OPTION FOR SOME PATIENTS AND THAT CLINICAL OUTCOMES MAY BE NONINFERIOR TO PLACEMENT BY MIDLINE STERNOTOMY WITH CPB NO FURTHER INFORMATION HAS BEEN PROVIDED | Heart Failure | Patient problem code
6818451 | 83630635 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED ELEVATION OF BG LEVELS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6818451 | 83630635 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED ELEVATION OF BG LEVELS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6818351 | 83633891 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TWENTY MINUTES AFTER THE EVENT AND THE PATIENTS BLOOD GLUCOSE LEVELS RETURNED TO NORMAL | Device Dislodged or Dislocated | Device problem code
6818351 | 83633891 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TWENTY MINUTES AFTER THE EVENT AND THE PATIENTS BLOOD GLUCOSE LEVELS RETURNED TO NORMAL | No Known Impact Or Consequence To Patient | Patient Problem Code
6818034 | 83637455 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD WAS HANGING OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT AN HOUR AFTER THE EVENT THE PATIENTS BLOOD GLUCOSE AND INSULIN HISTORY IS AS FOLLOWS TIME PM BGMMOLL BOLUSU PM HIGH POD DEACTIVATED | Device Dislodged or Dislocated | Device problem code
6818034 | 83637455 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD WAS HANGING OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ABOUT AN HOUR AFTER THE EVENT THE PATIENTS BLOOD GLUCOSE AND INSULIN HISTORY IS AS FOLLOWS TIME PM BGMMOLL BOLUSU PM HIGH POD DEACTIVATED | Hyperglycemia | Patient Problem Code
6818075 | 83645912 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6818075 | 83645912 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | Difficult to Insert | Device problem code
6818116 | 83646019 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE CUSTOMER REPORTED THE ADHESIVE FAILED NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED THREE MINUTES AFTER THE EVENT THE CUSTOMERS MOTHER ADMINISTERED CORRECTIONS TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BLOOD GLUCOSE LEVELS DECREASED | Device Dislodged or Dislocated | Device problem code
6818116 | 83646019 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE CUSTOMER REPORTED THE ADHESIVE FAILED NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED THREE MINUTES AFTER THE EVENT THE CUSTOMERS MOTHER ADMINISTERED CORRECTIONS TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BLOOD GLUCOSE LEVELS DECREASED | Hyperglycemia | Patient problem code
6818333 | 83648629 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ONE HOUR AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6818333 | 83648629 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED ONE HOUR AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6818439 | 83649483 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT AND INSULIN WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6818439 | 83649483 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT AND INSULIN WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6820695 | 83650746 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICE DID NOT RESPOND TO THE FRONT PANEL CONTROLS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6820695 | 83650746 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICE DID NOT RESPOND TO THE FRONT PANEL CONTROLS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Insufficient Information | Device Problem Code
6820643 | 83651039 | N | THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC B | Component Missing | Device problem code
6820643 | 83651039 | N | THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6820669 | 83652102 | N | B ECRW NLFF MFG DATE EXP DATE ADDITIONAL INFORMATION REQUESTED AND RECEIVED WAS THE BLEEDING INTRALUMINAL OR INTRAABDOMINAL ACCORDING TO THE PHYSICIAN INTRAABDOMINAL WAS A BLOOD TRANSFUSION PERFORMED DUE TO SYMPTOMS OR LOW HEMOGLOBIN YES WAS THE PATIENT STABILIZED WITH TRANSFUSION ALONE YES WHAT WAS THE PREOPERATIVE HEMOGLOBIN I WAS NOT INFORMED WAS THERE ANY STAPLE FORM ISSUE NOTED DURING INITIAL PROCEDURE NO WHAT SECTION OF STOMACH WAS THE BLUE AND WHITE CARTRIDGES USED UNINFORMED WAS THERE ANY STAPLE LINE INTERVENTION DONE DURING INITIAL SURGERY NO WAS BUTTRESSING MATERIAL USED NO WAS THE DEVICE DIFFICULT TO CLOSE OR FIRE NO DID NOT PRESENT ANY PROBLEMS DURING USE WHAT IS THE CURRENT PATIENT STATUS ACCORDING TO THE DOCTOR THE PATIENT IS WELL | No Device Problem Found | Medical Evaluation result code
6815978 | 83672060 | N | THE PATIENT WAS HOSPITALIZED FOR THE OCCLUSION OF THE SFA THE PATIENT WAS TREATED WITH MEDICATION AND REINTERVENTION WAS REQUIRED THUS RESULTING IN PROLONGED HOSPITALIZATION THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT J EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Findings Available | Medical Evaluation result code
6810310 | 83672950 | N | THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON AUG WHICH PERFORMED AS EXPECTED THE IMMUCOR LABORATORY ALSO TESTED A RETURNED BLOOD SAMPLE FROM THE CUSTOMER SITE ON AUG WHICH YIELDED AN EQUIVOCAL OUTCOME IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY NEGATIVE | Non-Functional Defect | Medical Evaluation result code
6817536 | 83676345 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE FROM A DIFFUSE BACKGROUND | Connection Problem | Device problem code
6817536 | 83676345 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE FROM A DIFFUSE BACKGROUND | Image Artifact | Medical Evaluation result code
6811439 | 83681588 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED NEGATIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE AUG TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND DETERMINED THAT THE INSTRUMENT WAS OPERATING AS EXPECTED | Connection Problem | Device problem code
6811439 | 83681588 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED NEGATIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE AUG TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND DETERMINED THAT THE INSTRUMENT WAS OPERATING AS EXPECTED | Image Artifact | Medical Evaluation result code
6820601 | 83686516 | N | THE PATIENT COMPLAINED OF LEFT LOWER EXTREMITY PAIN WITH SUSPECTED AMBULATION AND CLAUDICATION REVASCULARIZATION OF THE LESION USING ATHERECTOMY PTA AND DCB BALLOON WAS REQUIRED TO TREAT THE PATIENT STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT H EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Device Problem Found | Medical Evaluation result code
6821751 | 83686540 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FELT VERY SLEEPY THEY HAD URGENCY TO VOID AT AM BUT WAS UNABLE TO WHICH WAS VERY BLIZZARD THE NEXT DAY PATIENT SAID THEIR HEAD FELT WEIRD WHEN THEY INCREASED STIMULATION SO THEY TURNED IT BACK DOWN BECAUSE IT WAS MAKING THEIR TINNITUS WORSE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Micturation Urgency | Patient problem code
6821751 | 83686540 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FELT VERY SLEEPY THEY HAD URGENCY TO VOID AT AM BUT WAS UNABLE TO WHICH WAS VERY BLIZZARD THE NEXT DAY PATIENT SAID THEIR HEAD FELT WEIRD WHEN THEY INCREASED STIMULATION SO THEY TURNED IT BACK DOWN BECAUSE IT WAS MAKING THEIR TINNITUS WORSE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6821751 | 83686540 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FELT VERY SLEEPY THEY HAD URGENCY TO VOID AT AM BUT WAS UNABLE TO WHICH WAS VERY BLIZZARD THE NEXT DAY PATIENT SAID THEIR HEAD FELT WEIRD WHEN THEY INCREASED STIMULATION SO THEY TURNED IT BACK DOWN BECAUSE IT WAS MAKING THEIR TINNITUS WORSE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Tinnitus | Patient problem code
6821751 | 83686540 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FELT VERY SLEEPY THEY HAD URGENCY TO VOID AT AM BUT WAS UNABLE TO WHICH WAS VERY BLIZZARD THE NEXT DAY PATIENT SAID THEIR HEAD FELT WEIRD WHEN THEY INCREASED STIMULATION SO THEY TURNED IT BACK DOWN BECAUSE IT WAS MAKING THEIR TINNITUS WORSE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6821807 | 83693044 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF PAIN AND NODULES ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | No Findings Available | Medical Evaluation result code
6822557 | 83725556 | N | B SUTURE LOOPING OCCURS WHEN THE SURGEON INADVERTENTLY LOOPS A SUTURE OVER ONE OF THE COMMISSURAL POSTS THE SUTURE RESTRICTS THE LEAFLET MOTION PROHIBITING COAPTATION RESULTING IN CENTRAL LEAK WHETHER IDENTIFIED INTRAOPERATIVELY OR POSTOPERATIVELY CASES OF SUTURE LOOP WILL TYPICALLY REQUIRE REOPERATION THE IFU WERE REVIEWED AND NO INADEQUACIES WERE IDENTIFIED WITH REGARDS TO WARNINGS CONTRAINDICATIONS AND THE DIRECTIONSCONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL REPORTED EVENTS TRENDS ARE MONITORED ON A MONTHLY BASIS AND IF ACTION IS REQUIRED APPROPRIATE INVESTIGATION WILL BE PERFORMED | No Device Problem Found | Medical Evaluation result code
6822474 | 83726576 | N | STRICTURE OR NARROWING OF THE ESOPHAGUS IS A KNOWN AND ANTICIPATED ADVERSE EVENT AFTER AN ENDOSCOPIC THERAPY PROCEDURE FOR TREATMENT OF BARRETTS ESOPHAGUS THIS STRICTURE IS PROBABLY RELATED TO THE PROCEDURE | Material Deformation | Device problem code
6822474 | 83726576 | N | STRICTURE OR NARROWING OF THE ESOPHAGUS IS A KNOWN AND ANTICIPATED ADVERSE EVENT AFTER AN ENDOSCOPIC THERAPY PROCEDURE FOR TREATMENT OF BARRETTS ESOPHAGUS THIS STRICTURE IS PROBABLY RELATED TO THE PROCEDURE | Deformation Problem | Medical Evaluation result code
6822214 | 83728763 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6822214 | 83728763 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6822960 | 83734201 | N | REVASCULARIZATION OF THE TARGET VESSEL WAS PERFORMED BY PLACING A STENT DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT M EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Apparent Adverse Event | Device problem code
6822960 | 83734201 | N | REVASCULARIZATION OF THE TARGET VESSEL WAS PERFORMED BY PLACING A STENT DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT M EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Findings Available | Medical Evaluation result code
6820037 | 83734981 | D | THE CUSTOMER REPORTED A NEEDLE MECHANISM FAILURE THE BLUE CANNULA WAS VISIBLE AND THE STATUS OF THE PINK SLIDE WAS UNABLE TO BE DETERMINED HER BLOOD GLUCOSE REACHED OVER MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS | No Known Impact Or Consequence To Patient | Patient Problem Code
6820037 | 83734981 | D | THE CUSTOMER REPORTED A NEEDLE MECHANISM FAILURE THE BLUE CANNULA WAS VISIBLE AND THE STATUS OF THE PINK SLIDE WAS UNABLE TO BE DETERMINED HER BLOOD GLUCOSE REACHED OVER MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS | Defective Device | Device problem code
6819982 | 83735239 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED BG LEVELS RISING THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6819982 | 83735239 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED BG LEVELS RISING THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6815357 | 83735453 | N | THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED THAT THE DATE AND TIME WERE PREVIOUSLY SET ON IT IS LIKELY THAT THE CONTROLLER WAS NOT LEFT CONNECTED TO EXTERNAL POWER SOURCE ENOUGH TIME TO CHARGE THE INTERNAL BATTERY THE HVAD CONTROLLERS HAVE AN INTERNAL BATTERY BT WITH THE SOLE PURPOSE TO BACKUP THE DATE AND TIME ADDITIONALLY THE LOG FILES SHOWED A VAD DISCONNECT ALARM LOGGED ON THIS INDICATES DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO THE REAL TIME CLOCK WAS RESET TO ZERO MOST LIKELY BECAUSE CONTROLLER HAD NOT BEEN CONNECTED TO EXTERNAL POWER FOR EXTENDED PERIODS PER THE INSTRUCTIONS FOR USE IFU THERE ARE NO KNOWN CLINICAL FACTORS THAT COULD HAVE CONTRIBUTED TO THIS EVENT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6815357 | 83735453 | N | THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED THAT THE DATE AND TIME WERE PREVIOUSLY SET ON IT IS LIKELY THAT THE CONTROLLER WAS NOT LEFT CONNECTED TO EXTERNAL POWER SOURCE ENOUGH TIME TO CHARGE THE INTERNAL BATTERY THE HVAD CONTROLLERS HAVE AN INTERNAL BATTERY BT WITH THE SOLE PURPOSE TO BACKUP THE DATE AND TIME ADDITIONALLY THE LOG FILES SHOWED A VAD DISCONNECT ALARM LOGGED ON THIS INDICATES DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO THE REAL TIME CLOCK WAS RESET TO ZERO MOST LIKELY BECAUSE CONTROLLER HAD NOT BEEN CONNECTED TO EXTERNAL POWER FOR EXTENDED PERIODS PER THE INSTRUCTIONS FOR USE IFU THERE ARE NO KNOWN CLINICAL FACTORS THAT COULD HAVE CONTRIBUTED TO THIS EVENT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Non-Functional Defect | Medical Evaluation result code
6815369 | 83737540 | N | THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM ON B DUE TO A COMMUNICATION ISSUE BETWEEN THE PUMP MOTOR CONTROLLER AND THE USER INTERFACE CONTROLLER THE CONTROLLER WAS SENT TO THE SUPPLIER WHERE AN INVESTIGATION DETERMINED THAT THE MOST LIKELY CAUSE OF THE FAILURE WAS DUE TO A FAULTY PUMP MOTOR CONTROLLER POSSIBLY DUE TO A STATIC RANDOMACCESS MEMORY SRAM FAILURE PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Findings Available | Medical Evaluation result code
6818756 | 83739877 | N | B PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
6818756 | 83739877 | N | B PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
6820459 | 83742071 | D | PT CAME IN FOR MRI OF THE CSPINE PT WAS IN THE MRI MACHINE FOR ABOUT MINUTES AND STATED EXPERIENCING BURNING AND PAIN BEHIND THE LEFT EYE WHERE THE IMPLANT IS PT WAS TAKEN OUT OF THE MACHINE NO VISUAL DISTURBANCE | Burning Sensation | Patient problem code
6820459 | 83742071 | D | PT CAME IN FOR MRI OF THE CSPINE PT WAS IN THE MRI MACHINE FOR ABOUT MINUTES AND STATED EXPERIENCING BURNING AND PAIN BEHIND THE LEFT EYE WHERE THE IMPLANT IS PT WAS TAKEN OUT OF THE MACHINE NO VISUAL DISTURBANCE | Insufficient Information | Device problem code
6820459 | 83742071 | D | PT CAME IN FOR MRI OF THE CSPINE PT WAS IN THE MRI MACHINE FOR ABOUT MINUTES AND STATED EXPERIENCING BURNING AND PAIN BEHIND THE LEFT EYE WHERE THE IMPLANT IS PT WAS TAKEN OUT OF THE MACHINE NO VISUAL DISTURBANCE | Pain | Patient problem code
6822190 | 83742387 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING KNOT NODULES AND BACTERIAL INFECTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | Material Discolored | Device problem code
6813818 | 83743754 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER CON VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN ETERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6813818 | 83743754 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER CON VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN ETERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Findings Available | Medical Evaluation result code
6822586 | 83744370 | N | REPORTING EXEMPTION NUMBER B REGISTRATION NUMBER B B IS SUBMITTING ON BEHALF OF QMED AB MANUFACTURER ARISG CASE NUMBER B IMPORTERS REPORT NUMBER CAPA THE EVENTS ARE EXPECTED NO CORRECTIVEPREVENTIVE ACTION IS NEEDED MANUFACTURER NARRATIVE LOT NUMBER WAS NOT REPORTED AND TRENDING ON BATCH CANNOT BE PERFORMED PHARMACOVIGILANCE COMMENT THE SERIOUS EVENT OF ISCHEMIA AT IMPLANT SITE AND NONSERIOUS EVENTS OF IMPLANT SITE DISCOLORATION SWELLING BRUISING AND ERYTHEMA WERE CONSIDERED EXPECTED AND POSSIBLY RELATED TO THE TREATMENT POTENTIAL ETIOLOGIES INCLUDES VASCULAR COMPRESSION FROM THE FILLER OR INJECTION OF THE FILLER INTO THE VASCULATURE THE CORRECTIVE TREATMENT WITH HYALURONIDASE MIGHT HAVE CONTRIBUTED TO THE EVENTS OF SWELLING BRUISING AND ERYTHEMA THIS CASE MEETS THE CRITERIA FOR EXPEDITED REPORTING TO REGULATORY AUTHORITIES DUE TO INTERVENTION TO PREVENT A POTENTIAL PERMANENT IMPAIRMENT ADDITIONAL COMMENTS DEVICE NOT AVAILABLE TO MANUFACTURER | Non-Functional Defect | Medical Evaluation result code
6815364 | 83747193 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER CON VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN EXTERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS PER THE INSTRUCTIONS FOR USE IFU THE HVAD CONTROLLER IS A MICROPROCESSOR UNIT THAT CONTROLS AND MANAGES THE HEARTWARE SYSTEM OPERATION IT SENDS POWER AND OPERATING SIGNALS TO THE BLOOD PUMP AND COLLECTS INFORMATION FROM THE PUMP ACCORDING TO THE INSTRUCTIONS FOR USE IFU USERS ARE INSTRUCTED TO INPUT THE CURRENT DATE AND TIME WHEN SETTING UP BOTH THE PRIMARY AND BACKUP CONTROLLERS ACCORDING TO THE IFU CONTROLLERS ARE EXPECTED TO FUNCTION FOR AT LEAST ONE YEAR CONTROLLERS WITH REAL TIME CLOCK ERRORS WOULD NOT LIKELY CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY THE CONTROLLER WILL CONTINUE TO POWER THE PUMP THE REAL TIME CLOCK HAS A DIFFERENT POWER SOURCE FROM THE PUMP PARAMETERS DISPLAY AND ALARMS THIS WILL RESULT IN USER ANNOYANCE WITH NO AFFECT THE FUNCTION OF THE PUMP THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6815364 | 83747193 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER CON VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN EXTERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS PER THE INSTRUCTIONS FOR USE IFU THE HVAD CONTROLLER IS A MICROPROCESSOR UNIT THAT CONTROLS AND MANAGES THE HEARTWARE SYSTEM OPERATION IT SENDS POWER AND OPERATING SIGNALS TO THE BLOOD PUMP AND COLLECTS INFORMATION FROM THE PUMP ACCORDING TO THE INSTRUCTIONS FOR USE IFU USERS ARE INSTRUCTED TO INPUT THE CURRENT DATE AND TIME WHEN SETTING UP BOTH THE PRIMARY AND BACKUP CONTROLLERS ACCORDING TO THE IFU CONTROLLERS ARE EXPECTED TO FUNCTION FOR AT LEAST ONE YEAR CONTROLLERS WITH REAL TIME CLOCK ERRORS WOULD NOT LIKELY CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY THE CONTROLLER WILL CONTINUE TO POWER THE PUMP THE REAL TIME CLOCK HAS A DIFFERENT POWER SOURCE FROM THE PUMP PARAMETERS DISPLAY AND ALARMS THIS WILL RESULT IN USER ANNOYANCE WITH NO AFFECT THE FUNCTION OF THE PUMP THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6815145 | 83747240 | N | B METHOD THE COMPLAINT NEOPUFF WAS RETURNED TO FPH IN B FOR INVESTIGATION WHERE IT WAS VISUALLY INSPECTED RESULTS VISUAL INSPECTION REVEALED THAT THE UPPER AND LOWER MANOMETER NUTS WERE LOOSE AND CAUSING THE MANOMETER TO ROTATE FREELY INSIDE THE HOUSING CONCLUSION IT IS POSSIBLE THAT A PRODUCTION LINE OPERATOR MAY NOT HAVE TIGHTENED THE NUTS PROPERLY AND THAT THEY BECAME LOOSE DURING SHIPPING OR TRANSPORTATION | Detachment of Device or device Component | Device problem code
6815145 | 83747240 | N | B METHOD THE COMPLAINT NEOPUFF WAS RETURNED TO FPH IN B FOR INVESTIGATION WHERE IT WAS VISUALLY INSPECTED RESULTS VISUAL INSPECTION REVEALED THAT THE UPPER AND LOWER MANOMETER NUTS WERE LOOSE AND CAUSING THE MANOMETER TO ROTATE FREELY INSIDE THE HOUSING CONCLUSION IT IS POSSIBLE THAT A PRODUCTION LINE OPERATOR MAY NOT HAVE TIGHTENED THE NUTS PROPERLY AND THAT THEY BECAME LOOSE DURING SHIPPING OR TRANSPORTATION | Non-Functional Defect | Medical Evaluation result code
6817796 | 83747292 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE POD TO FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6817796 | 83747292 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE POD TO FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6817983 | 83747784 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BOTH OF WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | Connection Problem | Device problem code
6817983 | 83747784 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BOTH OF WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | No Findings Available | Medical Evaluation result code
6817986 | 83748071 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | Connection Problem | Device problem code
6817986 | 83748071 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | No Findings Available | Medical Evaluation result code
6817968 | 83749270 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6817968 | 83749270 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6817979 | 83749775 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | Connection Problem | Device problem code
6817979 | 83749775 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WITH SLIGHT RED BLOOD CELL ADHERENCE | No Findings Available | Medical Evaluation result code
6796822 | 83750059 | N | THE INVESTIGATION WAS COMPLETED BY CONDUCTING A THOROUGH EVALUATION OF THE PRODUCT AND THE REPORTED INFORMATION THE ISSUE WAS DETERMINED TO BE A DEFECT IN THE PRODUCT IF THE ERROR WAS TO OCCUR AND NOT BE DETECTED IT COULD RESULT IN SERIOUS INJURY OR SERIOUS ADVERSE HEALTH CONSEQUENCES IF THE CREATININE CLEARANCE IS CALCULATED USING A CREATININE VALUE PAST VALID AGE THE WRONG DOSE OF THE AUC DRUG COULD BE ORDERED FOR EXAMPLE IF THE ORIGINAL CREATININE WAS MOLL AND RESULTED IN A CREATININE CLEARANCE OF MLMIN J THE RESULTING CARBOPLATIN DOSE AT AUC IS MG IF THE CREATININE VALUE IS MOLL AND RESULTED IN A CREATININE CLEARANCE OF MLMIN J THE CARBOPLATIN DOSE AT AUC IS MG A DOSE DIFFERENCE OF THAT MAGNITUDE COULD RESULT IN SERIOUS INJURY IT IS UNLIKELY THAT THE PROBLEM WILL OCCUR IN CLINICAL USE WHEN ADJUSTMENTS ARE MADE TO ANY CHEMOTHERAPY DOSE PARTICULAR ATTENTION IS PAID TO MAKE SURE THE DOSE IS CORRECTLY ADJUSTED MULTIPLE CHECKS WOULD BE MADE BY THE PHYSICIAN THE PHARMACIST AND FINALLY BY THE INFUSION NURSE BEFORE ANY DRUG IS ADMINISTERED IN THIS CASE IT IS VERY LIKELY THE OLD DATE OF THE CREATININE WOULD BE DETECTED | Defective Device | Device problem code
6796822 | 83750059 | N | THE INVESTIGATION WAS COMPLETED BY CONDUCTING A THOROUGH EVALUATION OF THE PRODUCT AND THE REPORTED INFORMATION THE ISSUE WAS DETERMINED TO BE A DEFECT IN THE PRODUCT IF THE ERROR WAS TO OCCUR AND NOT BE DETECTED IT COULD RESULT IN SERIOUS INJURY OR SERIOUS ADVERSE HEALTH CONSEQUENCES IF THE CREATININE CLEARANCE IS CALCULATED USING A CREATININE VALUE PAST VALID AGE THE WRONG DOSE OF THE AUC DRUG COULD BE ORDERED FOR EXAMPLE IF THE ORIGINAL CREATININE WAS MOLL AND RESULTED IN A CREATININE CLEARANCE OF MLMIN J THE RESULTING CARBOPLATIN DOSE AT AUC IS MG IF THE CREATININE VALUE IS MOLL AND RESULTED IN A CREATININE CLEARANCE OF MLMIN J THE CARBOPLATIN DOSE AT AUC IS MG A DOSE DIFFERENCE OF THAT MAGNITUDE COULD RESULT IN SERIOUS INJURY IT IS UNLIKELY THAT THE PROBLEM WILL OCCUR IN CLINICAL USE WHEN ADJUSTMENTS ARE MADE TO ANY CHEMOTHERAPY DOSE PARTICULAR ATTENTION IS PAID TO MAKE SURE THE DOSE IS CORRECTLY ADJUSTED MULTIPLE CHECKS WOULD BE MADE BY THE PHYSICIAN THE PHARMACIST AND FINALLY BY THE INFUSION NURSE BEFORE ANY DRUG IS ADMINISTERED IN THIS CASE IT IS VERY LIKELY THE OLD DATE OF THE CREATININE WOULD BE DETECTED | No Findings Available | Medical Evaluation result code
6812090 | 83750447 | N | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CONTROLLER FAULT ALARM THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED FOUR CONTROLLER FAULT ALARMS LOGGED ON B THE CONTROLLER FAULT ALARMS WERE OF TYPE SYSUICAPSFAIL INDICATIVE OF A LEAKY DIODE ON THE ACTIVE POWER SELECTION APS CIRCUIT THE APS CIRCUIT MANAGES THE SELECTION OF THE ACTIVE POWER SOURCE WHEN THE MAIN INTEGRATED CHIP SENSES THAT THE CURRENT FROM THE APS CIRCUIT IS OUTSIDE THE NORMAL OR EXPECTED RANGE A CONTROLLER FAULT ALARM IS TRIGGERED THE ALARM DOES NOT AFFECT THE ELECTRICAL CONNECTION BETWEEN THE CONTROLLER AND POWER SOURCES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE IN THE APS CIRCUIT HEARTWARE IS INVESTIGATING LEAKY DIODES IN THE APS CIRCUIT PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6812090 | 83750447 | N | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CONTROLLER FAULT ALARM THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED FOUR CONTROLLER FAULT ALARMS LOGGED ON B THE CONTROLLER FAULT ALARMS WERE OF TYPE SYSUICAPSFAIL INDICATIVE OF A LEAKY DIODE ON THE ACTIVE POWER SELECTION APS CIRCUIT THE APS CIRCUIT MANAGES THE SELECTION OF THE ACTIVE POWER SOURCE WHEN THE MAIN INTEGRATED CHIP SENSES THAT THE CURRENT FROM THE APS CIRCUIT IS OUTSIDE THE NORMAL OR EXPECTED RANGE A CONTROLLER FAULT ALARM IS TRIGGERED THE ALARM DOES NOT AFFECT THE ELECTRICAL CONNECTION BETWEEN THE CONTROLLER AND POWER SOURCES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE IN THE APS CIRCUIT HEARTWARE IS INVESTIGATING LEAKY DIODES IN THE APS CIRCUIT PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device problem code
6812090 | 83750447 | N | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CONTROLLER FAULT ALARM THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED FOUR CONTROLLER FAULT ALARMS LOGGED ON B THE CONTROLLER FAULT ALARMS WERE OF TYPE SYSUICAPSFAIL INDICATIVE OF A LEAKY DIODE ON THE ACTIVE POWER SELECTION APS CIRCUIT THE APS CIRCUIT MANAGES THE SELECTION OF THE ACTIVE POWER SOURCE WHEN THE MAIN INTEGRATED CHIP SENSES THAT THE CURRENT FROM THE APS CIRCUIT IS OUTSIDE THE NORMAL OR EXPECTED RANGE A CONTROLLER FAULT ALARM IS TRIGGERED THE ALARM DOES NOT AFFECT THE ELECTRICAL CONNECTION BETWEEN THE CONTROLLER AND POWER SOURCES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE IN THE APS CIRCUIT HEARTWARE IS INVESTIGATING LEAKY DIODES IN THE APS CIRCUIT PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Operational Problem Identified | Medical Evaluation result code
6812829 | 83751182 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Difficult to Insert | Device problem code
6812829 | 83751182 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Failure to Deliver | Device problem code
6812829 | 83751182 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Assembly Problem Identified | Medical Evaluation result code
6823825 | 83766876 | D | A REPORT WAS RECEIVED THAT THE PATIENT WAS EXPERIENCING A BURNING SENSATION AROUND THE POCKET SITE THE PATIENT UNDERWENT AN EXPLANT PROCEDURE | Insufficient Information | Device problem code
6823825 | 83766876 | D | A REPORT WAS RECEIVED THAT THE PATIENT WAS EXPERIENCING A BURNING SENSATION AROUND THE POCKET SITE THE PATIENT UNDERWENT AN EXPLANT PROCEDURE | Burning Sensation | Patient problem code
6823524 | 83767828 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS WORSERETAINING MORE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6823524 | 83767828 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS WORSERETAINING MORE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6823671 | 83770048 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF NODULES PAIN GRANULOMAS SWELLING AND BONY NODULE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | No Findings Available | Medical Evaluation result code
6823591 | 83776096 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS IN A LOT OF PAIN AND WAS HAVING AN ALLERGIC REACTION TO THE ANTIBIOTICS AND GOT A RASH THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED AND NOT RESOLVED AT THE TIME OF THE REPORT TWO DAYS LATER PATIENTS FRIENDS TOLD THEM THEY WERE FLUSHED AND THE PATIENT ASKED IF THAT WAS A SIDE EFFECT OF THE EVALUATION PATIENT WILL FOLLOW UP WITH THEIR HEALTHCARE PROVIDER HCP ON B PATIENT WAS EXPERIENCING A SHORTNESS OF BREATH AND TOLD THEIR HCP AND SALES REP THE HCP DOESNT THINK THE SHORTNESS OF BREATH IS RELATED TO THE DEVICE BUT THEY TURNED THE CONTROLLER OFF THE PATIENT WILL TURN THE CONTROLLER BACK ON WHEN THE HCP THINKS IT IS APPROPRIATE THE NEXT DAY THE PATIENT WAS STILL EXPERIENCING A SHORTNESS OF BREATH NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Hypersensitivity | Patient problem code
6823591 | 83776096 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS IN A LOT OF PAIN AND WAS HAVING AN ALLERGIC REACTION TO THE ANTIBIOTICS AND GOT A RASH THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED AND NOT RESOLVED AT THE TIME OF THE REPORT TWO DAYS LATER PATIENTS FRIENDS TOLD THEM THEY WERE FLUSHED AND THE PATIENT ASKED IF THAT WAS A SIDE EFFECT OF THE EVALUATION PATIENT WILL FOLLOW UP WITH THEIR HEALTHCARE PROVIDER HCP ON B PATIENT WAS EXPERIENCING A SHORTNESS OF BREATH AND TOLD THEIR HCP AND SALES REP THE HCP DOESNT THINK THE SHORTNESS OF BREATH IS RELATED TO THE DEVICE BUT THEY TURNED THE CONTROLLER OFF THE PATIENT WILL TURN THE CONTROLLER BACK ON WHEN THE HCP THINKS IT IS APPROPRIATE THE NEXT DAY THE PATIENT WAS STILL EXPERIENCING A SHORTNESS OF BREATH NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6823591 | 83776096 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS IN A LOT OF PAIN AND WAS HAVING AN ALLERGIC REACTION TO THE ANTIBIOTICS AND GOT A RASH THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED AND NOT RESOLVED AT THE TIME OF THE REPORT TWO DAYS LATER PATIENTS FRIENDS TOLD THEM THEY WERE FLUSHED AND THE PATIENT ASKED IF THAT WAS A SIDE EFFECT OF THE EVALUATION PATIENT WILL FOLLOW UP WITH THEIR HEALTHCARE PROVIDER HCP ON B PATIENT WAS EXPERIENCING A SHORTNESS OF BREATH AND TOLD THEIR HCP AND SALES REP THE HCP DOESNT THINK THE SHORTNESS OF BREATH IS RELATED TO THE DEVICE BUT THEY TURNED THE CONTROLLER OFF THE PATIENT WILL TURN THE CONTROLLER BACK ON WHEN THE HCP THINKS IT IS APPROPRIATE THE NEXT DAY THE PATIENT WAS STILL EXPERIENCING A SHORTNESS OF BREATH NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Dyspnea | Patient problem code
6823591 | 83776096 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS IN A LOT OF PAIN AND WAS HAVING AN ALLERGIC REACTION TO THE ANTIBIOTICS AND GOT A RASH THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED AND NOT RESOLVED AT THE TIME OF THE REPORT TWO DAYS LATER PATIENTS FRIENDS TOLD THEM THEY WERE FLUSHED AND THE PATIENT ASKED IF THAT WAS A SIDE EFFECT OF THE EVALUATION PATIENT WILL FOLLOW UP WITH THEIR HEALTHCARE PROVIDER HCP ON B PATIENT WAS EXPERIENCING A SHORTNESS OF BREATH AND TOLD THEIR HCP AND SALES REP THE HCP DOESNT THINK THE SHORTNESS OF BREATH IS RELATED TO THE DEVICE BUT THEY TURNED THE CONTROLLER OFF THE PATIENT WILL TURN THE CONTROLLER BACK ON WHEN THE HCP THINKS IT IS APPROPRIATE THE NEXT DAY THE PATIENT WAS STILL EXPERIENCING A SHORTNESS OF BREATH NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Rash | Patient problem code
6823591 | 83776096 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS IN A LOT OF PAIN AND WAS HAVING AN ALLERGIC REACTION TO THE ANTIBIOTICS AND GOT A RASH THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED AND NOT RESOLVED AT THE TIME OF THE REPORT TWO DAYS LATER PATIENTS FRIENDS TOLD THEM THEY WERE FLUSHED AND THE PATIENT ASKED IF THAT WAS A SIDE EFFECT OF THE EVALUATION PATIENT WILL FOLLOW UP WITH THEIR HEALTHCARE PROVIDER HCP ON B PATIENT WAS EXPERIENCING A SHORTNESS OF BREATH AND TOLD THEIR HCP AND SALES REP THE HCP DOESNT THINK THE SHORTNESS OF BREATH IS RELATED TO THE DEVICE BUT THEY TURNED THE CONTROLLER OFF THE PATIENT WILL TURN THE CONTROLLER BACK ON WHEN THE HCP THINKS IT IS APPROPRIATE THE NEXT DAY THE PATIENT WAS STILL EXPERIENCING A SHORTNESS OF BREATH NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6823388 | 83786266 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR ENDOVASCULAR TREATMENT OF A CM IN DIAMETER ABDOMINAL AORTIC ANEURYSM THE AORTIC NECK WAS NARROW THROMBOTIC AND REVERSE TAPERED THE AORTIC NECK WAS MM AT THE RENAL ARTERY AND MM IN DIAMETER JUST ABOVE THE ANEURYSM IT WAS REPORTED THAT THE FINAL ANGIOGRAPHY INDICATED NOTABLE STENOSIS IN THE OSTIUM OF THE RIGHT RENAL ARTERY RA YET ALMOST NO DELAY WAS SEEN IN THE BLOOD FLOW INTO THE RA THE PHYSICIAN ELECTED TO IMPLANT THE STENT GRAFT PARTIALLY COVERING THE RIGHT RENAL ARTERY IN ORDER TO PREVENT THE OCCURRENCE OF A PROXIMAL TYPE I ENDOLEAK THE PHYSICIAN WAS SATISFIED WITH THE OUTCOME POST INDEX PROCEDURE HOWEVER A ONE WEEK POSTOPERATIVE CT SCAN SHOWED STENOSIS IN THE RIGHT RENAL ARTERY HAD PROGRESSED SIGNIFICANTLY THEREFORE PTRA PERCUTANEOUS TRANSLUMINAL RENAL ANGIOPLASTY WAS URGENTLY PERFORMED THE PHYSICIAN INSERTED VARIOUS CATHETERS AND GUIDEWIRES VIA THE BRACHIAL ARTERY APPROACH BUT CANNULATION TO THE RA WAS UNSUCCESSFUL AND THUS THE CATHETERS AND GUIDEWIRES USED WERE WITHDRAWN THE RA REMAINS OCCLUDED PER THE PHYSICIAN THE CAUSE OF THE EVENT IS RELATED TO THE PATIENTS ANATOMY NARROW THROMBOTIC AND REVERSE TAPERED AORTIC NECK NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT WILL BE MONITORED BY THE PHYSICIAN | Adverse Event Without Identified Device or Use Problem | Device problem code
6823388 | 83786266 | D | AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR ENDOVASCULAR TREATMENT OF A CM IN DIAMETER ABDOMINAL AORTIC ANEURYSM THE AORTIC NECK WAS NARROW THROMBOTIC AND REVERSE TAPERED THE AORTIC NECK WAS MM AT THE RENAL ARTERY AND MM IN DIAMETER JUST ABOVE THE ANEURYSM IT WAS REPORTED THAT THE FINAL ANGIOGRAPHY INDICATED NOTABLE STENOSIS IN THE OSTIUM OF THE RIGHT RENAL ARTERY RA YET ALMOST NO DELAY WAS SEEN IN THE BLOOD FLOW INTO THE RA THE PHYSICIAN ELECTED TO IMPLANT THE STENT GRAFT PARTIALLY COVERING THE RIGHT RENAL ARTERY IN ORDER TO PREVENT THE OCCURRENCE OF A PROXIMAL TYPE I ENDOLEAK THE PHYSICIAN WAS SATISFIED WITH THE OUTCOME POST INDEX PROCEDURE HOWEVER A ONE WEEK POSTOPERATIVE CT SCAN SHOWED STENOSIS IN THE RIGHT RENAL ARTERY HAD PROGRESSED SIGNIFICANTLY THEREFORE PTRA PERCUTANEOUS TRANSLUMINAL RENAL ANGIOPLASTY WAS URGENTLY PERFORMED THE PHYSICIAN INSERTED VARIOUS CATHETERS AND GUIDEWIRES VIA THE BRACHIAL ARTERY APPROACH BUT CANNULATION TO THE RA WAS UNSUCCESSFUL AND THUS THE CATHETERS AND GUIDEWIRES USED WERE WITHDRAWN THE RA REMAINS OCCLUDED PER THE PHYSICIAN THE CAUSE OF THE EVENT IS RELATED TO THE PATIENTS ANATOMY NARROW THROMBOTIC AND REVERSE TAPERED AORTIC NECK NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT WILL BE MONITORED BY THE PHYSICIAN | No Consequences Or Impact To Patient | Patient problem code
6823733 | 83790276 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6823733 | 83790276 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6823731 | 83790345 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE HIPBUTTOCKS | Device Dislodged or Dislocated | Device problem code
6823731 | 83790345 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6819976 | 83795961 | N | THE DETERMINATION COULD BE MADE THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE WAS BEING USED FOR TREATMENT WHEN THE REPORTED EVENT OCCURRED AND THERE IS A RELATIONSHIP OF THE DEVICE TO THE REPORTED PROBLEM CONFIRMED THE Q OF PM PMCA U OF MC PCBA AND THE CABLE BETWEEN THE FLOW SENSOR AND DA PCBA WERE BURNT | No Findings Available | Medical Evaluation result code
6824802 | 83799791 | N | B SERIAL NUMBER PROVIDED AS B IS MISSING ONE NUMBER | Insufficient Information | Device problem code
6823478 | 83809362 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPECTED TO FEEL STIMULATION ALL THE TIME THEY ALSO REPORTED A LEAD WAS HANGING FLUTTER SORENESS AND THE ANTIBIOTICS WAS CAUSING AN ALLERGIC REACTION THE NEXT DAY THEY WOKE UP RED LIKE A STRAWBERRY AND ITCHY THEY WENT TO THE HOSPITAL AND WAS PRESCRIBED PREDNISONE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Atrial Flutter | Patient problem code
6823478 | 83809362 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPECTED TO FEEL STIMULATION ALL THE TIME THEY ALSO REPORTED A LEAD WAS HANGING FLUTTER SORENESS AND THE ANTIBIOTICS WAS CAUSING AN ALLERGIC REACTION THE NEXT DAY THEY WOKE UP RED LIKE A STRAWBERRY AND ITCHY THEY WENT TO THE HOSPITAL AND WAS PRESCRIBED PREDNISONE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Itching | Patient problem code
6823478 | 83809362 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPECTED TO FEEL STIMULATION ALL THE TIME THEY ALSO REPORTED A LEAD WAS HANGING FLUTTER SORENESS AND THE ANTIBIOTICS WAS CAUSING AN ALLERGIC REACTION THE NEXT DAY THEY WOKE UP RED LIKE A STRAWBERRY AND ITCHY THEY WENT TO THE HOSPITAL AND WAS PRESCRIBED PREDNISONE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6823478 | 83809362 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPECTED TO FEEL STIMULATION ALL THE TIME THEY ALSO REPORTED A LEAD WAS HANGING FLUTTER SORENESS AND THE ANTIBIOTICS WAS CAUSING AN ALLERGIC REACTION THE NEXT DAY THEY WOKE UP RED LIKE A STRAWBERRY AND ITCHY THEY WENT TO THE HOSPITAL AND WAS PRESCRIBED PREDNISONE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6823478 | 83809362 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPECTED TO FEEL STIMULATION ALL THE TIME THEY ALSO REPORTED A LEAD WAS HANGING FLUTTER SORENESS AND THE ANTIBIOTICS WAS CAUSING AN ALLERGIC REACTION THE NEXT DAY THEY WOKE UP RED LIKE A STRAWBERRY AND ITCHY THEY WENT TO THE HOSPITAL AND WAS PRESCRIBED PREDNISONE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Hypersensitivity | Patient problem code
6797492 | 83833510 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6797492 | 83833510 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6825907 | 83837897 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF PALPABLE MASS AND SWELLING ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | No Findings Available | Medical Evaluation result code
6825843 | 83840329 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6825843 | 83840329 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6824362 | 83846034 | N | NATUS TECH SUPPORT ADVICE THE CUSTOMER TO RESEAT THE CONNECTION TO HIGHLOW INTENSITY SWITCH WHICH RESOLVED THE ISSUE WAS RESOLVED ONSITE AND NO FURTHER EVALUATION IS NEEDED | Connection Problem | Device problem code
6824362 | 83846034 | N | NATUS TECH SUPPORT ADVICE THE CUSTOMER TO RESEAT THE CONNECTION TO HIGHLOW INTENSITY SWITCH WHICH RESOLVED THE ISSUE WAS RESOLVED ONSITE AND NO FURTHER EVALUATION IS NEEDED | Results Pending Completion of Investigation | Medical Evaluation result code
6823324 | 83848566 | N | THE RETURNED CABLE WAS EVALUATED DURING EVALUATION THE CABLE PASSED VISUAL ANALYSIS AND SOFTWARE TESTING HOWEVER FAILED FUNCTIONAL ANALYSIS DUE TO OPEN IN CABLE ON THE GREEN AND WHITE CONDUCTORS ALONG THE LENGTH OF THE CABLE A SENSOR OFF PATIENT ERROR MESSAGE DISPLAYED WHICH WOULD HAVE PREVENTED THE UNIT FROM BEING ABLE TO ENGAGE IN MONITORING ACTIVITIES B | Device Displays Incorrect Message | Device problem code
6823324 | 83848566 | N | THE RETURNED CABLE WAS EVALUATED DURING EVALUATION THE CABLE PASSED VISUAL ANALYSIS AND SOFTWARE TESTING HOWEVER FAILED FUNCTIONAL ANALYSIS DUE TO OPEN IN CABLE ON THE GREEN AND WHITE CONDUCTORS ALONG THE LENGTH OF THE CABLE A SENSOR OFF PATIENT ERROR MESSAGE DISPLAYED WHICH WOULD HAVE PREVENTED THE UNIT FROM BEING ABLE TO ENGAGE IN MONITORING ACTIVITIES B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6809810 | 83853509 | N | REPORTABLE MALFUNCTION WITH INOMAX DSIR B WAS CREATED AS COMPLAINT B THE SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO OVERDELIVERY A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE FIELD THE ROOT CAUSE FOR THIS REPORT WAS PROPORTIONAL VALVE FAILURE THE DISTRIBUTOR WAS RECOMMENDED TO REPLACE THE PROPORTIONAL VALVE DUE TO THIS FINDING THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM THIS CASE DID NOT RESULT IN AN ADVERSE EVENTSERIOUS ADVERSE EVENT HOWEVER IT IS BEING SUBMITTED TO REGULATORY AUTHORITIES BECAUSE A SIMILAR FAILURE OCCURRED IN THE PAST WHICH RESULTED IN A SERIOUS ADVERSE EVENT MDR | Failure to Deliver | Device problem code
6809810 | 83853509 | N | REPORTABLE MALFUNCTION WITH INOMAX DSIR B WAS CREATED AS COMPLAINT B THE SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO OVERDELIVERY A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE FIELD THE ROOT CAUSE FOR THIS REPORT WAS PROPORTIONAL VALVE FAILURE THE DISTRIBUTOR WAS RECOMMENDED TO REPLACE THE PROPORTIONAL VALVE DUE TO THIS FINDING THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM THIS CASE DID NOT RESULT IN AN ADVERSE EVENTSERIOUS ADVERSE EVENT HOWEVER IT IS BEING SUBMITTED TO REGULATORY AUTHORITIES BECAUSE A SIMILAR FAILURE OCCURRED IN THE PAST WHICH RESULTED IN A SERIOUS ADVERSE EVENT MDR | Operational Problem Identified | Medical Evaluation result code
6826633 | 83858989 | D | COMPLAINANT ALLEGED THAT THE ASSOCIATED DEFIBRILLATOR WAS UNABLE TO DETECT THE ATTACHED ELECTRODES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6826633 | 83858989 | D | COMPLAINANT ALLEGED THAT THE ASSOCIATED DEFIBRILLATOR WAS UNABLE TO DETECT THE ATTACHED ELECTRODES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Defibrillation Problem | Device Problem Code
6820548 | 83862984 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE ANTIBODY SCREEN TEST WELL IMAGES IN QUESTION WHICH WERE VISUALLY NEGATIVE WITH A SLIGHT RED BLOOD CELL BACKGROUND THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON AUG WHICH PERFORMED AS EXPECTED B | Insufficient Information | Device problem code
6825923 | 83863266 | N | WE RECEIVED ONE SINGLE DPT KIT WITH AN IV SET AND PRESSURE TUBING FOR EXAMINATION THE REPORTED EVENT OF UNKNOWN MATERIAL WAS FOUND FLOATING INSIDE THE IV TUBING WAS CONFIRMED ONE UNKNOWN CLEAR MATERIAL WAS FOUND FLOATING WITHIN THE PRIMING SOLUTION INSIDE OF THE IV TUBE NEAR THE IV TUBING MALE CONNECTOR THE MATERIAL WAS APPROXIMATELY XMM IN SIZE THE MATERIAL DID NOT GET FLUSHED OUT OF THE KIT DURING MINUTES OF CONTINUOUS FLUSHING THE MATERIAL MOVED FROM THE IV TUBING TO THE DPT HOUSING AND GOT STUCK AT THE IV SIDE OF THE SNAP TAB IN THE FLUID PATH THE MATERIAL WAS REMOVED FROM THE DPT AND SENT TO CHEMISTRY FOR FURTHER ANALYSIS A SUPPLEMENTAL REPORT WILL BE FORTHCOMING WITH THE CHEMISTRY AND DEVICE HISTORY RESULTS ONCE RECEIVED | No Findings Available | Medical Evaluation result code
6827016 | 83872675 | N | THE CUSTOMER REQUESTED ASSISTANCE FROM GE HEALTHCARE TECHNICAL SUPPORT FOR THE REPORTED ISSUE THE CUSTOMER WAS ADVISED THAT A PRESSURE DIFFERENCE BETWEEN MANIFOLD AND PATIENT AIRWAY PRESSURE OF CM OR MORE WILL CAUSE THE ERROR AND SWITCH UNIT TO VOLUME MODE THE UNIT SWITCHED FROM PSV PRESSURE SUPPORT VENTILATION MODE TO VOLUME MODE AND CONTINUED TO VENTILATE THE CHANGE IN VENTILATION MODE IS NORMAL IN THE SITUATION THAT OCCURRED THE SYSTEM WAS REBOOTED WHICH CLEARED THE REPORTED ISSUE NO REPAIRS OR PARTS WERE NEEDED AND THE SYSTEM WAS RETURNED TO SERVICE THE CUSTOMER REQUESTED ASSISTANCE FROM GE HEALTHCARE TECHNICAL SUPPORT FOR THE REPORTED ISSUE THE CUSTOMER WAS ADVISED THAT A PRESSURE DIFFERENCE BETWEEN MANIFOLD AND PATIENT AIRWAY PRESSURE OF CM OR MORE WILL CAUSE THE ERROR AND SWITCH UNIT TO VOLUME MODE THE UNIT SWITCHED FROM PSV PRESSURE SUPPORT VENTILATION MODE TO VOLUME MODE AND CONTINUED TO VENTILATE THE CHANGE IN VENTILATION MODE IS NORMAL IN THE SITUATION THAT OCCURRED THE SYSTEM WAS REBOOTED WHICH CLEARED THE REPORTED ISSUE NO REPAIRS OR PARTS WERE NEEDED AND THE SYSTEM WAS RETURNED TO SERVICE | Computer Software Problem | Device problem code
6827016 | 83872675 | N | THE CUSTOMER REQUESTED ASSISTANCE FROM GE HEALTHCARE TECHNICAL SUPPORT FOR THE REPORTED ISSUE THE CUSTOMER WAS ADVISED THAT A PRESSURE DIFFERENCE BETWEEN MANIFOLD AND PATIENT AIRWAY PRESSURE OF CM OR MORE WILL CAUSE THE ERROR AND SWITCH UNIT TO VOLUME MODE THE UNIT SWITCHED FROM PSV PRESSURE SUPPORT VENTILATION MODE TO VOLUME MODE AND CONTINUED TO VENTILATE THE CHANGE IN VENTILATION MODE IS NORMAL IN THE SITUATION THAT OCCURRED THE SYSTEM WAS REBOOTED WHICH CLEARED THE REPORTED ISSUE NO REPAIRS OR PARTS WERE NEEDED AND THE SYSTEM WAS RETURNED TO SERVICE THE CUSTOMER REQUESTED ASSISTANCE FROM GE HEALTHCARE TECHNICAL SUPPORT FOR THE REPORTED ISSUE THE CUSTOMER WAS ADVISED THAT A PRESSURE DIFFERENCE BETWEEN MANIFOLD AND PATIENT AIRWAY PRESSURE OF CM OR MORE WILL CAUSE THE ERROR AND SWITCH UNIT TO VOLUME MODE THE UNIT SWITCHED FROM PSV PRESSURE SUPPORT VENTILATION MODE TO VOLUME MODE AND CONTINUED TO VENTILATE THE CHANGE IN VENTILATION MODE IS NORMAL IN THE SITUATION THAT OCCURRED THE SYSTEM WAS REBOOTED WHICH CLEARED THE REPORTED ISSUE NO REPAIRS OR PARTS WERE NEEDED AND THE SYSTEM WAS RETURNED TO SERVICE | Software Problem Identified | Medical Evaluation result code
6827099 | 83873087 | N | THE DENTIST REPORTED AN YEARS OLD PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AND EXPERIENCED INFECTION AT THE IMPLANT SITE AND DEVELOPED FISTULA HOWEVER NO HARM CAUSED TO THE PATIENT NO PATIENTS ETHNICITY AND RACE WERE PROVIDED THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES PUBLISHED STUDIES SHOW THAT FAILURE TO OSSEOINTEGRATE OCCURS IN TO OF ALL CASES AND IS CONSIDERED AN INHERENT RISK IN IMPLANT SURGERY ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR BONE QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | No Findings Available | Medical Evaluation result code
6821841 | 83873168 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND VERIFIED THE OCCURRENCE OF ERRORS TO FIX THE ISSUE THE FSE RELOADED THE SOFTWARE AND PERFORMED COMPLETE CALIBRATION AND FULL FUNCTION TEST TO FACTORY SPECIFICATIONS THE IABP WAS PUT TO RUN ON TRAINER FOR A SIGNIFICANT PERIOD OF TIME AND NO FURTHER ERRORS OCCURRED THE IABP WAS RETURNED TO CLINICAL USE | No Findings Available | Medical Evaluation result code
6819915 | 83880581 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT IT WAS REPORTED THAT IN THE FIRST SHOT THE DEVICE HAS A SMALL FAULT BUT IT FINALLY PERFORMED THE SHOT DOES THIS MEAN THAT THE DEVICE WAS DIFFICULT TO FIRE IF NO PLEASE EXPLAIN IT WAS REPORTED THAT ON THE SECOND FIRING THE DEVICE CUT BUT DID NOT STAPLE AND AND BOUNCING ALL THE HOOKS PLEASE CLARIFY WHAT IS MEANT BY AND BOUNCING ALL THE HOOKS WERE THERE ANY PATIENT CONSEQUENCES IF YES PLEASE DESCRIBE | Failure to Form Staple | Device problem code
6819915 | 83880581 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT IT WAS REPORTED THAT IN THE FIRST SHOT THE DEVICE HAS A SMALL FAULT BUT IT FINALLY PERFORMED THE SHOT DOES THIS MEAN THAT THE DEVICE WAS DIFFICULT TO FIRE IF NO PLEASE EXPLAIN IT WAS REPORTED THAT ON THE SECOND FIRING THE DEVICE CUT BUT DID NOT STAPLE AND AND BOUNCING ALL THE HOOKS PLEASE CLARIFY WHAT IS MEANT BY AND BOUNCING ALL THE HOOKS WERE THERE ANY PATIENT CONSEQUENCES IF YES PLEASE DESCRIBE | Complement Activation Problem | Medical Evaluation result code
6820565 | 83883765 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE ANTIBODY SCREEN TEST WELL IMAGES IN QUESTION WHICH WERE VISUALLY POSITIVE WITH A SLIGHT RED BLOOD CELL BACKGROUND AND FLARED EDGES THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON AUG WHICH PERFORMED AS EXPECTED THE FULL AND COMPLETE CUSTOMER TELEPHONE EXTENSION NUMBER IS B | Connection Problem | Device problem code
6820565 | 83883765 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE ANTIBODY SCREEN TEST WELL IMAGES IN QUESTION WHICH WERE VISUALLY POSITIVE WITH A SLIGHT RED BLOOD CELL BACKGROUND AND FLARED EDGES THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON AUG WHICH PERFORMED AS EXPECTED THE FULL AND COMPLETE CUSTOMER TELEPHONE EXTENSION NUMBER IS B | No Device Problem Found | Medical Evaluation result code
6826986 | 83886727 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER FELT PAIN AND NOTED THE ADHESIVE LET GO NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE MANUAL INSULIN INJECTIONS WERE ADMINISTERED TO TREAT THE HYPERGLYCEMIA AND A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6826986 | 83886727 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER FELT PAIN AND NOTED THE ADHESIVE LET GO NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE MANUAL INSULIN INJECTIONS WERE ADMINISTERED TO TREAT THE HYPERGLYCEMIA AND A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Hyperglycemia | Patient problem code
6826986 | 83886727 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER FELT PAIN AND NOTED THE ADHESIVE LET GO NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE MANUAL INSULIN INJECTIONS WERE ADMINISTERED TO TREAT THE HYPERGLYCEMIA AND A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Pain | Patient problem code
6827782 | 83893527 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6827782 | 83893527 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6823454 | 83898897 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS AFTER SEEING CANNULA THROUGH DEVICES VIEWING WINDOW PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6823454 | 83898897 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS AFTER SEEING CANNULA THROUGH DEVICES VIEWING WINDOW PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6824503 | 83900487 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT TREATED ELEVATED BG LEVELS BY PUTTING ON A NEW POD | Insufficient Information | Device problem code
6824503 | 83900487 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT TREATED ELEVATED BG LEVELS BY PUTTING ON A NEW POD | Hyperglycemia | Patient Problem Code
6828086 | 83903124 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A GYNECOLOGICAL SURGICAL PROCEDURE ON B AND UNKNOWN MESH WAS IMPLANTED IT WAS REPORTED THAT SHE EXPERIENCED UNDISCLOSED INJURIES NO ADDITIONAL INFORMATION WAS PROVIDED | No Consequences Or Impact To Patient | Patient problem code
6828086 | 83903124 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A GYNECOLOGICAL SURGICAL PROCEDURE ON B AND UNKNOWN MESH WAS IMPLANTED IT WAS REPORTED THAT SHE EXPERIENCED UNDISCLOSED INJURIES NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
6827934 | 83908470 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF BRUISING SWELLING AND POSTINFLAMMATORY HYPERPIGMENTATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF BRUISE INFLAMMATION SKIN DISCOLORATION AND EDEMA UNDESIRABLE EFFECTS THE PATIENTS MUST BE INFORMED THAT THERE ARE POTENTIAL SIDE EFFECTS ASSOCIATED LINKED TO THIS PROCEDURE WHICH CAN BE EITHER IMMEDIATE OR DELAYED THESE INCLUDE BUT ARE NOT LIMITED TO INFLAMMATORY REACTIONS REDNESS OEDEMA ERYTHEMA ETC WHICH MAY BE ASSOCIATED WITH ITCHING OR PAIN ON PRESSURE OR BOTH OCCURRING AFTER THE INJECTION THESE REACTIONS MAY LAST FOR A WEEK HAEMATOMAS STAINING OR DISCOLOURATION OF THE INJECTION SITE | No Findings Available | Medical Evaluation result code
6828228 | 83911618 | N | A FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO CUSTOMER SITE THE FSE REPORTED THE OIL RETURN VALVE HOSE CLAMP WAS LOOSE AND TIGHTENING THE VALVE RESOLVED THE SMOKEHAZEMIST ISSUE UNIT MEETS SPECIFICATIONS AND WAS RETURNED TO SERVICE B | Loose or Intermittent Connection | Device problem code
6828228 | 83911618 | N | A FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO CUSTOMER SITE THE FSE REPORTED THE OIL RETURN VALVE HOSE CLAMP WAS LOOSE AND TIGHTENING THE VALVE RESOLVED THE SMOKEHAZEMIST ISSUE UNIT MEETS SPECIFICATIONS AND WAS RETURNED TO SERVICE B | No Device Problem Found | Medical Evaluation result code
6824744 | 83913805 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6824744 | 83913805 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6829534 | 83933556 | N | DEVICE EVALUATION THE DEVICE WAS RETURNED FOR ANALYSIS AND TESTED OUT OF SPECIFICATION ANALYSIS FOUND THAT THE CYLINDERS HAD HOLES IN THE OUTER LAYER THAT WERE THE RESULT OF A SHARP INSTRUMENT BOTH CYLINDERS WERE FUNCTIONAL | Material Puncture / Hole | Device problem code
6829534 | 83933556 | N | DEVICE EVALUATION THE DEVICE WAS RETURNED FOR ANALYSIS AND TESTED OUT OF SPECIFICATION ANALYSIS FOUND THAT THE CYLINDERS HAD HOLES IN THE OUTER LAYER THAT WERE THE RESULT OF A SHARP INSTRUMENT BOTH CYLINDERS WERE FUNCTIONAL | No Device Problem Found | Medical Evaluation result code
6828864 | 83934623 | D | THIS INFORMATION WAS FOUND IN JOURNAL ARTICLE WASSERMANN B ET AL SURGICAL INFRARENAL NEONECK TECHNIQUE DURING ELECTIVE CONVERSION AFTER EVAR WITH SUPRARENAL FIXATION EUROPEAN SOCIETY FOR VASCULAR SURGERY VOL P THIS IS A RETROSPECTIVE REVIEW OF ALL CONSECUTIVE PATIENTS ADMITTED TO THE CLINIC OF VASCULAR AND ENDOVASCULAR SURGERY OF PADOVA UNIVERSITY WHO HAD UNDERGONE ENDOVASCULAR ANEURYSM REPAIR EVAR FOR INFRARENAL ABDOMINAL AORTIC ANEURYSMS AAA BETWEEN JANUARY AND JANUARY ONLY PATIENTS TREATED FOR LATE ELECTIVE OPEN CONVERSION OF SUPRARENAL FIXED EVAR WERE INCLUDED OF THE NINE PATIENTS INCLUDED IN THE STUDY FOUR OF THE STENTS WERE COOK ZENITH DEVICES THE MEAN AGE WAS YEARS WITH EIGHT MEN AND ONE WOMAN ONE PATIENT WITH A ZENITH FLEX STENT GRAFT HAD A TYPE III ENDOLEAK AND HAD ALREADY UNDERGONE AN UNSUCCESSFUL ENDOVASCULAR ATTEMPT AT ILIAC RELINING THE ZENITH FLEX STENT HAD BEEN IMPLANTED FOR MONTHS AT THE TIME OF CONVERSION SURGERY FOLLOWUP WAS PERFORMED AT MONTHS POSTOPERATIVELY THE CT ANGIOGRAM DEMONSTRATED THAT THE PROXIMAL ANASTOMOSIS WAS INTACT WITH NO DEGENERATION OF THE RESIDUAL INFRARENAL AORTIC NECK AND NO SIGNS OF ANASTOMOTIC PSEUDOANEURYSM PER THE ARTICLE NO PERI OR POSTOPERATIVE DEATHS WERE REGISTERED THERE WERE NO OBSERVED CASES OF MAJOR CARDIAC RENAL OR PULMONARY COMPLICATIONS THIS REPORT IS RELATED TO MFR REPORT NUMBERS AND ONE OF THE FOUR COOK ZENITH DEVICES LISTED IN THIS ARTICLE IS NOT SOLD IN THE UNITED STATES | Pseudoaneurysm | Patient problem code
6828864 | 83934623 | D | THIS INFORMATION WAS FOUND IN JOURNAL ARTICLE WASSERMANN B ET AL SURGICAL INFRARENAL NEONECK TECHNIQUE DURING ELECTIVE CONVERSION AFTER EVAR WITH SUPRARENAL FIXATION EUROPEAN SOCIETY FOR VASCULAR SURGERY VOL P THIS IS A RETROSPECTIVE REVIEW OF ALL CONSECUTIVE PATIENTS ADMITTED TO THE CLINIC OF VASCULAR AND ENDOVASCULAR SURGERY OF PADOVA UNIVERSITY WHO HAD UNDERGONE ENDOVASCULAR ANEURYSM REPAIR EVAR FOR INFRARENAL ABDOMINAL AORTIC ANEURYSMS AAA BETWEEN JANUARY AND JANUARY ONLY PATIENTS TREATED FOR LATE ELECTIVE OPEN CONVERSION OF SUPRARENAL FIXED EVAR WERE INCLUDED OF THE NINE PATIENTS INCLUDED IN THE STUDY FOUR OF THE STENTS WERE COOK ZENITH DEVICES THE MEAN AGE WAS YEARS WITH EIGHT MEN AND ONE WOMAN ONE PATIENT WITH A ZENITH FLEX STENT GRAFT HAD A TYPE III ENDOLEAK AND HAD ALREADY UNDERGONE AN UNSUCCESSFUL ENDOVASCULAR ATTEMPT AT ILIAC RELINING THE ZENITH FLEX STENT HAD BEEN IMPLANTED FOR MONTHS AT THE TIME OF CONVERSION SURGERY FOLLOWUP WAS PERFORMED AT MONTHS POSTOPERATIVELY THE CT ANGIOGRAM DEMONSTRATED THAT THE PROXIMAL ANASTOMOSIS WAS INTACT WITH NO DEGENERATION OF THE RESIDUAL INFRARENAL AORTIC NECK AND NO SIGNS OF ANASTOMOTIC PSEUDOANEURYSM PER THE ARTICLE NO PERI OR POSTOPERATIVE DEATHS WERE REGISTERED THERE WERE NO OBSERVED CASES OF MAJOR CARDIAC RENAL OR PULMONARY COMPLICATIONS THIS REPORT IS RELATED TO MFR REPORT NUMBERS AND ONE OF THE FOUR COOK ZENITH DEVICES LISTED IN THIS ARTICLE IS NOT SOLD IN THE UNITED STATES | Insufficient Information | Device problem code
6828864 | 83934623 | D | THIS INFORMATION WAS FOUND IN JOURNAL ARTICLE WASSERMANN B ET AL SURGICAL INFRARENAL NEONECK TECHNIQUE DURING ELECTIVE CONVERSION AFTER EVAR WITH SUPRARENAL FIXATION EUROPEAN SOCIETY FOR VASCULAR SURGERY VOL P THIS IS A RETROSPECTIVE REVIEW OF ALL CONSECUTIVE PATIENTS ADMITTED TO THE CLINIC OF VASCULAR AND ENDOVASCULAR SURGERY OF PADOVA UNIVERSITY WHO HAD UNDERGONE ENDOVASCULAR ANEURYSM REPAIR EVAR FOR INFRARENAL ABDOMINAL AORTIC ANEURYSMS AAA BETWEEN JANUARY AND JANUARY ONLY PATIENTS TREATED FOR LATE ELECTIVE OPEN CONVERSION OF SUPRARENAL FIXED EVAR WERE INCLUDED OF THE NINE PATIENTS INCLUDED IN THE STUDY FOUR OF THE STENTS WERE COOK ZENITH DEVICES THE MEAN AGE WAS YEARS WITH EIGHT MEN AND ONE WOMAN ONE PATIENT WITH A ZENITH FLEX STENT GRAFT HAD A TYPE III ENDOLEAK AND HAD ALREADY UNDERGONE AN UNSUCCESSFUL ENDOVASCULAR ATTEMPT AT ILIAC RELINING THE ZENITH FLEX STENT HAD BEEN IMPLANTED FOR MONTHS AT THE TIME OF CONVERSION SURGERY FOLLOWUP WAS PERFORMED AT MONTHS POSTOPERATIVELY THE CT ANGIOGRAM DEMONSTRATED THAT THE PROXIMAL ANASTOMOSIS WAS INTACT WITH NO DEGENERATION OF THE RESIDUAL INFRARENAL AORTIC NECK AND NO SIGNS OF ANASTOMOTIC PSEUDOANEURYSM PER THE ARTICLE NO PERI OR POSTOPERATIVE DEATHS WERE REGISTERED THERE WERE NO OBSERVED CASES OF MAJOR CARDIAC RENAL OR PULMONARY COMPLICATIONS THIS REPORT IS RELATED TO MFR REPORT NUMBERS AND ONE OF THE FOUR COOK ZENITH DEVICES LISTED IN THIS ARTICLE IS NOT SOLD IN THE UNITED STATES | Aneurysm | Patient problem code
6828864 | 83934623 | D | THIS INFORMATION WAS FOUND IN JOURNAL ARTICLE WASSERMANN B ET AL SURGICAL INFRARENAL NEONECK TECHNIQUE DURING ELECTIVE CONVERSION AFTER EVAR WITH SUPRARENAL FIXATION EUROPEAN SOCIETY FOR VASCULAR SURGERY VOL P THIS IS A RETROSPECTIVE REVIEW OF ALL CONSECUTIVE PATIENTS ADMITTED TO THE CLINIC OF VASCULAR AND ENDOVASCULAR SURGERY OF PADOVA UNIVERSITY WHO HAD UNDERGONE ENDOVASCULAR ANEURYSM REPAIR EVAR FOR INFRARENAL ABDOMINAL AORTIC ANEURYSMS AAA BETWEEN JANUARY AND JANUARY ONLY PATIENTS TREATED FOR LATE ELECTIVE OPEN CONVERSION OF SUPRARENAL FIXED EVAR WERE INCLUDED OF THE NINE PATIENTS INCLUDED IN THE STUDY FOUR OF THE STENTS WERE COOK ZENITH DEVICES THE MEAN AGE WAS YEARS WITH EIGHT MEN AND ONE WOMAN ONE PATIENT WITH A ZENITH FLEX STENT GRAFT HAD A TYPE III ENDOLEAK AND HAD ALREADY UNDERGONE AN UNSUCCESSFUL ENDOVASCULAR ATTEMPT AT ILIAC RELINING THE ZENITH FLEX STENT HAD BEEN IMPLANTED FOR MONTHS AT THE TIME OF CONVERSION SURGERY FOLLOWUP WAS PERFORMED AT MONTHS POSTOPERATIVELY THE CT ANGIOGRAM DEMONSTRATED THAT THE PROXIMAL ANASTOMOSIS WAS INTACT WITH NO DEGENERATION OF THE RESIDUAL INFRARENAL AORTIC NECK AND NO SIGNS OF ANASTOMOTIC PSEUDOANEURYSM PER THE ARTICLE NO PERI OR POSTOPERATIVE DEATHS WERE REGISTERED THERE WERE NO OBSERVED CASES OF MAJOR CARDIAC RENAL OR PULMONARY COMPLICATIONS THIS REPORT IS RELATED TO MFR REPORT NUMBERS AND ONE OF THE FOUR COOK ZENITH DEVICES LISTED IN THIS ARTICLE IS NOT SOLD IN THE UNITED STATES | Hemorrhage | Patient problem code
6828255 | 83935349 | D | IT WAS REPORTED THAT THE SAFETY WAS DIFFICULT TO ENGAGE ON A G X IN NEEDLE X IN TUBING BD SAFETYLOK BLOOD COLLECTION AND INFUSION SET WITH LUER ADAPTER CAUSING A NEEDLE STICK INJURY TO NURSE MEDICAL INTERVENTION | Injury | Patient problem code
6828255 | 83935349 | D | IT WAS REPORTED THAT THE SAFETY WAS DIFFICULT TO ENGAGE ON A G X IN NEEDLE X IN TUBING BD SAFETYLOK BLOOD COLLECTION AND INFUSION SET WITH LUER ADAPTER CAUSING A NEEDLE STICK INJURY TO NURSE MEDICAL INTERVENTION | Difficult to Insert | Device problem code
6828902 | 83937374 | D | IT WAS REPORTED THAT THE SAFETY MECHANISM OF THE G X IN BD ECLIPSE HYPODERMIC NEEDLE FAILED TO COVER THE NEEDLE RESULTING IN A NEEDLE STICK INJURY HEALTHCARE WORKER RECEIVED POST EXPOSURE LAB WORK NO ADDITIONAL TREATMENT REPORTED AT THIS TIME | Injury | Patient problem code
6828902 | 83937374 | D | IT WAS REPORTED THAT THE SAFETY MECHANISM OF THE G X IN BD ECLIPSE HYPODERMIC NEEDLE FAILED TO COVER THE NEEDLE RESULTING IN A NEEDLE STICK INJURY HEALTHCARE WORKER RECEIVED POST EXPOSURE LAB WORK NO ADDITIONAL TREATMENT REPORTED AT THIS TIME | Insufficient Information | Device problem code
6829052 | 83941145 | N | THIS REPORT IS RELATED TO PREVIOUSLY REPORTED COMPLAINT OF NOISE REPORTED ON B WITH MFR NUMBER | Noise, Audible | Device problem code
6829052 | 83941145 | N | THIS REPORT IS RELATED TO PREVIOUSLY REPORTED COMPLAINT OF NOISE REPORTED ON B WITH MFR NUMBER | No Device Problem Found | Medical Evaluation result code
6829467 | 83941480 | N | THE RETURNED CABLE WAS EVALUATED THE RETURNED UNIT PASSED VISUAL INSPECTION AND CONTINUITY TESTS BUT FAILED THE CHECK EEPROM CONTENTS VERIFICATION DURING FUNCTIONAL ANALYSIS THE CABLE WAS ABLE TO GENERATE A REPLACE CABLE ERROR MESSAGE B | Device Displays Incorrect Message | Device problem code
6829467 | 83941480 | N | THE RETURNED CABLE WAS EVALUATED THE RETURNED UNIT PASSED VISUAL INSPECTION AND CONTINUITY TESTS BUT FAILED THE CHECK EEPROM CONTENTS VERIFICATION DURING FUNCTIONAL ANALYSIS THE CABLE WAS ABLE TO GENERATE A REPLACE CABLE ERROR MESSAGE B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6829089 | 83942240 | N | UPON RECEIPT THE HELIX WAS EXTENDED AND STRETCHED OUT OF THE DISTAL TIP THIS STRETCHING WOULD HAVE OCCURRED DURING THE ATTEMPTED IMPLANT FURTHER HELIX TESTING COULD NOT BE PERFORMED DUE TO RECEIVED CONDITION OF THE HELIX | Stretched | Device problem code
6829089 | 83942240 | N | UPON RECEIPT THE HELIX WAS EXTENDED AND STRETCHED OUT OF THE DISTAL TIP THIS STRETCHING WOULD HAVE OCCURRED DURING THE ATTEMPTED IMPLANT FURTHER HELIX TESTING COULD NOT BE PERFORMED DUE TO RECEIVED CONDITION OF THE HELIX | No Findings Available | Medical Evaluation result code
6829679 | 83950078 | N | IT WAS REPORTED THAT THE PATIENT HAS POSSIBLE FEMORAL IMPLANT LOOSENING A REVISION SURGERY IS PLANNED TO RECEMENT THE FEMORAL IMPLANT AND EXCHANGE THE POLY INSERT | Loose or Intermittent Connection | Device problem code
6829679 | 83950078 | N | IT WAS REPORTED THAT THE PATIENT HAS POSSIBLE FEMORAL IMPLANT LOOSENING A REVISION SURGERY IS PLANNED TO RECEMENT THE FEMORAL IMPLANT AND EXCHANGE THE POLY INSERT | No Findings Available | Medical Evaluation result code
6829045 | 83952462 | N | NATUS MEDICAL HAS REQUESTED ADDITIONAL INFORMATION FROM CUSTOMER FOR FINAL DETERMINATION NATUS SERVICE REPAIR RECEIVE THE NEOBLUE RADIOMETER FOR CALIBRATION NATUS WILL FILE A SUPPLEMENTAL REPORT AS NEEDED | Calibration Problem | Device problem code
6829045 | 83952462 | N | NATUS MEDICAL HAS REQUESTED ADDITIONAL INFORMATION FROM CUSTOMER FOR FINAL DETERMINATION NATUS SERVICE REPAIR RECEIVE THE NEOBLUE RADIOMETER FOR CALIBRATION NATUS WILL FILE A SUPPLEMENTAL REPORT AS NEEDED | No Findings Available | Medical Evaluation result code
6824818 | 83954874 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6824818 | 83954874 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6826453 | 83961801 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS WALKING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A LONG ACTING INSULIN INJECTION WAS ADMINISTERED AND A NEW POD WAS APPLIED THE NEXT DAY | Device Dislodged or Dislocated | Device problem code
6826453 | 83961801 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER WAS WALKING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A LONG ACTING INSULIN INJECTION WAS ADMINISTERED AND A NEW POD WAS APPLIED THE NEXT DAY | No Known Impact Or Consequence To Patient | Patient Problem Code
6824939 | 83961810 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DID THE DEVICE LOCK OUT AND COULD NOT BE FIRED AT ALL OR WAS THE TRIGGER ADVANCED WITH NO STAPLES DEPLOYED THE TRIGGER ADVANCED WITH NO STAPLES DEPLOYED | Failure to Form Staple | Device problem code
6824939 | 83961810 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DID THE DEVICE LOCK OUT AND COULD NOT BE FIRED AT ALL OR WAS THE TRIGGER ADVANCED WITH NO STAPLES DEPLOYED THE TRIGGER ADVANCED WITH NO STAPLES DEPLOYED | Complement Activation Problem | Medical Evaluation result code
6830038 | 83961869 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICES TURNED COMPLETELY WHITE AND WAS ILLEGIBLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Display / Image | Device Problem Code
6830038 | 83961869 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICES TURNED COMPLETELY WHITE AND WAS ILLEGIBLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6830212 | 83969664 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER DECLINED TO REMOVE THE CANNULA AT THE TIME OF THE REPORT AND ALTHOUGH THE INFUSION SET SITE LOOKS GOOD THE CUSTOMER THOUGHT THAT THE CANNULA MAY HAVE BEEN KINKED REPORTEDLY AN INSULIN PEN WAS USED TO ADDRESS THE BG LEVEL UPON FOLLOWUP THE CUSTOMER REPORTED THE BG HAD INCREASED TO MGDL WITH A SMALL LEVEL OF KETONES THE CUSTOMER WAS UNABLE TO LOWER THE BG AFTER BLOUSING VIA INSULIN PEN AND WAS ADMITTED TO THE HOSPITAL ALTHOUGH THE CUSTOMER DID NOT OBSERVE ANY ISSUES WITH THE CARTRIDGE INFUSION SET TUBING OR CANNULA THE CUSTOMER THOUGHT THAT THE PUMP WAS AT FAULT THE CUSTOMER WAS TREATED WITH A SALINE DRIP AND INSULIN VIA AN INSULIN PEN THE HIGH BG WAS RESOLVED AND THE CUSTOMER WAS DISCHARGED THE SAME DAY | Hyperglycemia | Patient problem code
6830212 | 83969664 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER DECLINED TO REMOVE THE CANNULA AT THE TIME OF THE REPORT AND ALTHOUGH THE INFUSION SET SITE LOOKS GOOD THE CUSTOMER THOUGHT THAT THE CANNULA MAY HAVE BEEN KINKED REPORTEDLY AN INSULIN PEN WAS USED TO ADDRESS THE BG LEVEL UPON FOLLOWUP THE CUSTOMER REPORTED THE BG HAD INCREASED TO MGDL WITH A SMALL LEVEL OF KETONES THE CUSTOMER WAS UNABLE TO LOWER THE BG AFTER BLOUSING VIA INSULIN PEN AND WAS ADMITTED TO THE HOSPITAL ALTHOUGH THE CUSTOMER DID NOT OBSERVE ANY ISSUES WITH THE CARTRIDGE INFUSION SET TUBING OR CANNULA THE CUSTOMER THOUGHT THAT THE PUMP WAS AT FAULT THE CUSTOMER WAS TREATED WITH A SALINE DRIP AND INSULIN VIA AN INSULIN PEN THE HIGH BG WAS RESOLVED AND THE CUSTOMER WAS DISCHARGED THE SAME DAY | Adverse Event Without Identified Device or Use Problem | Device problem code
6828282 | 83986796 | N | IMMUCOR TECHNICAL SUPPORT COULD NOT USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BECAUSE OF A TESTING INSTRUMENT NETWORK ISSUE | Application Network Problem | Device problem code
6828282 | 83986796 | N | IMMUCOR TECHNICAL SUPPORT COULD NOT USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BECAUSE OF A TESTING INSTRUMENT NETWORK ISSUE | Network Communication Problem | Medical Evaluation result code
6831025 | 83991781 | N | DEVICE EVALUATION SUMMARY DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAS BEEN COMPLETED AS RECEIVED THE MONITOR AND BELT PASSED INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING A HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES FOLLOWED BY A HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM DURING THE TRANSITION TO THE HZ SIGNAL THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM AUDIO MESSAGING AND SIREN ALARMS AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY J BIPHASIC PULSES THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY ECG ACQUISITION DETECTION ALGORITHM PERFORMANCE AND PULSE DELIVERY FUNCTIONALITY DEVICE MANUFACTURE DATE MONITOR BELT | Vibration | Device problem code
6831025 | 83991781 | N | DEVICE EVALUATION SUMMARY DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAS BEEN COMPLETED AS RECEIVED THE MONITOR AND BELT PASSED INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING A HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES FOLLOWED BY A HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM DURING THE TRANSITION TO THE HZ SIGNAL THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM AUDIO MESSAGING AND SIREN ALARMS AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY J BIPHASIC PULSES THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY ECG ACQUISITION DETECTION ALGORITHM PERFORMANCE AND PULSE DELIVERY FUNCTIONALITY DEVICE MANUFACTURE DATE MONITOR BELT | Charging Problem | Device problem code
6831025 | 83991781 | N | DEVICE EVALUATION SUMMARY DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAS BEEN COMPLETED AS RECEIVED THE MONITOR AND BELT PASSED INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING A HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES FOLLOWED BY A HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM DURING THE TRANSITION TO THE HZ SIGNAL THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM AUDIO MESSAGING AND SIREN ALARMS AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY J BIPHASIC PULSES THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY ECG ACQUISITION DETECTION ALGORITHM PERFORMANCE AND PULSE DELIVERY FUNCTIONALITY DEVICE MANUFACTURE DATE MONITOR BELT | Energy Storage System Problem | Medical Evaluation result code
6827744 | 83996490 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE TWO SAME SAMPLE IDSSAME BLOOD TUBE TEN PATIENT SAMPLE RESULTS WERE SHOWN AS BEING RUN IN DUPLICATES WHICH DID NOT LEAD TO PATIENT SAMPLE MISIDENTIFICATION PATIENT SAMPLE MISIDENTIFICATION OCCURRED FOR ONLY THE LAST TUBE IN A CASSETTE AND ONLY WHEN THERE IS ANOTHER CASSETTE IN THE MULTILOADER OF THE INSTRUMENT THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT SEVEN ADDITIONAL INSTANCES WERE IDENTIFIED THAT LEAD TO ERRONEOUS RESULTS BASED ON THE SOFTWARE ENGINEERING ASSESSMENT THERE IS A MECHANISM BY WHICH THE AQUIOS SOFTWARE WILL GENERATE A DUPLICATE TEST WHEN A TEST QUERY RESPONSE IS RECEIVED FROM THE LIS IMMEDIATELY BEFORE THE QUERY TIMEOUT PERIOD EXPIRES A WINDOW IN THE ORDER OF MILLISECONDS IN SIZE ONCE THE DUPLICATE TEST IS CREATED THE SOFTWARE WILL CONTINUE TO CREATE DUPLICATE TESTS AS LONG AS THERE ARE TUBES AVAILABLE IN THE CASSETTE AND IN THE STACKER ONCE ALL TESTS HAVE BEEN COMPLETED THE SOFTWARE REVERTS TO NORMAL OPERATION THE SOFTWARE TEAM STATED THAT THE SOFTWARE DEFECT ONLY OCCURS UNDER THESE CONDITIONS A TEST QUERY RESPONSE TIMES OUT AND THE DEFAULT TEST OPTION IS ON AND WHEN THE RIGHT TINY TIME WINDOW IS HIT ALL PATTERNS DESCRIBED ABOVE ARE ALL SYMPTOMS OF THE SAME ROOT CAUSE PROBLEM WHEN THE ISSUE OCCURS THE SYSTEM WILL GET STUCK PRODUCING TWO REQUESTS FOR EACH TUBE THE PATTERN PRODUCED BY THE DUPLICATION DEPENDS ON THE POSITIONING OF THE TUBES IN THE CASSETTE WHICH IN TURN MAY CAUSE THE INSTRUMENT TO RESAMPLE THE SAME TUBE OR DRAW THE SAMPLE FROM ANOTHER TUBE BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION THERE HAVE BEEN NO REOCCURRENCE OF ERRONEOUS PATIENT RESULTS REPORTED TO DATE B RELATED EVENT BEC IDENTIFIER B | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6827744 | 83996490 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE TWO SAME SAMPLE IDSSAME BLOOD TUBE TEN PATIENT SAMPLE RESULTS WERE SHOWN AS BEING RUN IN DUPLICATES WHICH DID NOT LEAD TO PATIENT SAMPLE MISIDENTIFICATION PATIENT SAMPLE MISIDENTIFICATION OCCURRED FOR ONLY THE LAST TUBE IN A CASSETTE AND ONLY WHEN THERE IS ANOTHER CASSETTE IN THE MULTILOADER OF THE INSTRUMENT THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT SEVEN ADDITIONAL INSTANCES WERE IDENTIFIED THAT LEAD TO ERRONEOUS RESULTS BASED ON THE SOFTWARE ENGINEERING ASSESSMENT THERE IS A MECHANISM BY WHICH THE AQUIOS SOFTWARE WILL GENERATE A DUPLICATE TEST WHEN A TEST QUERY RESPONSE IS RECEIVED FROM THE LIS IMMEDIATELY BEFORE THE QUERY TIMEOUT PERIOD EXPIRES A WINDOW IN THE ORDER OF MILLISECONDS IN SIZE ONCE THE DUPLICATE TEST IS CREATED THE SOFTWARE WILL CONTINUE TO CREATE DUPLICATE TESTS AS LONG AS THERE ARE TUBES AVAILABLE IN THE CASSETTE AND IN THE STACKER ONCE ALL TESTS HAVE BEEN COMPLETED THE SOFTWARE REVERTS TO NORMAL OPERATION THE SOFTWARE TEAM STATED THAT THE SOFTWARE DEFECT ONLY OCCURS UNDER THESE CONDITIONS A TEST QUERY RESPONSE TIMES OUT AND THE DEFAULT TEST OPTION IS ON AND WHEN THE RIGHT TINY TIME WINDOW IS HIT ALL PATTERNS DESCRIBED ABOVE ARE ALL SYMPTOMS OF THE SAME ROOT CAUSE PROBLEM WHEN THE ISSUE OCCURS THE SYSTEM WILL GET STUCK PRODUCING TWO REQUESTS FOR EACH TUBE THE PATTERN PRODUCED BY THE DUPLICATION DEPENDS ON THE POSITIONING OF THE TUBES IN THE CASSETTE WHICH IN TURN MAY CAUSE THE INSTRUMENT TO RESAMPLE THE SAME TUBE OR DRAW THE SAMPLE FROM ANOTHER TUBE BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION THERE HAVE BEEN NO REOCCURRENCE OF ERRONEOUS PATIENT RESULTS REPORTED TO DATE B RELATED EVENT BEC IDENTIFIER B | Defective Device | Device problem code
6827744 | 83996490 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE TWO SAME SAMPLE IDSSAME BLOOD TUBE TEN PATIENT SAMPLE RESULTS WERE SHOWN AS BEING RUN IN DUPLICATES WHICH DID NOT LEAD TO PATIENT SAMPLE MISIDENTIFICATION PATIENT SAMPLE MISIDENTIFICATION OCCURRED FOR ONLY THE LAST TUBE IN A CASSETTE AND ONLY WHEN THERE IS ANOTHER CASSETTE IN THE MULTILOADER OF THE INSTRUMENT THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT SEVEN ADDITIONAL INSTANCES WERE IDENTIFIED THAT LEAD TO ERRONEOUS RESULTS BASED ON THE SOFTWARE ENGINEERING ASSESSMENT THERE IS A MECHANISM BY WHICH THE AQUIOS SOFTWARE WILL GENERATE A DUPLICATE TEST WHEN A TEST QUERY RESPONSE IS RECEIVED FROM THE LIS IMMEDIATELY BEFORE THE QUERY TIMEOUT PERIOD EXPIRES A WINDOW IN THE ORDER OF MILLISECONDS IN SIZE ONCE THE DUPLICATE TEST IS CREATED THE SOFTWARE WILL CONTINUE TO CREATE DUPLICATE TESTS AS LONG AS THERE ARE TUBES AVAILABLE IN THE CASSETTE AND IN THE STACKER ONCE ALL TESTS HAVE BEEN COMPLETED THE SOFTWARE REVERTS TO NORMAL OPERATION THE SOFTWARE TEAM STATED THAT THE SOFTWARE DEFECT ONLY OCCURS UNDER THESE CONDITIONS A TEST QUERY RESPONSE TIMES OUT AND THE DEFAULT TEST OPTION IS ON AND WHEN THE RIGHT TINY TIME WINDOW IS HIT ALL PATTERNS DESCRIBED ABOVE ARE ALL SYMPTOMS OF THE SAME ROOT CAUSE PROBLEM WHEN THE ISSUE OCCURS THE SYSTEM WILL GET STUCK PRODUCING TWO REQUESTS FOR EACH TUBE THE PATTERN PRODUCED BY THE DUPLICATION DEPENDS ON THE POSITIONING OF THE TUBES IN THE CASSETTE WHICH IN TURN MAY CAUSE THE INSTRUMENT TO RESAMPLE THE SAME TUBE OR DRAW THE SAMPLE FROM ANOTHER TUBE BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION THERE HAVE BEEN NO REOCCURRENCE OF ERRONEOUS PATIENT RESULTS REPORTED TO DATE B RELATED EVENT BEC IDENTIFIER B | Mechanical Jam | Device problem code
6827744 | 83996490 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE TWO SAME SAMPLE IDSSAME BLOOD TUBE TEN PATIENT SAMPLE RESULTS WERE SHOWN AS BEING RUN IN DUPLICATES WHICH DID NOT LEAD TO PATIENT SAMPLE MISIDENTIFICATION PATIENT SAMPLE MISIDENTIFICATION OCCURRED FOR ONLY THE LAST TUBE IN A CASSETTE AND ONLY WHEN THERE IS ANOTHER CASSETTE IN THE MULTILOADER OF THE INSTRUMENT THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT SEVEN ADDITIONAL INSTANCES WERE IDENTIFIED THAT LEAD TO ERRONEOUS RESULTS BASED ON THE SOFTWARE ENGINEERING ASSESSMENT THERE IS A MECHANISM BY WHICH THE AQUIOS SOFTWARE WILL GENERATE A DUPLICATE TEST WHEN A TEST QUERY RESPONSE IS RECEIVED FROM THE LIS IMMEDIATELY BEFORE THE QUERY TIMEOUT PERIOD EXPIRES A WINDOW IN THE ORDER OF MILLISECONDS IN SIZE ONCE THE DUPLICATE TEST IS CREATED THE SOFTWARE WILL CONTINUE TO CREATE DUPLICATE TESTS AS LONG AS THERE ARE TUBES AVAILABLE IN THE CASSETTE AND IN THE STACKER ONCE ALL TESTS HAVE BEEN COMPLETED THE SOFTWARE REVERTS TO NORMAL OPERATION THE SOFTWARE TEAM STATED THAT THE SOFTWARE DEFECT ONLY OCCURS UNDER THESE CONDITIONS A TEST QUERY RESPONSE TIMES OUT AND THE DEFAULT TEST OPTION IS ON AND WHEN THE RIGHT TINY TIME WINDOW IS HIT ALL PATTERNS DESCRIBED ABOVE ARE ALL SYMPTOMS OF THE SAME ROOT CAUSE PROBLEM WHEN THE ISSUE OCCURS THE SYSTEM WILL GET STUCK PRODUCING TWO REQUESTS FOR EACH TUBE THE PATTERN PRODUCED BY THE DUPLICATION DEPENDS ON THE POSITIONING OF THE TUBES IN THE CASSETTE WHICH IN TURN MAY CAUSE THE INSTRUMENT TO RESAMPLE THE SAME TUBE OR DRAW THE SAMPLE FROM ANOTHER TUBE BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION THERE HAVE BEEN NO REOCCURRENCE OF ERRONEOUS PATIENT RESULTS REPORTED TO DATE B RELATED EVENT BEC IDENTIFIER B | Mechanical Problem Identified | Medical Evaluation result code
6827744 | 83996490 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE TWO SAME SAMPLE IDSSAME BLOOD TUBE TEN PATIENT SAMPLE RESULTS WERE SHOWN AS BEING RUN IN DUPLICATES WHICH DID NOT LEAD TO PATIENT SAMPLE MISIDENTIFICATION PATIENT SAMPLE MISIDENTIFICATION OCCURRED FOR ONLY THE LAST TUBE IN A CASSETTE AND ONLY WHEN THERE IS ANOTHER CASSETTE IN THE MULTILOADER OF THE INSTRUMENT THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT SEVEN ADDITIONAL INSTANCES WERE IDENTIFIED THAT LEAD TO ERRONEOUS RESULTS BASED ON THE SOFTWARE ENGINEERING ASSESSMENT THERE IS A MECHANISM BY WHICH THE AQUIOS SOFTWARE WILL GENERATE A DUPLICATE TEST WHEN A TEST QUERY RESPONSE IS RECEIVED FROM THE LIS IMMEDIATELY BEFORE THE QUERY TIMEOUT PERIOD EXPIRES A WINDOW IN THE ORDER OF MILLISECONDS IN SIZE ONCE THE DUPLICATE TEST IS CREATED THE SOFTWARE WILL CONTINUE TO CREATE DUPLICATE TESTS AS LONG AS THERE ARE TUBES AVAILABLE IN THE CASSETTE AND IN THE STACKER ONCE ALL TESTS HAVE BEEN COMPLETED THE SOFTWARE REVERTS TO NORMAL OPERATION THE SOFTWARE TEAM STATED THAT THE SOFTWARE DEFECT ONLY OCCURS UNDER THESE CONDITIONS A TEST QUERY RESPONSE TIMES OUT AND THE DEFAULT TEST OPTION IS ON AND WHEN THE RIGHT TINY TIME WINDOW IS HIT ALL PATTERNS DESCRIBED ABOVE ARE ALL SYMPTOMS OF THE SAME ROOT CAUSE PROBLEM WHEN THE ISSUE OCCURS THE SYSTEM WILL GET STUCK PRODUCING TWO REQUESTS FOR EACH TUBE THE PATTERN PRODUCED BY THE DUPLICATION DEPENDS ON THE POSITIONING OF THE TUBES IN THE CASSETTE WHICH IN TURN MAY CAUSE THE INSTRUMENT TO RESAMPLE THE SAME TUBE OR DRAW THE SAMPLE FROM ANOTHER TUBE BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION THERE HAVE BEEN NO REOCCURRENCE OF ERRONEOUS PATIENT RESULTS REPORTED TO DATE B RELATED EVENT BEC IDENTIFIER B | Software Problem Identified | Medical Evaluation result code
6830118 | 83999465 | D | IT WAS REPORTED THAT THE PATIENT WAS AT SUBACUTE REHABILITATION FACILITY WHEN THE SITE RECEIVED A CALL ON B INDICATING THAT THE PATIENT WAS EXHIBITING HEMATURIA AND INCREASING LACTATE DEHYDROGENASE LDH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THERAPEUTIC THERE WAS NO REPORT OF ABNORMAL PUMP PARAMETERS OR CONTROLLER ALARMS CONTROLLER LOG FILES WERE NOT SENT IN DURING THE ADMISSION THE PATIENT WAS ADMITTED AND STARTED ON INTRAVENOUS IV HEPARIN THE CLINICAL TEAM SUSPECTED PUMP THROMBUS HOWEVER IT WAS NEVER CONFIRMED THE PATIENT DID NOT HAVE A RECENT ILLNESS THAT MAY HAVE PRECIPITATED PUMP THROMBUS ADDITIONALLY THE PATIENT HAD NO SIGNIFICANT HISTORY REGARDING PUMP THROMBUS BUT DID HAVE AN ISCHEMIC CEREBROVASCULAR ACCIDENT ON POSTOPERATIVE DAY THREE PREVIOUSLY REPORTED THE CLINICAL TEAM DID NOT ATTEMPT TISSUE PLASMINOGEN ACTIVATOR TPA DUE TO THE PATIENTS HISTORY OF STROKE CARDIOVASCULAR CV SURGERY WAS CONSULTED BUT THE PATIENT WAS REFUSING A PUMP EXCHANGE THE PATIENTS LDH BEGAN TO DECREASE WHILE ON HEPARIN AND WARFARIN WAS REINITIATED THE CLINICAL TEAM DISCHARGED THE PATIENT BACK TO THE SUBACUTE REHABILITATION FACILITY ON B FOR FURTHER REHAB WHEN INR WAS THERAPEUTIC GOAL IT WAS LAST REPORTED THAT THE PATIENT IS DOING WELL NO FURTHER INFORMATION HAS BEEN PROVIDED | No Consequences Or Impact To Patient | Patient problem code
6830118 | 83999465 | D | IT WAS REPORTED THAT THE PATIENT WAS AT SUBACUTE REHABILITATION FACILITY WHEN THE SITE RECEIVED A CALL ON B INDICATING THAT THE PATIENT WAS EXHIBITING HEMATURIA AND INCREASING LACTATE DEHYDROGENASE LDH THE PATIENTS INTERNATIONAL NORMALIZED RATIO INR WAS THERAPEUTIC THERE WAS NO REPORT OF ABNORMAL PUMP PARAMETERS OR CONTROLLER ALARMS CONTROLLER LOG FILES WERE NOT SENT IN DURING THE ADMISSION THE PATIENT WAS ADMITTED AND STARTED ON INTRAVENOUS IV HEPARIN THE CLINICAL TEAM SUSPECTED PUMP THROMBUS HOWEVER IT WAS NEVER CONFIRMED THE PATIENT DID NOT HAVE A RECENT ILLNESS THAT MAY HAVE PRECIPITATED PUMP THROMBUS ADDITIONALLY THE PATIENT HAD NO SIGNIFICANT HISTORY REGARDING PUMP THROMBUS BUT DID HAVE AN ISCHEMIC CEREBROVASCULAR ACCIDENT ON POSTOPERATIVE DAY THREE PREVIOUSLY REPORTED THE CLINICAL TEAM DID NOT ATTEMPT TISSUE PLASMINOGEN ACTIVATOR TPA DUE TO THE PATIENTS HISTORY OF STROKE CARDIOVASCULAR CV SURGERY WAS CONSULTED BUT THE PATIENT WAS REFUSING A PUMP EXCHANGE THE PATIENTS LDH BEGAN TO DECREASE WHILE ON HEPARIN AND WARFARIN WAS REINITIATED THE CLINICAL TEAM DISCHARGED THE PATIENT BACK TO THE SUBACUTE REHABILITATION FACILITY ON B FOR FURTHER REHAB WHEN INR WAS THERAPEUTIC GOAL IT WAS LAST REPORTED THAT THE PATIENT IS DOING WELL NO FURTHER INFORMATION HAS BEEN PROVIDED | Insufficient Information | Device problem code
6831529 | 84008612 | N | THE ROOT CAUSE OF THE PATIENTS INFECTION WAS DETERMINED TO BE THE SEROMA WHICH REQUIRED DRAINAGE THE SEROMA WAS THE LIKELY PRECURSOR TO THE COURSE OF EVENTS A SEROMA IS A KNOWN POTENTIAL COMPLICATION WITH THE USE OF ZENAPRO THEREFORE THE DEVICE CONTRIBUTED TO THE NEED FOR REOPERATION DUE TO THE DEVELOPMENT OF A SEROMA WHICH REQUIRED PERCUTANEOUS DRAINAGE WHICH LIKELY LED TO AN INFECTION THE REOPERATION WAS ALSO DETERMINED TO HAVE BEEN PERFORMED TO PRECLUDE SERIOUS INJURY AS THE ZENAPRO DEVICE IS CONTRAINDICATED FOR USE IN INFECTED FIELDS | Leak / Splash | Device problem code
6831529 | 84008612 | N | THE ROOT CAUSE OF THE PATIENTS INFECTION WAS DETERMINED TO BE THE SEROMA WHICH REQUIRED DRAINAGE THE SEROMA WAS THE LIKELY PRECURSOR TO THE COURSE OF EVENTS A SEROMA IS A KNOWN POTENTIAL COMPLICATION WITH THE USE OF ZENAPRO THEREFORE THE DEVICE CONTRIBUTED TO THE NEED FOR REOPERATION DUE TO THE DEVELOPMENT OF A SEROMA WHICH REQUIRED PERCUTANEOUS DRAINAGE WHICH LIKELY LED TO AN INFECTION THE REOPERATION WAS ALSO DETERMINED TO HAVE BEEN PERFORMED TO PRECLUDE SERIOUS INJURY AS THE ZENAPRO DEVICE IS CONTRAINDICATED FOR USE IN INFECTED FIELDS | Leakage/Seal | Medical Evaluation result code
6832199 | 84026806 | D | MEDTRONIC RECEIVED INFORMATION THAT FIVE DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS NOTED SIX DAYS FOLLOWING THE VALVE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6832199 | 84026806 | D | MEDTRONIC RECEIVED INFORMATION THAT FIVE DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS NOTED SIX DAYS FOLLOWING THE VALVE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6833100 | 84070845 | N | AS SOON AS ELLEX MEDICAL BECAME AWARE OF THE INCIDENT THE CALIBRATION OF THE LASER SYSTEM AND ALIGNMENT OF THE OPTICAL PATHS WAS VERIFIED BY ELLEX SERVICE THROUGH A PAF PRODUCT ACCEPTANCE FAULT RECORD ALL THE PARAMETERS IN ALIGNMENTS CHECKS INCLUDING OFFSET IN THE PAF ARE WITHIN THE ACCEPTANCE CRITERIA WHICH PROVES THAT THE DEVICE IS NOT AT FAULT AND IS PERFORMING WITHIN THE NORMAL SPECIFICATIONS THE DEVICE WAS TESTED IN MODEL EYE AND THERE WAS NO ISSUE WITH ALIGNMENT MECHANICAL OPERATION OR CALIBRATION THE GENERAL DEVICE ASPECTS WITH RESPECT TO USABILITY INSTRUCTIONS FOR USE AND THE RISK CONTROLS WERE ALSO EVALUATED ADEQUACY OF DEVICE DESIGN INDICATIONS OF PARAMETERS RELEVANT TO SAFETY OFFSET IS INDICATED ON LASER HEAD CONTROL KNOB BY AP NUMBERED COLOUR CODED INDICATIONSTHE LETTER A AND THE COLOUR ORANGE ON THE OFFSET CONTROL INDICATE ANTERIOR POSITIONS THE LETTER P AND THE COLOUR BLUE ON THE OFFSET CONTROLS INDICATE POSTERIOR POSITIONS ACCURACY OF CONTROLS MICRO M OFFSET RANGE WITH TOLERANCE OF MICRO M PRODUCT ACCEPTANCE FAULT RECORDS POST THE EVENT ALSO CONFIRMED THE OFFSET RESULTS WERE NORMAL ADEQUACY OF LABELLING OPERATOR MANUAL INCLUDES SUFFICIENT INSTRUCTIONS WARNINGS PRECAUTIONS FOR THE SAFE USE OF THE DEVICE AND THE LABELLING IS CONSIDERED ADEQUATE DO NOT AIM FOCUS OR FIRE THIS DEVICE IN ON OR NEAR THE FOVEA ON CORNEAL STRUCTURES ON THE PATIENTS LENS NATURAL OR ARTIFICIAL OR ON ANY OTHER STRUCTURE OF THE HUMAN BODY THAT IS NOT RELATED TO THE TREATMENT OF AN EYE DISORDER ALWAYS CONFIRM THE OFFSET DIRECTION ANTERIOR A OR POSTERIOR P AND OFFSET MEASUREMENT BEFORE YOU FIRE THE TREATMENT LASER WHEN TREATING CLOSE TO THE POSTERIOR LENS SURFACE A POSTERIOR OFFSET SHOULD BE CONSIDERED SET THE POSTERIOR P OFFSET TO AT LEAST MICROM WITH A MINIMUM ENERGY SETTING IF USING A HIGHER ENERGY SETTING CHOOSE A GREATER POSTERIOR P OFFSET ONLY FULLY TRAINED AND QUALIFIED PHYSICIANS MAY OPERATE THIS DEVICE THIS DEVICE IS A SAFE INSTRUMENT WHEN USED CORRECTLY HOWEVER LIKE ALL LASER SURGICAL EQUIPMENT IT CAN CAUSE INJURY IF NOT USED IN ACCORDANCE WITH THE CORRECT SAFETY PROCEDURES AND OPERATING INSTRUCTIONS RISK MANAGEMENT REVIEW INDICATED THAT THE RISK CONTROL MEASURES RISK ASSESSMENT RA TANGO ARE ADEQUATE AND WITHIN THE NORMAL USE RISK LEVELS ARE ACCEPTABLE CUSTOMER COMPLAINTS THE DEVICE HAS BEEN IN USE FOR LESS THAN A YEAR NO OTHER UNTOWARD INCIDENTS RELATED TO DEVICE USE WERE REPORTED | Calibration Problem | Device problem code
6833100 | 84070845 | N | AS SOON AS ELLEX MEDICAL BECAME AWARE OF THE INCIDENT THE CALIBRATION OF THE LASER SYSTEM AND ALIGNMENT OF THE OPTICAL PATHS WAS VERIFIED BY ELLEX SERVICE THROUGH A PAF PRODUCT ACCEPTANCE FAULT RECORD ALL THE PARAMETERS IN ALIGNMENTS CHECKS INCLUDING OFFSET IN THE PAF ARE WITHIN THE ACCEPTANCE CRITERIA WHICH PROVES THAT THE DEVICE IS NOT AT FAULT AND IS PERFORMING WITHIN THE NORMAL SPECIFICATIONS THE DEVICE WAS TESTED IN MODEL EYE AND THERE WAS NO ISSUE WITH ALIGNMENT MECHANICAL OPERATION OR CALIBRATION THE GENERAL DEVICE ASPECTS WITH RESPECT TO USABILITY INSTRUCTIONS FOR USE AND THE RISK CONTROLS WERE ALSO EVALUATED ADEQUACY OF DEVICE DESIGN INDICATIONS OF PARAMETERS RELEVANT TO SAFETY OFFSET IS INDICATED ON LASER HEAD CONTROL KNOB BY AP NUMBERED COLOUR CODED INDICATIONSTHE LETTER A AND THE COLOUR ORANGE ON THE OFFSET CONTROL INDICATE ANTERIOR POSITIONS THE LETTER P AND THE COLOUR BLUE ON THE OFFSET CONTROLS INDICATE POSTERIOR POSITIONS ACCURACY OF CONTROLS MICRO M OFFSET RANGE WITH TOLERANCE OF MICRO M PRODUCT ACCEPTANCE FAULT RECORDS POST THE EVENT ALSO CONFIRMED THE OFFSET RESULTS WERE NORMAL ADEQUACY OF LABELLING OPERATOR MANUAL INCLUDES SUFFICIENT INSTRUCTIONS WARNINGS PRECAUTIONS FOR THE SAFE USE OF THE DEVICE AND THE LABELLING IS CONSIDERED ADEQUATE DO NOT AIM FOCUS OR FIRE THIS DEVICE IN ON OR NEAR THE FOVEA ON CORNEAL STRUCTURES ON THE PATIENTS LENS NATURAL OR ARTIFICIAL OR ON ANY OTHER STRUCTURE OF THE HUMAN BODY THAT IS NOT RELATED TO THE TREATMENT OF AN EYE DISORDER ALWAYS CONFIRM THE OFFSET DIRECTION ANTERIOR A OR POSTERIOR P AND OFFSET MEASUREMENT BEFORE YOU FIRE THE TREATMENT LASER WHEN TREATING CLOSE TO THE POSTERIOR LENS SURFACE A POSTERIOR OFFSET SHOULD BE CONSIDERED SET THE POSTERIOR P OFFSET TO AT LEAST MICROM WITH A MINIMUM ENERGY SETTING IF USING A HIGHER ENERGY SETTING CHOOSE A GREATER POSTERIOR P OFFSET ONLY FULLY TRAINED AND QUALIFIED PHYSICIANS MAY OPERATE THIS DEVICE THIS DEVICE IS A SAFE INSTRUMENT WHEN USED CORRECTLY HOWEVER LIKE ALL LASER SURGICAL EQUIPMENT IT CAN CAUSE INJURY IF NOT USED IN ACCORDANCE WITH THE CORRECT SAFETY PROCEDURES AND OPERATING INSTRUCTIONS RISK MANAGEMENT REVIEW INDICATED THAT THE RISK CONTROL MEASURES RISK ASSESSMENT RA TANGO ARE ADEQUATE AND WITHIN THE NORMAL USE RISK LEVELS ARE ACCEPTABLE CUSTOMER COMPLAINTS THE DEVICE HAS BEEN IN USE FOR LESS THAN A YEAR NO OTHER UNTOWARD INCIDENTS RELATED TO DEVICE USE WERE REPORTED | Failure to Calibrate | Device Problem Code
6831616 | 84078434 | N | THE CUSTOMER REPORTED THAT THE POWER SUPPLY ON THE CENTRAL NURSES STATION CNS MONITOR FAILED WHILE IN PATIENT USE POWER SUPPLY PART NUMBER WAS EMAILED TO THE CUSTOMER MULTIPLE ATTEMPTS WERE MADE TO THE CUSTOMER TO OBTAIN THE SERIAL NUMBER CUSTOMER DID NOT RESPOND TO ANY CALLS OR EMAIL NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH CFR SECTION IF CUSTOMER PROVIDES THAT INFORMATION | No Display / Image | Device problem code
6831616 | 84078434 | N | THE CUSTOMER REPORTED THAT THE POWER SUPPLY ON THE CENTRAL NURSES STATION CNS MONITOR FAILED WHILE IN PATIENT USE POWER SUPPLY PART NUMBER WAS EMAILED TO THE CUSTOMER MULTIPLE ATTEMPTS WERE MADE TO THE CUSTOMER TO OBTAIN THE SERIAL NUMBER CUSTOMER DID NOT RESPOND TO ANY CALLS OR EMAIL NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH CFR SECTION IF CUSTOMER PROVIDES THAT INFORMATION | No Device Problem Found | Medical Evaluation result code
6833298 | 84078587 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF HAS SUFFERED BODILY INJURY AND RESULTING PAIN AND SUFFERING MENTAL ANGUISH DISABILITY DISFIGUREMENT AND LOSS OF THE CAPACITY FOR THE ENJOYMENT OF LIFE HAS INCURRED AND WILL INCUR IN THE FUTURE EXPENSE OF HOSPITALIZATION MEDICAL AND NURSING AND TREATMENT AND AGGRAVATION OF A PREVIOUSLY EXISTING CONDITION THESE LOSSES ARE PERMANENT OR CONTINUING IN NATURE AND PATIENT WILL SUFFER THEM IN THE FUTURE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENT DEVICE LABELING PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | No Device Problem Found | Medical Evaluation result code
6831654 | 84079065 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE CENTRAL NURSES STATION CNS WAS NOT ALARMING FOR ONE PATIENT THE NURSING STAFF HEARD THE BEDSIDE MONITOR BSM ALARM BUT NOT AT THE CENTRAL NURSES STATION CNS AFTER SOME TROUBLESHOOTING IT WAS DISCOVERED THAT THE ALARM CABLE WAS NOT PLUGGED IN PLUGGING THE CABLE BACK IN RESOLVED THE ISSUE | Connection Problem | Device problem code
6831654 | 84079065 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE CENTRAL NURSES STATION CNS WAS NOT ALARMING FOR ONE PATIENT THE NURSING STAFF HEARD THE BEDSIDE MONITOR BSM ALARM BUT NOT AT THE CENTRAL NURSES STATION CNS AFTER SOME TROUBLESHOOTING IT WAS DISCOVERED THAT THE ALARM CABLE WAS NOT PLUGGED IN PLUGGING THE CABLE BACK IN RESOLVED THE ISSUE | Wired Communication Problem | Medical Evaluation result code
6825803 | 84082628 | N | THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CRITICAL BATTERY ALARM ON B THAT WAS TRIGGERED BY B ON POWER PORT THE RETURNED BATTERY B WAS CONNECTED TO POWER PORT DURING THE REPORTED EVENT BUT DID NOT TRIGGER THE ALARM DURING THE EVENT THE BATTERY B WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY ANALYSIS OF B REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING HEARTWARE IS INVESTIGATING COMMUNICATION ERRORS PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communication or Transmission Problem | Device problem code
6825803 | 84082628 | N | THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CRITICAL BATTERY ALARM ON B THAT WAS TRIGGERED BY B ON POWER PORT THE RETURNED BATTERY B WAS CONNECTED TO POWER PORT DURING THE REPORTED EVENT BUT DID NOT TRIGGER THE ALARM DURING THE EVENT THE BATTERY B WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY ANALYSIS OF B REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING HEARTWARE IS INVESTIGATING COMMUNICATION ERRORS PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communications Problem Identified | Medical Evaluation result code
6833585 | 84084820 | D | IT WAS REPORTED THAT PATIENT WAS HOSPITALIZED FOR HYPERGLYCEMIA THE PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE ADHESIVE WAS LIFTING OFF THE SITE NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE THE POD WAS REMOVED BY THE PATIENT AT THE HOSPITAL AND THE PATIENT WAS TREATED WITH GRADUAL INSULIN DELIVERY VIA INTRAVENOUS IV ROUTE AS WELL AS IV FLUIDS THE PATIENTS BG LEVELS WHILE HOSPITALIZED DECREASED TO MMOLL MGDL | Device Dislodged or Dislocated | Device problem code
6833585 | 84084820 | D | IT WAS REPORTED THAT PATIENT WAS HOSPITALIZED FOR HYPERGLYCEMIA THE PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE ADHESIVE WAS LIFTING OFF THE SITE NEAR THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE THE POD WAS REMOVED BY THE PATIENT AT THE HOSPITAL AND THE PATIENT WAS TREATED WITH GRADUAL INSULIN DELIVERY VIA INTRAVENOUS IV ROUTE AS WELL AS IV FLUIDS THE PATIENTS BG LEVELS WHILE HOSPITALIZED DECREASED TO MMOLL MGDL | Hyperglycemia | Patient problem code
6830509 | 84108204 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Insufficient Information | Device problem code
6830509 | 84108204 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6833779 | 84109314 | N | BASED ON THE MANUFACTURER INVESTIGATION THE DIFFICULT PLATEHANDLE CONNECTION WAS THE RESULT OF MINOR OFFSET IN THE LOCATION OF PLATE RECESS INTENDED TO CONNECT THE PLATE TO THE INSERTION HANDLE IT IS NOTED THAT ALTHOUGH DIFFICULT CONNECTION WAS POSSIBLE THE PLATE WAS SATISFACTORILY IMPLANTED IT WAS CONCLUDED THAT THE REPORTED CASE WAS A RESULT OF A MANUFACTURING FAULT THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY DEVICE IS STILL IMPLANTED | Connection Problem | Device problem code
6833779 | 84109314 | N | BASED ON THE MANUFACTURER INVESTIGATION THE DIFFICULT PLATEHANDLE CONNECTION WAS THE RESULT OF MINOR OFFSET IN THE LOCATION OF PLATE RECESS INTENDED TO CONNECT THE PLATE TO THE INSERTION HANDLE IT IS NOTED THAT ALTHOUGH DIFFICULT CONNECTION WAS POSSIBLE THE PLATE WAS SATISFACTORILY IMPLANTED IT WAS CONCLUDED THAT THE REPORTED CASE WAS A RESULT OF A MANUFACTURING FAULT THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY DEVICE IS STILL IMPLANTED | Manufacturing Process Problem Identified | Medical Evaluation result code
6833780 | 84109905 | N | EXAMINATION OF THE INVOLVED PLATE REVEALED THAT THE RECESS INTENDED FOR HANDLE CONNECTION WAS LOCATED WITH A MM OFFSET THIS RESULTS IN A DIFFICULT ALTHOUGH POSSIBLE PLATEHANDLE CONNECTION FOLLOWING INVESTIGATION OF THE CASE IT WAS CONCLUDED THAT THE REPORTED CASE WAS A RESULT OF A MANUFACTURING FAULT THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY | Connection Problem | Device problem code
6833780 | 84109905 | N | EXAMINATION OF THE INVOLVED PLATE REVEALED THAT THE RECESS INTENDED FOR HANDLE CONNECTION WAS LOCATED WITH A MM OFFSET THIS RESULTS IN A DIFFICULT ALTHOUGH POSSIBLE PLATEHANDLE CONNECTION FOLLOWING INVESTIGATION OF THE CASE IT WAS CONCLUDED THAT THE REPORTED CASE WAS A RESULT OF A MANUFACTURING FAULT THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY | Manufacturing Process Problem Identified | Medical Evaluation result code
6831431 | 84120339 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6831431 | 84120339 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6831138 | 84122128 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT THE CUSTOMER CLAIMED THAT THEY WERE HAVING A HELIUM LEAK THE FSE COULD NOT RECREATE THE REPORTED ISSUE AND THERE WERE NO ERROR CODES ASSOCIATED WITH THE REPORTED MALFUNCTION THE FSE FOUND FLUID IN THE LINES CLEARED THE FLUID FROM THE LINES AND FOUND NO FURTHER ISSUES THE IABP PASSED ALL FUNCTIONAL AND SAFETY TEST PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | Leak / Splash | Device problem code
6831138 | 84122128 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT THE CUSTOMER CLAIMED THAT THEY WERE HAVING A HELIUM LEAK THE FSE COULD NOT RECREATE THE REPORTED ISSUE AND THERE WERE NO ERROR CODES ASSOCIATED WITH THE REPORTED MALFUNCTION THE FSE FOUND FLUID IN THE LINES CLEARED THE FLUID FROM THE LINES AND FOUND NO FURTHER ISSUES THE IABP PASSED ALL FUNCTIONAL AND SAFETY TEST PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | Component or Accessory Incompatibility | Device Problem Code
6831138 | 84122128 | N | B THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT THE CUSTOMER CLAIMED THAT THEY WERE HAVING A HELIUM LEAK THE FSE COULD NOT RECREATE THE REPORTED ISSUE AND THERE WERE NO ERROR CODES ASSOCIATED WITH THE REPORTED MALFUNCTION THE FSE FOUND FLUID IN THE LINES CLEARED THE FLUID FROM THE LINES AND FOUND NO FURTHER ISSUES THE IABP PASSED ALL FUNCTIONAL AND SAFETY TEST PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | Leakage/Seal | Medical Evaluation result code
6832547 | 84126472 | N | THE DENTIST REPORTED A B PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE BUT NO HARM CAUSED TO THE PATIENT PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Failure to Osseointegrate | Device problem code
6832547 | 84126472 | N | THE DENTIST REPORTED A B PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE BUT NO HARM CAUSED TO THE PATIENT PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Component or Accessory Incompatibility | Device problem code
6832547 | 84126472 | N | THE DENTIST REPORTED A B PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE BUT NO HARM CAUSED TO THE PATIENT PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | Results Pending Completion of Investigation | Medical Evaluation result code
6834115 | 84197331 | N | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING CRITICAL BATTERY ALARMS THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED CRITICAL BATTERY ALARM INVOLVING B IN THIS INSTANCE THE BATTERY WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY FAILURE ANALYSIS OF THE RETURNED BATTERY REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY THE MANUFACTURER HAS AN OPEN INTERNAL INVESTIGATION TO EVALUATE THE COMMUNICATION ANOMALIES BETWEEN THE CONTROLLER AND THE BATTERIES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Battery Problem | Device problem code
6834115 | 84197331 | N | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING CRITICAL BATTERY ALARMS THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED CRITICAL BATTERY ALARM INVOLVING B IN THIS INSTANCE THE BATTERY WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY FAILURE ANALYSIS OF THE RETURNED BATTERY REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY THE MANUFACTURER HAS AN OPEN INTERNAL INVESTIGATION TO EVALUATE THE COMMUNICATION ANOMALIES BETWEEN THE CONTROLLER AND THE BATTERIES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Communication or Transmission Problem | Device problem code
6834115 | 84197331 | N | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING CRITICAL BATTERY ALARMS THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED CRITICAL BATTERY ALARM INVOLVING B IN THIS INSTANCE THE BATTERY WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY FAILURE ANALYSIS OF THE RETURNED BATTERY REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY THE MANUFACTURER HAS AN OPEN INTERNAL INVESTIGATION TO EVALUATE THE COMMUNICATION ANOMALIES BETWEEN THE CONTROLLER AND THE BATTERIES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Communications Problem Identified | Medical Evaluation result code
6834115 | 84197331 | N | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING CRITICAL BATTERY ALARMS THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED CRITICAL BATTERY ALARM INVOLVING B IN THIS INSTANCE THE BATTERY WENT FROM A HIGH RELATIVE STATE OF CHARGE RSOC TO RSOC AND BACK TO PREVIOUS RSOC VALUES THIS OBSERVATION IS INDICATIVE OF A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY FAILURE ANALYSIS OF THE RETURNED BATTERY REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERY THE MANUFACTURER HAS AN OPEN INTERNAL INVESTIGATION TO EVALUATE THE COMMUNICATION ANOMALIES BETWEEN THE CONTROLLER AND THE BATTERIES THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Energy Storage System Problem | Medical Evaluation result code
6834718 | 84243228 | N | THE DENTIST REPORTED A PATIENT WHO WAS A SMOKER HAD IMPLANT LOSS OF OSSEOINTEGRATE WITH INFECTION AND BONE LOST OCCURRED AT THE IMPLANT SITE THE IMPLANT WAS REMOVED AND NO HARM WAS CAUSED TO THE PATIENT NO PATIENTS RACE WAS AVAILABLE THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES LOSS OF OSSEOINTEGRATION IS CAUSED BY MULTIPLE FACTORS INCLUDING OCCLUSAL OVERLOAD AND PERIMPLANTITIS THE BONE VOLUME AND QUANTITY ALSO PLAY A ROLE IN OSSEOINTEGRATION AND IMPLANTS IN GRAFTED SITES ARE MORE LIKELY TO FAIL COMPARED TO THOSE PLACED IN NATURAL BONE THE PATIENT WAS A SMOKER AND SMOKING MAY HAVE CONTRIBUTED TO THE LATE FAILURE OF THE IMPLANT ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR BONE QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION | No Device Problem Found | Medical Evaluation result code
6835290 | 84258114 | N | THE REPORTED EVENTS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | No Device Problem Found | Medical Evaluation result code
6824738 | 84258631 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT TREATED ELEVATED BG LEVELS WITH AN INJECTION SHOT OF INSULIN AND PUTTING ON A NEW POD THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | Device Dislodged or Dislocated | Device problem code
6824738 | 84258631 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT TREATED ELEVATED BG LEVELS WITH AN INJECTION SHOT OF INSULIN AND PUTTING ON A NEW POD THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | Device Fell | Device Problem Code
6824738 | 84258631 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT TREATED ELEVATED BG LEVELS WITH AN INJECTION SHOT OF INSULIN AND PUTTING ON A NEW POD THE ADHESIVE PAD FAILED WHICH RESULTED IN THE POD FALLING OFF AND CANNULA DISLODGING FROM THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6834947 | 84259066 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6834947 | 84259066 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6835334 | 84262188 | N | THE REPORTED EVENTS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | No Device Problem Found | Medical Evaluation result code
6835387 | 84262427 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT HAD A STROKE ON B THE PATIENT WAS DISCHARGED TO A REHABILITATION CENTER THE PATIENTS FAMILY DECIDED TO WITHDRAW CARE THE PATIENT EXPIRED ON B | Cerebrovascular Accident | Patient problem code
6835387 | 84262427 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT HAD A STROKE ON B THE PATIENT WAS DISCHARGED TO A REHABILITATION CENTER THE PATIENTS FAMILY DECIDED TO WITHDRAW CARE THE PATIENT EXPIRED ON B | Death | Patient problem code
6835387 | 84262427 | D | THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE LVAD ON B IT WAS REPORTED THAT THE PATIENT HAD A STROKE ON B THE PATIENT WAS DISCHARGED TO A REHABILITATION CENTER THE PATIENTS FAMILY DECIDED TO WITHDRAW CARE THE PATIENT EXPIRED ON B | Insufficient Information | Device problem code
6835655 | 84267877 | D | IT WAS REPORTED THE PATIENT WAS EXPERIENCING A BURNING PAIN AT THE IPG INCISION SITE THERE WAS A SIZABLE LUMP OVER THE IPG SITE THE PHYSICIAN ASPIRATED AND DRAINED MORE THAN CC OF BLOOD THE PATIENT IS BEING TREATED WITH ANTIBIOTICS | Insufficient Information | Device problem code
6835655 | 84267877 | D | IT WAS REPORTED THE PATIENT WAS EXPERIENCING A BURNING PAIN AT THE IPG INCISION SITE THERE WAS A SIZABLE LUMP OVER THE IPG SITE THE PHYSICIAN ASPIRATED AND DRAINED MORE THAN CC OF BLOOD THE PATIENT IS BEING TREATED WITH ANTIBIOTICS | Pain | Patient problem code
6835655 | 84267877 | D | IT WAS REPORTED THE PATIENT WAS EXPERIENCING A BURNING PAIN AT THE IPG INCISION SITE THERE WAS A SIZABLE LUMP OVER THE IPG SITE THE PHYSICIAN ASPIRATED AND DRAINED MORE THAN CC OF BLOOD THE PATIENT IS BEING TREATED WITH ANTIBIOTICS | Burning Sensation | Patient problem code
6836416 | 84301256 | N | THE DENTIST REPORTED A PATIENT EXPERIENCED IMPLANT LOSS OF OSSEOINTEGRATE WITHIN WEEKS OF IMPLANT PLACEMENT AND HAD BONE LOST AND INFECTION AT THE IMPLANT SITE NO HARM CAUSED TO THE PATIENT THE RETURNED PRODUCT AND DHR WERE REVIEWED NO DEFECT AND NONCONFORMITY WERE OBSERVED WITH THE PRODUCT THE DHR SHOWED NO DISCREPANCIES OR RELATED ISSUES PUBLISHED STUDIES SHOW THAT FAILURE TO OSSEOINTEGRATE OCCURS IN TO OF ALL CASES AND IS CONSIDERED AN INHERENT RISK IN IMPLANT SURGERY ALTHOUGH THE ROOT CAUSE FOR THE FAILED IMPLANT COMPLAINT IS INCONCLUSIVE AND SPECIFIC TO EACH CASE PROBABLE CAUSES COULD BE THE LOSS OF PRIMARY STABILITY AT THE OSTEOTOMY SITE DUE TO INSUFFICIENT BONE OR POOR BONE QUALITY EITHER THE BONE WAS TOO SOFT OR THE OPERATOR ERRED IN CREATING AN OSTEOTOMY BIGGER THAN THE SIZE OF THE IMPLANT DIAMETER PREMATURE LOADING PATIENTS HEALTH PERIIMPLANTITIS SMOKING AND LACK OF ORAL HYGIENE MAY ALSO BE CONTRIBUTING FACTORS TO THE LACK OF OSSEOINTEGRATION HOWEVER ESTABLISHED BIOCOMPATIBILITY STUDIES HAVE SHOWN THAT IMPLANT MATERIALS OR MANUFACTURING TECHNIQUES ARE NOT CAUSE FOR IMPLANT FAILURE OR INFECTION CURRENTLY NO PATIENTS AGE WEIGHT AND ETHNICITYRACE WERE GIVEN IF INFORMATION ON THE PATIENT BECOMES AVAILABLE A FOLLOWUP REPORT WILL BE PROVIDED | No Device Problem Found | Medical Evaluation result code
6836635 | 84311008 | N | A TEMPORAL ASSOCIATION EXISTS BETWEEN PD THERAPY WITH FRESENIUS PRODUCTS AND THE PATIENT EXPERIENCING DRAIN COMPLICATIONS ABDOMINAL PAIN CLOUDY PD EFFLUENT AND SUBSEQUENT DIAGNOSIS OF ESCHERICHIA COLI PERITONITIS WITHOUT REPORTED HOSPITALIZATION AT THE TIME OF THIS CLINICAL INVESTIGATION THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE THAT INDICATES A CAUSAL RELATIONSHIP BETWEEN FRESENIUS PRODUCTS AND THE PATIENTS PERITONITIS EVENT THE ETIOLOGY OF PERITONITIS IS LIKELY RELATED TO A BREACH IN ASEPTIC TECHNIQUE DURING CCPD TREATMENT AT HOME REUSING CCPD TREATMENT SUPPLIES AS REPORTED BY THE PD NURSE PD THERAPY WAS CONTINUING WITHOUT REPORTED FURTHER ISSUES AT THE TIME OF THIS REPORT A FOLLOW UP WILL BE SUBMITTED FOLLOWING EVALUATION | No Findings Available | Medical Evaluation result code
6836394 | 84314284 | N | ALLERGAN IS UNABLE TO CONFIRM WITH THE HEALTHCARE PROFESSIONAL THEREFORE ADDITIONAL EVENT PRODUCT OR PATIENT DETAILS ARE NOT ATTAINABLE THE EVENTS OF CYST TUMOR MUCOCELE AND PAIN ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | No Device Problem Found | Medical Evaluation result code
6832025 | 84317660 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE SERVICE TERRITORY MANAGER STM EVALUATED THE IABP FOR A HELIUM LEAK THE STM PERFORMED THE DESIGNATED TEST PER THE SERVICE MANUAL AND ISOLATED THE LEAK TO INSIDE THE IABP NOT THE CART THEN DISASSEMBLED THE IABP AND TRACED THE HELIUM LINE UP TO THE FILL MANIFOLD THE STM TIGHTENED DOWN THE FITTINGS ON THE LINE AND STILL DETECTED A LEAK HE THEN REEVALUATED THE IABP AND FOUND THAT THE PLUG ON THE BACK OF THE FILL MANIFOLD WAS LOOSE UPON TIGHTENING THE PLUG THE HELIUM PRESSURE STOPPED DROPPING THE STM WATCHED THE PRESSURE FOR ANOTHER MINUTES AND OBSERVED NO HELIUM LOSS THE STM PUT THE IABP BACK TOGETHER THE IABP THEN PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Device Problem Found | Medical Evaluation result code
6837504 | 84317761 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTEDREBOOTED ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6837504 | 84317761 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTEDREBOOTED ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6837498 | 84317799 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RESTARTREBOOT ITSELF MULTIPLE TIMES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6837498 | 84317799 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RESTARTREBOOT ITSELF MULTIPLE TIMES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6837468 | 84319200 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT COULDNT EMPTY THEIR BLADDER LAST NIGHT HAD BLADDER SPASMS AND FELT A BURNING SENSATION THEY WERE ONLY ABLE TO SELF VOID ON WARM WATER BUT HAD A HORRIBLE NIGHT THE PATIENT STATES THAT THEY ARENT ABLE TO HAVE A STREAM AND THAT IS NEW FOR THEM THEY SPOKE WITH THE SALES REP AND THEY WERE ADVISED TO STAY ON PROGRAM AT V THE NEXT DAY PATIENT SAID THEY HAVE ABSOLUTELY NO STREAM ANYMORE WHEN THEY VOID THEY ONLY DRIBBLE IN VERY SMALL AMOUNTS THE PATIENT SAID THINGS ARE WORSE THAN PRIOR TO EVALUATION BECAUSE THEY NEVER REALLY HAD TO CATHERIZE MUCH BEFORE SOMETIMES THE PATIENT FEELS LIKE THEY NEED TO VOID BUT IS UNABLE TO ON B PATIENT SLEPT FROM PM TO AM AND THAT WAS THE BEST NIGHT OF THEIR LIFE THEY TOOK SOMETHING TO HELP THEM SLEEP AND THEY ARENT SURE IF THIS HAD ANYTHING TO DO WITH NOT WAKING UP LAST NIGHT THE PATIENT CATHED ONCE CCS SEVEN DAYS LATER PATIENT SAID THE ENS BATTERIES ARE GOING DEAD FOR THE SECOND TIME THEY ARE ONLY LASTING A TOTAL OF ONE WEEK THE PATIENT HAS A FOLLOW UP ON WEDNESDAY AND THE PATIENT SAID THEY WONT LAST UNTIL THEN AS THEY ARE ALREADY ON RED THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Insufficient Information | Device problem code
6837468 | 84319200 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT COULDNT EMPTY THEIR BLADDER LAST NIGHT HAD BLADDER SPASMS AND FELT A BURNING SENSATION THEY WERE ONLY ABLE TO SELF VOID ON WARM WATER BUT HAD A HORRIBLE NIGHT THE PATIENT STATES THAT THEY ARENT ABLE TO HAVE A STREAM AND THAT IS NEW FOR THEM THEY SPOKE WITH THE SALES REP AND THEY WERE ADVISED TO STAY ON PROGRAM AT V THE NEXT DAY PATIENT SAID THEY HAVE ABSOLUTELY NO STREAM ANYMORE WHEN THEY VOID THEY ONLY DRIBBLE IN VERY SMALL AMOUNTS THE PATIENT SAID THINGS ARE WORSE THAN PRIOR TO EVALUATION BECAUSE THEY NEVER REALLY HAD TO CATHERIZE MUCH BEFORE SOMETIMES THE PATIENT FEELS LIKE THEY NEED TO VOID BUT IS UNABLE TO ON B PATIENT SLEPT FROM PM TO AM AND THAT WAS THE BEST NIGHT OF THEIR LIFE THEY TOOK SOMETHING TO HELP THEM SLEEP AND THEY ARENT SURE IF THIS HAD ANYTHING TO DO WITH NOT WAKING UP LAST NIGHT THE PATIENT CATHED ONCE CCS SEVEN DAYS LATER PATIENT SAID THE ENS BATTERIES ARE GOING DEAD FOR THE SECOND TIME THEY ARE ONLY LASTING A TOTAL OF ONE WEEK THE PATIENT HAS A FOLLOW UP ON WEDNESDAY AND THE PATIENT SAID THEY WONT LAST UNTIL THEN AS THEY ARE ALREADY ON RED THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Muscle Spasm(s) | Patient problem code
6837468 | 84319200 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT COULDNT EMPTY THEIR BLADDER LAST NIGHT HAD BLADDER SPASMS AND FELT A BURNING SENSATION THEY WERE ONLY ABLE TO SELF VOID ON WARM WATER BUT HAD A HORRIBLE NIGHT THE PATIENT STATES THAT THEY ARENT ABLE TO HAVE A STREAM AND THAT IS NEW FOR THEM THEY SPOKE WITH THE SALES REP AND THEY WERE ADVISED TO STAY ON PROGRAM AT V THE NEXT DAY PATIENT SAID THEY HAVE ABSOLUTELY NO STREAM ANYMORE WHEN THEY VOID THEY ONLY DRIBBLE IN VERY SMALL AMOUNTS THE PATIENT SAID THINGS ARE WORSE THAN PRIOR TO EVALUATION BECAUSE THEY NEVER REALLY HAD TO CATHERIZE MUCH BEFORE SOMETIMES THE PATIENT FEELS LIKE THEY NEED TO VOID BUT IS UNABLE TO ON B PATIENT SLEPT FROM PM TO AM AND THAT WAS THE BEST NIGHT OF THEIR LIFE THEY TOOK SOMETHING TO HELP THEM SLEEP AND THEY ARENT SURE IF THIS HAD ANYTHING TO DO WITH NOT WAKING UP LAST NIGHT THE PATIENT CATHED ONCE CCS SEVEN DAYS LATER PATIENT SAID THE ENS BATTERIES ARE GOING DEAD FOR THE SECOND TIME THEY ARE ONLY LASTING A TOTAL OF ONE WEEK THE PATIENT HAS A FOLLOW UP ON WEDNESDAY AND THE PATIENT SAID THEY WONT LAST UNTIL THEN AS THEY ARE ALREADY ON RED THE EVENT WAS STAMPED AS SALES ATTENTION NEEDED NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Burning Sensation | Patient problem code
6836756 | 84323478 | N | B A UNIQUE DEVICE IDENTIFIERS WAS NOT KNOWN AS THE LOT NUMBERS WERE NOT AVAILABLE B USERS ARE MADE AWARE OF THE RISKS ASSOCIATED WITH TYPE II ENDOLEAKS IN THE IFU AND ARE INSTRUCTED TO CONSIDER THE RISKS AND BENEFITS DISCUSSED IN THE IFU FOR EACH PATIENT BEFORE USING THE DEVICES | Insufficient Information | Device problem code
6838097 | 84351224 | N | B PER IFU USERS ARE MADE AWARE OF THE RISKS ASSOCIATED WITH TYPE II ENDOLEAKS AND ARE INSTRUCTED TO CONSIDER THE RISKS AND BENEFITS DISCUSSED IN THE IFU FOR EACH PATIENT BEFORE USING THE DEVICES | Insufficient Information | Device problem code
6836121 | 84360908 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER B VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN ETERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN EXTERNAL POWER SOURCE THE CONTROLLER WAS ABLE TO RETAIN THE DATE AND TIME AFTER ALLOWING THE REAL TIME CLOCK TO FULLY CHARGE PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS AND ADAPTERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6836121 | 84360908 | N | THE REPORTED EVENT OF A REAL TIME CLOCK RESET WAS CONFIRMED ON THE RETURNED CONTROLLER B VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO THIS CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK THIS IS MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN ETERNAL POWER SOURCE HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A DEPLETED REAL TIME CLOCK MOST LIKELY DUE TO AN EXTENDED PERIOD OF TIME WHERE THE CONTROLLER WAS NOT CONNECTED TO AN EXTERNAL POWER SOURCE THE CONTROLLER WAS ABLE TO RETAIN THE DATE AND TIME AFTER ALLOWING THE REAL TIME CLOCK TO FULLY CHARGE PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS AND ADAPTERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6838359 | 84366136 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS AND HALF LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6838359 | 84366136 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS AND HALF LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6838732 | 84376758 | N | THE RETURNED PUMP WAS EVALUATED FOR EVIDENCE OF ALLEGATION THE RETURNED AMEDA PURELY YOURS BREAST PUMP MET AMEDA SPECIFICATIONS FOR BOTH SUCTION AND SPEED AND PASSED VISUAL INSPECTION STANDARDS NO EVIDENCE OF MALFUNCTION WAS OBSERVED | No Device Problem Found | Medical Evaluation result code
6836558 | 84391224 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6836558 | 84391224 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6841111 | 84619021 | N | B ATTEMPT HAS BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED PARENT INFORMATION | Insufficient Information | Device problem code
6841102 | 84619992 | N | ATTEMPT HAS BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | Insufficient Information | Device problem code
6841271 | 84623184 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B HOSPITAL FACILITY AND WAS TREATED WITH HUMALOG INSULIN THROUGH IV | Insufficient Information | Device problem code
6841271 | 84623184 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B HOSPITAL FACILITY AND WAS TREATED WITH HUMALOG INSULIN THROUGH IV | No Known Impact Or Consequence To Patient | Patient Problem Code
6841279 | 84624153 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B DUE TO HIGH BG LEVELS AND EXPERIENCING SYMPTOMS OF VOMITING PATIENT WAS TREATED WITH INSULIN THROUGH IV AND WAS GIVEN LIQUID POTASSIUM IN ADDITION PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Hyperglycemia | Patient problem code
6841279 | 84624153 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B DUE TO HIGH BG LEVELS AND EXPERIENCING SYMPTOMS OF VOMITING PATIENT WAS TREATED WITH INSULIN THROUGH IV AND WAS GIVEN LIQUID POTASSIUM IN ADDITION PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6841279 | 84624153 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B DUE TO HIGH BG LEVELS AND EXPERIENCING SYMPTOMS OF VOMITING PATIENT WAS TREATED WITH INSULIN THROUGH IV AND WAS GIVEN LIQUID POTASSIUM IN ADDITION PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Vomiting | Patient problem code
6841116 | 84629359 | N | VYAIRE MEDICAL COMPLAINT NUMBER B RESULTS OF INVESTIGATION THE VYAIRE MEDICAL FACTORY SERVICE CENTER RECEIVED THE SUSPECT DEVICE A SIPAP VENTILATOR AND EVALUATED THE DEVICE AN EVALUATION OF THE DEVICE CONFIRMED THE REPORTED ISSUE AND ISOLATED THE ISSUE TO THE BLENDER BEING OUT OF CALIBRATION | Calibration Problem | Device problem code
6841116 | 84629359 | N | VYAIRE MEDICAL COMPLAINT NUMBER B RESULTS OF INVESTIGATION THE VYAIRE MEDICAL FACTORY SERVICE CENTER RECEIVED THE SUSPECT DEVICE A SIPAP VENTILATOR AND EVALUATED THE DEVICE AN EVALUATION OF THE DEVICE CONFIRMED THE REPORTED ISSUE AND ISOLATED THE ISSUE TO THE BLENDER BEING OUT OF CALIBRATION | Failure to Calibrate | Device Problem Code
6841508 | 84640054 | N | THE EVENTS OF LYMPHOMA AND SEROMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL SEROMA POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY AFTER BREAST IMPLANT SURGERY THE FOLLOWING MAY OCCUR ANDOR PERSIST WITH VARYING INTENSITY ANDOR FOR A VARYING LENGTH OF TIME HEMATOMASEROMA IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING AND BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS CAN FORM IN THE TISSUE CAPSULE SURROUNDING THE IMPLANT WITH SYMPTOMS THAT MAY INCLUDE PAIN AND FIRMNESS LYMPHADENOPATHY HAS ALSO BEEN REPORTED IN SOME WOMEN WITH IMPLANTS | No Findings Available | Medical Evaluation result code
6841577 | 84641614 | D | IT WAS REPORTED THAT A PATIENT RECEIVED A SKIN IRRITATION AFTER WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT ADVISED THIS TO BE AN ONGOING ISSUE SUBSEQUENTLY A PHYSICIAN PRESCRIBED TRIAMCINOLONE CREAM AS TREATMENT NO SPECIFIC DIAGNOSIS WAS PROVIDED | Skin Irritation | Patient problem code
6841577 | 84641614 | D | IT WAS REPORTED THAT A PATIENT RECEIVED A SKIN IRRITATION AFTER WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT ADVISED THIS TO BE AN ONGOING ISSUE SUBSEQUENTLY A PHYSICIAN PRESCRIBED TRIAMCINOLONE CREAM AS TREATMENT NO SPECIFIC DIAGNOSIS WAS PROVIDED | Insufficient Information | Device problem code
6841577 | 84641614 | D | IT WAS REPORTED THAT A PATIENT RECEIVED A SKIN IRRITATION AFTER WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT ADVISED THIS TO BE AN ONGOING ISSUE SUBSEQUENTLY A PHYSICIAN PRESCRIBED TRIAMCINOLONE CREAM AS TREATMENT NO SPECIFIC DIAGNOSIS WAS PROVIDED | Irritation | Patient problem code
6836807 | 84642987 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6836807 | 84642987 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6837703 | 84643145 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Hyperglycemia | Patient Problem Code
6837703 | 84643145 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Insufficient Information | Device problem code
6837461 | 84645318 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS POD WAS NOT STICKING WELL TO THE SKIN AND THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6837461 | 84645318 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS POD WAS NOT STICKING WELL TO THE SKIN AND THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6839252 | 84647698 | N | B CAPA THE EVENT OF VISUAL IMPAIRMENT IS EXPECTED NO CORRECTIVE OR PREVENTIVE ACTION IS NEEDED MANUFACTURING NARRATIVE LOT NUMBER WAS NOT REPORTED PHARMACOVIGILANCE COMMENTS THE SERIOUS EXPECTED EVENT OF VISUAL IMPAIRMENT WAS CONSIDERED POSSIBLY RELATED TO THE TREATMENT PROCEDURE SERIOUS CRITERIA INCLUDE POTENTIAL DISABILITY POTENTIAL CONTRIBUTORY FACTOR TO THE EVENT INCLUDE INJECTION TECHNIQUE THIS CASE MEETS THE CRITERIA OF SERIOUSNESS AND CAUSALITY FOR EXPEDITED REPORTING TO THE REGULATORY AUTHORITIES FURTHER INFORMATION HAS BEEN REQUESTED ADDITIONAL COMMENTSNO INFORMATION WAS MADE AVAILABLE FOR DEVICE EVALUATION BY THE MANUFACTURER | Insufficient Information | Device problem code
6839252 | 84647698 | N | B CAPA THE EVENT OF VISUAL IMPAIRMENT IS EXPECTED NO CORRECTIVE OR PREVENTIVE ACTION IS NEEDED MANUFACTURING NARRATIVE LOT NUMBER WAS NOT REPORTED PHARMACOVIGILANCE COMMENTS THE SERIOUS EXPECTED EVENT OF VISUAL IMPAIRMENT WAS CONSIDERED POSSIBLY RELATED TO THE TREATMENT PROCEDURE SERIOUS CRITERIA INCLUDE POTENTIAL DISABILITY POTENTIAL CONTRIBUTORY FACTOR TO THE EVENT INCLUDE INJECTION TECHNIQUE THIS CASE MEETS THE CRITERIA OF SERIOUSNESS AND CAUSALITY FOR EXPEDITED REPORTING TO THE REGULATORY AUTHORITIES FURTHER INFORMATION HAS BEEN REQUESTED ADDITIONAL COMMENTSNO INFORMATION WAS MADE AVAILABLE FOR DEVICE EVALUATION BY THE MANUFACTURER | Results Pending Completion of Investigation | Medical Evaluation result code
6841747 | 84648258 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURN MARKS ON THE AC POWER ADAPTER AND CIRCUIT BOARD WHICH CAUSED THE REPORTED INABILITY TO POWER THE TRANSMITTER | Power Problem | Device problem code
6841747 | 84648258 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURN MARKS ON THE AC POWER ADAPTER AND CIRCUIT BOARD WHICH CAUSED THE REPORTED INABILITY TO POWER THE TRANSMITTER | Power Source Problem Identified | Medical Evaluation result code
6838407 | 84648713 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6838407 | 84648713 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6835331 | 84648906 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED AS VISUALLY NEGATIVE | No Device Problem Found | Medical Evaluation result code
6841812 | 84653516 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | No Findings Available | Medical Evaluation result code
6838374 | 84655375 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6838374 | 84655375 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6837628 | 84655525 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6837628 | 84655525 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | Insufficient Information | Device problem code
6837368 | 84660465 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing Information | Device problem code
6837368 | 84660465 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6842024 | 84708069 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT STARTED FEELING HIGH THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6842024 | 84708069 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT STARTED FEELING HIGH THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Hyperglycemia | Patient Problem Code
6841653 | 84764178 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Insufficient Information | Device problem code
6841653 | 84764178 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6829789 | 84772465 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WAS FIRED IN THE USUAL MANNER BUT DID NOT CRUNCH AS USUAL SURGEON WAS ONLY ABLE TO FIND ONE DONUT WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT | No Device Problem Found | Medical Evaluation result code
6832550 | 84805708 | N | A SPACELABS FSE ARRIVED ONSITE FOR FURTHER INVESTIGATION FINDINGS SHOW THAT THE CONNECTION WAS LOST BETWEEN THE CENTRAL MONITOR AND TELEMETRY RECEIVER THE CUSTOMER BIOMED ELECTED TO REPLACE THE SIXTEEN YEAR OLD DEVICE WITH A SPARE SPACELABS TECHNICAL SUPPORTED ASSISTED THE BIOMED WITH CONFIGURING THE NETWORK SETTINGS FOLLOWING INSTALLATION THERE HAVE BEEN NO REPORTS OF RECURRENCE THIS REPORT IS CONSIDERED COMPLETE AND THE MATTER CLOSED | Connection Problem | Device problem code
6832550 | 84805708 | N | A SPACELABS FSE ARRIVED ONSITE FOR FURTHER INVESTIGATION FINDINGS SHOW THAT THE CONNECTION WAS LOST BETWEEN THE CENTRAL MONITOR AND TELEMETRY RECEIVER THE CUSTOMER BIOMED ELECTED TO REPLACE THE SIXTEEN YEAR OLD DEVICE WITH A SPARE SPACELABS TECHNICAL SUPPORTED ASSISTED THE BIOMED WITH CONFIGURING THE NETWORK SETTINGS FOLLOWING INSTALLATION THERE HAVE BEEN NO REPORTS OF RECURRENCE THIS REPORT IS CONSIDERED COMPLETE AND THE MATTER CLOSED | Telemetry Discrepancy | Device problem code
6832550 | 84805708 | N | A SPACELABS FSE ARRIVED ONSITE FOR FURTHER INVESTIGATION FINDINGS SHOW THAT THE CONNECTION WAS LOST BETWEEN THE CENTRAL MONITOR AND TELEMETRY RECEIVER THE CUSTOMER BIOMED ELECTED TO REPLACE THE SIXTEEN YEAR OLD DEVICE WITH A SPARE SPACELABS TECHNICAL SUPPORTED ASSISTED THE BIOMED WITH CONFIGURING THE NETWORK SETTINGS FOLLOWING INSTALLATION THERE HAVE BEEN NO REPORTS OF RECURRENCE THIS REPORT IS CONSIDERED COMPLETE AND THE MATTER CLOSED | Communications Problem Identified | Medical Evaluation result code
6836817 | 84848017 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6836817 | 84848017 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6842366 | 84848085 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6842366 | 84848085 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6842569 | 84855720 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS EXPERIENCING PAIN IN THEIR RIGHT BUTTOCKS WHEN THEY MOVE AND IT IS HARD TO MOVE AROUND IT WAS SUGGESTED TO DECREASE STIMULATION TO SEE IF THAT ALLEVIATES THE PAIN TWO DAYS LATER PATIENT SAID THEY WERE IN THE HOSPITAL AS A RESULT OF EXCESSIVE BLEEDING THEY WERE STILL IN THE HOSPITAL A DAY LATER AND WAS COMPLAINING OF PAIN IN THEIR BUTTOCKS ON B PATIENT MADE ANOTHER TRIP TO THE ER BECAUSE OF INCREASING PAIN IN THEIR BUTTOCK THEY THOUGHT THEY WERE BLEEDING INTERNALLY BUT A CAT SCAN SHOWED EVERYTHING WAS FINE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6842569 | 84855720 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS EXPERIENCING PAIN IN THEIR RIGHT BUTTOCKS WHEN THEY MOVE AND IT IS HARD TO MOVE AROUND IT WAS SUGGESTED TO DECREASE STIMULATION TO SEE IF THAT ALLEVIATES THE PAIN TWO DAYS LATER PATIENT SAID THEY WERE IN THE HOSPITAL AS A RESULT OF EXCESSIVE BLEEDING THEY WERE STILL IN THE HOSPITAL A DAY LATER AND WAS COMPLAINING OF PAIN IN THEIR BUTTOCKS ON B PATIENT MADE ANOTHER TRIP TO THE ER BECAUSE OF INCREASING PAIN IN THEIR BUTTOCK THEY THOUGHT THEY WERE BLEEDING INTERNALLY BUT A CAT SCAN SHOWED EVERYTHING WAS FINE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient Problem Code
6842733 | 84864331 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHERE WAS THE LEAK WHAT COLOR CARTRIDGES WERE USED AND THE ORDER WAS BUTTRESSING MATERIAL USED WERE ANY ISSUES NOTED WITH DEVICE PERFORMANCE DURING PROCEDURE HOW WAS THE LEAK IDENTIFIED WHAT IS THE CURRENT PATIENT STATUS ADDITIONAL INFORMATION RECEIVED THE SURGEON WANTED TO NOTE THAT THE PATIENT HAD BEEN VAPING DURING RECOVERY ASSUMING TO MEAN USING AN ECIGARETTE | Leak / Splash | Device problem code
6842733 | 84864331 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHERE WAS THE LEAK WHAT COLOR CARTRIDGES WERE USED AND THE ORDER WAS BUTTRESSING MATERIAL USED WERE ANY ISSUES NOTED WITH DEVICE PERFORMANCE DURING PROCEDURE HOW WAS THE LEAK IDENTIFIED WHAT IS THE CURRENT PATIENT STATUS ADDITIONAL INFORMATION RECEIVED THE SURGEON WANTED TO NOTE THAT THE PATIENT HAD BEEN VAPING DURING RECOVERY ASSUMING TO MEAN USING AN ECIGARETTE | Leakage/Seal | Medical Evaluation result code
6843551 | 84915458 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF GRANULOMAS BUMPS AND HARDENED ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | Material Discolored | Device problem code
6843551 | 84915458 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF GRANULOMAS BUMPS AND HARDENED ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION | No Findings Available | Medical Evaluation result code
6843722 | 84931996 | N | EXPLANT DATE IF EXPLANTED GIVE DATE NOT APPLICABLE TO DATE THERE IS NO INDICATION THAT THE LENS HAS BEEN EXPLANTED B ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Apparent Adverse Event | Device problem code
6843722 | 84931996 | N | EXPLANT DATE IF EXPLANTED GIVE DATE NOT APPLICABLE TO DATE THERE IS NO INDICATION THAT THE LENS HAS BEEN EXPLANTED B ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Findings Available | Medical Evaluation result code
6843632 | 84934802 | N | THE RETURNED PRODUCT WAS EVALUATED FOR EVIDENCE OF ALLEGATION THE RETURNED AMEDA PURELY YOURS BREAST PUMP MET AMEDA SPECIFICATIONS FOR BOTH SUCTION AND SPEED AND PASSED VISUAL INSPECTION STANDARDS NO EVIDENCE OF MALFUNCTION WAS OBSERVED BASED ON THE INVESTIGATION RESULTS IT CANNOT BE DEFINITELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS | No Device Problem Found | Medical Evaluation result code
6843180 | 84937167 | D | COMPLAINANT ALLEGED THAT THE DEVICE DID NOT RESPOND TO THE FRONT PANEL CONTROLS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6843180 | 84937167 | D | COMPLAINANT ALLEGED THAT THE DEVICE DID NOT RESPOND TO THE FRONT PANEL CONTROLS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Insufficient Information | Device Problem Code
6843283 | 84937380 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RESTARTREBOOT ITSELF MULTIPLE TIMES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6843283 | 84937380 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE RESTARTREBOOT ITSELF MULTIPLE TIMES COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6843703 | 84937563 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE HALTS AFTER PARTIAL BOOT CYCLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6843703 | 84937563 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE HALTS AFTER PARTIAL BOOT CYCLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6840674 | 84941682 | N | AS PART OF A REMEDIAL ACTIVITY LUMENIS CONDUCTED A RETROSPECTIVE REVIEW OF ALL ITS SAFETY COMPLAINT FILES FROM JANUARY TO JULY AN INVESTIGATION OF THE REPORTED EVENT FOUND THAT ALTHOUGH THERE WAS NO PATIENT INVOLVEMENT AND NO ACTUAL HARM HAD OCCURED THE MALFUNCTION MIGHT LEAD TO SERIOUS INJURY SHOULD IT RECUR THE FOOTSWITCH WAS RETURNED TO LUMENIS FOR EVALUATION WHERE A LUMENIS QUALITY ENGINEER INVESTIGATED THE FOOTSWITCH AND WAS ABLE TO DUPLICATE THE ISSUE THE FOOTSWITCH WOULD STICK WHEN RELEASING THE SWITCH SLOWLY THE ISSUE WAS DETERMINED TO BE A MECHANICAL MALFUNCTION THE MALFUNCTION WAS MOST PROBABLY CAUSED BY A MECHANICAL DIMENSIONAL INCONSISTENCY CAUSED BY A MANUFACTURER QUALITY ISSUE IN RESPONSE LUMENIS HAD OPENED A SUPPLIER CORRECTIVE ACTION REQUEST SCAR TO THE FOOTSWITCH MANUFACTURER B THE RISK HAS BEEN QUANTIFIED AND FOUND TO BE NEGLIGIBLY SMALL AND THE RISK HAS BEEN CHARACTERIZED AND DOCUMENTED AS ACCEPTABLE WITHIN FULL RISK ASSESSMENT NO CORRECTIVE ACTION OR REMEDIAL ACTIONS WERE DEEMED NECESSARY ALTHOUGH THE EVENT HAPPENED IN A CLOSED LAB ENVIRONMENT WITH NO PATIENT INVOLVEMENT THIS MALFUNCTION MIGHT LEAD TO SERIOUS INJURY SHOULD IT RECUR AND IN AN ABUNDANCE OF CAUTION LUMENIS IS REPORTING THIS MALFUNCTION | Product Quality Problem | Device problem code
6840674 | 84941682 | N | AS PART OF A REMEDIAL ACTIVITY LUMENIS CONDUCTED A RETROSPECTIVE REVIEW OF ALL ITS SAFETY COMPLAINT FILES FROM JANUARY TO JULY AN INVESTIGATION OF THE REPORTED EVENT FOUND THAT ALTHOUGH THERE WAS NO PATIENT INVOLVEMENT AND NO ACTUAL HARM HAD OCCURED THE MALFUNCTION MIGHT LEAD TO SERIOUS INJURY SHOULD IT RECUR THE FOOTSWITCH WAS RETURNED TO LUMENIS FOR EVALUATION WHERE A LUMENIS QUALITY ENGINEER INVESTIGATED THE FOOTSWITCH AND WAS ABLE TO DUPLICATE THE ISSUE THE FOOTSWITCH WOULD STICK WHEN RELEASING THE SWITCH SLOWLY THE ISSUE WAS DETERMINED TO BE A MECHANICAL MALFUNCTION THE MALFUNCTION WAS MOST PROBABLY CAUSED BY A MECHANICAL DIMENSIONAL INCONSISTENCY CAUSED BY A MANUFACTURER QUALITY ISSUE IN RESPONSE LUMENIS HAD OPENED A SUPPLIER CORRECTIVE ACTION REQUEST SCAR TO THE FOOTSWITCH MANUFACTURER B THE RISK HAS BEEN QUANTIFIED AND FOUND TO BE NEGLIGIBLY SMALL AND THE RISK HAS BEEN CHARACTERIZED AND DOCUMENTED AS ACCEPTABLE WITHIN FULL RISK ASSESSMENT NO CORRECTIVE ACTION OR REMEDIAL ACTIONS WERE DEEMED NECESSARY ALTHOUGH THE EVENT HAPPENED IN A CLOSED LAB ENVIRONMENT WITH NO PATIENT INVOLVEMENT THIS MALFUNCTION MIGHT LEAD TO SERIOUS INJURY SHOULD IT RECUR AND IN AN ABUNDANCE OF CAUTION LUMENIS IS REPORTING THIS MALFUNCTION | No Findings Available | Medical Evaluation result code
6827897 | 84952146 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION THE KNOWN C ANTIGEN POSITIVE TEST WELL COMBINED AS KNOWN E ANTIGEN POSITIVE TEST WELL ALSO APPEARED VISUALLY WEAK POSITIVE AND DIFFUSE RED BLOOD CELL ADHESION THE KNOWN E ANTIGEN POSITIVE TEST WELL COMBINED AS KNOWN C ANTIGEN NEGATIVE TEST WELL ALSO APPEARED VISUALLY WEAK POSITIVE AND DIFFUSE RED BLOOD CELL ADHESION | Connection Problem | Device problem code
6827897 | 84952146 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION THE KNOWN C ANTIGEN POSITIVE TEST WELL COMBINED AS KNOWN E ANTIGEN POSITIVE TEST WELL ALSO APPEARED VISUALLY WEAK POSITIVE AND DIFFUSE RED BLOOD CELL ADHESION THE KNOWN E ANTIGEN POSITIVE TEST WELL COMBINED AS KNOWN C ANTIGEN NEGATIVE TEST WELL ALSO APPEARED VISUALLY WEAK POSITIVE AND DIFFUSE RED BLOOD CELL ADHESION | No Device Problem Found | Medical Evaluation result code
6841210 | 84957245 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6841210 | 84957245 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS SWIMMING WHEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6841200 | 84957471 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6841200 | 84957471 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6841205 | 84957832 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE ADHESIVE FAILED AROUND THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED AND A CORRECTION BOLUS WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6841205 | 84957832 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE ADHESIVE FAILED AROUND THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED AND A CORRECTION BOLUS WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6841206 | 84958539 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM CORRECTION DOSES WERE ATTEMPTED WITH THE POD BUT THEY WERE UNSUCCESSFUL THE PATIENT NOTED THE ADHESIVE WAS NOT SECURE WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6841206 | 84958539 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM CORRECTION DOSES WERE ATTEMPTED WITH THE POD BUT THEY WERE UNSUCCESSFUL THE PATIENT NOTED THE ADHESIVE WAS NOT SECURE WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6838365 | 84966595 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6838365 | 84966595 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6844613 | 84968510 | D | DURING A SURGICAL PROCEDURE THE PATIENT EXPERIENCE A BURNING SENSATION FROM THE RENEW HANDPIECE | Burning Sensation | Patient problem code
6844613 | 84968510 | D | DURING A SURGICAL PROCEDURE THE PATIENT EXPERIENCE A BURNING SENSATION FROM THE RENEW HANDPIECE | Insufficient Information | Device problem code
6840957 | 84978792 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6840957 | 84978792 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6840958 | 84985771 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6840958 | 84985771 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6840961 | 84987647 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6840961 | 84987647 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6840981 | 84987685 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT NOTED THE BLUE CANNULA WAS NOT VISIBLE INDICATING A NEEDLE MECHANISM FAILURE | Defective Device | Device problem code
6840981 | 84987685 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT NOTED THE BLUE CANNULA WAS NOT VISIBLE INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6839546 | 84991152 | N | THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT OF WOULD NOT HOLD DATE AND TIME THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF LOG FILES WHICH REVEALED A SERIES OF POWER UP EVENTS WHERE THE TIME REVERTED TO AND IMPLIES THAT THE REAL TIME CLOCK LOST BATTERY POWER WHILE THE CONTROLLER WAS POWER OFF WHEN POWERED ON THE REAL TIME CLOCK WAS RESET BUT THE DEVICE WAS NOT POWERED ON LONG ENOUGH TO RECHARGE THE REAL TIME CLOCK BATTERY ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE SYSTEM DID NOT INDICATE ANY ABNORMAL OPERATION OF THE CONTROLLER WHEN THE INTERNAL BATTERY WAS ADEQUATELY RECHARGED THE CONTROLLER WAS ABLE TO KEEP ACCURATE DATE AND TIME THE REAL TIME CLOCK RTC WAS MOST LIKELY RESET TO ZERO BECAUSE THE CONTROLLER HAD NOT BEEN CONNECTED TO EXTERNAL POWER FOR EXTENDED PERIODS AND ITS INTERNAL COIN CELL BATTERY HAD RUN DOWN PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Insufficient Information | Device problem code
6845215 | 84993211 | D | OUS MDR AN EXIT BLOCK WAS DIAGNOSED OF THE LV LEAD BY CRT CONTROL THE PATIENT WAS ASYMPTOMATIC THIS LV LEAD WAS REPOSITIONED SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS FILE WILL BE UPDATED | Positioning Problem | Device problem code
6845215 | 84993211 | D | OUS MDR AN EXIT BLOCK WAS DIAGNOSED OF THE LV LEAD BY CRT CONTROL THE PATIENT WAS ASYMPTOMATIC THIS LV LEAD WAS REPOSITIONED SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS FILE WILL BE UPDATED | Exit Block | Patient problem code
6841371 | 84994516 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR LONGER THAN HOURS TREATMENT INCLUDED A NEW POD THE DEVICE WAS RETURNED FOR EVALUATION | Insufficient Information | Device problem code
6841371 | 84994516 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR LONGER THAN HOURS TREATMENT INCLUDED A NEW POD THE DEVICE WAS RETURNED FOR EVALUATION | No Known Impact Or Consequence To Patient | Patient Problem Code
6845077 | 84998071 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AT MULTIPLE SITES AND THERE WAS GENERAL BONE LOSS HOWEVER NO HARM WAS CAUSED TO THE PATIENT NO PATIENTS ETHNICITY AND RACE WERE PROVIDED PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | No Findings Available | Medical Evaluation result code
6845337 | 84998182 | N | THE DENTIST REPORTED A PATIENT HAD IMPLANT FAILURE TO OSSEOINTEGRATE AT MULTIPLE SITES AND THERE WAS GENERAL BONE LOSS HOWEVER NO HARM WAS CAUSED TO THE PATIENT NO PATIENTS ETHNICITY AND RACE WERE PROVIDED PRODUCT INVESTIGATION AND DHR REVIEW ARE PENDING A FOLLOWUP REPORT WILL BE COMPLETED WHEN THE RESULTS BECOME AVAILABLE | No Findings Available | Medical Evaluation result code
6840539 | 84998227 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT THEIR ONETOUCH VERIO IQ METER WAS READING INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED INACCURACY OCCURRED AT AM ON B AT THIS TIME THE PATIENT OBTAINED BLOOD GLUCOSE RESULTS OF AND MGDL WITH THE SUBJECT METER WITHIN MINUTES OF ONE ANOTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE OF THESE GLUCOSE RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES THEIR DIABETES WITH AN UN SPECIFIED DOSAGE OF METFORMIN AND TOOK AN INCREASED DOSAGE OF THIS MEDICATION DOSAGE UNSPECIFIED IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT DID NOT DEVELOP ANY PHYSICAL SYMPTOMS OR REQUIRE ANY FURTHER TREATMENT AS A RESULT OF THIS AT THE TIME OF TROUBLESHOOTING THE CCA NOTED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER AND AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE PATIENT DID NOT HAVE CONTROL SOLUTION AVAILABLE TO WALK THROUGH A CONTROL SOLUTION TEST THE PATIENTS PRODUCTS WERE REPLACED BASED ON THE INFORMATION PROVIDED THERE IS NO INDICATION THAT THE SUBJECT METER CAUSED OR CONTRIBUTED TO A SERIOUS INJURY ADVERSE EVENT THE PATIENT DID NOT DEVELOP SIGNSSYMPTOMS SUGGESTIVE OF A SERIOUS INJURY NOR DID THEY RECEIVE MEDICAL INTERVENTION FOR AN ACUTE BLOOD GLUCOSE EXCURSION HOWEVER THIS COMPLAINT IS BEING REPORTED AS A MALFUNCTION BECAUSE THE CALCULATED DIFFERENCE BETWEEN THE REPORTED RESULTS FALLS OUTSIDE LFS PRECISION CRITERIA AND THE ALLEGED INACCURACY WAS NOT RESOLVED DURING TROUBLESHOOTING | Incorrect, Inadequate or Imprecise Result or Readings | Device Problem Code
6840539 | 84998227 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT THEIR ONETOUCH VERIO IQ METER WAS READING INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED INACCURACY OCCURRED AT AM ON B AT THIS TIME THE PATIENT OBTAINED BLOOD GLUCOSE RESULTS OF AND MGDL WITH THE SUBJECT METER WITHIN MINUTES OF ONE ANOTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE OF THESE GLUCOSE RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES THEIR DIABETES WITH AN UN SPECIFIED DOSAGE OF METFORMIN AND TOOK AN INCREASED DOSAGE OF THIS MEDICATION DOSAGE UNSPECIFIED IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT DID NOT DEVELOP ANY PHYSICAL SYMPTOMS OR REQUIRE ANY FURTHER TREATMENT AS A RESULT OF THIS AT THE TIME OF TROUBLESHOOTING THE CCA NOTED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER AND AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE PATIENT DID NOT HAVE CONTROL SOLUTION AVAILABLE TO WALK THROUGH A CONTROL SOLUTION TEST THE PATIENTS PRODUCTS WERE REPLACED BASED ON THE INFORMATION PROVIDED THERE IS NO INDICATION THAT THE SUBJECT METER CAUSED OR CONTRIBUTED TO A SERIOUS INJURY ADVERSE EVENT THE PATIENT DID NOT DEVELOP SIGNSSYMPTOMS SUGGESTIVE OF A SERIOUS INJURY NOR DID THEY RECEIVE MEDICAL INTERVENTION FOR AN ACUTE BLOOD GLUCOSE EXCURSION HOWEVER THIS COMPLAINT IS BEING REPORTED AS A MALFUNCTION BECAUSE THE CALCULATED DIFFERENCE BETWEEN THE REPORTED RESULTS FALLS OUTSIDE LFS PRECISION CRITERIA AND THE ALLEGED INACCURACY WAS NOT RESOLVED DURING TROUBLESHOOTING | No Known Impact Or Consequence To Patient | Patient problem code
6845616 | 85002915 | D | DURING TESTING THE DEVICE HAD DAMAGEDBENT SCOPE THERE WAS NO PATIENT INVOLVEMENT | No Patient Involvement | Patient problem code
6845616 | 85002915 | D | DURING TESTING THE DEVICE HAD DAMAGEDBENT SCOPE THERE WAS NO PATIENT INVOLVEMENT | Break | Device problem code
6845616 | 85002915 | D | DURING TESTING THE DEVICE HAD DAMAGEDBENT SCOPE THERE WAS NO PATIENT INVOLVEMENT | Material Twisted / Bent | Device problem code
6841286 | 85011373 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS PLAYING ON THE SWINGS WHEN THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED LESS THAN AN HOUR AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6841286 | 85011373 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS PLAYING ON THE SWINGS WHEN THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED LESS THAN AN HOUR AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6846007 | 85014328 | N | A SERVICE TECHNICIAN EVALUATED THE UNIT ON SITE THE PROBLEM WAS DUPLICATED BY PRESSING EVERY FUNCTION UNTIL THE CASSETTE REJECTED THE ERROR MESSAGE THE AIR PRESSURE OUTPUT IS LOWER THAN COMMANDED APPEARED SEVERAL TIMES THE UNIT HAS BEEN RECEIVED AND WILL UNDERGO FURTHER TESTING | Device Displays Incorrect Message | Device problem code
6846007 | 85014328 | N | A SERVICE TECHNICIAN EVALUATED THE UNIT ON SITE THE PROBLEM WAS DUPLICATED BY PRESSING EVERY FUNCTION UNTIL THE CASSETTE REJECTED THE ERROR MESSAGE THE AIR PRESSURE OUTPUT IS LOWER THAN COMMANDED APPEARED SEVERAL TIMES THE UNIT HAS BEEN RECEIVED AND WILL UNDERGO FURTHER TESTING | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6845961 | 85015321 | N | PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT B | Component Missing | Device problem code
6845961 | 85015321 | N | PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6838981 | 85019491 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | Failure to Analyze Signal | Device Problem Code
6838981 | 85019491 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | No Consequences Or Impact To Patient | Patient problem code
6846924 | 85078282 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | Insufficient Information | Device problem code
6846924 | 85078282 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | No Consequences Or Impact To Patient | Patient problem code
6846938 | 85078665 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | Insufficient Information | Device problem code
6846938 | 85078665 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | No Consequences Or Impact To Patient | Patient problem code
6846947 | 85078673 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | No Consequences Or Impact To Patient | Patient problem code
6846947 | 85078673 | D | IT WAS REPORTED THAT THE ASYMPTOMATIC PATIENT DEVELOPED PNEUMOTHORAX DISCOVERED BY CHEST XRAY TAKEN THE DAY AFTER IMPLANT PROCEDURE THE PATIENT WAS GIVEN OXYGEN VIA A NASAL CANNULA AND WAS MONITORED OVERNIGHT THE PATIENT WAS DISCHARGED WITH CLOSE FOLLOWUP | Insufficient Information | Device problem code
6846949 | 85084629 | D | THE PATIENT REPORTED THAT SHE WAS HOSPITALIZED AND DIAGNOSED WITH HIGH BLOOD GLUCOSE AND GASTROPARESIS HER BLOOD GLUCOSE HAD REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS TREATMENT INCLUDED BEING GIVEN AZITHROMYCIN AND INSULIN | Hyperglycemia | Patient Problem Code
6846949 | 85084629 | D | THE PATIENT REPORTED THAT SHE WAS HOSPITALIZED AND DIAGNOSED WITH HIGH BLOOD GLUCOSE AND GASTROPARESIS HER BLOOD GLUCOSE HAD REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS TREATMENT INCLUDED BEING GIVEN AZITHROMYCIN AND INSULIN | Insufficient Information | Device problem code
6846909 | 85084691 | D | IT WAS REPORTED THAT A PATIENT WITH A MM TFX AORTIC VALVE DEVELOPED SEVERE AORTIC INSUFFICIENCY AND WAS TREATED WITH TAVR AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS AND MONTHS THE PATIENT TOLERATED THE PROCEDURE WELL AND WAS REPORTED TO BE DISCHARGED | Aortic Insufficiency | Patient problem code
6846909 | 85084691 | D | IT WAS REPORTED THAT A PATIENT WITH A MM TFX AORTIC VALVE DEVELOPED SEVERE AORTIC INSUFFICIENCY AND WAS TREATED WITH TAVR AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS AND MONTHS THE PATIENT TOLERATED THE PROCEDURE WELL AND WAS REPORTED TO BE DISCHARGED | Insufficient Information | Device problem code
6846920 | 85087089 | N | IN FOLLOW UP WITH A COMPLAINT HANDLER ON B THE CUSTOMER REPORTED THAT OR YEARS AGO SHE DEVELOPED MASTITIS WHILE PUMPING WITH THE PUMP IN STYLE BREAST PUMP AND WAS PRESCRIBED ANTIBIOTICS SHE STATED THAT WHEN SHE STOPPED PUMPING HER ISSUE WENT AWAY SHE DID NOT REPORT THIS TO MEDELA AT THE TIME SHE JUST STOPPED PUMPING BASED ON THE RESULTS OF B IT CANNOT BE DEFINITIVELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS MASTITIS IS USUALLY A BENIGN SELFLIMITING INFECTION WITH FEW CONSEQUENCES FOR THE SUCKLING INFANT THE RISK OF MASTITIS IS HIGHER AMONG WOMEN WHO HAVE BREASTFED PREVIOUSLY ESPECIALLY THOSE WITH A HISTORY OR MASTITIS RIORDAN WAMBACH TH ED P BREASTFEEDING AND HUMAN LACTATION MASTITIS REQUIRES PROMPT MEDICAL ATTENTION FOR THE MOTHER FOR PAIN RELIEF AND PRESCRIPTION ANTIBIOTICS TO AVOID PROGRESSION TO OVERWHELMING SEPSIS | Component or Accessory Incompatibility | Device Problem Code
6846920 | 85087089 | N | IN FOLLOW UP WITH A COMPLAINT HANDLER ON B THE CUSTOMER REPORTED THAT OR YEARS AGO SHE DEVELOPED MASTITIS WHILE PUMPING WITH THE PUMP IN STYLE BREAST PUMP AND WAS PRESCRIBED ANTIBIOTICS SHE STATED THAT WHEN SHE STOPPED PUMPING HER ISSUE WENT AWAY SHE DID NOT REPORT THIS TO MEDELA AT THE TIME SHE JUST STOPPED PUMPING BASED ON THE RESULTS OF B IT CANNOT BE DEFINITIVELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS MASTITIS IS USUALLY A BENIGN SELFLIMITING INFECTION WITH FEW CONSEQUENCES FOR THE SUCKLING INFANT THE RISK OF MASTITIS IS HIGHER AMONG WOMEN WHO HAVE BREASTFED PREVIOUSLY ESPECIALLY THOSE WITH A HISTORY OR MASTITIS RIORDAN WAMBACH TH ED P BREASTFEEDING AND HUMAN LACTATION MASTITIS REQUIRES PROMPT MEDICAL ATTENTION FOR THE MOTHER FOR PAIN RELIEF AND PRESCRIPTION ANTIBIOTICS TO AVOID PROGRESSION TO OVERWHELMING SEPSIS | Pumping Stopped | Device problem code
6846920 | 85087089 | N | IN FOLLOW UP WITH A COMPLAINT HANDLER ON B THE CUSTOMER REPORTED THAT OR YEARS AGO SHE DEVELOPED MASTITIS WHILE PUMPING WITH THE PUMP IN STYLE BREAST PUMP AND WAS PRESCRIBED ANTIBIOTICS SHE STATED THAT WHEN SHE STOPPED PUMPING HER ISSUE WENT AWAY SHE DID NOT REPORT THIS TO MEDELA AT THE TIME SHE JUST STOPPED PUMPING BASED ON THE RESULTS OF B IT CANNOT BE DEFINITIVELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS MASTITIS IS USUALLY A BENIGN SELFLIMITING INFECTION WITH FEW CONSEQUENCES FOR THE SUCKLING INFANT THE RISK OF MASTITIS IS HIGHER AMONG WOMEN WHO HAVE BREASTFED PREVIOUSLY ESPECIALLY THOSE WITH A HISTORY OR MASTITIS RIORDAN WAMBACH TH ED P BREASTFEEDING AND HUMAN LACTATION MASTITIS REQUIRES PROMPT MEDICAL ATTENTION FOR THE MOTHER FOR PAIN RELIEF AND PRESCRIPTION ANTIBIOTICS TO AVOID PROGRESSION TO OVERWHELMING SEPSIS | Complement Activation Problem | Medical Evaluation result code
6841280 | 85098980 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6841280 | 85098980 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6841207 | 85103850 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD FOR HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED AROUND THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE HALF WAY A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6841207 | 85103850 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD FOR HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED AROUND THE CANNULA WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE HALF WAY A NEW POD WAS APPLIED MINUTES AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6840551 | 85103987 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6840551 | 85103987 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6847901 | 85104082 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAD THE URGE TO GO BUT WAS UNABLE TO VOID AS IF THEY HAD RETENTION THE NEXT DAY PATIENT WAS CONCERNED THAT THEY HAD TO VOID HELD IT WAS UNABLE TO GO AND FELT RETENTION THEY ARE SCHEDULED FOR LEAD REMOVAL ON B BUT WOULD LIKE TO EXTEND THE TRIAL ON B PATIENT REPORTS HAVING CONTROLLER ISSUE THEY TRIED TROUBLESHOOTING WITH THE SALES REP AND ISNT SURE WHAT TO DO NEXT THEY HAD TO PUT A FOLEY BACK IN TO ASSIST WITH EMPTYING THEIR BLADDER NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6847901 | 85104082 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAD THE URGE TO GO BUT WAS UNABLE TO VOID AS IF THEY HAD RETENTION THE NEXT DAY PATIENT WAS CONCERNED THAT THEY HAD TO VOID HELD IT WAS UNABLE TO GO AND FELT RETENTION THEY ARE SCHEDULED FOR LEAD REMOVAL ON B BUT WOULD LIKE TO EXTEND THE TRIAL ON B PATIENT REPORTS HAVING CONTROLLER ISSUE THEY TRIED TROUBLESHOOTING WITH THE SALES REP AND ISNT SURE WHAT TO DO NEXT THEY HAD TO PUT A FOLEY BACK IN TO ASSIST WITH EMPTYING THEIR BLADDER NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6847929 | 85105360 | D | MEDTRONIC RECEIVED INFORMATION THAT IMMEDIATELY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB OCCURRED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6847929 | 85105360 | D | MEDTRONIC RECEIVED INFORMATION THAT IMMEDIATELY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB OCCURRED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6837765 | 85115438 | N | CONTINUED B PENTAX VIDEO COLONOSCOPE MODEL ECIF IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | Insufficient Information | Device problem code
6836547 | 85119011 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6836547 | 85119011 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6840199 | 85121575 | N | PATIENT WEIGHT IS NOT AVAILABLE FROM THE FACILITY THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING REVEALED A TWIST IN THE WORKING LENGTH OF THE DEVICE INCHES FROM THE BOTTOM OF THE STRAIN RELIEF ALONG WITH A SECOND TWIST INCHES FROM STRAIN RELIEF A GUIDEWIRE COULD BE PASSED NORMALLY ALTHOUGH RESISTANCE WAS FELT WHEN PASSING A GUIDEWIRE | Material Twisted / Bent | Device problem code
6840199 | 85121575 | N | PATIENT WEIGHT IS NOT AVAILABLE FROM THE FACILITY THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING REVEALED A TWIST IN THE WORKING LENGTH OF THE DEVICE INCHES FROM THE BOTTOM OF THE STRAIN RELIEF ALONG WITH A SECOND TWIST INCHES FROM STRAIN RELIEF A GUIDEWIRE COULD BE PASSED NORMALLY ALTHOUGH RESISTANCE WAS FELT WHEN PASSING A GUIDEWIRE | Stress Problem Identified | Medical Evaluation result code
6838390 | 85123246 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6838390 | 85123246 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6846626 | 85124067 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6848512 | 85134204 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6848512 | 85134204 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6848550 | 85134223 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Component or Accessory Incompatibility | Device problem code
6848550 | 85134223 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Incompatible Component/ Accessory | Medical Evaluation result code
6848532 | 85134411 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Component or Accessory Incompatibility | Device problem code
6848532 | 85134411 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Incompatible Component/ Accessory | Medical Evaluation result code
6848498 | 85134637 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6848498 | 85134637 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6844859 | 85134852 | N | B BATCH NR WAS RECEIVED IN THE STERILE PACKAGE UPON INSPECTION OF THE INSTRUMENT WITHIN THE PACKAGE IT WAS NOTED TO HAVE A LOOSE COMPONENT THE LOOSE COMPONENT WAS DEFINED AS AN INTERNAL CLICKER THE INSTRUMENT WAS OPENED AND TESTED AND WAS FOUND TO BE FUNCTIONAL THE INSTRUMENT WAS DISASSEMBLED AND IT WAS NOTED TO HAVE A CLICKER IN PLACE THE CLICKER FOUND IN THE STERILE PACKAGE WAS AN EXTRA COMPONENT THAT WAS INADVERTENTLY PLACED IN THE PACKAGE DURING OUR PACKAGING PROCESS THE BATCH HISTORY RECORD WAS REVIEWED AND THERE WERE NO DEFECTS PROTOCOLS OR NCRS FOUND DURING THE MANUFACTURING PROCESS RELATED TO THIS COMPLAINT | Detachment of Device or device Component | Device problem code
6844859 | 85134852 | N | B BATCH NR WAS RECEIVED IN THE STERILE PACKAGE UPON INSPECTION OF THE INSTRUMENT WITHIN THE PACKAGE IT WAS NOTED TO HAVE A LOOSE COMPONENT THE LOOSE COMPONENT WAS DEFINED AS AN INTERNAL CLICKER THE INSTRUMENT WAS OPENED AND TESTED AND WAS FOUND TO BE FUNCTIONAL THE INSTRUMENT WAS DISASSEMBLED AND IT WAS NOTED TO HAVE A CLICKER IN PLACE THE CLICKER FOUND IN THE STERILE PACKAGE WAS AN EXTRA COMPONENT THAT WAS INADVERTENTLY PLACED IN THE PACKAGE DURING OUR PACKAGING PROCESS THE BATCH HISTORY RECORD WAS REVIEWED AND THERE WERE NO DEFECTS PROTOCOLS OR NCRS FOUND DURING THE MANUFACTURING PROCESS RELATED TO THIS COMPLAINT | No Findings Available | Medical Evaluation result code
6848651 | 85134980 | N | REPORT SOURCE B | Insufficient Information | Device Problem Code
6849531 | 85162540 | N | NARROWING OF THE ESOPHAGUS AFTER ENDOSCOPIC ABLATION IS AN EXPECTED AND KNOWN ADVERSE EVENT AND IS TYPICALLY PROCEDURE RELATED HOWEVER IN THE CASE OF THIS PATIENT THE SITE REPORTED THAT IT IS PROCEDURERELATED AND MAY ALSO BE DEVICERELATED IT IS KNOWN THAT ENDOSCOPE ABLATION THERAPY SUCH AS RADIOFREQUENCY OR CRYOTHERAPY ABLATION CAN CAUSE NARROWING OF THE ESOPHAGUS THE PATIENTS CONDITION IMPROVED AFTER DILATION OF THE NARROWED AREA OF THE ESOPHAGUS | Stress Problem Identified | Medical Evaluation result code
6849833 | 85264468 | N | DEVICE EVALUATION LENS WAS RETURNED IN A SPECIMEN CUP VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE B | Component Missing | Device problem code
6849833 | 85264468 | N | DEVICE EVALUATION LENS WAS RETURNED IN A SPECIMEN CUP VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6850962 | 85266254 | N | ANY ADDITIONAL INFORMATION PROVIDED BY THE CUSTOMER WILL BE INCLUDED IN A FOLLOW UP REPORT RESULTS OF INVESTIGATION VYAIRE CERTIFIED SERVICE TECHNICIAN WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE THE DEVICE FAILED THE LEAK TEST WITH MESSAGE STATING FAILED FLOW BENCH TESTING REVEALED THAT THE BLOWER GASKET WAS NOT SEATED PROPERLY ON THE MUFFLER COVER THE SERVICE TECHNICIAN RESEATED THE BLOWER GASKET AND THE REPORTED ISSUE WAS RESOLVED THE DEVICE PASSED THE LEAK TEST FIFTY TIMES THE DEVICE PASSED ALL TESTING AND MET ALL VYAIRE MANUFACTURER SPECIFICATIONS | No Device Problem Found | Medical Evaluation result code
6850705 | 85267380 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT HER DAUGHTER WAS HOSPITALIZED ON B DUE TO HIGH BLOOD GLUCOSE BLOOD GLUCOSE AT THE TIME OF THE ADMISSION WAS MGDL THE CUSTOMERS MOTHER STATED HAD JUST PUT ON HER TRANSMITTER AND SENSOR SHE HAD A PROBLEM AND WAS IN THE WARM UP THE CUSTOMER WAS REFUSING TREATMENT WE GOT OUT AND SHE STARTED WALKING DOWN THE STREET THE CUSTOMER HAD A SEIZURE WHILE FLOATING DOWN THE RIVER THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION CUSTOMER WAS WEARING PUMP WITHIN HOURS OF HOSPITALIZATION TROUBLESHOOTING WAS DECLINED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6850705 | 85267380 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT HER DAUGHTER WAS HOSPITALIZED ON B DUE TO HIGH BLOOD GLUCOSE BLOOD GLUCOSE AT THE TIME OF THE ADMISSION WAS MGDL THE CUSTOMERS MOTHER STATED HAD JUST PUT ON HER TRANSMITTER AND SENSOR SHE HAD A PROBLEM AND WAS IN THE WARM UP THE CUSTOMER WAS REFUSING TREATMENT WE GOT OUT AND SHE STARTED WALKING DOWN THE STREET THE CUSTOMER HAD A SEIZURE WHILE FLOATING DOWN THE RIVER THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION CUSTOMER WAS WEARING PUMP WITHIN HOURS OF HOSPITALIZATION TROUBLESHOOTING WAS DECLINED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6850705 | 85267380 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT HER DAUGHTER WAS HOSPITALIZED ON B DUE TO HIGH BLOOD GLUCOSE BLOOD GLUCOSE AT THE TIME OF THE ADMISSION WAS MGDL THE CUSTOMERS MOTHER STATED HAD JUST PUT ON HER TRANSMITTER AND SENSOR SHE HAD A PROBLEM AND WAS IN THE WARM UP THE CUSTOMER WAS REFUSING TREATMENT WE GOT OUT AND SHE STARTED WALKING DOWN THE STREET THE CUSTOMER HAD A SEIZURE WHILE FLOATING DOWN THE RIVER THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION CUSTOMER WAS WEARING PUMP WITHIN HOURS OF HOSPITALIZATION TROUBLESHOOTING WAS DECLINED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Seizures | Patient problem code
6851029 | 85268899 | D | AN ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN THE PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM IN DIAMETER ABDOMINAL AORTIC ANEURYSM THE PROXIMAL AORTIC NECK MEASURED MM IN DIAMETER TO MM IN DIAMETER ALONG A MM LENGTH THE PROXIMAL AORTIC NECK ANGLE MEASURED DEGREES AND THE SUPRA TO INFRARENAL ANGULATION MEASURED DEGREES THERE WAS NO THROMBUS OR CALCIUM OBSERVED AT THE SEAL ZONE DURING THE INDEX PROCEDURE A SLUGGISH PROXIMAL TYPE I ENDOLEAK WAS OBSERVED THE PHYSICIAN ELECTED TO IMPLANT SIX HELIFX ENDOANCHORS SO THAT THE PROXIMAL END OF THE STENT GRAFT WAS ANCHORED ONTO THE AORTA HOWEVER THE PROXIMAL TYPE I ENDOLEAK PERSISTED AND IT WAS DETERMINED THAT THE OVERLAP OF THE STENT GRAFT TO THE AORTIC WALL ON THE LEFT SIDE WAS SUB OPTIMAL THE PHYSICIAN THEN ELECTED TO IMPLANT AN ENDURANT PROXIMAL AORTIC EXTENSION TO GAIN BETTER APPOSITION OF THE ENDOGRAFT TO THE AORTA IT WAS DELIVERED THROUGH THE RIGHT COMMON FEMORAL ARTERY AND DEPLOYED SO THAT AN ADDITIONAL MM OF GRAFT APPOSITION TO THE LEFT SIDE OF THE AORTA WAS GAINED BALLOONING WAS PERFORMED AND FOUR HELIFX ENDOANCHORS WERE IMPLANTED ON THE LEFT SIDE OF THE ENDURANT AORTIC EXTENSION THE PROXIMAL TYPE I ENDOLEAK WAS DRAMATICALLY IMPROVED AND THE PHYSICIAN WAS SATISFIED WITH THE RESULT NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | No Consequences Or Impact To Patient | Patient problem code
6851029 | 85268899 | D | AN ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN THE PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM IN DIAMETER ABDOMINAL AORTIC ANEURYSM THE PROXIMAL AORTIC NECK MEASURED MM IN DIAMETER TO MM IN DIAMETER ALONG A MM LENGTH THE PROXIMAL AORTIC NECK ANGLE MEASURED DEGREES AND THE SUPRA TO INFRARENAL ANGULATION MEASURED DEGREES THERE WAS NO THROMBUS OR CALCIUM OBSERVED AT THE SEAL ZONE DURING THE INDEX PROCEDURE A SLUGGISH PROXIMAL TYPE I ENDOLEAK WAS OBSERVED THE PHYSICIAN ELECTED TO IMPLANT SIX HELIFX ENDOANCHORS SO THAT THE PROXIMAL END OF THE STENT GRAFT WAS ANCHORED ONTO THE AORTA HOWEVER THE PROXIMAL TYPE I ENDOLEAK PERSISTED AND IT WAS DETERMINED THAT THE OVERLAP OF THE STENT GRAFT TO THE AORTIC WALL ON THE LEFT SIDE WAS SUB OPTIMAL THE PHYSICIAN THEN ELECTED TO IMPLANT AN ENDURANT PROXIMAL AORTIC EXTENSION TO GAIN BETTER APPOSITION OF THE ENDOGRAFT TO THE AORTA IT WAS DELIVERED THROUGH THE RIGHT COMMON FEMORAL ARTERY AND DEPLOYED SO THAT AN ADDITIONAL MM OF GRAFT APPOSITION TO THE LEFT SIDE OF THE AORTA WAS GAINED BALLOONING WAS PERFORMED AND FOUR HELIFX ENDOANCHORS WERE IMPLANTED ON THE LEFT SIDE OF THE ENDURANT AORTIC EXTENSION THE PROXIMAL TYPE I ENDOLEAK WAS DRAMATICALLY IMPROVED AND THE PHYSICIAN WAS SATISFIED WITH THE RESULT NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Insufficient Information | Device problem code
6851131 | 85274613 | N | REPORT SOURCE B | Insufficient Information | Device Problem Code
6851052 | 85275681 | D | A PATIENTS BLOOD GLUCOSE REACHED AND MGDL ON THE PDM METER WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AT THE B MEDICAL CENTER AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV AND INSULIN | Insufficient Information | Device problem code
6851052 | 85275681 | D | A PATIENTS BLOOD GLUCOSE REACHED AND MGDL ON THE PDM METER WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AT THE B MEDICAL CENTER AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV AND INSULIN | Diabetic Ketoacidosis | Patient problem code
6832006 | 85282190 | N | IMMUCOR TECHNICAL SUPPORT COULD NOT USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BECAUSE ACCESS WAS NOT GIVEN ADDITIONALLY NO CUSTOMER BLOOD SAMPLE WAS SENT TO THE IMMUCOR LABORATORY FOR TESTING BECAUSE THE CUSTOMER STATED THAT THE BLOOD SAMPLE WAS PREVIOUSLY DRAWN ON EITHER B OR B | Connection Problem | Device problem code
6832006 | 85282190 | N | IMMUCOR TECHNICAL SUPPORT COULD NOT USE A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION BECAUSE ACCESS WAS NOT GIVEN ADDITIONALLY NO CUSTOMER BLOOD SAMPLE WAS SENT TO THE IMMUCOR LABORATORY FOR TESTING BECAUSE THE CUSTOMER STATED THAT THE BLOOD SAMPLE WAS PREVIOUSLY DRAWN ON EITHER B OR B | Interoperability Problem Identified | Medical Evaluation result code
6846685 | 85287727 | N | B PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO GRAY DISPLAY OR DISPLAY ANOMALY NOTED PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE PUMP WAS MONITORED FOR HOURS AND NO UNEXPECTED GREY DISPLAY MISSING SEGMENTS PARTIAL DISPLAY GRAINY DISPLAY OR ADDITIONAL DISPLAY ANOMALY NOTED PUMP RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Scratched Material | Device problem code
6846685 | 85287727 | N | B PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO GRAY DISPLAY OR DISPLAY ANOMALY NOTED PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE PUMP WAS MONITORED FOR HOURS AND NO UNEXPECTED GREY DISPLAY MISSING SEGMENTS PARTIAL DISPLAY GRAINY DISPLAY OR ADDITIONAL DISPLAY ANOMALY NOTED PUMP RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Display or Visual Feedback Problem | Device problem code
6846685 | 85287727 | N | B PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO GRAY DISPLAY OR DISPLAY ANOMALY NOTED PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE PUMP WAS MONITORED FOR HOURS AND NO UNEXPECTED GREY DISPLAY MISSING SEGMENTS PARTIAL DISPLAY GRAINY DISPLAY OR ADDITIONAL DISPLAY ANOMALY NOTED PUMP RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Component Missing | Device problem code
6846685 | 85287727 | N | B PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO GRAY DISPLAY OR DISPLAY ANOMALY NOTED PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE PUMP WAS MONITORED FOR HOURS AND NO UNEXPECTED GREY DISPLAY MISSING SEGMENTS PARTIAL DISPLAY GRAINY DISPLAY OR ADDITIONAL DISPLAY ANOMALY NOTED PUMP RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY | Missing or Inadequate Safety Measures | Medical Evaluation result code
6851428 | 85295470 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | No Findings Available | Medical Evaluation result code
6846934 | 85296207 | N | UNIT DID NOT POWER UP AND NO RF PROBLEM ISOLATED TO THE PCB BOARD PER B B | Failure to Power Up | Device problem code
6846934 | 85296207 | N | UNIT DID NOT POWER UP AND NO RF PROBLEM ISOLATED TO THE PCB BOARD PER B B | Power Source Problem Identified | Medical Evaluation result code
6843289 | 85298577 | D | IT WAS REPORTED THAT WHILE LOWERING THE COLON DURING SURGERY THE DEVICE DID NOT WORK PROPERLY ACCORDING TO THE SURGEON IT CUT AND DID NOT CLIP IT STAPLED THE RECTUM AND NOT THE WARHEAD THE STAPLER WAS CHANGED WITHOUT PROBLEMS THERE WAS NO HARM TO THE PATIENT | Insufficient Information | Device Problem Code
6843289 | 85298577 | D | IT WAS REPORTED THAT WHILE LOWERING THE COLON DURING SURGERY THE DEVICE DID NOT WORK PROPERLY ACCORDING TO THE SURGEON IT CUT AND DID NOT CLIP IT STAPLED THE RECTUM AND NOT THE WARHEAD THE STAPLER WAS CHANGED WITHOUT PROBLEMS THERE WAS NO HARM TO THE PATIENT | No Consequences Or Impact To Patient | Patient problem code
6851200 | 85298673 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE POD WAS NOT STICKING WELL TO THE SKIN AND THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6851200 | 85298673 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR THE POD WAS NOT STICKING WELL TO THE SKIN AND THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6837567 | 85306613 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL AT PM AFTER PATIENT HAD EATEN WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6837567 | 85306613 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL AT PM AFTER PATIENT HAD EATEN WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6837318 | 85306661 | D | THE CUSTOMER REPORTED THAT HIS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS LEG FOR BETWEEN AND HOURS HE STATED THAT HE DID NOT FEEL THAT THE CANNULA WENT ALL THE WAY IN TREATMENT INCLUDED CHANGING THE POD | No Known Impact Or Consequence To Patient | Patient Problem Code
6837318 | 85306661 | D | THE CUSTOMER REPORTED THAT HIS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS LEG FOR BETWEEN AND HOURS HE STATED THAT HE DID NOT FEEL THAT THE CANNULA WENT ALL THE WAY IN TREATMENT INCLUDED CHANGING THE POD | Component or Accessory Incompatibility | Device problem code
6852318 | 85405230 | D | IT WAS REPORTED THAT THE PATIENT AND PATIENT CAREGIVER RECEIVED A NEEDLE STICK INJURY WHEN USING A G X MM BD ULTRA FINE INSULIN PEN NEEDLE NO MEDICAL INTERVENTION WAS SOUGHT | No Consequences Or Impact To Patient | Patient problem code
6852318 | 85405230 | D | IT WAS REPORTED THAT THE PATIENT AND PATIENT CAREGIVER RECEIVED A NEEDLE STICK INJURY WHEN USING A G X MM BD ULTRA FINE INSULIN PEN NEEDLE NO MEDICAL INTERVENTION WAS SOUGHT | Physical Resistance / Sticking | Device problem code
6852441 | 85410704 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURN MARKS ON THE AC POWER ADAPTER AND MAIN PRINTED CIRCUIT BOARD WHICH CAUSED THE REPORTED INABILITY TO POWER THE TRANSMITTER | Power Problem | Device problem code
6852441 | 85410704 | N | NO CONCLUSION CODE AVAILABLE FINAL ANALYSIS FOUND BURN MARKS ON THE AC POWER ADAPTER AND MAIN PRINTED CIRCUIT BOARD WHICH CAUSED THE REPORTED INABILITY TO POWER THE TRANSMITTER | Power Source Problem Identified | Medical Evaluation result code
6853169 | 85450771 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF NODULES IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF SKIN IRRITATION PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES AFTER TREATMENT WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING | No Findings Available | Medical Evaluation result code
6853168 | 85459090 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF NODULES IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF SKIN IRRITATION PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY SEVERITY AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS | No Findings Available | Medical Evaluation result code
6853930 | 85573790 | N | ON B A SERVICE ENGINEER DETERMINED THERE WERE NO ANALYZER DEFICIENCIES THE ANALYZER ALERTED THE OPERATOR TO POSSIBLE SAMPLE ABNORMALITY REQUIRING FURTHER REVIEW PRIOR TO REPORTING USERS ARE INFORMED OF THE SITUATIONS IN WHICH RESULTS MAY BE AFFECTED IN CHAPTER POSSIBLE SAMPLE INTERFERENCES FOR PLATELETS THE WARNING IS NOTED WHERE THE FOLLOWING ARE PRESENT THE PLATELET COUNT MAY BE REPORTED FALSELY LOW PLATELET AGGREGATION GIANT PLATELETS PSEUDOTHROMBOCYTOPENIA PLT CLUMPS CAN BE CAUSED BY ABNORMAL PROTEINS IN THE PATIENTS PLASMA WHEN EXPOSED TO EDTA ANTICOAGULANT NO ANALYZER MALFUNCTION OCCURRED ROOT CAUSE IS A PATIENT SPECIFIC ABNORMALITY DUE TO THE PRESENCE OF PLATELET CLUMPING | Clumping in Device or Device Ingredient | Device problem code
6853930 | 85573790 | N | ON B A SERVICE ENGINEER DETERMINED THERE WERE NO ANALYZER DEFICIENCIES THE ANALYZER ALERTED THE OPERATOR TO POSSIBLE SAMPLE ABNORMALITY REQUIRING FURTHER REVIEW PRIOR TO REPORTING USERS ARE INFORMED OF THE SITUATIONS IN WHICH RESULTS MAY BE AFFECTED IN CHAPTER POSSIBLE SAMPLE INTERFERENCES FOR PLATELETS THE WARNING IS NOTED WHERE THE FOLLOWING ARE PRESENT THE PLATELET COUNT MAY BE REPORTED FALSELY LOW PLATELET AGGREGATION GIANT PLATELETS PSEUDOTHROMBOCYTOPENIA PLT CLUMPS CAN BE CAUSED BY ABNORMAL PROTEINS IN THE PATIENTS PLASMA WHEN EXPOSED TO EDTA ANTICOAGULANT NO ANALYZER MALFUNCTION OCCURRED ROOT CAUSE IS A PATIENT SPECIFIC ABNORMALITY DUE TO THE PRESENCE OF PLATELET CLUMPING | No Findings Available | Medical Evaluation result code
6853609 | 85574368 | D | EDWARDS REVIEWED THE LITERATURE ARTICLE LONGTERM FOLLOWUP AFTER TRANSCATHETER TRICUSPID VALVEINVALVE REPLACEMENT WITH BALLOONEXPANDABLE AORTIC VALVES AUTHOR FLAVIO RIBICHINI ET AL INTERNATIONAL JOURNAL OF CARDIOLOGY BACKGROUND TRANSCATHETER TRICUSPID VIV REPLACEMENT HAS BEEN PROPOSED AS A FEASIBLE OPTION FOR HIGHRISK PATIENTS WITH PREVIOUS TRICUSPID VALVE REPLACEMENT THAT UNDERGO VALVE DEGENERATION CAUSING REFRACTORY HEART FAILURE HOWEVER LITTLE IS KNOWN ABOUT THE LONGTERM OUTCOME OF PATIENTS TREATED WITH TRANSCATHETER TRICUSPID VIV WE EVALUATE THE SAFETY OF TRANSCATHETER TRICUSPID VALVEINVALVE VIV REPLACEMENT BY USING BALLOONEXPANDABLE AORTIC VALVE STENTS AND THE LONGTERM FOLLOWUP CONCLUSIONS IN OUR EXPERIENCE TRANSCATHETER TRICUSPID VIV DEMONSTRATED GOOD LONGTERM RESULTS IN TERMS OF VALVE PERFORMANCE AND FUNCTIONAL CLASS IMPROVEMENT AT MONTHS FROM INTERVENTION IN THE CONTEXT OF THIS ARTICLE THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE ONE B MALE PATIENT WITH A MM PERICARDIAL MITRAL VALVE IMPLANTED IN TRICUSPID POSITION UNDERWENT SURGERY FOR VALVE IN VALVE REPLACEMENT AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS DUE TO DEGENERATION LEADING TO SEVERE TRICUSPID STENOSIS AND INSUFFICIENCY A MM TRANSCATHETER VALVE WAS IMPLANTED AS A REPLACEMENT PATIENT WAS NOTED AS TO BE DISCHARGED | Stenosis | Patient problem code
6853609 | 85574368 | D | EDWARDS REVIEWED THE LITERATURE ARTICLE LONGTERM FOLLOWUP AFTER TRANSCATHETER TRICUSPID VALVEINVALVE REPLACEMENT WITH BALLOONEXPANDABLE AORTIC VALVES AUTHOR FLAVIO RIBICHINI ET AL INTERNATIONAL JOURNAL OF CARDIOLOGY BACKGROUND TRANSCATHETER TRICUSPID VIV REPLACEMENT HAS BEEN PROPOSED AS A FEASIBLE OPTION FOR HIGHRISK PATIENTS WITH PREVIOUS TRICUSPID VALVE REPLACEMENT THAT UNDERGO VALVE DEGENERATION CAUSING REFRACTORY HEART FAILURE HOWEVER LITTLE IS KNOWN ABOUT THE LONGTERM OUTCOME OF PATIENTS TREATED WITH TRANSCATHETER TRICUSPID VIV WE EVALUATE THE SAFETY OF TRANSCATHETER TRICUSPID VALVEINVALVE VIV REPLACEMENT BY USING BALLOONEXPANDABLE AORTIC VALVE STENTS AND THE LONGTERM FOLLOWUP CONCLUSIONS IN OUR EXPERIENCE TRANSCATHETER TRICUSPID VIV DEMONSTRATED GOOD LONGTERM RESULTS IN TERMS OF VALVE PERFORMANCE AND FUNCTIONAL CLASS IMPROVEMENT AT MONTHS FROM INTERVENTION IN THE CONTEXT OF THIS ARTICLE THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE ONE B MALE PATIENT WITH A MM PERICARDIAL MITRAL VALVE IMPLANTED IN TRICUSPID POSITION UNDERWENT SURGERY FOR VALVE IN VALVE REPLACEMENT AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS DUE TO DEGENERATION LEADING TO SEVERE TRICUSPID STENOSIS AND INSUFFICIENCY A MM TRANSCATHETER VALVE WAS IMPLANTED AS A REPLACEMENT PATIENT WAS NOTED AS TO BE DISCHARGED | Insufficient Information | Device problem code
6853609 | 85574368 | D | EDWARDS REVIEWED THE LITERATURE ARTICLE LONGTERM FOLLOWUP AFTER TRANSCATHETER TRICUSPID VALVEINVALVE REPLACEMENT WITH BALLOONEXPANDABLE AORTIC VALVES AUTHOR FLAVIO RIBICHINI ET AL INTERNATIONAL JOURNAL OF CARDIOLOGY BACKGROUND TRANSCATHETER TRICUSPID VIV REPLACEMENT HAS BEEN PROPOSED AS A FEASIBLE OPTION FOR HIGHRISK PATIENTS WITH PREVIOUS TRICUSPID VALVE REPLACEMENT THAT UNDERGO VALVE DEGENERATION CAUSING REFRACTORY HEART FAILURE HOWEVER LITTLE IS KNOWN ABOUT THE LONGTERM OUTCOME OF PATIENTS TREATED WITH TRANSCATHETER TRICUSPID VIV WE EVALUATE THE SAFETY OF TRANSCATHETER TRICUSPID VALVEINVALVE VIV REPLACEMENT BY USING BALLOONEXPANDABLE AORTIC VALVE STENTS AND THE LONGTERM FOLLOWUP CONCLUSIONS IN OUR EXPERIENCE TRANSCATHETER TRICUSPID VIV DEMONSTRATED GOOD LONGTERM RESULTS IN TERMS OF VALVE PERFORMANCE AND FUNCTIONAL CLASS IMPROVEMENT AT MONTHS FROM INTERVENTION IN THE CONTEXT OF THIS ARTICLE THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE ONE B MALE PATIENT WITH A MM PERICARDIAL MITRAL VALVE IMPLANTED IN TRICUSPID POSITION UNDERWENT SURGERY FOR VALVE IN VALVE REPLACEMENT AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS DUE TO DEGENERATION LEADING TO SEVERE TRICUSPID STENOSIS AND INSUFFICIENCY A MM TRANSCATHETER VALVE WAS IMPLANTED AS A REPLACEMENT PATIENT WAS NOTED AS TO BE DISCHARGED | Heart Failure | Patient problem code
6853609 | 85574368 | D | EDWARDS REVIEWED THE LITERATURE ARTICLE LONGTERM FOLLOWUP AFTER TRANSCATHETER TRICUSPID VALVEINVALVE REPLACEMENT WITH BALLOONEXPANDABLE AORTIC VALVES AUTHOR FLAVIO RIBICHINI ET AL INTERNATIONAL JOURNAL OF CARDIOLOGY BACKGROUND TRANSCATHETER TRICUSPID VIV REPLACEMENT HAS BEEN PROPOSED AS A FEASIBLE OPTION FOR HIGHRISK PATIENTS WITH PREVIOUS TRICUSPID VALVE REPLACEMENT THAT UNDERGO VALVE DEGENERATION CAUSING REFRACTORY HEART FAILURE HOWEVER LITTLE IS KNOWN ABOUT THE LONGTERM OUTCOME OF PATIENTS TREATED WITH TRANSCATHETER TRICUSPID VIV WE EVALUATE THE SAFETY OF TRANSCATHETER TRICUSPID VALVEINVALVE VIV REPLACEMENT BY USING BALLOONEXPANDABLE AORTIC VALVE STENTS AND THE LONGTERM FOLLOWUP CONCLUSIONS IN OUR EXPERIENCE TRANSCATHETER TRICUSPID VIV DEMONSTRATED GOOD LONGTERM RESULTS IN TERMS OF VALVE PERFORMANCE AND FUNCTIONAL CLASS IMPROVEMENT AT MONTHS FROM INTERVENTION IN THE CONTEXT OF THIS ARTICLE THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE ONE B MALE PATIENT WITH A MM PERICARDIAL MITRAL VALVE IMPLANTED IN TRICUSPID POSITION UNDERWENT SURGERY FOR VALVE IN VALVE REPLACEMENT AFTER AN IMPLANT DURATION OF APPROXIMATELY YEARS DUE TO DEGENERATION LEADING TO SEVERE TRICUSPID STENOSIS AND INSUFFICIENCY A MM TRANSCATHETER VALVE WAS IMPLANTED AS A REPLACEMENT PATIENT WAS NOTED AS TO BE DISCHARGED | Tricuspid Stenosis | Patient problem code
6847521 | 85577692 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT HE TESTED THE IABP AND FOUND WATER IN THE LINES THE FSE COULD NOT RECREATE THE LOW VACUUM ALARM AND ALL TESTS CAME BACK NORMAL THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE B | Expiration Date Error | Device problem code
6847521 | 85577692 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT HE TESTED THE IABP AND FOUND WATER IN THE LINES THE FSE COULD NOT RECREATE THE LOW VACUUM ALARM AND ALL TESTS CAME BACK NORMAL THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE B | No Device Problem Found | Medical Evaluation result code
6850722 | 85578415 | N | THE DEVICE WAS RECEIVED BY RESMED FOR AN ENGINEERING INVESTIGATION REVIEW OF THE DEVICE LOGS CONFIRMED THE OCCURRENCE OF SYSTEM FAULT SF PERFORMANCE TESTING WAS NOT ABLE TO REPRODUCE THE REPORTED DEVICE FAULT AND THE DEVICE AND BATTERY PERFORMED PER SPECIFICATIONS BASED ON ALL EVIDENCE IT IS POSSIBLE THAT THE SYSTEM FAULT WAS DUE TO THE DEVICE BEING OPERATED AT TEMPERATURES OUTSIDE ITS SPECIFICATIONS THE DEVICE WAS SERVICED CLEANED CALIBRATED AND TESTED BEFORE IT WAS RETURNED TO THE CUSTOMER RESMEDS RISK ANALYSIS FOR THESE FAILURE MODES CONCLUDES THAT THE RISK IS ACCEPTABLE B | Defective Device | Device problem code
6850722 | 85578415 | N | THE DEVICE WAS RECEIVED BY RESMED FOR AN ENGINEERING INVESTIGATION REVIEW OF THE DEVICE LOGS CONFIRMED THE OCCURRENCE OF SYSTEM FAULT SF PERFORMANCE TESTING WAS NOT ABLE TO REPRODUCE THE REPORTED DEVICE FAULT AND THE DEVICE AND BATTERY PERFORMED PER SPECIFICATIONS BASED ON ALL EVIDENCE IT IS POSSIBLE THAT THE SYSTEM FAULT WAS DUE TO THE DEVICE BEING OPERATED AT TEMPERATURES OUTSIDE ITS SPECIFICATIONS THE DEVICE WAS SERVICED CLEANED CALIBRATED AND TESTED BEFORE IT WAS RETURNED TO THE CUSTOMER RESMEDS RISK ANALYSIS FOR THESE FAILURE MODES CONCLUDES THAT THE RISK IS ACCEPTABLE B | Operational Problem Identified | Medical Evaluation result code
6853496 | 85578525 | D | THE CUSTOMER REPORTED THAT DURING A PROCEDURE THE CLEAR INNER INSULATION WAS PULLED OFF OF THE ELECTROCAUTERY BLADE DURING CLEANING THE INSULATION DID NOT FALL WITHIN THE PATIENT CAVITY AND THERE WAS NO PATIENT INJURY | Detachment of Device or device Component | Device problem code
6853496 | 85578525 | D | THE CUSTOMER REPORTED THAT DURING A PROCEDURE THE CLEAR INNER INSULATION WAS PULLED OFF OF THE ELECTROCAUTERY BLADE DURING CLEANING THE INSULATION DID NOT FALL WITHIN THE PATIENT CAVITY AND THERE WAS NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
6854302 | 85590170 | N | AGESDATE OF BIRTH UNKNOWN NOT PROVIDED GENDERSSEX UNKNOWNNOT PROVIDED DATE OF EVENT UNKNOWNNOT PROVIDED SERIAL NUMBER UNKNOWNNOT PROVIDED CATALOG NUMBER A COMPLETE CATALOG NUMBER IS UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED EXPIRATION DATE AND UDI NUMBER ARE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED IF IMPLANTED GIVE DATE UNKNOWN NOT PROVIDED IF EXPLANTED GIVE DATE UNKNOWN NOT PROVIDED MANUFACTURING DATE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Findings Available | Medical Evaluation result code
6854306 | 85591047 | N | AGESDATE OF BIRTH UNKNOWN NOT PROVIDED GENDERSSEX UNKNOWNNOT PROVIDED DATE OF EVENT UNKNOWNNOT PROVIDED SERIAL NUMBER UNKNOWNNOT PROVIDED CATALOG NUMBER A COMPLETE CATALOG NUMBER IS UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED EXPIRATION DATE AND UDI NUMBER ARE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED IF IMPLANTED GIVE DATE UNKNOWN NOT PROVIDED IF EXPLANTED GIVE DATE UNKNOWN NOT PROVIDED MANUFACTURING DATE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Findings Available | Medical Evaluation result code
6854309 | 85593009 | N | SERIAL NUMBER UNKNOWNNOT PROVIDED CATALOG NUMBER A COMPLETE CATALOG NUMBER IS UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED EXPIRATION DATE AND UDI NUMBER ARE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED MANUFACTURING DATES UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Findings Available | Medical Evaluation result code
6851789 | 85593768 | D | ON B THE REPORTER CONTACTED ANIMAS AND REPORTED A POWERBATTERY LIFE ISSUE THERE WAS NO INDICATION THAT THE PRODUCT CAUSED OR CONTRIBUTED TO AN ADVERSE EVENT THIS COMPLAINT IS BEING REPORTED AS THE ALLEGED MALFUNCTION HAS THE ABILITY TO RESULT IN A DELAY IN TREATMENT OR LONG TERM CESSATION IN DELIVERY | Battery Problem | Device Problem Code
6851789 | 85593768 | D | ON B THE REPORTER CONTACTED ANIMAS AND REPORTED A POWERBATTERY LIFE ISSUE THERE WAS NO INDICATION THAT THE PRODUCT CAUSED OR CONTRIBUTED TO AN ADVERSE EVENT THIS COMPLAINT IS BEING REPORTED AS THE ALLEGED MALFUNCTION HAS THE ABILITY TO RESULT IN A DELAY IN TREATMENT OR LONG TERM CESSATION IN DELIVERY | No Consequences Or Impact To Patient | Patient problem code
6854313 | 85594902 | N | CATALOG NUMBER COMPLETE CATALOG NUMBER IS UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED EXPIRATION DATE AND UDI NUMBER ARE UNKNOWN SINCE THE SERIAL NUMBER WAS NOT PROVIDED MANUFACTURING DATE UNKNOWN SINCE SERIAL NUMBER WAS NOT PROVIDED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO THE MANUFACTURER HAS BEEN SUBMITTED | No Findings Available | Medical Evaluation result code
6854334 | 85597222 | N | NO INFORMATION HAS BEEN PROVIDED BY THE INITIAL REPORTER INDICATING THAT THE SEROMA WAS DIRECTLY RESULTANT OF THE OVITEX DEVICE OR THE PATIENT OUTCOME | Insufficient Information | Device problem code
6851268 | 85599277 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG A MANULA INSULIN INJECTION WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Insufficient Information | Device problem code
6851268 | 85599277 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG A MANULA INSULIN INJECTION WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6826465 | 85602198 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6826465 | 85602198 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE CUSTOMER NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6838397 | 85605886 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6838397 | 85605886 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6853302 | 85608107 | N | WILLIAM COOK B INITIALLY REPORTED THIS EVENT BASED ON AVAILABLE INFORMATION AT THE TIME OF COMPLAINT REGISTRATION UNDER NEW INFORMATION WAS PROVIDED IDENTIFYING THAT THE PRODUCT WAS ACTUALLY A COOK INC MANUFACTURED DEVICE B EVALUATION IT HAS NOT BEEN POSSIBLE TO FURTHER INVESTIGATE OR EVALUATE THIS ALLEGED EVENT BASED ON THE LIMITED INFORMATION PROVIDED TO DATE VIA THE OPERATIVE NOTE STATING TULIP MIGRATION COOK WILL REOPEN ITS INVESTIGATION IF FURTHER INFORMATION IS RECEIVED WARRANTING SUPPLEMENTATION IN ACCORDANCE WITH CFR MANIPULATION IN THE AREA OF THE FILTER IMPLANT MAY CAUSE MIGRATION OR CONTRIBUTE TO CHANGES IN THE FILTER CONFIGURATION AND PLACEMENT NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS AND NOTHING INDICATES THAT THE FILTER DID NOT PERFORM AS INTENDED EG INTENDED FOR THE PREVENTION OF RECURRENT PULMONARY EMBOLISM PE VIA PLACEMENT IN THE VENA CAVA | Migration or Expulsion of Device | Device problem code
6853302 | 85608107 | N | WILLIAM COOK B INITIALLY REPORTED THIS EVENT BASED ON AVAILABLE INFORMATION AT THE TIME OF COMPLAINT REGISTRATION UNDER NEW INFORMATION WAS PROVIDED IDENTIFYING THAT THE PRODUCT WAS ACTUALLY A COOK INC MANUFACTURED DEVICE B EVALUATION IT HAS NOT BEEN POSSIBLE TO FURTHER INVESTIGATE OR EVALUATE THIS ALLEGED EVENT BASED ON THE LIMITED INFORMATION PROVIDED TO DATE VIA THE OPERATIVE NOTE STATING TULIP MIGRATION COOK WILL REOPEN ITS INVESTIGATION IF FURTHER INFORMATION IS RECEIVED WARRANTING SUPPLEMENTATION IN ACCORDANCE WITH CFR MANIPULATION IN THE AREA OF THE FILTER IMPLANT MAY CAUSE MIGRATION OR CONTRIBUTE TO CHANGES IN THE FILTER CONFIGURATION AND PLACEMENT NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS AND NOTHING INDICATES THAT THE FILTER DID NOT PERFORM AS INTENDED EG INTENDED FOR THE PREVENTION OF RECURRENT PULMONARY EMBOLISM PE VIA PLACEMENT IN THE VENA CAVA | Device Migration | Medical Evaluation result code
6833843 | 85608146 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE | Connection Problem | Device problem code
6833843 | 85608146 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE | No Findings Available | Medical Evaluation result code
6836795 | 85608354 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6836795 | 85608354 | D | PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6848071 | 85609176 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK IT WAS NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6848071 | 85609176 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK IT WAS NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | Positioning Problem | Device problem code
6837322 | 85612803 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER BACK FOR LESS THAN AN HOUR DUE TO THE CANNULA POKING HER AND BEING REALLY UNCOMFORTABLE IT WAS NOT BELIEVED THAT IT WAS INSERTED PROPERLY | Insufficient Information | Device problem code
6837322 | 85612803 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER BACK FOR LESS THAN AN HOUR DUE TO THE CANNULA POKING HER AND BEING REALLY UNCOMFORTABLE IT WAS NOT BELIEVED THAT IT WAS INSERTED PROPERLY | Discomfort | Patient problem code
6846762 | 85613766 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING CONFIRMED THE REPORTED ISSUE WITH THE DEVICE FINDING A TWIST IN THE TUBING INCHES FROM THE STRAIN RELIEF | Material Twisted / Bent | Device problem code
6846762 | 85613766 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING CONFIRMED THE REPORTED ISSUE WITH THE DEVICE FINDING A TWIST IN THE TUBING INCHES FROM THE STRAIN RELIEF | Stress Problem Identified | Medical Evaluation result code
6840088 | 85614726 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING REVEALED THAT THE HUB WAS BLOCKED AND A PREP KIT GUIDEWIRE COULD NOT BE PASSED A MANDREL WAS ABLE TO BE PASSED HOWEVER WITH NO RESISTANCE | Partial Blockage | Device Problem Code
6840088 | 85614726 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION VISUAL INSPECTION AND SIMULATED USE TESTING REVEALED THAT THE HUB WAS BLOCKED AND A PREP KIT GUIDEWIRE COULD NOT BE PASSED A MANDREL WAS ABLE TO BE PASSED HOWEVER WITH NO RESISTANCE | No Device Problem Found | Medical Evaluation result code
6831261 | 85617416 | N | THE INVESTIGATION HAS DETERMINED THAT A LOWER THAN EXPECTED VITROS CREA RESULT WAS OBTAINED WHEN TESTING ONE PATIENT SAMPLE ON A VITROS INTEGRATED SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT IS LIKELY A SAMPLE RELATED ISSUE DUE TO INAPPROPRIATE PREANALYTICAL SAMPLE HANDLING THE CUSTOMER WAS NOT PERFORMING THE PREANALYTICAL SAMPLE PREPARATION IN ACCORDANCE WITH THE SAMPLE COLLECTION DEVICE MANUFACTURERS RECOMMENDATIONS UPON VISUAL INSPECTION FIBRIN WAS CONFIRMED TO BE PRESENT IN THE AFFECTED PATIENT SAMPLE BEFORE BEING PROCESSED ON THE VITROS SYSTEM A VITROS SYSTEM ISSUE CANNOT BE COMPLETELY RULED OUT AS THE CUSTOMER OBSERVED INTERMITTENT MICROSLIDE METERING CONDITION CODES AROUND THE TIME OF THE EVENT BASED ON HISTORICAL QUALITY CONTROL PERFORMANCE THERE IS NO INDICATION OF AN ATYPICAL VITROS CREA REAGENT PERFORMANCE ISSUE AT THE TIME OF THE EVENT | Low Test Results | Device problem code
6831261 | 85617416 | N | THE INVESTIGATION HAS DETERMINED THAT A LOWER THAN EXPECTED VITROS CREA RESULT WAS OBTAINED WHEN TESTING ONE PATIENT SAMPLE ON A VITROS INTEGRATED SYSTEM THE ASSIGNABLE CAUSE OF THE EVENT IS LIKELY A SAMPLE RELATED ISSUE DUE TO INAPPROPRIATE PREANALYTICAL SAMPLE HANDLING THE CUSTOMER WAS NOT PERFORMING THE PREANALYTICAL SAMPLE PREPARATION IN ACCORDANCE WITH THE SAMPLE COLLECTION DEVICE MANUFACTURERS RECOMMENDATIONS UPON VISUAL INSPECTION FIBRIN WAS CONFIRMED TO BE PRESENT IN THE AFFECTED PATIENT SAMPLE BEFORE BEING PROCESSED ON THE VITROS SYSTEM A VITROS SYSTEM ISSUE CANNOT BE COMPLETELY RULED OUT AS THE CUSTOMER OBSERVED INTERMITTENT MICROSLIDE METERING CONDITION CODES AROUND THE TIME OF THE EVENT BASED ON HISTORICAL QUALITY CONTROL PERFORMANCE THERE IS NO INDICATION OF AN ATYPICAL VITROS CREA REAGENT PERFORMANCE ISSUE AT THE TIME OF THE EVENT | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6849073 | 85621625 | N | B DEVICE EMITS ODOR WAS FOR THE BURNING SMELL PMAK NUMBER K K K THE DEVICE WILL NOT BE RETURNING TO ZIMMER BIOMET SINCE IT IS GOING TO BE EVALUATED BY AN EXTERNAL CONTRACTOR ONCE THE INVESTIGATION HAS BEEN COMPLETED A FOLLOW UP MDR WILL BE SUBMITTED | Device Emits Odor | Device problem code
6849073 | 85621625 | N | B DEVICE EMITS ODOR WAS FOR THE BURNING SMELL PMAK NUMBER K K K THE DEVICE WILL NOT BE RETURNING TO ZIMMER BIOMET SINCE IT IS GOING TO BE EVALUATED BY AN EXTERNAL CONTRACTOR ONCE THE INVESTIGATION HAS BEEN COMPLETED A FOLLOW UP MDR WILL BE SUBMITTED | Results Pending Completion of Investigation | Medical Evaluation result code
6848768 | 85625457 | N | THE DEVICE WAS RETURNED TO MMDG FOR EVALUATION OF THE FREE FLOW COMPLAINT DURING THE INVESTIGATION MMDG WAS NOT ABLE TO CONFIRM OR REPLICATE THE COMPLAINT THE PUMP OPERATED WITHIN PRODUCT SPECIFICATIONS THE DEVICE WAS SERVICED AND RETURNED TO THE CUSTOMER | No Device Problem Found | Medical Evaluation result code
6851353 | 85628280 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6851353 | 85628280 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6850308 | 85628765 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED PERFORMING A COMPLETE PREVENTIVE MAINTENANCE PM WITH FULL CALIBRATION FUNCTIONAL TESTING AND SAFETY CHECK TO FACTORY SPECIFICATIONS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE THE SAFETY DISK WAS EXPIRED BY TIME AND A NEW SAFETY DISK WAS INSTALLED THE FULL NAME OF THE INITIAL REPORTER LISTED IN IS B AN ABBREVIATION WAS USED AS THE FULL NAME EXCEEDED THE MAXIMUM CHARACTER LIMIT FOR THAT FIELD | Calibration Problem | Device problem code
6850308 | 85628765 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED PERFORMING A COMPLETE PREVENTIVE MAINTENANCE PM WITH FULL CALIBRATION FUNCTIONAL TESTING AND SAFETY CHECK TO FACTORY SPECIFICATIONS THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE THE SAFETY DISK WAS EXPIRED BY TIME AND A NEW SAFETY DISK WAS INSTALLED THE FULL NAME OF THE INITIAL REPORTER LISTED IN IS B AN ABBREVIATION WAS USED AS THE FULL NAME EXCEEDED THE MAXIMUM CHARACTER LIMIT FOR THAT FIELD | Failure to Calibrate | Device Problem Code
6843435 | 85637252 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE WITH A CLEAR HALO AROUND THE BUTTON | Connection Problem | Device problem code
6843435 | 85637252 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE WITH A CLEAR HALO AROUND THE BUTTON | No Findings Available | Medical Evaluation result code
6844883 | 85637290 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON AUG WHICH APPEARED VISUALLY POSITIVE | Connection Problem | Device problem code
6844883 | 85637290 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON AUG WHICH APPEARED VISUALLY POSITIVE | No Findings Available | Medical Evaluation result code
6852062 | 85637651 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH VISUALLY APPEARED WITH A SLIGHT ADHERENCE | Connection Problem | Device problem code
6852062 | 85637651 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH VISUALLY APPEARED WITH A SLIGHT ADHERENCE | No Findings Available | Medical Evaluation result code
6855095 | 85639029 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED ARE MULTIFACTORIAL AND MAY BE ATTRIBUTED TO MEDICAL TREATMENT PROGRESSION OF DISEASE ANTICOAGULATION THERAPY AND PATIENT COMORBIDITIES IN ADDITION LVAD PUMP CANDIDATES OFTEN POSSESS RISK FACTORS FOR INTRAVASCULAR COAGULATION WHICH MAY INCLUDE ARTERIAL ANDOR VENOUS VASCULAR DISEASE CHRONIC LOW FLOW STATE AND DECREASED IMMOBILITY THERE ARE PATIENT PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Coagulation in Device or device Ingredient | Device Problem Code
6855095 | 85639029 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED ARE MULTIFACTORIAL AND MAY BE ATTRIBUTED TO MEDICAL TREATMENT PROGRESSION OF DISEASE ANTICOAGULATION THERAPY AND PATIENT COMORBIDITIES IN ADDITION LVAD PUMP CANDIDATES OFTEN POSSESS RISK FACTORS FOR INTRAVASCULAR COAGULATION WHICH MAY INCLUDE ARTERIAL ANDOR VENOUS VASCULAR DISEASE CHRONIC LOW FLOW STATE AND DECREASED IMMOBILITY THERE ARE PATIENT PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Infusion or Flow Problem | Device problem code
6855095 | 85639029 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED ARE MULTIFACTORIAL AND MAY BE ATTRIBUTED TO MEDICAL TREATMENT PROGRESSION OF DISEASE ANTICOAGULATION THERAPY AND PATIENT COMORBIDITIES IN ADDITION LVAD PUMP CANDIDATES OFTEN POSSESS RISK FACTORS FOR INTRAVASCULAR COAGULATION WHICH MAY INCLUDE ARTERIAL ANDOR VENOUS VASCULAR DISEASE CHRONIC LOW FLOW STATE AND DECREASED IMMOBILITY THERE ARE PATIENT PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Operational Problem Identified | Medical Evaluation result code
6855635 | 85664840 | D | THE PATIENT REPORTED AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM SHE NOTICED HER SITE WAS RED AND IRRITATED SHE WENT TO SEE HER DOCTOR WHO PROVIDED CORTISONE FOR HER SITE IRRITATION | Insufficient Information | Device problem code
6855635 | 85664840 | D | THE PATIENT REPORTED AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM SHE NOTICED HER SITE WAS RED AND IRRITATED SHE WENT TO SEE HER DOCTOR WHO PROVIDED CORTISONE FOR HER SITE IRRITATION | Irritation | Patient problem code
6856876 | 85730855 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Infection | Patient problem code
6856876 | 85730855 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Venipuncture | Patient problem code
6856876 | 85730855 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Insufficient Information | Device problem code
6856840 | 85734600 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Infection | Patient problem code
6856840 | 85734600 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Insufficient Information | Device problem code
6856840 | 85734600 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Venipuncture | Patient problem code
6856820 | 85735836 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Insufficient Information | Device problem code
6856820 | 85735836 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Infection | Patient problem code
6856820 | 85735836 | D | IT WAS REPORTED THAT TWO PATIENTS DEVELOPED INFECTIONS POSSIBILITY RELATED TO VENIPUNCTURE PROCEDURES USING THE G X IN BD ECLIPSE BLOOD COLLECTION NEEDLE WITH LUER ADAPTER THE PATIENTS WERE PLACED ON ANTIBIOTICS BUT NO OTHER MEDICAL INTERVENTIONS WERE MENTIONED | Venipuncture | Patient problem code
6855566 | 85740519 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TARGET LESION AND TARGET VESSEL THE PATIENT WAS HOSPITALIZED FOR DAYS THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT J EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Device Problem Found | Medical Evaluation result code
6857280 | 85741813 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6857280 | 85741813 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6851715 | 85751399 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | Connection Problem | Device problem code
6851715 | 85751399 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Findings Available | Medical Evaluation result code
6854421 | 85752786 | N | A RISK TO THE PATIENTS HEALTH COULD NOT BE EXCLUDED FOR THESE SPECIFIC CIRCUMSTANCES SINCE A BIOPSY WAS APPLIED IN A DIFFERENT LOCATION IN THE BRAIN THAN ANTICIPATED WITH THE BRAINLAB DEVICE INVOLVED AND THE DESIRED DIAGNOSTIC SAMPLE WAS NOT RETRIEVED DESPITE ACCORDING TO THE SURGEON THE SURGERY WAS TO RETRIEVE A DIAGNOSTIC SAMPLE NOT TO REMOVE OR TREAT THE LESION THE BURR HOLE WAS IN THE INTENDED LOCATION AND THERE WERE NO FURTHER UNPLANNED BIOPSY PATHS APPLIED AT THIS SURGERY THERE WAS NO DIRECT OR INCREASED RISK TO HARM A CRITICAL BRAIN STRUCTURE EG BLOOD VESSELS OR IMPORTANT BRAIN TISSUE DUE TO THE APPLIED BIOPSY OR PATH THERE WERE NO NEGATIVE EFFECTS TO THE PATIENT NEITHER DUE TO BIOPSIES NOR SURGERYANESTHESIA TIME OF CA MIN OTHER THAN PLANNED REVISION BIOPSY SURGERY THERE ARE NO REMEDIAL ACTIONS NECESSARY DONE OR PLANNED FOR THIS PATIENT HOSPITALIZATION DID NOT NEED TO BE PROLONGED EITHER ACCORDING TO THE RESULTS OF THE BRAINLAB INVESTIGATION AND THE INFORMATION PROVIDED BY THE HOSPITAL IT CAN BE CONCLUDED THAT THE ROOT CAUSE FOR THE BIOPSIES TAKEN IN A LOCATION DIFFERENT THAN ANTICIPATED WITH NAVIGATION INVOLVED IS THAT IS THAT THE PATIENTS HEAD IN THE HEAD HOLDER MOST LIKELY MOVED DURING THE SURGERY AS A FURTHER CONTRIBUTING FACTOR A SHIFT OF THE PATIENTS BRAIN MIGHT HAVE OCCURRED IN BETWEEN THE PREOP MRI AND THE ANATOMICAL SITUATION DURING THE SURGERY EG DUE TO THE BONE FLAP AND OR LOSS OF CEREBROSPINAL FLUID APPARENTLY THE RESULTING DEVIATION BETWEEN DISPLAYED NAVIGATION INFORMATION AND THE ACTUAL PATIENT ANATOMY WAS NOT DETECTED WITH THE NECESSARY ACCURACY VERIFICATION TO BE PERFORMED BY THE USER THERE IS NO INDICATION OF A SYSTEMATIC ERROR OR MALFUNCTION OF THE BRAINLAB DEVICE NAVIGATION CORRESPONDING BRAINLAB MEASURES TO MINIMIZE THIS ANTICIPATED RISK AS LOW AS REASONABLY PRACTICABLE ARE ALREADY IN PLACE BRAINLAB INTENDS TO REITERATE THE RELEVANT TOPICS REGARDING THE USE OF THE DEVICE TO THIS CUSTOMER | No Device Problem Found | Medical Evaluation result code
6854035 | 85758040 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHES ON LCD WINDOW B | Scratched Material | Device problem code
6854035 | 85758040 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHES ON LCD WINDOW B | Component Missing | Device problem code
6854035 | 85758040 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHES ON LCD WINDOW B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6857560 | 85758072 | N | B THE INTERVENTION PERFORMED ON B WAS REPORTED UNDER MANUFACTURER REPORT NUMBER THEREFORE THIS REPORT WILL CAPTURE THE PROXIMAL TYPE I ENDOLEAK IDENTIFIED ON B AND THE INTERVENTION PERFORMED ON B TO TREAT THE ENDOLEAK B | Insufficient Information | Device problem code
6840926 | 85758940 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6840926 | 85758940 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6840929 | 85759111 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6840929 | 85759111 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6853426 | 85762806 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW | Scratched Material | Device problem code
6853426 | 85762806 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW | Component Missing | Device problem code
6853426 | 85762806 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
6853426 | 85762806 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
6851622 | 85764117 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6851622 | 85764117 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6852433 | 85766073 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6852433 | 85766073 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6848008 | 85775148 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT FELT INSULIN ON THEIR SKIN AND NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | Positioning Problem | Device problem code
6848008 | 85775148 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT FELT INSULIN ON THEIR SKIN AND NOTED THE CANNULA WAS NOT INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6852918 | 85775551 | D | THE USER FACILITY REPORTED AFTER A COMPLETED STERILIZATION CYCLE AN EMPLOYEE REMOVED AN INSTRUMENT TRAY FOR USE A SECOND EMPLOYEE HANDLED THE INSTRUMENT TRAY AND FELT A BURNING SENSATION ON THEIR HANDS THE EMPLOYEE SOUGHT AND RECEIVED MEDICAL TREATMENT | Burning Sensation | Patient problem code
6852918 | 85775551 | D | THE USER FACILITY REPORTED AFTER A COMPLETED STERILIZATION CYCLE AN EMPLOYEE REMOVED AN INSTRUMENT TRAY FOR USE A SECOND EMPLOYEE HANDLED THE INSTRUMENT TRAY AND FELT A BURNING SENSATION ON THEIR HANDS THE EMPLOYEE SOUGHT AND RECEIVED MEDICAL TREATMENT | Insufficient Information | Device problem code
6839488 | 85778320 | N | PATIENT WEIGHT IS NOT AVAILABLE FROM THE FACILITY THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION UPON VISUAL INSPECTION AND SIMULATED USE TESTING IT WAS FOUND THAT THERE WAS A BLOCKAGE AT THE HUB WHEN USING A GUIDEWIRE A GUIDEWIRE PASSED THE HUB BUT RESISTANCE WAS FELT IN THE WORKING LENGTH DUE TO A BUNCHING OF THE DEVICE THAT OCCURRED DURING DEVICE RETURN | Partial Blockage | Device problem code
6839488 | 85778320 | N | PATIENT WEIGHT IS NOT AVAILABLE FROM THE FACILITY THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION UPON VISUAL INSPECTION AND SIMULATED USE TESTING IT WAS FOUND THAT THERE WAS A BLOCKAGE AT THE HUB WHEN USING A GUIDEWIRE A GUIDEWIRE PASSED THE HUB BUT RESISTANCE WAS FELT IN THE WORKING LENGTH DUE TO A BUNCHING OF THE DEVICE THAT OCCURRED DURING DEVICE RETURN | Operational Problem Identified | Medical Evaluation result code
6852869 | 85780825 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Defective Alarm | Device Problem Code
6852869 | 85780825 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Mechanical Problem Identified | Medical Evaluation result code
6858178 | 85785668 | N | DEVICE EVALUATIONS OF MONITOR SN B AND BELT SN B HAVE BEEN COMPLETED DURING THE INCOMING FUNCTIONAL TESTING OF THE ELECTRODE BELT THE ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WERE VERIFIED THE MONITOR FAILED INCOMING TESTING UPON EVALUATION THE R RESISTOR WAS OPEN THE OPEN RESISTOR IS CONSISTENT WITH THE REPORTED EXTERNAL DEFIBRILLATIONS THERE IS NO INDICATION THAT THE OPEN RESISTOR CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH DEVICE MANUFACTURE DATE MONITOR B BELT | No Display / Image | Device problem code
6858178 | 85785668 | N | DEVICE EVALUATIONS OF MONITOR SN B AND BELT SN B HAVE BEEN COMPLETED DURING THE INCOMING FUNCTIONAL TESTING OF THE ELECTRODE BELT THE ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WERE VERIFIED THE MONITOR FAILED INCOMING TESTING UPON EVALUATION THE R RESISTOR WAS OPEN THE OPEN RESISTOR IS CONSISTENT WITH THE REPORTED EXTERNAL DEFIBRILLATIONS THERE IS NO INDICATION THAT THE OPEN RESISTOR CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH DEVICE MANUFACTURE DATE MONITOR B BELT | No Findings Available | Medical Evaluation result code
6858229 | 85786629 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A HEAT RASH UNDER THE LIFEVEST THE PATIENTS PHYSICIAN RECOMMENDED A CREAM FOR THE RASH FOLLOWUP INDICATED THAT THE RASH WAS IMPROVING | Rash | Patient problem code
6858229 | 85786629 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A HEAT RASH UNDER THE LIFEVEST THE PATIENTS PHYSICIAN RECOMMENDED A CREAM FOR THE RASH FOLLOWUP INDICATED THAT THE RASH WAS IMPROVING | Insufficient Information | Device problem code
6841087 | 85787699 | D | IT WAS REPORTED THAT DURING A CONTROLLER CHECK IT WAS NOTED THAT THE TIMEDATE HAD RESET TO THE DEFAULT SETTING IT WAS PRESUMED THE INTERNAL BATTERY HAD DISCHARGED THE DEVICE WAS EXCHANGED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6841087 | 85787699 | D | IT WAS REPORTED THAT DURING A CONTROLLER CHECK IT WAS NOTED THAT THE TIMEDATE HAD RESET TO THE DEFAULT SETTING IT WAS PRESUMED THE INTERNAL BATTERY HAD DISCHARGED THE DEVICE WAS EXCHANGED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Insufficient Information | Device problem code
6851708 | 85788811 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | Connection Problem | Device problem code
6851708 | 85788811 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON AUG TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Findings Available | Medical Evaluation result code
6841283 | 85789563 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A FEW HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6841283 | 85789563 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE CUSTOMER NOTED THE ADHESIVE FAILED WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A FEW HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6853439 | 85794158 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE | Connection Problem | Device problem code
6853439 | 85794158 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE | No Findings Available | Medical Evaluation result code
6858499 | 85797269 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B CUSTOMER WAS NOT SURE WHAT DAY DUE TO HIGH BLOOD GLUCOSE THE CUSTOMER EXPERIENCED NO SYMPTOMS BLOOD GLUCOSE AT THE TIME OF THE ADMISSION WAS MGDL THE CUSTOMER WAS TREATED WITH MANUAL INJECTION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT WAS WITHOUT PUMP LESS THAN HOURS CUSTOMER ALSO REPORTS AIR BUBBLES THAT WERE REMOVED WITH TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Appropriate Term/Code Not Available | Device problem code
6858499 | 85797269 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED ON B CUSTOMER WAS NOT SURE WHAT DAY DUE TO HIGH BLOOD GLUCOSE THE CUSTOMER EXPERIENCED NO SYMPTOMS BLOOD GLUCOSE AT THE TIME OF THE ADMISSION WAS MGDL THE CUSTOMER WAS TREATED WITH MANUAL INJECTION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION BUT WAS WITHOUT PUMP LESS THAN HOURS CUSTOMER ALSO REPORTS AIR BUBBLES THAT WERE REMOVED WITH TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6857421 | 85853918 | N | THE A ANESTHESIA SYSTEM UNIT IN OPERATING ROOM AND OTHER A ANESTHESIA SYSTEM UNITS WERE TESTED ON THE SITE BY MINDRAYS FIELD SERVICE REPRESENTATIVE THE UNITS AND VAPORIZERS PASSED ALL LEAK AND FUNCTIONAL TESTS PER SPECIFICATIONS NO LEAKS OR OTHER MALFUNCTION WERE IDENTIFIED HOWEVER THE A ANESTHESIA SYSTEMS IN OPERATING ROOM AND OPERATING ROOM WERE FOUND TO HAVE THE SCAVENGER VALVE TURNED OFF THE INVESTIGATION FINDINGS AND THE LOGS OF THE FOUR A ANESTHESIA SYSTEM UNITS WERE COLLECTED AND SENT TO THE MANUFACTURER FOR FURTHER EVALUATION THE MANUFACTURER DID NOT IDENTIFY ANY ANOMALY IN THE LOGS AND CONCLUDED THAT DUE TO THE SCAVENGER VALVE BEING TURNED OFF DURING USE IN OPERATING ROOM WASTE GAS WAS EMITTED INTO THE OPERATING ROOM ON THE DAY OF THE REPORTED EVENT THE OPERATING MANUAL INSTRUCTS THE USER TO INSPECT THE SCAVENGER SYSTEM PRIOR TO USE THEREFORE THE LACK OF THE PREOPERATIVE TESTS APPEAR TO HAVE CONTRIBUTED TO THE REPORTED EVENT B | Leak / Splash | Device problem code
6857421 | 85853918 | N | THE A ANESTHESIA SYSTEM UNIT IN OPERATING ROOM AND OTHER A ANESTHESIA SYSTEM UNITS WERE TESTED ON THE SITE BY MINDRAYS FIELD SERVICE REPRESENTATIVE THE UNITS AND VAPORIZERS PASSED ALL LEAK AND FUNCTIONAL TESTS PER SPECIFICATIONS NO LEAKS OR OTHER MALFUNCTION WERE IDENTIFIED HOWEVER THE A ANESTHESIA SYSTEMS IN OPERATING ROOM AND OPERATING ROOM WERE FOUND TO HAVE THE SCAVENGER VALVE TURNED OFF THE INVESTIGATION FINDINGS AND THE LOGS OF THE FOUR A ANESTHESIA SYSTEM UNITS WERE COLLECTED AND SENT TO THE MANUFACTURER FOR FURTHER EVALUATION THE MANUFACTURER DID NOT IDENTIFY ANY ANOMALY IN THE LOGS AND CONCLUDED THAT DUE TO THE SCAVENGER VALVE BEING TURNED OFF DURING USE IN OPERATING ROOM WASTE GAS WAS EMITTED INTO THE OPERATING ROOM ON THE DAY OF THE REPORTED EVENT THE OPERATING MANUAL INSTRUCTS THE USER TO INSPECT THE SCAVENGER SYSTEM PRIOR TO USE THEREFORE THE LACK OF THE PREOPERATIVE TESTS APPEAR TO HAVE CONTRIBUTED TO THE REPORTED EVENT B | Leakage/Seal | Medical Evaluation result code
6859671 | 85854421 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6859468 | 85859747 | N | THE HILLROM TECHNICIAN FOUND THERE WERE TOO MANY BEADS AND THE BEADS NEEDED TO BE DECONTAMINATED THE TECHNICIAN SWAPPED THE BED AND DECONTAMINATED THE BEADS TO RESOLVE THE ISSUE THE PATIENT HAD A STAGE SACROCOCCYGEAL WOUND THAT PROGRESSED TO UNSTAGEABLE THE PATIENT IS BEING TREATED WITH DAKINS SOLUTION DEVELOPMENT OF PRESSURE ULCERS IS MULTIFACTORIAL AND CANNOT BE ONLY ATTRIBUTED TO PERFORMANCE OF THE SURFACE RISK FACTORS INCLUDE PROTEINCALORIE MALNUTRITION MICROCLIMATE SKIN WETNESS CAUSED BY SWEATING OR INCONTINENCE DISEASES THAT REDUCE BLOOD FLOW TO THE SKIN SUCH AS ARTERIOSCLEROSIS OR DISEASES THAT REDUCE THE SENSATION IN THE SKIN SUCH AS PARALYSIS OR NEUROPATHY POSITION CHANGES ARE KEY TO PRESSURE SORE PREVENTION AND TREATMENT THESE CHANGES NEED TO BE FREQUENT REPOSITIONING NEEDS TO AVOID STRESS ON THE SKIN AND BODY POSITIONS NEED TO MINIMIZE THE RISK OF PRESSURE ON VULNERABLE AREAS PER THE HILLROM USER MANUAL POOR FLUIDIZATION IF FLUIDIZATION IS SLUGGISH OR UNEVEN NOTIFY YOUR HILLROM REPRESENTATIVE FLUIDIZATION IS AFFECTED BY THE FOLLOWING ROOM TEMPERATURE HUMIDITY THE AMOUNT OF MATERIALS SUCH AS FLUID CELLS OR CELLULAR DEBRIS WHICH HAS ESCAPED FROM THE BLOOD VESSELS AND HAS BEEN DEPOSITED IN TISSUES OR ON TISSUE SURFACES RESTRICTED AIR CIRCULATION FROM BLANKETS ON THE BED A SEARCH OF THE HILLROM MAINTENANCE RECORDS SHOWED HILLROM PERFORMED PREVENTATIVE MAINTENANCE ON THIS BED IN IT IS UNKNOWN IF THE FACILITY PERFORMED ANY OTHER PREVENTATIVE MAINTENANCE ON THIS BED BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | Temperature Problem | Device Problem Code
6859468 | 85859747 | N | THE HILLROM TECHNICIAN FOUND THERE WERE TOO MANY BEADS AND THE BEADS NEEDED TO BE DECONTAMINATED THE TECHNICIAN SWAPPED THE BED AND DECONTAMINATED THE BEADS TO RESOLVE THE ISSUE THE PATIENT HAD A STAGE SACROCOCCYGEAL WOUND THAT PROGRESSED TO UNSTAGEABLE THE PATIENT IS BEING TREATED WITH DAKINS SOLUTION DEVELOPMENT OF PRESSURE ULCERS IS MULTIFACTORIAL AND CANNOT BE ONLY ATTRIBUTED TO PERFORMANCE OF THE SURFACE RISK FACTORS INCLUDE PROTEINCALORIE MALNUTRITION MICROCLIMATE SKIN WETNESS CAUSED BY SWEATING OR INCONTINENCE DISEASES THAT REDUCE BLOOD FLOW TO THE SKIN SUCH AS ARTERIOSCLEROSIS OR DISEASES THAT REDUCE THE SENSATION IN THE SKIN SUCH AS PARALYSIS OR NEUROPATHY POSITION CHANGES ARE KEY TO PRESSURE SORE PREVENTION AND TREATMENT THESE CHANGES NEED TO BE FREQUENT REPOSITIONING NEEDS TO AVOID STRESS ON THE SKIN AND BODY POSITIONS NEED TO MINIMIZE THE RISK OF PRESSURE ON VULNERABLE AREAS PER THE HILLROM USER MANUAL POOR FLUIDIZATION IF FLUIDIZATION IS SLUGGISH OR UNEVEN NOTIFY YOUR HILLROM REPRESENTATIVE FLUIDIZATION IS AFFECTED BY THE FOLLOWING ROOM TEMPERATURE HUMIDITY THE AMOUNT OF MATERIALS SUCH AS FLUID CELLS OR CELLULAR DEBRIS WHICH HAS ESCAPED FROM THE BLOOD VESSELS AND HAS BEEN DEPOSITED IN TISSUES OR ON TISSUE SURFACES RESTRICTED AIR CIRCULATION FROM BLANKETS ON THE BED A SEARCH OF THE HILLROM MAINTENANCE RECORDS SHOWED HILLROM PERFORMED PREVENTATIVE MAINTENANCE ON THIS BED IN IT IS UNKNOWN IF THE FACILITY PERFORMED ANY OTHER PREVENTATIVE MAINTENANCE ON THIS BED BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | Contamination of Device Ingredient or Reagent | Device problem code
6859468 | 85859747 | N | THE HILLROM TECHNICIAN FOUND THERE WERE TOO MANY BEADS AND THE BEADS NEEDED TO BE DECONTAMINATED THE TECHNICIAN SWAPPED THE BED AND DECONTAMINATED THE BEADS TO RESOLVE THE ISSUE THE PATIENT HAD A STAGE SACROCOCCYGEAL WOUND THAT PROGRESSED TO UNSTAGEABLE THE PATIENT IS BEING TREATED WITH DAKINS SOLUTION DEVELOPMENT OF PRESSURE ULCERS IS MULTIFACTORIAL AND CANNOT BE ONLY ATTRIBUTED TO PERFORMANCE OF THE SURFACE RISK FACTORS INCLUDE PROTEINCALORIE MALNUTRITION MICROCLIMATE SKIN WETNESS CAUSED BY SWEATING OR INCONTINENCE DISEASES THAT REDUCE BLOOD FLOW TO THE SKIN SUCH AS ARTERIOSCLEROSIS OR DISEASES THAT REDUCE THE SENSATION IN THE SKIN SUCH AS PARALYSIS OR NEUROPATHY POSITION CHANGES ARE KEY TO PRESSURE SORE PREVENTION AND TREATMENT THESE CHANGES NEED TO BE FREQUENT REPOSITIONING NEEDS TO AVOID STRESS ON THE SKIN AND BODY POSITIONS NEED TO MINIMIZE THE RISK OF PRESSURE ON VULNERABLE AREAS PER THE HILLROM USER MANUAL POOR FLUIDIZATION IF FLUIDIZATION IS SLUGGISH OR UNEVEN NOTIFY YOUR HILLROM REPRESENTATIVE FLUIDIZATION IS AFFECTED BY THE FOLLOWING ROOM TEMPERATURE HUMIDITY THE AMOUNT OF MATERIALS SUCH AS FLUID CELLS OR CELLULAR DEBRIS WHICH HAS ESCAPED FROM THE BLOOD VESSELS AND HAS BEEN DEPOSITED IN TISSUES OR ON TISSUE SURFACES RESTRICTED AIR CIRCULATION FROM BLANKETS ON THE BED A SEARCH OF THE HILLROM MAINTENANCE RECORDS SHOWED HILLROM PERFORMED PREVENTATIVE MAINTENANCE ON THIS BED IN IT IS UNKNOWN IF THE FACILITY PERFORMED ANY OTHER PREVENTATIVE MAINTENANCE ON THIS BED BASED ON THIS INFORMATION NO FURTHER ACTION IS REQUIRED | Environmental Temperature Problem Identified | Medical Evaluation result code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Scratched Material | Device problem code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Deformation Due to Compressive Stress | Device problem code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6859597 | 85860498 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF TEMPORARY PARALYSIS AND THE LIP CORNER TURNED UP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS ADVERSE EVENTS TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Stress Problem Identified | Medical Evaluation result code
6859975 | 85861984 | N | DEVICE EVALUATION PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS AND A PIECE OF THE HAPTIC MISSING B | Component Missing | Device problem code
6859975 | 85861984 | N | DEVICE EVALUATION PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS AND A PIECE OF THE HAPTIC MISSING B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6860711 | 85903964 | D | ENDOLOGIX WAS MADE AWARE OF A POSSIBLE ENDOLEAK TYPE IIIA OF THE PROXIMAL AND MAIN BODY STENT PATIENT WAS INITIALLY SCHEDULED FOR A SECONDARY ON B BUT SUFFERED A STROKE SECONDARY WAS CANCELED PATIENT IS CURRENTLY BEING MONITORED | Hemorrhage | Patient problem code
6860711 | 85903964 | D | ENDOLOGIX WAS MADE AWARE OF A POSSIBLE ENDOLEAK TYPE IIIA OF THE PROXIMAL AND MAIN BODY STENT PATIENT WAS INITIALLY SCHEDULED FOR A SECONDARY ON B BUT SUFFERED A STROKE SECONDARY WAS CANCELED PATIENT IS CURRENTLY BEING MONITORED | Insufficient Information | Device problem code
6860711 | 85903964 | D | ENDOLOGIX WAS MADE AWARE OF A POSSIBLE ENDOLEAK TYPE IIIA OF THE PROXIMAL AND MAIN BODY STENT PATIENT WAS INITIALLY SCHEDULED FOR A SECONDARY ON B BUT SUFFERED A STROKE SECONDARY WAS CANCELED PATIENT IS CURRENTLY BEING MONITORED | Cerebrovascular Accident | Patient problem code
6857302 | 85904563 | N | EXAMINATION OF THE ITEM REVEALED A MICROSCOPIC DEFECT IN THE ORING OF THE PUMP EXAMINATION OF THE PRODUCTION RECORDS OF THE INVOLVED DEVICE DID NOT REVEAL A DEVIATION THAT MAY HAVE CONTRIBUTED TO THE REPORTED CASE THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY NOTE THIS PRODUCT IS NO LONGER MARKETED IN THE USA | Defective Device | Device problem code
6857302 | 85904563 | N | EXAMINATION OF THE ITEM REVEALED A MICROSCOPIC DEFECT IN THE ORING OF THE PUMP EXAMINATION OF THE PRODUCTION RECORDS OF THE INVOLVED DEVICE DID NOT REVEAL A DEVIATION THAT MAY HAVE CONTRIBUTED TO THE REPORTED CASE THIS REPORT IS SUBMITTED FOLLOWING AN FDA INSPECTION AT THE MANUFACTURER FACILITY NOTE THIS PRODUCT IS NO LONGER MARKETED IN THE USA | No Findings Available | Medical Evaluation result code
6857601 | 85905629 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6857601 | 85905629 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6860979 | 85906091 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A PATIENT REGARDING AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR PATIENT EXPERIENCED PAIN AROUND THEIR TAILBONE AREA AND REPROGRAMMING DID NOT RESOLVE IT AS A RESULT OF THE EVENT THE HEALTHCARE PROVIDER HCP DECIDED TO REPLACE THE LEAD NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6860979 | 85906091 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A PATIENT REGARDING AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR PATIENT EXPERIENCED PAIN AROUND THEIR TAILBONE AREA AND REPROGRAMMING DID NOT RESOLVE IT AS A RESULT OF THE EVENT THE HEALTHCARE PROVIDER HCP DECIDED TO REPLACE THE LEAD NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6860979 | 85906091 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A PATIENT REGARDING AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR PATIENT EXPERIENCED PAIN AROUND THEIR TAILBONE AREA AND REPROGRAMMING DID NOT RESOLVE IT AS A RESULT OF THE EVENT THE HEALTHCARE PROVIDER HCP DECIDED TO REPLACE THE LEAD NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6856963 | 85914897 | N | B THE INSULIN PUMP WAS RECEIVED WITH A DETACHED END CAP THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE DISPLAY WINDOW AND CASE HOWEVER THE INSULIN PUMP PASSED THE DISPLACEMENT TEST | Scratched Material | Device problem code
6856963 | 85914897 | N | B THE INSULIN PUMP WAS RECEIVED WITH A DETACHED END CAP THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE DISPLAY WINDOW AND CASE HOWEVER THE INSULIN PUMP PASSED THE DISPLACEMENT TEST | Detachment of Device or device Component | Device problem code
6856963 | 85914897 | N | B THE INSULIN PUMP WAS RECEIVED WITH A DETACHED END CAP THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE DISPLAY WINDOW AND CASE HOWEVER THE INSULIN PUMP PASSED THE DISPLACEMENT TEST | Stress Problem Identified | Medical Evaluation result code
6857583 | 85918921 | N | THE MANUFACTURER IS WAITING FOR THE ARRIVAL OF THE AFFECTED ADMINISTRATION SET THE FOREIGN OBJECT IDENTIFIED COULD BE BURNT PLASTIC DURING MANUFACTURING PROCESS WHICH WILL NEED FURTHER INVESTIGATION ZYNO MEDICAL HAS SENT A QUALITY ALERT TO ITS CONTRACT SUPPLIER ON REGARDING THIS COMPLAINT A SCAR B IS ISSUED TO ADDRESS THIS COMPLAINT | Power Problem | Device problem code
6857583 | 85918921 | N | THE MANUFACTURER IS WAITING FOR THE ARRIVAL OF THE AFFECTED ADMINISTRATION SET THE FOREIGN OBJECT IDENTIFIED COULD BE BURNT PLASTIC DURING MANUFACTURING PROCESS WHICH WILL NEED FURTHER INVESTIGATION ZYNO MEDICAL HAS SENT A QUALITY ALERT TO ITS CONTRACT SUPPLIER ON REGARDING THIS COMPLAINT A SCAR B IS ISSUED TO ADDRESS THIS COMPLAINT | Power Source Problem Identified | Medical Evaluation result code
6857392 | 85927453 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BLOOD GLUCOSE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK | Hyperglycemia | Patient problem code
6857392 | 85927453 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BLOOD GLUCOSE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK | Positioning Problem | Device problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Nausea | Patient problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Headache | Patient problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Pain | Patient problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Sweating | Patient problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Dizziness | Patient problem code
6857662 | 85931338 | D | I HAD A PAIN IN MY LEFT FALLOPIAN TUBE WHERE THE ESSURE DEVICE WAS IMPLANTED I HAD NAUSEA SEVERE DIZZINESS AND LIGHTHEADEDNESS ALONG WITH SWEATING | Insufficient Information | Device problem code
6861362 | 85936551 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | Defective Device | Device Problem Code
6861362 | 85936551 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | No Consequences Or Impact To Patient | Patient problem code
6861369 | 85936610 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | No Consequences Or Impact To Patient | Patient problem code
6861369 | 85936610 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION PLEASE REFERENCE MEDWATCH REPORT FOR A SIMILAR EVENT REPORTED FROM THE SAME CUSTOMER | Defective Device | Device Problem Code
6840087 | 85943032 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THE PATIENT WAS EXPERIENCING INTERMITTENT BEEPING ONE CONTROLLER B AND TWO BATTERIES B WERE RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICES REVEALED THAT THE CONTROLLER AND BATTERIES PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING LOG FILE ANALYSIS REVEALED THE CONTROLLER B DID NOT CONTAIN THE SMR SOFTWARE ANALYSIS OF THE DATA LOG FILES REVEALED SEVERAL PREMATURE POWER SWITCHING EVENTS THAT WERE DUE TO COMMUNICATION ERRORS ANDOR MOMENTARY DISCONNECTIONS INVOLVING B MOMENTARY DISCONNECTIONS WILL RESULT IN AN AUDIBLE TONE THE MOST LIKELY ROOT CAUSE OF THE POWER SWITCHING EVENTS ARE MOST OFTEN ATTRIBUTED TO COMMUNICATION ERRORS ANDOR MOMENTARY DISCONNECTIONS HOWEVER THE REPORTED BEEPS ARE MOST LIKELY ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND THE BATTERY TWO INTERNAL INVESTIGATIONS HAVE BEEN OPENED TO INVESTIGATE COMMUNICATION ERRORS ON NONSMR CONTROLLERS AND MOMENTARY DISCONNECTIONS ADDITIONAL DEVICES B BATTERY CATALOG NUMBER EXPIRATION DATE FEB B YES PRODUCT RECEIVED ON AUG B B BATTERY CATALOG NUMBER EXPIRATION DATE FEB B YES PRODUCT RECEIVED ON AUG B HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Noise, Audible | Device problem code
6840087 | 85943032 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THE PATIENT WAS EXPERIENCING INTERMITTENT BEEPING ONE CONTROLLER B AND TWO BATTERIES B WERE RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICES REVEALED THAT THE CONTROLLER AND BATTERIES PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING LOG FILE ANALYSIS REVEALED THE CONTROLLER B DID NOT CONTAIN THE SMR SOFTWARE ANALYSIS OF THE DATA LOG FILES REVEALED SEVERAL PREMATURE POWER SWITCHING EVENTS THAT WERE DUE TO COMMUNICATION ERRORS ANDOR MOMENTARY DISCONNECTIONS INVOLVING B MOMENTARY DISCONNECTIONS WILL RESULT IN AN AUDIBLE TONE THE MOST LIKELY ROOT CAUSE OF THE POWER SWITCHING EVENTS ARE MOST OFTEN ATTRIBUTED TO COMMUNICATION ERRORS ANDOR MOMENTARY DISCONNECTIONS HOWEVER THE REPORTED BEEPS ARE MOST LIKELY ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND THE BATTERY TWO INTERNAL INVESTIGATIONS HAVE BEEN OPENED TO INVESTIGATE COMMUNICATION ERRORS ON NONSMR CONTROLLERS AND MOMENTARY DISCONNECTIONS ADDITIONAL DEVICES B BATTERY CATALOG NUMBER EXPIRATION DATE FEB B YES PRODUCT RECEIVED ON AUG B B BATTERY CATALOG NUMBER EXPIRATION DATE FEB B YES PRODUCT RECEIVED ON AUG B HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Non-Functional Defect | Medical Evaluation result code
6861550 | 85949370 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT PLEASE CONFIRM THAT THE SURGEON BELIEVED THE STAPLING WAS ADEQUATE WAS THE PATIENTS HOSPITALIZATION EXTENDED DUE TO DEVICE ISSUE WHAT IS THE CURRENT PATIENT STATUS | No Device Problem Found | Medical Evaluation result code
6861672 | 85952588 | D | A REPORT WAS RECEIVED THAT DURING TRIAL PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED TRIAL LEAD THE BSN SALES REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED BUT WAS USED DURING THE PROCEDURE | No Information | Patient problem code
6861672 | 85952588 | D | A REPORT WAS RECEIVED THAT DURING TRIAL PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED TRIAL LEAD THE BSN SALES REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED BUT WAS USED DURING THE PROCEDURE | Defective Component | Device problem code
6861670 | 85953400 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6861670 | 85953400 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6860651 | 85954883 | N | B ANALYSIS WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO A MISSING RESERVOIR RETAINER THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE LCD WINDOW A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND A MISSING RESERVOIR TUBE ORING | Scratched Material | Device problem code
6860651 | 85954883 | N | B ANALYSIS WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO A MISSING RESERVOIR RETAINER THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE LCD WINDOW A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND A MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
6860651 | 85954883 | N | B ANALYSIS WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO A MISSING RESERVOIR RETAINER THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHES ON THE LCD WINDOW A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND A MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
6858816 | 85956729 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6858816 | 85956729 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6857388 | 85958164 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON SITE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN THE POD IS NOT AVAILABLE FOR RETURN AS IT HAS BEEN DISCARDED BY THE PATIENT | Positioning Problem | Device problem code
6857388 | 85958164 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON SITE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN THE POD IS NOT AVAILABLE FOR RETURN AS IT HAS BEEN DISCARDED BY THE PATIENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6862081 | 85995510 | N | AS THE EXACT DATE OF THE EVENT IS UNKNOWN B THE DATE OF CONFERENCE PRESENTATION WAS ESTIMATED TO BE THE DATE OF EVENT B AS LOT NUMBERS WERE UNAVAILABLE A UDI WAS NOT PROVIDED THE INSTRUCTIONS FOR USE IFU FOR THE GORE EXCLUDER AAA ENDOPROSTHESIS STATES ADVERSE EVENTS THAT MAY OCCUR AND OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK | Insufficient Information | Device problem code
6859337 | 86004607 | N | B FROM THE PRELIMINARY INVESTIGATION THE DESCRIPTION OF THE ISSUE IS SHEATH BODY SEPARATION | Material Separation | Device problem code
6859337 | 86004607 | N | B FROM THE PRELIMINARY INVESTIGATION THE DESCRIPTION OF THE ISSUE IS SHEATH BODY SEPARATION | Fracture Problem | Medical Evaluation result code
6851824 | 86007919 | N | PATIENT AGE DATE OF BIRTH GENDER AND WEIGHT ARE UNAVAILABLE THE DEVICE AND LEAD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION THE DEVICE WAS STILL LOCKED WITHIN THE LEAD UPON RECEIPT TO THE MANUFACTURER VISUAL INSPECTION OF THE DEVICE FOUND EVIDENCE OF THE LEAD BEING COMPROMISED IT IS PROBABLE THAT THE AGE OF THE LEAD CONTRIBUTED TO THE DEGRADATION IN INTEGRITY NO EVIDENCE WAS OBSERVED TO INDICATE THAT THE LLD DEVICE MALFUNCTIONED TO CAUSE THIS EVENT | No Device Problem Found | Medical Evaluation result code
6851636 | 86008588 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6851636 | 86008588 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6851632 | 86008592 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6851632 | 86008592 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6840457 | 86009548 | D | OLYMPUS WAS INFORMED THAT DURING PREPARATION FOR USE THE SCOPE WAS FOUND WITH A CUT ON THE BENDING SECTION RUBBER THERE WAS NO PATIENT INVOLVEMENT WITH THE REPORTED SCOPE | Material Twisted / Bent | Device problem code
6840457 | 86009548 | D | OLYMPUS WAS INFORMED THAT DURING PREPARATION FOR USE THE SCOPE WAS FOUND WITH A CUT ON THE BENDING SECTION RUBBER THERE WAS NO PATIENT INVOLVEMENT WITH THE REPORTED SCOPE | No Patient Involvement | Patient problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Thrombosis | Patient problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Ischemia | Patient problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Insufficient Information | Device problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Thrombus | Patient problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Occlusion | Patient problem code
6862615 | 86071736 | D | THE PATIENT HAD PREVIOUSLY BEEN TREATED FOR AN ABDOMINAL AORTIC ANEURYSM AAA AND A RIGHT INTERNAL ILIAC OCCLUSION USING A SELF OR BALLOON EXPANDABLE UNCOVERED STENT THAT EXTENDED FROM THE PATIENTS AORTIC BIFURCATION ALL THE WAY INTO THE EXTERNAL ILIAC IN B A FENESTRATED ENDOVASCULAR ANEURYSM REPAIR FEVAR WAS PERFORMED USING A ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE THIS WAS PERFORMED BY A DIFFERENT PHYSICIAN AND AT ANOTHER FACILITY THE PATIENTS RT INTERNAL ILIAC WAS ALL READY OCCLUDED PRIOR TO THIS ZFEN PROCEDURE AS REPORTED TO COOK MEDICAL BY THE PATIENTS NEW PHYSICIAN THE PREVIOUSLY IMPLANTED ZENITH FLEX WITH SPIRALZ TECHNOLOGY AAA ENDOVASCULAR GRAFT ILIAC LEG ZSLE WAS THROMBOSED AND COMPLETELY OCCLUDED THE PATIENT REQUIRED EMERGENT INTERVENTION TO TREAT CRITICAL LIMB ISCHEMIA WITH AN OCCLUDED RIGHT ILIAC LIMB DURING THE INTERVENTION PROCEDURE THE PREVIOUSLY IMPLANTED ZSLE WAS REALIGNED WITH A LIMB FROM A DIFFERENT COMPANY THE PHYSICIAN USED BALLOON EXPANDABLE COVERED STENTS PLACED JUST PROXIMAL TO THE FLOW DIVIDER IN THE AORTOILIAC SEGMENTS AND INFLATED TO ATMOSPHERES AS THERE WERE STILL RESIDUAL THROMBUS A X AND A X OVERLAPPED COVERED BALLOON EXPANDABLE STENT WAS PLACED INTO THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC AN EXCELLENT ANGIOGRAPHIC RESULT CONFIRMED BRISK FLOW DOWN BOTH ILIAC LIMBS AND INTO THE COMMON FEMORAL BILATERALLY THE PATIENT WAS DISCHARGED AT A LATER DATE BUT DID PRESENT BACK IN THE HOSPITAL IT HAS NOT BEEN CLARIFIED YET IF THE PATIENTS LAST VISIT TO THE HOSPITAL WAS IN ANY WAY RELATED TO COMPLICATIONS FROM THIS REINTERVENTION | Aneurysm | Patient problem code
6863292 | 86072498 | N | THE WARNINGS IN THE PACKAGE INSERT STATE THIS TYPE OF EVENT CAN OCCUR WITHOUT A PRODUCT RETURN NO PRODUCT EVALUATION IS ABLE TO BE CONDUCTED NO XRAYS SCANS PICTURES OR PHYSICIANS REPORTS WERE PROVIDED BASED ON THE INFORMATION AVAILABLE THE MOST LIKELY CAUSE OF THE SCREWS BACKING OUT IS THE PATIENTS CONDITION IF ADDITIONAL INFORMATION IS OBTAINED THAT ADDS VALUE TO THE RELEVANT CONTENT OF THIS REPORT ANDOR A CONCLUSION CAN BE DRAWN A FOLLOWUP REPORT WILL BE SENT THIS IS REPORT FOUR OF SEVEN FOR THE SAME EVENT REFERENCE REPORTS THROUGH | Detachment of Device or device Component | Device problem code
6863292 | 86072498 | N | THE WARNINGS IN THE PACKAGE INSERT STATE THIS TYPE OF EVENT CAN OCCUR WITHOUT A PRODUCT RETURN NO PRODUCT EVALUATION IS ABLE TO BE CONDUCTED NO XRAYS SCANS PICTURES OR PHYSICIANS REPORTS WERE PROVIDED BASED ON THE INFORMATION AVAILABLE THE MOST LIKELY CAUSE OF THE SCREWS BACKING OUT IS THE PATIENTS CONDITION IF ADDITIONAL INFORMATION IS OBTAINED THAT ADDS VALUE TO THE RELEVANT CONTENT OF THIS REPORT ANDOR A CONCLUSION CAN BE DRAWN A FOLLOWUP REPORT WILL BE SENT THIS IS REPORT FOUR OF SEVEN FOR THE SAME EVENT REFERENCE REPORTS THROUGH | Fracture Problem | Medical Evaluation result code
6862989 | 86076814 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Material Discolored | Device problem code
6862989 | 86076814 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Findings Available | Medical Evaluation result code
6863298 | 86077223 | N | THE WARNINGS IN THE PACKAGE INSERT STATE THIS TYPE OF EVENT CAN OCCUR WITHOUT A PRODUCT RETURN NO PRODUCT EVALUATION IS ABLE TO BE CONDUCTED NO XRAYS SCANS PICTURES OR PHYSICIANS REPORTS WERE PROVIDED BASED ON THE INFORMATION AVAILABLE THE MOST LIKELY CAUSE OF THE SCREWS BACKING OUT IS THE PATIENTS CONDITION IF ADDITIONAL INFORMATION IS OBTAINED THAT ADDS VALUE TO THE RELEVANT CONTENT OF THIS REPORT ANDOR A CONCLUSION CAN BE DRAWN A FOLLOWUP REPORT WILL BE SENT THIS IS REPORT ONE OF SEVEN FOR THE SAME EVENT REFERENCE REPORTS THROUGH | No Findings Available | Medical Evaluation result code
6863309 | 86077601 | N | THE WARNINGS IN THE PACKAGE INSERT STATE THIS TYPE OF EVENT CAN OCCUR WITHOUT A PRODUCT RETURN NO PRODUCT EVALUATION IS ABLE TO BE CONDUCTED NO XRAYS SCANS PICTURES OR PHYSICIANS REPORTS WERE PROVIDED BASED ON THE INFORMATION AVAILABLE THE MOST LIKELY CAUSE OF THE SCREWS BACKING OUT IS THE PATIENTS CONDITION IF ADDITIONAL INFORMATION IS OBTAINED THAT ADDS VALUE TO THE RELEVANT CONTENT OF THIS REPORT ANDOR A CONCLUSION CAN BE DRAWN A FOLLOWUP REPORT WILL BE SENT THIS IS REPORT TWO OF SEVEN FOR THE SAME EVENT REFERENCE REPORTS THROUGH | No Findings Available | Medical Evaluation result code
6863307 | 86081012 | N | THE WARNINGS IN THE PACKAGE INSERT STATE THIS TYPE OF EVENT CAN OCCUR WITHOUT A PRODUCT RETURN NO PRODUCT EVALUATION IS ABLE TO BE CONDUCTED NO XRAYS SCANS PICTURES OR PHYSICIANS REPORTS WERE PROVIDED BASED ON THE INFORMATION AVAILABLE THE MOST LIKELY CAUSE OF THE SCREWS BACKING OUT IS THE PATIENTS CONDITION IF ADDITIONAL INFORMATION IS OBTAINED THAT ADDS VALUE TO THE RELEVANT CONTENT OF THIS REPORT ANDOR A CONCLUSION CAN BE DRAWN A FOLLOWUP REPORT WILL BE SENT THIS IS REPORT THREE OF SEVEN FOR THE SAME EVENT REFERENCE REPORTS THROUGH | No Findings Available | Medical Evaluation result code
6863260 | 86081979 | D | B THIS SPONTANEOUS CASE REPORTED BY A CONSUMER WHO CONTACTED THE COMPANY TO REPORT A PRODUCT COMPLAINT AND ADVERSE EVENTS CONCERNS A B FEMALE PATIENT OF UNKNOWN ETHNICITY MEDICAL HISTORY AND CONCOMITANT MEDICATIONS WERE NOT PROVIDED THE PATIENT RECEIVED INSULIN LISPRO HUMALOG CARTRIDGE UNKNOWN DOSE ROUTE FREQUENCY INDICATION AND START DATE SINCE UNKNOWN DATE UNSPECIFIED TIME AFTER BEGINNING TREATMENT WITH INSULIN LISPRO VIA HUMAPEN SAVVIO GRAY LOT UNKNOWN INSULIN LISPRO WAS NOT GETTING OUT OF DEVICE WHICH WAS NOT WORKING B HENCE INSULIN LISPRO HAD NOT BEEN INJECTED AND FOR THIS REASON PATIENT S GLUCOSE WAS VERY HIGH ON B BLOOD GLUCOSE INCREASED LED PATIENT TO HOSPITALIZATION DURING ONE WEEK ON B PATIENT WAS DISCHARGED INFORMATION REGARDING LABORATORIAL EXAMS CORRECTIVE TREATMENT AND OUTCOME WAS NOT PROVIDED AS OF AUG INSULIN LISPRO HAD BEEN DELIVERED VIA SYRINGE TREATMENT WITH INSULIN LISPRO WAS ONGOING IT WAS UNKNOWN WHO OPERATED THE DEVICE AND IF THE OPERATOR WAS TRAINED THE DEVICE MODEL AND THE REPORTED DEVICE HAD BEEN USED FOR UNSPECIFIED TIME DEVICE RETURN STATUS WAS UNKNOWN AND IT WAS DISCONTINUED THE REPORTING CONSUMER DID NOT PROVIDE ANY RELATEDNESS OPINION BETWEEN THE EVENTS AND INSULIN LISPRO UPDATE SEP UPDATED MEDWATCH FIELDS FOR EXPEDITED DEVICE REPORTING NO NEW INFORMATION ADDED | Insufficient Information | Device problem code
6863260 | 86081979 | D | B THIS SPONTANEOUS CASE REPORTED BY A CONSUMER WHO CONTACTED THE COMPANY TO REPORT A PRODUCT COMPLAINT AND ADVERSE EVENTS CONCERNS A B FEMALE PATIENT OF UNKNOWN ETHNICITY MEDICAL HISTORY AND CONCOMITANT MEDICATIONS WERE NOT PROVIDED THE PATIENT RECEIVED INSULIN LISPRO HUMALOG CARTRIDGE UNKNOWN DOSE ROUTE FREQUENCY INDICATION AND START DATE SINCE UNKNOWN DATE UNSPECIFIED TIME AFTER BEGINNING TREATMENT WITH INSULIN LISPRO VIA HUMAPEN SAVVIO GRAY LOT UNKNOWN INSULIN LISPRO WAS NOT GETTING OUT OF DEVICE WHICH WAS NOT WORKING B HENCE INSULIN LISPRO HAD NOT BEEN INJECTED AND FOR THIS REASON PATIENT S GLUCOSE WAS VERY HIGH ON B BLOOD GLUCOSE INCREASED LED PATIENT TO HOSPITALIZATION DURING ONE WEEK ON B PATIENT WAS DISCHARGED INFORMATION REGARDING LABORATORIAL EXAMS CORRECTIVE TREATMENT AND OUTCOME WAS NOT PROVIDED AS OF AUG INSULIN LISPRO HAD BEEN DELIVERED VIA SYRINGE TREATMENT WITH INSULIN LISPRO WAS ONGOING IT WAS UNKNOWN WHO OPERATED THE DEVICE AND IF THE OPERATOR WAS TRAINED THE DEVICE MODEL AND THE REPORTED DEVICE HAD BEEN USED FOR UNSPECIFIED TIME DEVICE RETURN STATUS WAS UNKNOWN AND IT WAS DISCONTINUED THE REPORTING CONSUMER DID NOT PROVIDE ANY RELATEDNESS OPINION BETWEEN THE EVENTS AND INSULIN LISPRO UPDATE SEP UPDATED MEDWATCH FIELDS FOR EXPEDITED DEVICE REPORTING NO NEW INFORMATION ADDED | No Known Impact Or Consequence To Patient | Patient Problem Code
6859164 | 86086623 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW NO TRACES OF MOISTURE WERE NOTED AT ELECTRONIC OR MOTOR ASSEMBLIES PER VISUAL INSPECTION | Scratched Material | Device problem code
6859164 | 86086623 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW NO TRACES OF MOISTURE WERE NOTED AT ELECTRONIC OR MOTOR ASSEMBLIES PER VISUAL INSPECTION | Component Missing | Device problem code
6859164 | 86086623 | N | B UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED MISSING RESERVOIR TUBE ORING MINOR SCRATCHES ON CASE AND MINOR SCRATCHES ON LCD WINDOW NO TRACES OF MOISTURE WERE NOTED AT ELECTRONIC OR MOTOR ASSEMBLIES PER VISUAL INSPECTION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6860965 | 86087392 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6860965 | 86087392 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6860954 | 86087661 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Positioning Failure | Device Problem Code
6860954 | 86087661 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Assembly Problem Identified | Medical Evaluation result code
6860967 | 86089277 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6860967 | 86089277 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6860973 | 86089345 | N | THE REVOLVE STEREOTACTIC PROBE IS A STERILE SINGLEPATIENT USE DEVICE THAT MAY BE USED WITH IMAGING GUIDANCE TO EXCISE A TISSUE SAMPLE FOR DIAGNOSIS ONE MST WAS RECEIVED FOR INVESTIGATION ON AND ANALYZED ON DEVICE WAS FOUND TO BE IN GOOD CONDITION EVIDENCE OF USE IN A PROCEDURE WAS PRESENT ON THE NEEDLE AND INSIDE THE VACUUM TUBES CUTTER WAS IN THE CLOSED POSITION THE DEVICE WAS CONNECTED TO A HOLSTER FOR FUNCTIONAL EVALUATION UPON INITIALIZATION TISSUE WAS FOUND CLOGGING THE CUTTER THE LATERAL VACUUM LINE IS NOT FUNCTIONING UPON FURTHER SAMPLING LATERAL LINE SEEMS TO HAVE BEEN UNCLOGGED DUE TO THE DISCOVERY OF THE TISSUE WITHIN THE CUTTER EVENT WAS REASSESSED FOR REPORTABILITY AGAINST THE RESULT OF THE INVESTIGATION FOLLOWING CONSULTATION WITH OUR MEDICAL DIRECTOR DUE TO THE POTENTIAL TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AS A RESULT OF POTENTIAL MISSED OR LOST TISSUE SAMPLES PURSUANT TO CFR THIS FAILURE MODE WAS DETERMINED TO BE A REPORTABLE MALFUNCTION THUS WE ARE SUBMITTING THIS MEDWATCH REPORT | Complete Blockage | Device problem code
6860973 | 86089345 | N | THE REVOLVE STEREOTACTIC PROBE IS A STERILE SINGLEPATIENT USE DEVICE THAT MAY BE USED WITH IMAGING GUIDANCE TO EXCISE A TISSUE SAMPLE FOR DIAGNOSIS ONE MST WAS RECEIVED FOR INVESTIGATION ON AND ANALYZED ON DEVICE WAS FOUND TO BE IN GOOD CONDITION EVIDENCE OF USE IN A PROCEDURE WAS PRESENT ON THE NEEDLE AND INSIDE THE VACUUM TUBES CUTTER WAS IN THE CLOSED POSITION THE DEVICE WAS CONNECTED TO A HOLSTER FOR FUNCTIONAL EVALUATION UPON INITIALIZATION TISSUE WAS FOUND CLOGGING THE CUTTER THE LATERAL VACUUM LINE IS NOT FUNCTIONING UPON FURTHER SAMPLING LATERAL LINE SEEMS TO HAVE BEEN UNCLOGGED DUE TO THE DISCOVERY OF THE TISSUE WITHIN THE CUTTER EVENT WAS REASSESSED FOR REPORTABILITY AGAINST THE RESULT OF THE INVESTIGATION FOLLOWING CONSULTATION WITH OUR MEDICAL DIRECTOR DUE TO THE POTENTIAL TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AS A RESULT OF POTENTIAL MISSED OR LOST TISSUE SAMPLES PURSUANT TO CFR THIS FAILURE MODE WAS DETERMINED TO BE A REPORTABLE MALFUNCTION THUS WE ARE SUBMITTING THIS MEDWATCH REPORT | Operational Problem Identified | Medical Evaluation result code
6860952 | 86089460 | N | ACCORDING TO THE PRACTITIONER THE TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE THE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE TWO IMPLANTS FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6860952 | 86089460 | N | ACCORDING TO THE PRACTITIONER THE TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE THE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE TWO IMPLANTS FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6861046 | 86089493 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6861046 | 86089493 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Assembly Problem Identified | Medical Evaluation result code
6860955 | 86091588 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6860955 | 86091588 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6861257 | 86094055 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6861257 | 86094055 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6859319 | 86101489 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6859319 | 86101489 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6860906 | 86108398 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6860906 | 86108398 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6862802 | 86113075 | N | B THE INSULIN PUMP WAS RECEIVED WITH AN OPEN BOOK CRITICAL PUMP ERROR AND A PUMP ERROR ALARM IN THE LONG TRACE THE PROBLEM WAS ISOLATED TO THE SOFTWARE ERROR ANALYSIS WAS UNABLE TO PERFORM THE FUNCTIONAL TESTS INCLUDING THE SELF TEST THE SLEEP CURRENT MEASUREMENT TEST THE ACTIVE CURRENT MEASUREMENT TEST THE REWIND TEST THE PRIME TEST THE BASIC OCCLUSION TEST THE OCCLUSION TEST THE FORCE SENSOR TEST AND THE DISPLACEMENT TEST DUE TO THE OPEN BOOK CRITICAL PUMP ERROR THE INSULIN PUMP WAS RECEIVED WITH A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Component or Accessory Incompatibility | Device problem code
6862802 | 86113075 | N | B THE INSULIN PUMP WAS RECEIVED WITH AN OPEN BOOK CRITICAL PUMP ERROR AND A PUMP ERROR ALARM IN THE LONG TRACE THE PROBLEM WAS ISOLATED TO THE SOFTWARE ERROR ANALYSIS WAS UNABLE TO PERFORM THE FUNCTIONAL TESTS INCLUDING THE SELF TEST THE SLEEP CURRENT MEASUREMENT TEST THE ACTIVE CURRENT MEASUREMENT TEST THE REWIND TEST THE PRIME TEST THE BASIC OCCLUSION TEST THE OCCLUSION TEST THE FORCE SENSOR TEST AND THE DISPLACEMENT TEST DUE TO THE OPEN BOOK CRITICAL PUMP ERROR THE INSULIN PUMP WAS RECEIVED WITH A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Scratched Material | Device problem code
6862802 | 86113075 | N | B THE INSULIN PUMP WAS RECEIVED WITH AN OPEN BOOK CRITICAL PUMP ERROR AND A PUMP ERROR ALARM IN THE LONG TRACE THE PROBLEM WAS ISOLATED TO THE SOFTWARE ERROR ANALYSIS WAS UNABLE TO PERFORM THE FUNCTIONAL TESTS INCLUDING THE SELF TEST THE SLEEP CURRENT MEASUREMENT TEST THE ACTIVE CURRENT MEASUREMENT TEST THE REWIND TEST THE PRIME TEST THE BASIC OCCLUSION TEST THE OCCLUSION TEST THE FORCE SENSOR TEST AND THE DISPLACEMENT TEST DUE TO THE OPEN BOOK CRITICAL PUMP ERROR THE INSULIN PUMP WAS RECEIVED WITH A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Computer Software Problem | Device problem code
6862802 | 86113075 | N | B THE INSULIN PUMP WAS RECEIVED WITH AN OPEN BOOK CRITICAL PUMP ERROR AND A PUMP ERROR ALARM IN THE LONG TRACE THE PROBLEM WAS ISOLATED TO THE SOFTWARE ERROR ANALYSIS WAS UNABLE TO PERFORM THE FUNCTIONAL TESTS INCLUDING THE SELF TEST THE SLEEP CURRENT MEASUREMENT TEST THE ACTIVE CURRENT MEASUREMENT TEST THE REWIND TEST THE PRIME TEST THE BASIC OCCLUSION TEST THE OCCLUSION TEST THE FORCE SENSOR TEST AND THE DISPLACEMENT TEST DUE TO THE OPEN BOOK CRITICAL PUMP ERROR THE INSULIN PUMP WAS RECEIVED WITH A SCRATCHED CASE A PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHES ON THE LCD WINDOW | Software Problem Identified | Medical Evaluation result code
6864545 | 86114025 | D | A REPORT WAS RECEIVED THAT THE PATIENT WAS EXPERIENCING BURNING SENSATION AT THE LEAD SITE IT WAS ALSO REPORTED THAT THE PATIENT WAS RECEIVING SHOTS NEAR THE IMPLANT SITE BY THE PHYSICIAN NO FURTHER COURSE OF ACTION WILL BE TAKEN AT THIS TIME | Burning Sensation | Patient problem code
6864545 | 86114025 | D | A REPORT WAS RECEIVED THAT THE PATIENT WAS EXPERIENCING BURNING SENSATION AT THE LEAD SITE IT WAS ALSO REPORTED THAT THE PATIENT WAS RECEIVING SHOTS NEAR THE IMPLANT SITE BY THE PHYSICIAN NO FURTHER COURSE OF ACTION WILL BE TAKEN AT THIS TIME | Insufficient Information | Device problem code
6864553 | 86114940 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION RECEIVED DURING CONVERSATION WITH SURGEON HE COMMENTED THAT THE EVENTS REPORTED WERE MORE HIS OBSERVATION AND NOT SPECIFICALLY A DEVICE ISSUE THE BLOOD CLOT WOULD FORM ON THE JJ ANASTOMOSIS AND THE PATIENT WOULD DEVELOP A BOWEL OBSTRUCTION A FEW DAYS LATER HE HAS NOT SEEN THIS ISSUE IN ANY ROUX EN Y PROCEDURE PREVIOUSLY WHILE USING RELOADS PRIOR TO THE GRIPPING SURFACE TECHNOLOGY NO STAPLE FORMATION ISSUES WERE OBSERVED DURING ANY OF THE PROCEDURES AS WELL AS THE REOPERATIONS WHITE RELOADS WERE UTILIZED WITH THE TRIPLE STAPLE TECHNIQUE USED THE SURGEON STATED THAT HE ALWAYS WAITS SECONDS PRIOR TO FIRING AND USES THE SAME TECHNIQUE IS ALL CASES | Obstruction of Flow | Device problem code
6864553 | 86114940 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION RECEIVED DURING CONVERSATION WITH SURGEON HE COMMENTED THAT THE EVENTS REPORTED WERE MORE HIS OBSERVATION AND NOT SPECIFICALLY A DEVICE ISSUE THE BLOOD CLOT WOULD FORM ON THE JJ ANASTOMOSIS AND THE PATIENT WOULD DEVELOP A BOWEL OBSTRUCTION A FEW DAYS LATER HE HAS NOT SEEN THIS ISSUE IN ANY ROUX EN Y PROCEDURE PREVIOUSLY WHILE USING RELOADS PRIOR TO THE GRIPPING SURFACE TECHNOLOGY NO STAPLE FORMATION ISSUES WERE OBSERVED DURING ANY OF THE PROCEDURES AS WELL AS THE REOPERATIONS WHITE RELOADS WERE UTILIZED WITH THE TRIPLE STAPLE TECHNIQUE USED THE SURGEON STATED THAT HE ALWAYS WAITS SECONDS PRIOR TO FIRING AND USES THE SAME TECHNIQUE IS ALL CASES | Operational Problem Identified | Medical Evaluation result code
6860993 | 86121231 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WOULD NOT OPEN WAS IT STUCK ON TISSUE IF YES HOW WAS THE DEVICE REMOVED WAS THE DEVICE EVENTUALLY ABLE TO BE OPENED | Entrapment of Device | Device Problem Code
6860993 | 86121231 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WOULD NOT OPEN WAS IT STUCK ON TISSUE IF YES HOW WAS THE DEVICE REMOVED WAS THE DEVICE EVENTUALLY ABLE TO BE OPENED | Difficult to Open or Close | Device problem code
6860993 | 86121231 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WOULD NOT OPEN WAS IT STUCK ON TISSUE IF YES HOW WAS THE DEVICE REMOVED WAS THE DEVICE EVENTUALLY ABLE TO BE OPENED | Operational Problem Identified | Medical Evaluation result code
6860993 | 86121231 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WOULD NOT OPEN WAS IT STUCK ON TISSUE IF YES HOW WAS THE DEVICE REMOVED WAS THE DEVICE EVENTUALLY ABLE TO BE OPENED | Mechanical Jam | Device problem code
6860993 | 86121231 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE DEVICE WOULD NOT OPEN WAS IT STUCK ON TISSUE IF YES HOW WAS THE DEVICE REMOVED WAS THE DEVICE EVENTUALLY ABLE TO BE OPENED | Mechanical Problem Identified | Medical Evaluation result code
6860996 | 86121983 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD WAS NOT STICKING WELL TO THE SKIN AND PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Physical Resistance / Sticking | Device problem code
6860996 | 86121983 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD WAS NOT STICKING WELL TO THE SKIN AND PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6860996 | 86121983 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD WAS NOT STICKING WELL TO THE SKIN AND PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6861211 | 86125949 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B HOSPITAL ON B AND WAS TREATED WITH INSULIN THROUGH IV | No Known Impact Or Consequence To Patient | Patient Problem Code
6861211 | 86125949 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B HOSPITAL ON B AND WAS TREATED WITH INSULIN THROUGH IV | Insufficient Information | Device problem code
6864870 | 86128072 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6864870 | 86128072 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6864929 | 86130810 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS ARE LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANTS HAVE BEEN RESTORED TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864929 | 86130810 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS ARE LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANTS HAVE BEEN RESTORED TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6864821 | 86132385 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864821 | 86132385 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6864879 | 86133066 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864879 | 86133066 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6864887 | 86133297 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6864887 | 86133297 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6861529 | 86133616 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE DURING THE VIDEO TEST THERE WAS RIGHT IMAGE LOSS THE ISSUE WAS CAUSED BY THE RIGHT CAMERA CONTROL UNIT CCU THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Data Problem | Device problem code
6861529 | 86133616 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE DURING THE VIDEO TEST THERE WAS RIGHT IMAGE LOSS THE ISSUE WAS CAUSED BY THE RIGHT CAMERA CONTROL UNIT CCU THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Data Storage or Loss of Data | Medical Evaluation result code
6864818 | 86133803 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864818 | 86133803 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6864840 | 86134055 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864840 | 86134055 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6864830 | 86134212 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864830 | 86134212 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6864907 | 86134290 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6864907 | 86134290 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Operational Problem Identified | Medical Evaluation result code
6864838 | 86134649 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864838 | 86134649 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6864834 | 86135254 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6864834 | 86135254 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6852055 | 86135571 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | Connection Problem | Device problem code
6852055 | 86135571 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION | No Findings Available | Medical Evaluation result code
6864826 | 86135650 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6864826 | 86135650 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6864823 | 86135899 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6864823 | 86135899 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6861000 | 86140959 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6861000 | 86140959 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6860764 | 86145350 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM A MANUAL INSULIN INJECTION WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA THE POD ALARMED WITH A HAZARD ALARM | Insufficient Information | Device problem code
6860764 | 86145350 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM A MANUAL INSULIN INJECTION WAS ADMINISTERED TO TREAT THE HYPERGLYCEMIA THE POD ALARMED WITH A HAZARD ALARM | Hyperglycemia | Patient problem code
6851156 | 86163616 | D | THE CUSTOMER REPORTED HER BLOOD GLUCOSE REACHED MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THERE WAS WETNESS AND INSULIN NOTED AT THE SITE | Insufficient Information | Device problem code
6851156 | 86163616 | D | THE CUSTOMER REPORTED HER BLOOD GLUCOSE REACHED MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THERE WAS WETNESS AND INSULIN NOTED AT THE SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6866412 | 86241368 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6866458 | 86241524 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6866479 | 86242160 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6866446 | 86242234 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6866374 | 86243227 | D | THE PATIENT REPORTED AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED PAINFUL AND SORE SHE WENT TO SEE HER GENERAL PRACTITIONER WHO PRESCRIBED FLUCLOXACILLIN MG ONE TABLET PER DAY FOR DAYS FOR HER SITE INFECTION | Infection | Patient problem code
6866374 | 86243227 | D | THE PATIENT REPORTED AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED PAINFUL AND SORE SHE WENT TO SEE HER GENERAL PRACTITIONER WHO PRESCRIBED FLUCLOXACILLIN MG ONE TABLET PER DAY FOR DAYS FOR HER SITE INFECTION | Pain | Patient problem code
6866374 | 86243227 | D | THE PATIENT REPORTED AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED PAINFUL AND SORE SHE WENT TO SEE HER GENERAL PRACTITIONER WHO PRESCRIBED FLUCLOXACILLIN MG ONE TABLET PER DAY FOR DAYS FOR HER SITE INFECTION | Insufficient Information | Device problem code
6866236 | 86243753 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6867945 | 86255595 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B MEMORIAL ON B AND WAS TREATED WITH IV FLUIDS ANTIBIOTICS AND POTASSIUM | No Known Impact Or Consequence To Patient | Patient Problem Code
6867945 | 86255595 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT THE B MEMORIAL ON B AND WAS TREATED WITH IV FLUIDS ANTIBIOTICS AND POTASSIUM | Insufficient Information | Device problem code
6864415 | 86259043 | D | THE CUSTOMER REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR LESS THAN AN HOUR SHE SAW THE CANNULA WAS NOT INSERTED AND ALSO FELT THE INSULIN | Positioning Failure | Device problem code
6864415 | 86259043 | D | THE CUSTOMER REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR LESS THAN AN HOUR SHE SAW THE CANNULA WAS NOT INSERTED AND ALSO FELT THE INSULIN | No Known Impact Or Consequence To Patient | Patient Problem Code
6866068 | 86260221 | N | VYAIRE COMPLAINT B RESULTS OF INVESTIGATION THE VYAIRE FAILURE ANALYSIS LAB RECEIVED THE SUSPECTED DEVICECOMPONENT AND PERFORMED A FAILURE INVESTIGATION THE REPORTED ISSUE WAS DUPLICATED IN THE LABORATORY SETTING THE ROOT CAUSE OF THE REPORTED ISSUE WAS DUE TO LOW BATTERY VOLTAGE | Battery Problem | Device problem code
6866068 | 86260221 | N | VYAIRE COMPLAINT B RESULTS OF INVESTIGATION THE VYAIRE FAILURE ANALYSIS LAB RECEIVED THE SUSPECTED DEVICECOMPONENT AND PERFORMED A FAILURE INVESTIGATION THE REPORTED ISSUE WAS DUPLICATED IN THE LABORATORY SETTING THE ROOT CAUSE OF THE REPORTED ISSUE WAS DUE TO LOW BATTERY VOLTAGE | Power Source Problem Identified | Medical Evaluation result code
6866052 | 86260735 | N | RESULTS OF THE EVALUATION SHOWED THE CT PARAMETERS DID NOT MATCH RECOMMENDATIONS AND NUMBER OF SLICES WAS BORDERLINE A D TREE COULD NOT BE CREATED BY THE PLANNING APPLICATION AND WAS GENERATED MANUALLY BY THE PHYSICIAN TWO POINTS IN THE MANUAL REGISTRATION PLAN HAD DIFFERENT NAMES BUT THE SAME VALUES CAUSING THE SYSTEM TO CRASH WHEN THE MANUAL REGISTRATION PLAN IS LOADED THE COMPLAINT ANALYSIS REVEALED A SOFTWARE DESIGN ANOMALY | Inadequacy of Device Shape and/or Size | Device problem code
6866052 | 86260735 | N | RESULTS OF THE EVALUATION SHOWED THE CT PARAMETERS DID NOT MATCH RECOMMENDATIONS AND NUMBER OF SLICES WAS BORDERLINE A D TREE COULD NOT BE CREATED BY THE PLANNING APPLICATION AND WAS GENERATED MANUALLY BY THE PHYSICIAN TWO POINTS IN THE MANUAL REGISTRATION PLAN HAD DIFFERENT NAMES BUT THE SAME VALUES CAUSING THE SYSTEM TO CRASH WHEN THE MANUAL REGISTRATION PLAN IS LOADED THE COMPLAINT ANALYSIS REVEALED A SOFTWARE DESIGN ANOMALY | Deformation Problem | Medical Evaluation result code
6867097 | 86261595 | N | THE CUSTOMER REPORTED THAT HIS MULTI GAS UNIT STOPS WORKING WHILE MONITORING A PATIENT AND WONT PRODUCE GAS READINGS THE GAS UNIT DOES NOT DISPLAY AN ERROR MESSAGE WHEN THE UNIT STOPS READING EVEN AFTER WARMING THE DEVICE FOR AN HOUR BEFORE USE CUSTOMER STATES THAT THEIR GAS UNIT IS WORKING AS IT SHOULD NO EVALUATION ON THE UNIT IS NEEDED CUSTOMER IS SENDING BACK THE LOANER | No Display / Image | Device problem code
6867097 | 86261595 | N | THE CUSTOMER REPORTED THAT HIS MULTI GAS UNIT STOPS WORKING WHILE MONITORING A PATIENT AND WONT PRODUCE GAS READINGS THE GAS UNIT DOES NOT DISPLAY AN ERROR MESSAGE WHEN THE UNIT STOPS READING EVEN AFTER WARMING THE DEVICE FOR AN HOUR BEFORE USE CUSTOMER STATES THAT THEIR GAS UNIT IS WORKING AS IT SHOULD NO EVALUATION ON THE UNIT IS NEEDED CUSTOMER IS SENDING BACK THE LOANER | Device Displays Incorrect Message | Device problem code
6867097 | 86261595 | N | THE CUSTOMER REPORTED THAT HIS MULTI GAS UNIT STOPS WORKING WHILE MONITORING A PATIENT AND WONT PRODUCE GAS READINGS THE GAS UNIT DOES NOT DISPLAY AN ERROR MESSAGE WHEN THE UNIT STOPS READING EVEN AFTER WARMING THE DEVICE FOR AN HOUR BEFORE USE CUSTOMER STATES THAT THEIR GAS UNIT IS WORKING AS IT SHOULD NO EVALUATION ON THE UNIT IS NEEDED CUSTOMER IS SENDING BACK THE LOANER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6867919 | 86261914 | D | THERE WAS NO PATIENT INVOLVED IN THIS EVENT LOW BATTERY WARNING | Battery Problem: Low Impedance | Device problem code
6867919 | 86261914 | D | THERE WAS NO PATIENT INVOLVED IN THIS EVENT LOW BATTERY WARNING | No Patient Involvement | Patient problem code
6867070 | 86267532 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE NON INVASIVE BLOOD PRESSURE NIBP SETTING ON THE CORE UNIT G MONITOR WAS NOT APPROPRIATELY DEFAULTING TO THE NICU INFLATION PRESSURE RATE WHEN IN NICU MODE IT WOULD DEFAULT TO THE ADULT INFLATION PRESSURE RATE A TEMPORARY SOLUTION WAS TO MANUALLY SET THE UNIT TO THE NICU SETTING THE SETTING WOULD HOLD UNTIL THE UNIT WAS POWERED DOWN AND THEN DEFAULT BACK TO ADULT THIS ISSUE WAS DISCOVERED DURING A GO LIVE AND NO PATIENT HARM WAS REPORTED AN INVESTIGATION WAS CONDUCTED AT NIHON KOHDEN THE INITIAL CUFF PRESSURE OF THE DETAILED SETTINGS TAB IS A SETTING THAT CHANGES THE INITIAL CUFF PRESSURE OF EACH MODE ADULTCHILDNEONATE AND CAN BE OPERATED FREELY IRRESPECTIVE OF THE TYPE OF AIR HOSE CONNECTED TO THE DEVICE WE CONFIRMED THAT WHEN THE NEONATAL AIR HOSE WAS CONNECTED IT OPERATED ACCORDING TO PRODUCT SPECIFICATION IE INFLATED TO THE INITIAL NEONATAL PRESSURE | Inflation Problem | Device problem code
6867070 | 86267532 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE NON INVASIVE BLOOD PRESSURE NIBP SETTING ON THE CORE UNIT G MONITOR WAS NOT APPROPRIATELY DEFAULTING TO THE NICU INFLATION PRESSURE RATE WHEN IN NICU MODE IT WOULD DEFAULT TO THE ADULT INFLATION PRESSURE RATE A TEMPORARY SOLUTION WAS TO MANUALLY SET THE UNIT TO THE NICU SETTING THE SETTING WOULD HOLD UNTIL THE UNIT WAS POWERED DOWN AND THEN DEFAULT BACK TO ADULT THIS ISSUE WAS DISCOVERED DURING A GO LIVE AND NO PATIENT HARM WAS REPORTED AN INVESTIGATION WAS CONDUCTED AT NIHON KOHDEN THE INITIAL CUFF PRESSURE OF THE DETAILED SETTINGS TAB IS A SETTING THAT CHANGES THE INITIAL CUFF PRESSURE OF EACH MODE ADULTCHILDNEONATE AND CAN BE OPERATED FREELY IRRESPECTIVE OF THE TYPE OF AIR HOSE CONNECTED TO THE DEVICE WE CONFIRMED THAT WHEN THE NEONATAL AIR HOSE WAS CONNECTED IT OPERATED ACCORDING TO PRODUCT SPECIFICATION IE INFLATED TO THE INITIAL NEONATAL PRESSURE | No Device Problem Found | Medical Evaluation result code
6857922 | 86279793 | N | INVESTIGATION WITH THE DONORS TBV OF ML THE DONOR WOULD HAVE QUALIFIED FOR THE COLLECTION PROCEDURE THAT WAS COMPLETED A TERUMO BCTS SYSTEM ENGINEER PERFORMED A CALCULATION FOR THE POTENTIAL AC OVERINFUSION AND HYPOVOLEMIA FOR THIS EVENT THE CALCULATED AC INFUSION RATE FOR BOTH THE ENTERED AND CORRECT PATIENT DATA ARE LESS THAN THE MAXIMUM ALLOWED AC INFUSION RATE PER THE VERSION TRIMA ACCEL SYSTEM ADMINISTRATORS GUIDE THEREFORE IT IS UNLIKELY THAT AN AC OVERINFUSION OCCURRED DUE TO THE INCORRECTLY ENTERED DONOR INFORMATION BASED ON THE TRIMAACCEL AUTOMATED BLOOD COLLECTION SYSTEM ADMINISTRATORS GUIDE THE TOTAL VOLUME REMOVED FROM THE MALE DONOR FOR BOTH THE ENTERED AND CORRECT PATIENT DATA RESULTED IN LESS THAN TBV BEING REMOVED INVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | Excess Flow or Over-Infusion | Device problem code
6857922 | 86279793 | N | INVESTIGATION WITH THE DONORS TBV OF ML THE DONOR WOULD HAVE QUALIFIED FOR THE COLLECTION PROCEDURE THAT WAS COMPLETED A TERUMO BCTS SYSTEM ENGINEER PERFORMED A CALCULATION FOR THE POTENTIAL AC OVERINFUSION AND HYPOVOLEMIA FOR THIS EVENT THE CALCULATED AC INFUSION RATE FOR BOTH THE ENTERED AND CORRECT PATIENT DATA ARE LESS THAN THE MAXIMUM ALLOWED AC INFUSION RATE PER THE VERSION TRIMA ACCEL SYSTEM ADMINISTRATORS GUIDE THEREFORE IT IS UNLIKELY THAT AN AC OVERINFUSION OCCURRED DUE TO THE INCORRECTLY ENTERED DONOR INFORMATION BASED ON THE TRIMAACCEL AUTOMATED BLOOD COLLECTION SYSTEM ADMINISTRATORS GUIDE THE TOTAL VOLUME REMOVED FROM THE MALE DONOR FOR BOTH THE ENTERED AND CORRECT PATIENT DATA RESULTED IN LESS THAN TBV BEING REMOVED INVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | No Device Problem Found | Medical Evaluation result code
6855336 | 86283346 | D | A PATIENTS BLOOD GLUCOSE REACHED ALMOST MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS FEELING PAIN DURING INSULIN DELIVERY AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6855336 | 86283346 | D | A PATIENTS BLOOD GLUCOSE REACHED ALMOST MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS FEELING PAIN DURING INSULIN DELIVERY AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Disconnection | Device problem code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Component or Accessory Incompatibility | Device Problem Code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Scratched Material | Device problem code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Battery Problem | Device Problem Code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Loss of Power | Device problem code
6863598 | 86286236 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED THE SELF REWIND PRIMESEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TESTS THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR THE PUMP ALSO ALARMED POWER LOSS FAILED BATTERY TEST AND REPLACE BATTERY NOW ALARMS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Power Source Problem Identified | Medical Evaluation result code
6865851 | 86289331 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS IN ADDITION TO ELEVATED BG LEVELS ADHESIVE WAS WET THUS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6865851 | 86289331 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS IN ADDITION TO ELEVATED BG LEVELS ADHESIVE WAS WET THUS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6866245 | 86291770 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6866245 | 86291770 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6868638 | 86302441 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WA MGDL AND A BOLUS VIA THE PUMP WAS DELIVERED TO ADDRESS THE BG LEVEL THE CUSTOMER WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG THE CUSTOMER WAS NOT IN DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT SURE WHAT CAUSED THE HIGH BG AND THOUGHT THAT SCAR TISSUE MAY BE AFFECTING THE INSULIN ABSORPTION THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND GLUCOSE THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG RESOLVED NEW PUMP SUPPLIES WERE PUT ON AND A PUMP SYSTEM CHECK WAS PERFORMED NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMERS CURRENT BG WAS MGDL | Adverse Event Without Identified Device or Use Problem | Device problem code
6868638 | 86302441 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WA MGDL AND A BOLUS VIA THE PUMP WAS DELIVERED TO ADDRESS THE BG LEVEL THE CUSTOMER WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG THE CUSTOMER WAS NOT IN DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT SURE WHAT CAUSED THE HIGH BG AND THOUGHT THAT SCAR TISSUE MAY BE AFFECTING THE INSULIN ABSORPTION THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND GLUCOSE THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG RESOLVED NEW PUMP SUPPLIES WERE PUT ON AND A PUMP SYSTEM CHECK WAS PERFORMED NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMERS CURRENT BG WAS MGDL | Scar Tissue | Patient problem code
6868638 | 86302441 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WA MGDL AND A BOLUS VIA THE PUMP WAS DELIVERED TO ADDRESS THE BG LEVEL THE CUSTOMER WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG THE CUSTOMER WAS NOT IN DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT SURE WHAT CAUSED THE HIGH BG AND THOUGHT THAT SCAR TISSUE MAY BE AFFECTING THE INSULIN ABSORPTION THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND GLUCOSE THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG RESOLVED NEW PUMP SUPPLIES WERE PUT ON AND A PUMP SYSTEM CHECK WAS PERFORMED NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMERS CURRENT BG WAS MGDL | Diabetic Ketoacidosis | Patient problem code
6868638 | 86302441 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WA MGDL AND A BOLUS VIA THE PUMP WAS DELIVERED TO ADDRESS THE BG LEVEL THE CUSTOMER WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG THE CUSTOMER WAS NOT IN DIABETIC KETOACIDOSIS THE CUSTOMER WAS NOT SURE WHAT CAUSED THE HIGH BG AND THOUGHT THAT SCAR TISSUE MAY BE AFFECTING THE INSULIN ABSORPTION THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND GLUCOSE THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG RESOLVED NEW PUMP SUPPLIES WERE PUT ON AND A PUMP SYSTEM CHECK WAS PERFORMED NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMERS CURRENT BG WAS MGDL | Hyperglycemia | Patient Problem Code
6868592 | 86303444 | N | INITIAL REPORTER POSTAL CODE B THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED AT THE TIME OF COMPLAINT ENTRY NO DEVICE SPECIFIC INFORMATION IE CATALOGUELOT NUMBER IS AVAILABLE THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS ARTICLE ATTACHED TO THE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UDI UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Insufficient Information | Device problem code
6868592 | 86303444 | N | INITIAL REPORTER POSTAL CODE B THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED AT THE TIME OF COMPLAINT ENTRY NO DEVICE SPECIFIC INFORMATION IE CATALOGUELOT NUMBER IS AVAILABLE THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS ARTICLE ATTACHED TO THE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UDI UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Results Pending Completion of Investigation | Medical Evaluation result code
6868698 | 86305240 | N | B THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED ARTICLE ATTACHED TO THE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Insufficient Information | Device problem code
6868698 | 86305240 | N | B THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED ARTICLE ATTACHED TO THE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Results Pending Completion of Investigation | Medical Evaluation result code
6862382 | 86306509 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO EVALUATE THE UNIT THE FSE REMOVED THE COMPRESSOR AND RESEATED ALL GASKETS AND HOSES THE FSE ENSURED ALL HOSES AND FITTING WERE CORRECTLY SEATED SCREWS WERE TIGHTENED TO IN LBS AND ENSURED CORRECT VACUUM HAD RETURNED TO PUMP THE FSE PERFORMED A FULL CALIBRATION AND FUNCTIONAL TEST AND IABP WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Calibration Problem | Device problem code
6862382 | 86306509 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO EVALUATE THE UNIT THE FSE REMOVED THE COMPRESSOR AND RESEATED ALL GASKETS AND HOSES THE FSE ENSURED ALL HOSES AND FITTING WERE CORRECTLY SEATED SCREWS WERE TIGHTENED TO IN LBS AND ENSURED CORRECT VACUUM HAD RETURNED TO PUMP THE FSE PERFORMED A FULL CALIBRATION AND FUNCTIONAL TEST AND IABP WAS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Failure to Calibrate | Device problem code
6840455 | 86312455 | D | OLYMPUS WAS INFORMED THAT A CUT WAS FOUND ON THE BENDING SECTION COVER THERE WAS NO PATIENT INVOLVEMENT WITH THE REPORTED SCOPE | Material Split, Cut or Torn | Device problem code
6840455 | 86312455 | D | OLYMPUS WAS INFORMED THAT A CUT WAS FOUND ON THE BENDING SECTION COVER THERE WAS NO PATIENT INVOLVEMENT WITH THE REPORTED SCOPE | No Patient Involvement | Patient problem code
6868980 | 86318773 | N | THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED ARTICLE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UDI UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Insufficient Information | Device problem code
6868980 | 86318773 | N | THIS MDR INVOLVED ENTERPRISE STENTS NO DEVICE SPECIFIC INFORMATION COULD BE OBTAINED ARTICLE MDR RHO MH PARK HJ CHUNG EC ET AL VARIOUS TECHNIQUES OF STENTASSISTED COIL EMBOLIZATION OF WIDENECKED OR FUSIFORM ARTHEROSCLEROTIC AND DISSECTING UNRUPTURED VERTEBROBASILAR ARTERY ANEURYSMS FOR REDUCING RECANALIZATION MIDTERM RESULTS ACTA NEUROCHIR DOI SY THIS SUBMISSION IS RELATED TO A LITERATURE ARTICLE DISCOVERED TO SUPPORT THE CER SUBMISSION PROCESS AS SUCH THE ASSOCIATED TIME FRAME OF EVENT DATES INCLUDES BUT IS NOT LIMITED TO YEARS THESE ARTICLES ARE BEING REVIEWED MONTHLY FOR SAFETY SIGNALS AND WILL BE FOLLOWED BY MONTHLY TRENDING ASSESSMENTS AS WELL AS PMS REVIEWS DATE OF EVENT PRODUCT CODE AND LOT NUMBER COULD NOT BE OBTAINED FROM THE AUTHOR UDI UNKNOWN PART NUMBER ATTEMPTS TO OBTAIN PRODUCT PART NUMBER WERE UNSUCCESSFUL UDI UNAVAILABLE THERE WILL BE THREE MDRS SUBMITTED FOR THIS LITERATURE ARTICLE ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN DAYS OF RECEIPT | Results Pending Completion of Investigation | Medical Evaluation result code
6869329 | 86331555 | N | A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT WHEN PERFORMING A DEMO ON A RESIN HEAD IN THE PRONE POSITION THE REPORTED IMPRECISION COULD BE REPLICATED THE HARDWARE SOFTWARE AND INSTRUMENTS PASSED THE SYSTEM CHECKOUT THE SYSTEM WAS FOUND TO BE FULLY FUNCTIONAL | Imprecision | Device problem code
6869329 | 86331555 | N | A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT WHEN PERFORMING A DEMO ON A RESIN HEAD IN THE PRONE POSITION THE REPORTED IMPRECISION COULD BE REPLICATED THE HARDWARE SOFTWARE AND INSTRUMENTS PASSED THE SYSTEM CHECKOUT THE SYSTEM WAS FOUND TO BE FULLY FUNCTIONAL | No Device Problem Found | Medical Evaluation result code
6869416 | 86333750 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A B OLD MALE PATIENT DURING FLIGHT TRANSPORT THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6869416 | 86333750 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A B OLD MALE PATIENT DURING FLIGHT TRANSPORT THE DEVICE RESTARTEDREBOOTED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6869475 | 86335369 | N | NO REPORT OF PATIENT INVOLVEMENT THE CUSTOMER REQUESTED TECHNICAL SUPPORT FROM GE HEALTHCARE IT WAS RECOMMENDED TO REPLACE THE COMPACT FLASH CARD | Flare or Flash | Device problem code
6869475 | 86335369 | N | NO REPORT OF PATIENT INVOLVEMENT THE CUSTOMER REQUESTED TECHNICAL SUPPORT FROM GE HEALTHCARE IT WAS RECOMMENDED TO REPLACE THE COMPACT FLASH CARD | No Device Problem Found | Medical Evaluation result code
6867409 | 86419853 | N | ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE CUSTOMER REPORTED FAILURE MODE VISUAL INSPECTION FOUND ONE BAD FUSE AND DUSTY FANS THE ILLUMINATOR WAS INSTALLED AND TESTED ON AN INHOUSE TEST SYSTEM AND FAILED TO POWER ON WITH ERRORS THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6867409 | 86419853 | N | ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE CUSTOMER REPORTED FAILURE MODE VISUAL INSPECTION FOUND ONE BAD FUSE AND DUSTY FANS THE ILLUMINATOR WAS INSTALLED AND TESTED ON AN INHOUSE TEST SYSTEM AND FAILED TO POWER ON WITH ERRORS THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | No Device Problem Found | Medical Evaluation result code
6867409 | 86419853 | N | ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE CUSTOMER REPORTED FAILURE MODE VISUAL INSPECTION FOUND ONE BAD FUSE AND DUSTY FANS THE ILLUMINATOR WAS INSTALLED AND TESTED ON AN INHOUSE TEST SYSTEM AND FAILED TO POWER ON WITH ERRORS THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6867409 | 86419853 | N | ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE CUSTOMER REPORTED FAILURE MODE VISUAL INSPECTION FOUND ONE BAD FUSE AND DUSTY FANS THE ILLUMINATOR WAS INSTALLED AND TESTED ON AN INHOUSE TEST SYSTEM AND FAILED TO POWER ON WITH ERRORS THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Power Source Problem Identified | Medical Evaluation result code
6869973 | 86421132 | N | PER THE INSTRUCTIONS FOR USE PARAVALVULAR LEAK PVL IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE CARDIAC TISSUE AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE SOME PVL IS NOT UNCOMMON POST DEPLOYMENT MANY CASES ARE MILD TO MODERATE AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING IN THIS CASE AS REPORTED CALCIUM WAS A LIKELY CONTRIBUTING FACTOR FOR THE PVL COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW NO CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED | Leak / Splash | Device problem code
6869973 | 86421132 | N | PER THE INSTRUCTIONS FOR USE PARAVALVULAR LEAK PVL IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE CARDIAC TISSUE AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE SOME PVL IS NOT UNCOMMON POST DEPLOYMENT MANY CASES ARE MILD TO MODERATE AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING IN THIS CASE AS REPORTED CALCIUM WAS A LIKELY CONTRIBUTING FACTOR FOR THE PVL COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW NO CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED | Activation, Positioning or Separation Problem | Device problem code
6869973 | 86421132 | N | PER THE INSTRUCTIONS FOR USE PARAVALVULAR LEAK PVL IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE CARDIAC TISSUE AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE SOME PVL IS NOT UNCOMMON POST DEPLOYMENT MANY CASES ARE MILD TO MODERATE AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING IN THIS CASE AS REPORTED CALCIUM WAS A LIKELY CONTRIBUTING FACTOR FOR THE PVL COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW NO CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED | Assembly Problem Identified | Medical Evaluation result code
6868521 | 86426643 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6868521 | 86426643 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6864101 | 86427753 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6864101 | 86427753 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Operational Problem Identified | Medical Evaluation result code
6869141 | 86429738 | N | THIS DEVICE WAS RETURNED TO MMDG FOR EVALUATION DURING THE INVESTIGATION MMDG WAS ABLE TO CONFIRM THAT THE SET WAS LEAKING AT SEAL OF THE ADDITION PORT THE SET WAS SENT TO THE SUPPLIER FOR FURTHER EVALUATION THIS REPORT WILL BE UPDATED WHEN THOSE RESULTS ARE RECEIVED BY MMDG | Leak / Splash | Device problem code
6869141 | 86429738 | N | THIS DEVICE WAS RETURNED TO MMDG FOR EVALUATION DURING THE INVESTIGATION MMDG WAS ABLE TO CONFIRM THAT THE SET WAS LEAKING AT SEAL OF THE ADDITION PORT THE SET WAS SENT TO THE SUPPLIER FOR FURTHER EVALUATION THIS REPORT WILL BE UPDATED WHEN THOSE RESULTS ARE RECEIVED BY MMDG | Leakage/Seal | Medical Evaluation result code
6871671 | 86435905 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV PUMP AND FLUID PUMP | Insufficient Information | Device problem code
6871671 | 86435905 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL ON B AND WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS PATIENT WAS TREATED WITH IV PUMP AND FLUID PUMP | Diabetic Ketoacidosis | Patient problem code
6851653 | 86437192 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6851653 | 86437192 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Operational Problem Identified | Medical Evaluation result code
6866255 | 86447965 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6866255 | 86447965 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6872340 | 86468178 | N | LOT HISTORY RECORDS REVIEWED FOR LOT NUMBERS WITH NO NONCONFORMITIES NOTED FOR EITHER LOT PRODUCT IFU AND SURGICAL TECHNIQUE WERE REVIEWED AND NO DISCREPANCIES FOUND IN EITHER DOCUMENT OSTEOPOROSIS IS LISTED AS A RELATIVE CONTRAINDICATION WITHIN THE IFU XM REV B AND SURGICAL TECHNIQUE XMMKT REV B WITHIN THE FOLLOWING STATEMENT RAPID JOINT DISEASE BONE ABSORPTION OSTEOPENIA OSTEOPOROSIS RELATIVE CONTRAINDICATION MAY LIMIT THE DEGREE OF OBTAINABLE CORRECTION STABILIZATION ANDOR THE AMOUNT OF MECHANICAL FIXATION PATIENT BONE WAS NOTED AS OSTEOPOROTIC AND MAY HAVE CONTRIBUTED TO DECREASE IN BONE PURCHASE AND SCREW BACK OUT THE EVENT IS BEING REPORTED AFTER DAYS DUE TO ISSUES SETTING UP WEBTRADER ACCOUNT | Detachment of Device or device Component | Device problem code
6872340 | 86468178 | N | LOT HISTORY RECORDS REVIEWED FOR LOT NUMBERS WITH NO NONCONFORMITIES NOTED FOR EITHER LOT PRODUCT IFU AND SURGICAL TECHNIQUE WERE REVIEWED AND NO DISCREPANCIES FOUND IN EITHER DOCUMENT OSTEOPOROSIS IS LISTED AS A RELATIVE CONTRAINDICATION WITHIN THE IFU XM REV B AND SURGICAL TECHNIQUE XMMKT REV B WITHIN THE FOLLOWING STATEMENT RAPID JOINT DISEASE BONE ABSORPTION OSTEOPENIA OSTEOPOROSIS RELATIVE CONTRAINDICATION MAY LIMIT THE DEGREE OF OBTAINABLE CORRECTION STABILIZATION ANDOR THE AMOUNT OF MECHANICAL FIXATION PATIENT BONE WAS NOTED AS OSTEOPOROTIC AND MAY HAVE CONTRIBUTED TO DECREASE IN BONE PURCHASE AND SCREW BACK OUT THE EVENT IS BEING REPORTED AFTER DAYS DUE TO ISSUES SETTING UP WEBTRADER ACCOUNT | Operational Problem Identified | Medical Evaluation result code
6872680 | 86479091 | D | IT WAS REPORTED TO NEVRO THAT A TRIAL PATIENT DEVELOPED AN INFECTION THE PATIENT WAS HOSPITALIZED AND WAS GIVEN IV ANTIBIOTICS THE DEVICE WAS EXPLANTED THE PATIENT WAS DISCHARGED AND THERE WERE NO REPORTS OF FURTHER COMPLICATIONS | Infection | Patient problem code
6872680 | 86479091 | D | IT WAS REPORTED TO NEVRO THAT A TRIAL PATIENT DEVELOPED AN INFECTION THE PATIENT WAS HOSPITALIZED AND WAS GIVEN IV ANTIBIOTICS THE DEVICE WAS EXPLANTED THE PATIENT WAS DISCHARGED AND THERE WERE NO REPORTS OF FURTHER COMPLICATIONS | Insufficient Information | Device problem code
6872892 | 86485448 | D | MEDTRONIC RECEIVED INFORMATION THAT THREE DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6872892 | 86485448 | D | MEDTRONIC RECEIVED INFORMATION THAT THREE DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6851646 | 86506872 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6851646 | 86506872 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6861041 | 86514783 | N | B BATCH UNK THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE JAW COULD NOT BE OPENED HOW WAS THE DEVICE REMOVED FROM TISSUE WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED DID THE DEVICE STAPLE IF THE DEVICE STAPLED WAS THE STAPLE LINE COMPLETE IF THE DEVICE STAPLED WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT DID THE DEVICE CUT IF THE DEVICE CUT WAS THE CUT LINE COMPLETE WERE THE CUT LINE AND STAPLE LINES EQUAL | Failure to Form Staple | Device problem code
6861041 | 86514783 | N | B BATCH UNK THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHEN THE JAW COULD NOT BE OPENED HOW WAS THE DEVICE REMOVED FROM TISSUE WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED DID THE DEVICE STAPLE IF THE DEVICE STAPLED WAS THE STAPLE LINE COMPLETE IF THE DEVICE STAPLED WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT DID THE DEVICE CUT IF THE DEVICE CUT WAS THE CUT LINE COMPLETE WERE THE CUT LINE AND STAPLE LINES EQUAL | Operational Problem Identified | Medical Evaluation result code
6874401 | 86553282 | N | B THE EVENTS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS CANNOT BE OBTAINED AS PATIENT IS NOT WILLING TO PROVIDE FURTHER INFORMATION DOES NOT WISH TO MOVE FORWARD AT THIS TIME NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Deflation Problem | Device problem code
6874401 | 86553282 | N | B THE EVENTS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS CANNOT BE OBTAINED AS PATIENT IS NOT WILLING TO PROVIDE FURTHER INFORMATION DOES NOT WISH TO MOVE FORWARD AT THIS TIME NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Calcified | Device problem code
6874401 | 86553282 | N | B THE EVENTS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS CANNOT BE OBTAINED AS PATIENT IS NOT WILLING TO PROVIDE FURTHER INFORMATION DOES NOT WISH TO MOVE FORWARD AT THIS TIME NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Operational Problem Identified | Medical Evaluation result code
6874790 | 86553974 | N | SC SN B DEVICE EVALUATION INDICATED THAT THE COMPLAINT OF A KINK IN THE PADDLE END OF THE LEAD WAS NOT CONFIRMED THERE WAS NO KINK OR DEFORMATION PRESENT IN THE LEAD LEAD EXHIBITED NORMAL DEVICE CHARACTERISTICS | Material Deformation | Device problem code
6874790 | 86553974 | N | SC SN B DEVICE EVALUATION INDICATED THAT THE COMPLAINT OF A KINK IN THE PADDLE END OF THE LEAD WAS NOT CONFIRMED THERE WAS NO KINK OR DEFORMATION PRESENT IN THE LEAD LEAD EXHIBITED NORMAL DEVICE CHARACTERISTICS | Stress Problem Identified | Medical Evaluation result code
6874394 | 86555594 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE AT THE LEVEL OF MGDL CUSTOMER STATED HE WAS NOT WEARING THE INSULIN PUMP CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE CUSTOMER WAS HE CAN INSERT THE SERTER MANUALLY THE INSULIN PUMP WAS NOT RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6874394 | 86555594 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE AT THE LEVEL OF MGDL CUSTOMER STATED HE WAS NOT WEARING THE INSULIN PUMP CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE CUSTOMER WAS HE CAN INSERT THE SERTER MANUALLY THE INSULIN PUMP WAS NOT RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6874039 | 86556211 | N | B CONCLUSION CODETRAUMA CAUSED BY FIRWORK NO CONNECTION TO SURGERY B | No Device Problem Found | Medical Evaluation result code
6874222 | 86563098 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF NODULES IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE AND TABLE INJECTION SITE RESPONSES BY SEVERITY AND DURATION AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS N FOR POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLBELLA XC INCLUDE SWELLING TENDERNESS FIRMNESS BRUISING LUMPSBUMPS REDNESS PAIN DISCOLORATION ITCHING AND DRYNESS INJECTION SITE RESPONSES BY SEVERITY AND DURATION AFTER REPEAT TREATMENT WITH JUVÉDERM VOLBELLA XC OCCURRING IN OF TREATED SUBJECTS N INCLUDE SWELLING TENDERNESS FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE CLINICAL STUDY SUBJECTS HAD LUMPSBUMPS OR SWELLING THAT OCCURRED WEEKS TO MONTHS AFTER TREATMENT ALL OF THESE AES WERE MILD OR MODERATE SWELLING WAS TREATED WITH ACETAMINOPHEN OR DOXYCYCLINE AND NO TREATMENT WAS GIVEN FOR THE LUMPSBUMPS ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE | No Findings Available | Medical Evaluation result code
6850090 | 86575905 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6850090 | 86575905 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6844482 | 86581434 | N | A BIOMÉRIEUX INTERNAL INVESTIGATION WAS PERFORMED A CUSTOMER REPORTED LOW RESULTS WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH IDAS CMV IGM LOT VT VT VT VT AND VT COMPARED TO THE SPECIFICATION THE CUSTOMER PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS A REVIEW OF QUALITY RECORDS CONFIRMED THERE WAS NO ISSUE IN THE PRODUCTION OF LOT AN ANALYSIS OF THE CONTROL CHART FOR INTERNAL SAMPLES TARGET TV ON DIFFERENT BATCHES OF VIDAS CMVM SHOWED THAT VIDAS CMV IGM LOT WAS IN THE TREND OF THE OTHER BATCHES THE QUALITY PRODUCT LABORATORY TESTED THESE INTERNAL SAMPLES ON THE RETAINED KIT OF VIDAS CMV IGM LOT AND OBTAINED THE FOLLOWING RESULTS CM TV CM TV CM TV CM TV CM TV ALL THE RESULTS WERE WITHIN THE SPECIFICATIONS AND SIMILAR TO THOSE OBTAINED BEFORE THE RELEASE OF THE BATCH VIDAS CMV IGM LOT THERE WAS NO DRIFT OF LOT THE QUALITY PRODUCT LABORATORY HAD TESTED THE IQC BIORAD VIROTROL WITH THE VIDAS CMVM LOT RETAIN KIT AND LOT REFERENCE KIT THE RESULTS WERE POSITIVE WITH THE BATCHES BATCH TV BATCH TV THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS THE IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM THE QUALITY PRODUCT LABORATORY RECOMPUTED THE FOLLOWING RANGE BASED ON THE CUSTOMERS DATA DISCREPANT VALUES OUT OF VALUES NOT TAKEN INTO ACCOUNT BECAUSE VISIBLY OUT OF TREND TV TV TV MEAN TV ET TV ET TV ACCORDING TO THIS NEW STATISTICS CUSTOMERS VALUES EXCLUDING DISCREPANT VALUES CONFORM WITHIN RANGE AT ET EXCEPT FOR ONE VALUE AT VT CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | False Negative Result | Device problem code
6845830 | 86583679 | N | THE CUSTOMER COMPLAINED ABOUT A LOW RESULT OBTAINED WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH VIDAS CMV IGM LOT TV COMPARED TO THE SPECIFICATION HE HAD PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS THE ANALYSIS OF THE BATCH HISTORY RECORDS SHOWED NO ANOMALY THERE IS NO OTHER COMPLAINT RECORDED FOR THE SAME ISSUE ON THE KIT VIDAS CMV IGM LOT THERE IS NO NONCONFORMITY LINKED TO THE CUSTOMERS ANOMALY ON THIS LOT AND NO CAPA ON THE PRODUCT REFERENCE B THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS FURTHERMORE THIS IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM B CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN THE QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | False Negative Result | Device problem code
6845830 | 86583679 | N | THE CUSTOMER COMPLAINED ABOUT A LOW RESULT OBTAINED WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH VIDAS CMV IGM LOT TV COMPARED TO THE SPECIFICATION HE HAD PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS THE ANALYSIS OF THE BATCH HISTORY RECORDS SHOWED NO ANOMALY THERE IS NO OTHER COMPLAINT RECORDED FOR THE SAME ISSUE ON THE KIT VIDAS CMV IGM LOT THERE IS NO NONCONFORMITY LINKED TO THE CUSTOMERS ANOMALY ON THIS LOT AND NO CAPA ON THE PRODUCT REFERENCE B THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS FURTHERMORE THIS IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM B CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN THE QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | Low Test Results | Device problem code
6845830 | 86583679 | N | THE CUSTOMER COMPLAINED ABOUT A LOW RESULT OBTAINED WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH VIDAS CMV IGM LOT TV COMPARED TO THE SPECIFICATION HE HAD PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS THE ANALYSIS OF THE BATCH HISTORY RECORDS SHOWED NO ANOMALY THERE IS NO OTHER COMPLAINT RECORDED FOR THE SAME ISSUE ON THE KIT VIDAS CMV IGM LOT THERE IS NO NONCONFORMITY LINKED TO THE CUSTOMERS ANOMALY ON THIS LOT AND NO CAPA ON THE PRODUCT REFERENCE B THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS FURTHERMORE THIS IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM B CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN THE QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6847396 | 86584670 | N | THE FOLLOWING THINGS WERE CONFIRMED BY A LOCAL ENGINEER THERE WAS NOT ANY ABNORMALITY WITH EACH CAMERAHEAD OTVSPROHHDLE WHICH WERE USED FOR THIS EVENT THE SUBJECT OTVSV WAS USED FOR THIS EVENT THE SUBJECT OTVSV IS NOT COMPATIBLE WITH THE OTVSPROHHDLE THE SUBJECT OTVSV DOES NOT HAVE A FUNCTION OF NBI A CABLE WHICH WAS CONNECTED WITH THE SUBJECT OTVS WAS MISSING FROM THE ENDOSCOPY SYSTEM TOWER THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED THE OTVSV AND THE OTVSPROHHDLE INSTRUCTION MANUAL STATES THE NORMAL COMBINATION RANGE OF THE RELEVANT PRODUCT AND RELATED EQUIPMENT | Compatibility Problem | Device problem code
6847396 | 86584670 | N | THE FOLLOWING THINGS WERE CONFIRMED BY A LOCAL ENGINEER THERE WAS NOT ANY ABNORMALITY WITH EACH CAMERAHEAD OTVSPROHHDLE WHICH WERE USED FOR THIS EVENT THE SUBJECT OTVSV WAS USED FOR THIS EVENT THE SUBJECT OTVSV IS NOT COMPATIBLE WITH THE OTVSPROHHDLE THE SUBJECT OTVSV DOES NOT HAVE A FUNCTION OF NBI A CABLE WHICH WAS CONNECTED WITH THE SUBJECT OTVS WAS MISSING FROM THE ENDOSCOPY SYSTEM TOWER THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED THE OTVSV AND THE OTVSPROHHDLE INSTRUCTION MANUAL STATES THE NORMAL COMBINATION RANGE OF THE RELEVANT PRODUCT AND RELATED EQUIPMENT | Component Missing | Device problem code
6847396 | 86584670 | N | THE FOLLOWING THINGS WERE CONFIRMED BY A LOCAL ENGINEER THERE WAS NOT ANY ABNORMALITY WITH EACH CAMERAHEAD OTVSPROHHDLE WHICH WERE USED FOR THIS EVENT THE SUBJECT OTVSV WAS USED FOR THIS EVENT THE SUBJECT OTVSV IS NOT COMPATIBLE WITH THE OTVSPROHHDLE THE SUBJECT OTVSV DOES NOT HAVE A FUNCTION OF NBI A CABLE WHICH WAS CONNECTED WITH THE SUBJECT OTVS WAS MISSING FROM THE ENDOSCOPY SYSTEM TOWER THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED THE OTVSV AND THE OTVSPROHHDLE INSTRUCTION MANUAL STATES THE NORMAL COMBINATION RANGE OF THE RELEVANT PRODUCT AND RELATED EQUIPMENT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6871608 | 86585185 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE READ HIGH MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS IT WAS ALSO STATED THAT THE CANNULA APPEARED TO HAVE A CREASE BEND WHEN THE POD WAS REMOVED TREATMENT INCLUDED MANUAL SHOTS | Hyperglycemia | Patient Problem Code
6871608 | 86585185 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE READ HIGH MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS IT WAS ALSO STATED THAT THE CANNULA APPEARED TO HAVE A CREASE BEND WHEN THE POD WAS REMOVED TREATMENT INCLUDED MANUAL SHOTS | Material Twisted / Bent | Device problem code
6874245 | 86591302 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF FIRMNESS IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS WARNINGS INJECTION SITE RESPONSES CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND LASTING DAYS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATEONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLLURE XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE AND TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY AND DURATION AFTER INITIAL TREATMENT OCCURRING IN OF TREATED SUBJECTS N FOR POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION PER TABLE INJECTION SITE RESPONSES BY SEVERITY AND DURATION AFTER REPEAT TREATMENT WITH JUVÉDERM VOLLURE XC OCCURRING IN OF TREATED SUBJECTS N INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH REDNESS LUMPSBUMPS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION AES AFTER INITIALTOUCHUP TREATMENT OCCURRING IN OF NLFS INCLUDED INJECTION SITE BRUISING ERYTHEMA PAIN DISCOLORATION PRURITUS REACTION AND FACIAL ASYMMETRY | No Device Problem Found | Medical Evaluation result code
6845855 | 86594898 | N | A BIOMÉRIEUX INTERNAL INVESTIGATION WAS PERFORMED THE CUSTOMER REPORTED LOW RESULTS WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH IDAS CMV IGM LOT VT VT VT VT AND VT COMPARED TO THE SPECIFICATION THE CUSTOMER PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS A REVIEW OF QUALITY RECORDS CONFIRMED THERE WAS NO ISSUE IN THE PRODUCTION OF LOT AN ANALYSIS OF THE CONTROL CHART FOR INTERNAL SAMPLES TARGET TV ON DIFFERENT BATCHES OF VIDAS CMVM SHOWED THAT VIDAS CMV IGM LOT WAS IN THE TREND OF THE OTHER BATCHES THE QUALITY PRODUCT LABORATORY TESTED THESE INTERNAL SAMPLES ON THE RETAINED KIT OF VIDAS CMV IGM LOT AND OBTAINED THE FOLLOWING RESULTS CM TV CM TV CM TV CM TV CM TV ALL THE RESULTS WERE WITHIN THE SPECIFICATIONS AND SIMILAR TO THOSE OBTAINED BEFORE THE RELEASE OF THE BATCH VIDAS CMV IGM LOT THERE WAS NO DRIFT OF LOT THE QUALITY PRODUCT LABORATORY HAD TESTED THE IQC BIORAD VIROTROL WITH THE VIDAS CMVM LOT RETAIN KIT AND LOT REFERENCE KIT THE RESULTS WERE POSITIVE WITH THE BATCHES BATCH TV BATCH TV THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS THE IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM THE QUALITY PRODUCT LABORATORY RECOMPUTED THE FOLLOWING RANGE BASED ON THE CUSTOMERS DATA DISCREPANT VALUES OUT OF VALUES NOT TAKEN INTO ACCOUNT BECAUSE VISIBLY OUT OF TREND TV TV TV MEAN TV ET TV ET TV ACCORDING TO THIS NEW STATISTICS CUSTOMERS VALUES EXCLUDING DISCREPANT VALUES CONFORM WITHIN RANGE AT ET EXCEPT FOR ONE VALUE AT VT CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | False Negative Result | Device problem code
6845855 | 86594898 | N | A BIOMÉRIEUX INTERNAL INVESTIGATION WAS PERFORMED THE CUSTOMER REPORTED LOW RESULTS WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH IDAS CMV IGM LOT VT VT VT VT AND VT COMPARED TO THE SPECIFICATION THE CUSTOMER PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS A REVIEW OF QUALITY RECORDS CONFIRMED THERE WAS NO ISSUE IN THE PRODUCTION OF LOT AN ANALYSIS OF THE CONTROL CHART FOR INTERNAL SAMPLES TARGET TV ON DIFFERENT BATCHES OF VIDAS CMVM SHOWED THAT VIDAS CMV IGM LOT WAS IN THE TREND OF THE OTHER BATCHES THE QUALITY PRODUCT LABORATORY TESTED THESE INTERNAL SAMPLES ON THE RETAINED KIT OF VIDAS CMV IGM LOT AND OBTAINED THE FOLLOWING RESULTS CM TV CM TV CM TV CM TV CM TV ALL THE RESULTS WERE WITHIN THE SPECIFICATIONS AND SIMILAR TO THOSE OBTAINED BEFORE THE RELEASE OF THE BATCH VIDAS CMV IGM LOT THERE WAS NO DRIFT OF LOT THE QUALITY PRODUCT LABORATORY HAD TESTED THE IQC BIORAD VIROTROL WITH THE VIDAS CMVM LOT RETAIN KIT AND LOT REFERENCE KIT THE RESULTS WERE POSITIVE WITH THE BATCHES BATCH TV BATCH TV THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS THE IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM THE QUALITY PRODUCT LABORATORY RECOMPUTED THE FOLLOWING RANGE BASED ON THE CUSTOMERS DATA DISCREPANT VALUES OUT OF VALUES NOT TAKEN INTO ACCOUNT BECAUSE VISIBLY OUT OF TREND TV TV TV MEAN TV ET TV ET TV ACCORDING TO THIS NEW STATISTICS CUSTOMERS VALUES EXCLUDING DISCREPANT VALUES CONFORM WITHIN RANGE AT ET EXCEPT FOR ONE VALUE AT VT CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6869631 | 86600886 | N | B THE CUSTOMER HAD TUBING INSTALLED DURING CALIBRATION THEY TRIED MULTIPLE TIMES TO CLOSE THE OCCLUDER USING THE OCCLUSION KNOB AND THE CLOSE BUTTON BUT WAS UNABLE TO FULLY CLOSE THE TUBING THERE WAS A SLIGHT FLOW OF THE PRIMING LIQUID | Calibration Problem | Device problem code
6869631 | 86600886 | N | B THE CUSTOMER HAD TUBING INSTALLED DURING CALIBRATION THEY TRIED MULTIPLE TIMES TO CLOSE THE OCCLUDER USING THE OCCLUSION KNOB AND THE CLOSE BUTTON BUT WAS UNABLE TO FULLY CLOSE THE TUBING THERE WAS A SLIGHT FLOW OF THE PRIMING LIQUID | Failure to Calibrate | Device Problem Code
6872007 | 86604811 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6872007 | 86604811 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6845876 | 86613457 | N | A BIOMÉRIEUX INTERNAL INVESTIGATION WAS PERFORMED THE CUSTOMER REPORTED LOW RESULTS WHEN TESTING THE IQC VIROTROL FROM BIORAD WITH IDAS CMV IGM LOT VT VT VT VT AND VT COMPARED TO THE SPECIFICATION THE CUSTOMER PUT IN PLACE BASED ON RESULTS OBTAINED ON PREVIOUS LOTS A REVIEW OF QUALITY RECORDS CONFIRMED THERE WAS NO ISSUE IN THE PRODUCTION OF LOT AN ANALYSIS OF THE CONTROL CHART FOR INTERNAL SAMPLES TARGET TV ON DIFFERENT BATCHES OF VIDAS CMVM SHOWED THAT VIDAS CMV IGM LOT WAS IN THE TREND OF THE OTHER BATCHES THE QUALITY PRODUCT LABORATORY TESTED THESE INTERNAL SAMPLES ON THE RETAINED KIT OF VIDAS CMV IGM LOT AND OBTAINED THE FOLLOWING RESULTS CM TV CM TV CM TV CM TV CM TV ALL THE RESULTS WERE WITHIN THE SPECIFICATIONS AND SIMILAR TO THOSE OBTAINED BEFORE THE RELEASE OF THE BATCH VIDAS CMV IGM LOT THERE WAS NO DRIFT OF LOT THE QUALITY PRODUCT LABORATORY HAD TESTED THE IQC BIORAD VIROTROL WITH THE VIDAS CMVM LOT RETAIN KIT AND LOT REFERENCE KIT THE RESULTS WERE POSITIVE WITH THE BATCHES BATCH TV BATCH TV THE DIFFERENCES OBSERVED ARE LINKED TO THE NORMAL INTERLOT VARIABILITY WHICH HAS NO IMPACT ON PATIENT RESULT OR INTERPRETATION STATUS THE IQC VIROTROL IS MANUFACTURED NOT NATURAL SAMPLE AND MANUFACTURING PROCESS CAN AFFECT SAMPLE MATRIX SO THE HYPOTHESIS IS THAT THIS SAMPLE COULD BEHAVE DIFFERENTLY FOLLOWING SOME LOTS OF VIDAS CMVM IT IS RECOMMENDED THAT EACH LABORATORY ESTABLISHES ITS OWN RANGE FOR EACH ANALYTE EXTRACT FROM PACKAGE INSERT VIROTROL TORCHM THE QUALITY PRODUCT LABORATORY RECOMPUTED THE FOLLOWING RANGE BASED ON THE CUSTOMERS DATA DISCREPANT VALUES OUT OF VALUES NOT TAKEN INTO ACCOUNT BECAUSE VISIBLY OUT OF TREND TV TV TV MEAN TV ET TV ET TV ACCORDING TO THIS NEW STATISTICS CUSTOMERS VALUES EXCLUDING DISCREPANT VALUES CONFORM WITHIN RANGE AT ET EXCEPT FOR ONE VALUE AT VT CHECK WITH THE CUSTOMER HOW STATISTICS ARE COMPUTED AND SUGGEST THE CUSTOMER THAT IS ALSO POSSIBLE TO WORK WITH A MOBILE MEAN QUALITY PRODUCT LABORATORY ANALYZED COMPLAINTS ABOUT SENSITIVITY PROBLEM POTENTIAL FALSE NEGATIVE RESULT SINCE AND OBSERVED THAT THERE IS NO INCREASE OF COMPLAINTS BASED ON COMPLAINTS THE RELATIVE SENSITIVITY IS ABOUT AND CONFORM TO PACKAGE INSERT PERFORMANCES SENSITIVITY ACCORDING TO THE DATA MENTIONED ABOVE VIDAS CMV IGM LOT IS WITHIN THE EXPECTED SPECIFICATIONS | No Device Problem Found | Medical Evaluation result code
6876425 | 86618856 | D | INFORMATION WAS RECEIVED FROM A FRIEND OF A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR CONSUMER REPORTED THE PATIENT HAS BLADDER RETENTION AND FURTHER STATES THEIR BLADDER IS NOT DRAINING IT WAS CLARIFIED THAT THE PATIENT CURRENTLY HAS THIS SYMPTOM THE PATIENTS HEALTHCARE PROVIDER HCP IS NO LONGER IN THE PATIENTS INSURANCE NETWORK SO THE PATIENTS HCP REFERRED THEM TO ANOTHER UROLOGIST THE CONSUMER SAID THE PATIENT MADE AN APPOINTMENT WITH THE HCP BUT DOES NOT KNOW IF THE HCP IS ON THE HCP LISTINGS THE MANUFACTURING COMPANY SENT TO THEM IT WAS REVIEWED THAT THE LIST IS VOLUNTARY AND NOT ALL HCPS WILL BE ON IT AT THE TIME OF THE INITIAL REPORT THE CONSUMER DID NOT HAVE ACCESS TO THE PATIENT PROGRAMMER PP SO NO TROUBLESHOOTING COULD OCCUR DUE TO THE LACK OF ACCESS TO THE PRODUCT THE CONSUMER NOTED THAT IT HAS BEEN A LONG TIME SINCE THE INS WAS CHECKED AND THE PATIENT HAS DEMENTIA NOT RELATED TO DEVICETHERAPY IT IS UNKNOWN IF THE INS IS ON CONSUMER FURTHER CLARIFIED THAT THEY HAVE NOT BEEN REALLY GOOD WITH THE PP AND FURTHER NOTED THERE IS NOTHING WRONG WITH ANYTHING THE PATIENT HAS BEEN GOING IN THEIR PAD FOR YEARS SINCE GOING INTO ASSISTED LIVING THEY HAVE BEEN MAKING ADJUSTMENTS TO THEIR SETTINGS THAT THEY THOUGHT WOULD WORK FOR THEM WITHIN DOING THAT TIME WENT ON AND THE PATIENT GOT DEMENTIA SO THEY NO LONGER GO TO THE BATHROOM ON THEIR OWN AND ONLY GO IN THEIR PAD THE PATIENT HAS FOUND AN HCP THAT IS FAMILIAR WITH THE MANUFACTURING COMPANYS INS AND WILL GO SEE THEM NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6876425 | 86618856 | D | INFORMATION WAS RECEIVED FROM A FRIEND OF A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR CONSUMER REPORTED THE PATIENT HAS BLADDER RETENTION AND FURTHER STATES THEIR BLADDER IS NOT DRAINING IT WAS CLARIFIED THAT THE PATIENT CURRENTLY HAS THIS SYMPTOM THE PATIENTS HEALTHCARE PROVIDER HCP IS NO LONGER IN THE PATIENTS INSURANCE NETWORK SO THE PATIENTS HCP REFERRED THEM TO ANOTHER UROLOGIST THE CONSUMER SAID THE PATIENT MADE AN APPOINTMENT WITH THE HCP BUT DOES NOT KNOW IF THE HCP IS ON THE HCP LISTINGS THE MANUFACTURING COMPANY SENT TO THEM IT WAS REVIEWED THAT THE LIST IS VOLUNTARY AND NOT ALL HCPS WILL BE ON IT AT THE TIME OF THE INITIAL REPORT THE CONSUMER DID NOT HAVE ACCESS TO THE PATIENT PROGRAMMER PP SO NO TROUBLESHOOTING COULD OCCUR DUE TO THE LACK OF ACCESS TO THE PRODUCT THE CONSUMER NOTED THAT IT HAS BEEN A LONG TIME SINCE THE INS WAS CHECKED AND THE PATIENT HAS DEMENTIA NOT RELATED TO DEVICETHERAPY IT IS UNKNOWN IF THE INS IS ON CONSUMER FURTHER CLARIFIED THAT THEY HAVE NOT BEEN REALLY GOOD WITH THE PP AND FURTHER NOTED THERE IS NOTHING WRONG WITH ANYTHING THE PATIENT HAS BEEN GOING IN THEIR PAD FOR YEARS SINCE GOING INTO ASSISTED LIVING THEY HAVE BEEN MAKING ADJUSTMENTS TO THEIR SETTINGS THAT THEY THOUGHT WOULD WORK FOR THEM WITHIN DOING THAT TIME WENT ON AND THE PATIENT GOT DEMENTIA SO THEY NO LONGER GO TO THE BATHROOM ON THEIR OWN AND ONLY GO IN THEIR PAD THE PATIENT HAS FOUND AN HCP THAT IS FAMILIAR WITH THE MANUFACTURING COMPANYS INS AND WILL GO SEE THEM NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6876425 | 86618856 | D | INFORMATION WAS RECEIVED FROM A FRIEND OF A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR CONSUMER REPORTED THE PATIENT HAS BLADDER RETENTION AND FURTHER STATES THEIR BLADDER IS NOT DRAINING IT WAS CLARIFIED THAT THE PATIENT CURRENTLY HAS THIS SYMPTOM THE PATIENTS HEALTHCARE PROVIDER HCP IS NO LONGER IN THE PATIENTS INSURANCE NETWORK SO THE PATIENTS HCP REFERRED THEM TO ANOTHER UROLOGIST THE CONSUMER SAID THE PATIENT MADE AN APPOINTMENT WITH THE HCP BUT DOES NOT KNOW IF THE HCP IS ON THE HCP LISTINGS THE MANUFACTURING COMPANY SENT TO THEM IT WAS REVIEWED THAT THE LIST IS VOLUNTARY AND NOT ALL HCPS WILL BE ON IT AT THE TIME OF THE INITIAL REPORT THE CONSUMER DID NOT HAVE ACCESS TO THE PATIENT PROGRAMMER PP SO NO TROUBLESHOOTING COULD OCCUR DUE TO THE LACK OF ACCESS TO THE PRODUCT THE CONSUMER NOTED THAT IT HAS BEEN A LONG TIME SINCE THE INS WAS CHECKED AND THE PATIENT HAS DEMENTIA NOT RELATED TO DEVICETHERAPY IT IS UNKNOWN IF THE INS IS ON CONSUMER FURTHER CLARIFIED THAT THEY HAVE NOT BEEN REALLY GOOD WITH THE PP AND FURTHER NOTED THERE IS NOTHING WRONG WITH ANYTHING THE PATIENT HAS BEEN GOING IN THEIR PAD FOR YEARS SINCE GOING INTO ASSISTED LIVING THEY HAVE BEEN MAKING ADJUSTMENTS TO THEIR SETTINGS THAT THEY THOUGHT WOULD WORK FOR THEM WITHIN DOING THAT TIME WENT ON AND THE PATIENT GOT DEMENTIA SO THEY NO LONGER GO TO THE BATHROOM ON THEIR OWN AND ONLY GO IN THEIR PAD THE PATIENT HAS FOUND AN HCP THAT IS FAMILIAR WITH THE MANUFACTURING COMPANYS INS AND WILL GO SEE THEM NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Dementia | Patient problem code
6876413 | 86620285 | N | THIS MULTIPACK DEVICE CONSISTS OF SET SCREWS WITH PART WHICH IS NOT APPROVED FOR USE IN THE US HOWEVER A LIKE DEVICE WITH PART KK AN UPN B IS MARKETED IN THE US PRODUCT ANALYSIS VISUAL OPTICAL AND MICROSCOPIC EXAMINATION OF THE SET SCREW WHICH DISPLAYED HEAVY WITNESS MARKS ON THE PERIPHERAL FACE OF THE SET SCREW NODE HEIGHT AND MORPHOLOGY EVIDENCE OF THREAD OVERLOAD STARTING AT THE BEGINNING OF THE THREAD IS NOTED CONSISTENT WITH INCOMPLETE SEATING OF THE ROD IN THE MAS SADDLE AT FINAL TIGHTENING SET SCREW DISPLAYED EVIDENCE OF SIGNIFICANT DEFORMATION AT THE HEXALOBE INTERFACE ALSO CONSISTENT WITH INCOMPLETE SEATING OF THE ROD IN THE MAS SADDLE AT FINAL TIGHTENING IT NOTED THAT NO MENTION WAS MADE WITH THE EVENT DESCRIPTION REGARDING UTILIZATION OF REDUCTION INSTRUMENTATION PRIOR TO FINAL TIGHTENING THE ABOVE OBSERVATIONS ARE CONSISTENT WITH USING RMAS AND SET SCREW IMPLANTS TO REDUCE THE ROD IMMEDIATELY PRIOR TO FINAL TIGHTENING | No Device Problem Found | Medical Evaluation result code
6876124 | 86620692 | N | SYSTEM ANALYSIS ON B BSC SERVICE ENGINEER OBSERVED ERROR CODE DATA POINT OUT OF SPEC ATTEMPTED TO RECALIBRATE LASER FOUND RF AND FLOW WERE GOOD SUSPECTS FAULTY RESONATOR SERVICE ACTIVITY HAS NOT CONCLUDED | Device Displays Incorrect Message | Device problem code
6876124 | 86620692 | N | SYSTEM ANALYSIS ON B BSC SERVICE ENGINEER OBSERVED ERROR CODE DATA POINT OUT OF SPEC ATTEMPTED TO RECALIBRATE LASER FOUND RF AND FLOW WERE GOOD SUSPECTS FAULTY RESONATOR SERVICE ACTIVITY HAS NOT CONCLUDED | Calibration Problem | Device problem code
6876124 | 86620692 | N | SYSTEM ANALYSIS ON B BSC SERVICE ENGINEER OBSERVED ERROR CODE DATA POINT OUT OF SPEC ATTEMPTED TO RECALIBRATE LASER FOUND RF AND FLOW WERE GOOD SUSPECTS FAULTY RESONATOR SERVICE ACTIVITY HAS NOT CONCLUDED | Failure to Calibrate | Device Problem Code
6867076 | 86644605 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE CONTROLLER PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING HOWEVER REVIEW OF THE CONTROLLER LOG FILES SHOW A POWER DISCONNECT ALARM BETWEEN B AND CONTROLLER THIS WAS MOST LIKELY DUE TO A COMMUNICATION ANOMALY BETWEEN BATTERY AND CONTROLLER HEARTWARE IS CURRENTLY INVESTIGATING THE ANOMALIES IN COMMUNICATION BETWEEN BATTERIES AND CONTROLLER THERE WERE NO REPORTED ADVERSE PATIENT EFFECTS IN RELATION TO THE REPORTED EVENT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6877314 | 86689832 | N | INFORMATION CONTAINED IN THIS REPORT WAS PREVIOUSLY SUBMITTED THROUGH ASR ON APR THE EVENT OF PAIN STIFFNESS AND TIGHTNESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Operational Problem Identified | Medical Evaluation result code
6877486 | 86690063 | D | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO PAIN CAUSED BY A PATIENT FALL | Pain | Patient problem code
6877486 | 86690063 | D | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO PAIN CAUSED BY A PATIENT FALL | Insufficient Information | Device problem code
6877486 | 86690063 | D | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO PAIN CAUSED BY A PATIENT FALL | Fall | Patient problem code
6875082 | 86697785 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR ON RIGHT SIDE OF STOMACH THE NEEDLE HAD DEPLOYED EARLY INDICATING A NEEDLE MECHANISM FAILURE | Premature Activation | Device problem code
6875082 | 86697785 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LESS THAN ONE HOUR ON RIGHT SIDE OF STOMACH THE NEEDLE HAD DEPLOYED EARLY INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6877905 | 86713113 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF FILLER MIGRATION IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF MIGRATION OF DEVICE OR DEVICE COMPONENT POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Migration or Expulsion of Device | Device problem code
6877905 | 86713113 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF FILLER MIGRATION IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF MIGRATION OF DEVICE OR DEVICE COMPONENT POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Leak / Splash | Device problem code
6877905 | 86713113 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF FILLER MIGRATION IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF MIGRATION OF DEVICE OR DEVICE COMPONENT POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Leakage/Seal | Medical Evaluation result code
6813505 | 86724990 | N | WE HAD AN ORIGINAL REPORT OF A DEATH FROM OUR DISTRIBUTOR IN B IT DID NOT CONTAIN COMPLETE INFORMATION SUCH AS DATE OF DEATH WE REQUESTED ADDITIONAL INFORMATION WE RECEIVED THIS FURTHER INFORMATION LATE YESTERDAY FROM THE DISTRIBUTOR STATING PLEASE DISREGARD THE INFORMATION ABOUT A DEAD PATIENT THE PATIENT WAS NOT CONNECTED TO ANY EVENT VENTILATOR | Connection Problem | Device problem code
6879901 | 86765762 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN UPON POWER UP THE DEVICES DISPLAY TURNED COMPLETELY GRAY WITH VERTICAL LINES AND NO CLINICAL INFORMATION COULD BE SEEN COMPLAINANT INDICATED THAT THE CLINICIAN USED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6879901 | 86765762 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN UPON POWER UP THE DEVICES DISPLAY TURNED COMPLETELY GRAY WITH VERTICAL LINES AND NO CLINICAL INFORMATION COULD BE SEEN COMPLAINANT INDICATED THAT THE CLINICIAN USED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Erratic or Intermittent Display | Device Problem Code
6879957 | 86765850 | D | COMPLAINANT ALLEGED THAT THE DEVICE WAS UNABLE TO RECOGNIZE A BATTERY WAS INSTALLED AND INAPPROPRIATELY INDICATED AC POWER COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Battery Problem | Device Problem Code
6879957 | 86765850 | D | COMPLAINANT ALLEGED THAT THE DEVICE WAS UNABLE TO RECOGNIZE A BATTERY WAS INSTALLED AND INAPPROPRIATELY INDICATED AC POWER COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6880396 | 86791741 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS MGDL AND A CORRECTIVE BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL DUE TO GOING INTO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND ELECTROLYTES THE CUSTOMER WAS DISCHARGED THE NEXT DAY WITH THE HIGH BG AND DKA RESOLVED THE CAUSE OF THE HIGH BG WAS NOT KNOWN A PUMP SYSTEM CHECK WAS PERFORMED USING THE CURRENT SUPPLIES ON THE PUMP AND NO DEVICE ISSUES WERE IDENTIFIED | Adverse Event Without Identified Device or Use Problem | Device problem code
6880396 | 86791741 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS MGDL AND A CORRECTIVE BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL DUE TO GOING INTO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND ELECTROLYTES THE CUSTOMER WAS DISCHARGED THE NEXT DAY WITH THE HIGH BG AND DKA RESOLVED THE CAUSE OF THE HIGH BG WAS NOT KNOWN A PUMP SYSTEM CHECK WAS PERFORMED USING THE CURRENT SUPPLIES ON THE PUMP AND NO DEVICE ISSUES WERE IDENTIFIED | Hyperglycemia | Patient problem code
6880396 | 86791741 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS MGDL AND A CORRECTIVE BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL DUE TO GOING INTO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP AND ELECTROLYTES THE CUSTOMER WAS DISCHARGED THE NEXT DAY WITH THE HIGH BG AND DKA RESOLVED THE CAUSE OF THE HIGH BG WAS NOT KNOWN A PUMP SYSTEM CHECK WAS PERFORMED USING THE CURRENT SUPPLIES ON THE PUMP AND NO DEVICE ISSUES WERE IDENTIFIED | Diabetic Ketoacidosis | Patient problem code
6883604 | 86989536 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENTS BG LEVELS PRIOR TO BED WERE MMOLL MGDL WHEN THE PATIENT WOKE UP THE SMELL OF INSULIN WAS STRONG THE ADHESIVE PAD WAS WET AND BG LEVELS HAD RISEN TO MMOLL MGDL THE PATIENT WAS TAKEN TO THE HOSPITAL AND DIAGNOSED WITH DIABETIC KETOACIDOSIS DKA THE PATIENT WAS VOMITING AND KETONE LEVELS WERE AT AT THE HOSPITAL THE PATIENT WAS SEDATED WITH ATIVAN AND HALDO THE PATIENT WAS THEN TREATED WITH A POTASSIUM DRIP INSULIN DRIP AND SALINE DRIP | Insufficient Information | Device problem code
6883604 | 86989536 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENTS BG LEVELS PRIOR TO BED WERE MMOLL MGDL WHEN THE PATIENT WOKE UP THE SMELL OF INSULIN WAS STRONG THE ADHESIVE PAD WAS WET AND BG LEVELS HAD RISEN TO MMOLL MGDL THE PATIENT WAS TAKEN TO THE HOSPITAL AND DIAGNOSED WITH DIABETIC KETOACIDOSIS DKA THE PATIENT WAS VOMITING AND KETONE LEVELS WERE AT AT THE HOSPITAL THE PATIENT WAS SEDATED WITH ATIVAN AND HALDO THE PATIENT WAS THEN TREATED WITH A POTASSIUM DRIP INSULIN DRIP AND SALINE DRIP | Vomiting | Patient problem code
6881865 | 86989773 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON HIS LEG PATIENT HAD EXPERIENCED HEADACHE VOMITING AND STOMACH ACHE PATIENT WAS SUBSEQUENTLY SEEN BY A HEALTH CARE PROFESSIONAL AT B IN B ON B AND WAS TREATED WITH INSULIN DROP AND IV | Insufficient Information | Device problem code
6881865 | 86989773 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON HIS LEG PATIENT HAD EXPERIENCED HEADACHE VOMITING AND STOMACH ACHE PATIENT WAS SUBSEQUENTLY SEEN BY A HEALTH CARE PROFESSIONAL AT B IN B ON B AND WAS TREATED WITH INSULIN DROP AND IV | Vomiting | Patient problem code
6881865 | 86989773 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON HIS LEG PATIENT HAD EXPERIENCED HEADACHE VOMITING AND STOMACH ACHE PATIENT WAS SUBSEQUENTLY SEEN BY A HEALTH CARE PROFESSIONAL AT B IN B ON B AND WAS TREATED WITH INSULIN DROP AND IV | Headache | Patient problem code
6887943 | 87155931 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS AT THEIR LEAD SITE SO THEY WERE REDIRECTED TO THEIR HEALTHCARE PROVIDER HCP SIX DAYS LATER PATIENT HAD A URINARY TRACT INFECTION UTI AND SAID THINGS ARE UNCOMFORTABLE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient Problem Code
6887943 | 87155931 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS AT THEIR LEAD SITE SO THEY WERE REDIRECTED TO THEIR HEALTHCARE PROVIDER HCP SIX DAYS LATER PATIENT HAD A URINARY TRACT INFECTION UTI AND SAID THINGS ARE UNCOMFORTABLE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6887943 | 87155931 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS AT THEIR LEAD SITE SO THEY WERE REDIRECTED TO THEIR HEALTHCARE PROVIDER HCP SIX DAYS LATER PATIENT HAD A URINARY TRACT INFECTION UTI AND SAID THINGS ARE UNCOMFORTABLE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6887943 | 87155931 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS AT THEIR LEAD SITE SO THEY WERE REDIRECTED TO THEIR HEALTHCARE PROVIDER HCP SIX DAYS LATER PATIENT HAD A URINARY TRACT INFECTION UTI AND SAID THINGS ARE UNCOMFORTABLE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Discomfort | Patient problem code
6887943 | 87155931 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS AT THEIR LEAD SITE SO THEY WERE REDIRECTED TO THEIR HEALTHCARE PROVIDER HCP SIX DAYS LATER PATIENT HAD A URINARY TRACT INFECTION UTI AND SAID THINGS ARE UNCOMFORTABLE NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Urinary Tract | Patient problem code
6884329 | 87156826 | D | UPON CLINICAL ASSESSMENT OF THIS EVENT IT WAS DISCOVERED THAT A TYPE IIIB ENDOLEAK OF THE BIFURCATED STENT GRAFT WAS ALSO PRESENT | Hemorrhage | Patient problem code
6884329 | 87156826 | D | UPON CLINICAL ASSESSMENT OF THIS EVENT IT WAS DISCOVERED THAT A TYPE IIIB ENDOLEAK OF THE BIFURCATED STENT GRAFT WAS ALSO PRESENT | Insufficient Information | Device problem code
6882268 | 87158358 | D | A CENTER DIRECTOR REPORTED A PATIENT PRESENTED WITH OCULAR IRRITATION AND LIGHT SENSITIVITY TWO WEEKS POST LASIK THE PATIENT WAS NOTED TO HAVE A CORNEAL ABRASION ADDL INFORMATION REQUESTED | Irritation | Patient problem code
6882268 | 87158358 | D | A CENTER DIRECTOR REPORTED A PATIENT PRESENTED WITH OCULAR IRRITATION AND LIGHT SENSITIVITY TWO WEEKS POST LASIK THE PATIENT WAS NOTED TO HAVE A CORNEAL ABRASION ADDL INFORMATION REQUESTED | Corneal Abrasion | Patient problem code
6882268 | 87158358 | D | A CENTER DIRECTOR REPORTED A PATIENT PRESENTED WITH OCULAR IRRITATION AND LIGHT SENSITIVITY TWO WEEKS POST LASIK THE PATIENT WAS NOTED TO HAVE A CORNEAL ABRASION ADDL INFORMATION REQUESTED | Insufficient Information | Device problem code
6883940 | 87158562 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A B PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENTS PHYSICIAN IMPLANTED AN ICD EARLIER THAN PLANNED DUE TO THE ALLERGY | Hypersensitivity | Patient problem code
6883940 | 87158562 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A B PATIENT DEVELOPED AN ALLERGIC REACTION WHILE WEARING THE LIFEVEST THE PATIENTS PHYSICIAN IMPLANTED AN ICD EARLIER THAN PLANNED DUE TO THE ALLERGY | Insufficient Information | Device problem code
6882246 | 87159210 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD FOR HOURS PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B ON B AND WAS TREATED WITH FLUIDS AND INSULIN | Insufficient Information | Device problem code
6882246 | 87159210 | D | A PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD FOR HOURS PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B ON B AND WAS TREATED WITH FLUIDS AND INSULIN | No Known Impact Or Consequence To Patient | Patient Problem Code
6886184 | 87181469 | D | THE PATIENT REPORTED AFTER WEARING THE POD FOR HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED AND PAINFUL SHE WENT TO SEE HER DOCTOR WHO PRESCRIBED HER ANTIBIOTICS FOR HER SITE INFECTION | Pain | Patient problem code
6886184 | 87181469 | D | THE PATIENT REPORTED AFTER WEARING THE POD FOR HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED AND PAINFUL SHE WENT TO SEE HER DOCTOR WHO PRESCRIBED HER ANTIBIOTICS FOR HER SITE INFECTION | Insufficient Information | Device problem code
6886184 | 87181469 | D | THE PATIENT REPORTED AFTER WEARING THE POD FOR HOURS ON THE ARM SHE NOTICED HER SITE WAS INFECTED AND PAINFUL SHE WENT TO SEE HER DOCTOR WHO PRESCRIBED HER ANTIBIOTICS FOR HER SITE INFECTION | Infection | Patient problem code
6886827 | 87182888 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT WAS OUT OF IT WHEN THE VOIDING TRACKER WAS EXPLAINED TO THEM AND IS CONFUSED ON HOW TO USE IT TWO DAYS LATER PATIENT RATED THE EVALUATION A ON B PATIENT SOMETIMES FEELS A HIGH SENSE OF URGENCY BUT IS UNABLE TO VOID AT OTHER TIMES THE PATIENT HAS A VOID WITHOUT FEELING THE URGE TO GO NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Micturation Urgency | Patient problem code
6886827 | 87182888 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT WAS OUT OF IT WHEN THE VOIDING TRACKER WAS EXPLAINED TO THEM AND IS CONFUSED ON HOW TO USE IT TWO DAYS LATER PATIENT RATED THE EVALUATION A ON B PATIENT SOMETIMES FEELS A HIGH SENSE OF URGENCY BUT IS UNABLE TO VOID AT OTHER TIMES THE PATIENT HAS A VOID WITHOUT FEELING THE URGE TO GO NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6885685 | 87189390 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH VERIO IQ METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B SHE OBTAINED RESULTS OF AND MGDL ON THE SUBJECT METER PERFORMED MORE THAN MINUTES APART WHICH SHE FELT WERE INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND STATED THAT ON B SHE ADMINISTERED TWO DOSES OF INSULIN IN RESPONSE TO THE ALLEGEDLY HIGH RESULTS THE PATIENT REPORTED THAT ONE AND A HALF HOURS AFTER THE ALLEGED ISSUE OCCURRED SHE FAINTED AND THAT AT TO AM THE EMERGENCY MEDICAL SERVICES EMS WERE CALLED WHO PERFORMED A BLOOD GLUCOSE TEST ON AN EMS METER WHICH GAVE A RESULT OF MGDL AND SHE WAS TAKEN TO AN URGENT CARE CLINIC WHERE SHE WAS ADMINISTERED A GLUCAGON INJECTION AND GIVEN FOOD DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
6885685 | 87189390 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH VERIO IQ METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B SHE OBTAINED RESULTS OF AND MGDL ON THE SUBJECT METER PERFORMED MORE THAN MINUTES APART WHICH SHE FELT WERE INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND STATED THAT ON B SHE ADMINISTERED TWO DOSES OF INSULIN IN RESPONSE TO THE ALLEGEDLY HIGH RESULTS THE PATIENT REPORTED THAT ONE AND A HALF HOURS AFTER THE ALLEGED ISSUE OCCURRED SHE FAINTED AND THAT AT TO AM THE EMERGENCY MEDICAL SERVICES EMS WERE CALLED WHO PERFORMED A BLOOD GLUCOSE TEST ON AN EMS METER WHICH GAVE A RESULT OF MGDL AND SHE WAS TAKEN TO AN URGENT CARE CLINIC WHERE SHE WAS ADMINISTERED A GLUCAGON INJECTION AND GIVEN FOOD DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Insufficient Information | Device problem code
6885685 | 87189390 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT HER ONETOUCH VERIO IQ METER READ INACCURATELY HIGH IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT REPORTED THAT ON B SHE OBTAINED RESULTS OF AND MGDL ON THE SUBJECT METER PERFORMED MORE THAN MINUTES APART WHICH SHE FELT WERE INACCURATELY HIGH METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES BY SELFADJUSTING HER INSULIN DOSES AND STATED THAT ON B SHE ADMINISTERED TWO DOSES OF INSULIN IN RESPONSE TO THE ALLEGEDLY HIGH RESULTS THE PATIENT REPORTED THAT ONE AND A HALF HOURS AFTER THE ALLEGED ISSUE OCCURRED SHE FAINTED AND THAT AT TO AM THE EMERGENCY MEDICAL SERVICES EMS WERE CALLED WHO PERFORMED A BLOOD GLUCOSE TEST ON AN EMS METER WHICH GAVE A RESULT OF MGDL AND SHE WAS TAKEN TO AN URGENT CARE CLINIC WHERE SHE WAS ADMINISTERED A GLUCAGON INJECTION AND GIVEN FOOD DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND AN APPROVED SAMPLE SITE WAS USED THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Hyperglycemia | Patient Problem Code
6878707 | 87198768 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6878707 | 87198768 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6890058 | 87199585 | D | MEDTRONIC RECEIVED INFORMATION THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6890058 | 87199585 | D | MEDTRONIC RECEIVED INFORMATION THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6886143 | 87206496 | D | A USER FACILITY REPORTED THAT ONE PATIENT SUSTAINED BURNS TO AN UNKNOWN ANATOMICAL AREA FOLLOWING TREATMENT WITH A LUMENIS LIGHTSHEER DESIRE LASER NO REPORTS OF SERIOUS INJURY OR MEDICAL INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT WERE RECEIVED | No Known Impact Or Consequence To Patient | Patient problem code
6886143 | 87206496 | D | A USER FACILITY REPORTED THAT ONE PATIENT SUSTAINED BURNS TO AN UNKNOWN ANATOMICAL AREA FOLLOWING TREATMENT WITH A LUMENIS LIGHTSHEER DESIRE LASER NO REPORTS OF SERIOUS INJURY OR MEDICAL INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT WERE RECEIVED | Insufficient Information | Device problem code
6882025 | 87210745 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6882025 | 87210745 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6882033 | 87210986 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6882033 | 87210986 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6882024 | 87211505 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED RISING BG LEVELS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6882024 | 87211505 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED RISING BG LEVELS THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6881892 | 87212562 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6881892 | 87212562 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | Positioning Problem | Device problem code
6890405 | 87216705 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6890405 | 87216705 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6889846 | 87217869 | D | ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LIVING DONOR NEPHRECTOMY PROCEDURE THE DEVICE CUT TISSUE POORLY AND DID NOT GRASP PROPERLY DU TO TISSUE PARTICLES ADHERED TO THE HINGE OF THE JAWS A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE NO PATIENT HARM | Misfocusing | Device Problem Code
6889846 | 87217869 | D | ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LIVING DONOR NEPHRECTOMY PROCEDURE THE DEVICE CUT TISSUE POORLY AND DID NOT GRASP PROPERLY DU TO TISSUE PARTICLES ADHERED TO THE HINGE OF THE JAWS A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE NO PATIENT HARM | No Consequences Or Impact To Patient | Patient problem code
6889547 | 87220071 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS AND PATIENT REPORTING CANNULA DID NOT FEEL THE SAME AS THE OTHERS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6889547 | 87220071 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS AND PATIENT REPORTING CANNULA DID NOT FEEL THE SAME AS THE OTHERS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6889809 | 87231006 | D | ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LIVING DONOR NEPHRECTOMY PROCEDURE THE DEVICE CUT TISSUE POORLY AND DID NOT GRASP PROPERLY DU TO TISSUE PARTICLES ADHERED TO THE HINGE OF THE JAWS A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE NO PATIENT HARM | Misfocusing | Device Problem Code
6889809 | 87231006 | D | ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LIVING DONOR NEPHRECTOMY PROCEDURE THE DEVICE CUT TISSUE POORLY AND DID NOT GRASP PROPERLY DU TO TISSUE PARTICLES ADHERED TO THE HINGE OF THE JAWS A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE NO PATIENT HARM | No Consequences Or Impact To Patient | Patient problem code
6882564 | 87231966 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER LEG FOR BETWEEN AND HOURS SHE STATED THAT THE PINK SLIDE WAS OFF TO THE LEFT OF THE WHITE CIRCLE TREATMENT INCLUDED REPLACING THE POD | Component or Accessory Incompatibility | Device problem code
6882564 | 87231966 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER LEG FOR BETWEEN AND HOURS SHE STATED THAT THE PINK SLIDE WAS OFF TO THE LEFT OF THE WHITE CIRCLE TREATMENT INCLUDED REPLACING THE POD | No Known Impact Or Consequence To Patient | Patient Problem Code
6890843 | 87249214 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6890843 | 87249214 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6893051 | 87296292 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS AND HOSPITALIZED TREATMENT INCLUDED INTRAVENOUS DELIVERY OF INSULIN | Diabetic Ketoacidosis | Patient problem code
6893051 | 87296292 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS AND HOSPITALIZED TREATMENT INCLUDED INTRAVENOUS DELIVERY OF INSULIN | Insufficient Information | Device problem code
6893077 | 87299708 | D | ORIGINAL EVAR PERFORMED B IDENTIFIED A TYPE A ENDOLEAK DURING PATIENT FOLLOWUP ON B REINTERVENTION PLANNED AND PERFORMED ON B | Hemorrhage | Patient problem code
6893077 | 87299708 | D | ORIGINAL EVAR PERFORMED B IDENTIFIED A TYPE A ENDOLEAK DURING PATIENT FOLLOWUP ON B REINTERVENTION PLANNED AND PERFORMED ON B | Insufficient Information | Device problem code
6893076 | 87307453 | D | IT WAS REPORTED THE PATIENTS BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN HYPERGLYCEMIA TREATED BY PATIENT WITH AN INSULIN PEN | Insufficient Information | Device problem code
6893076 | 87307453 | D | IT WAS REPORTED THE PATIENTS BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN HYPERGLYCEMIA TREATED BY PATIENT WITH AN INSULIN PEN | Hyperglycemia | Patient problem code
6882017 | 87309672 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED BUBBLES AND MOISTURE AT SITE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6882017 | 87309672 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED BUBBLES AND MOISTURE AT SITE THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6881880 | 87310062 | D | A PATIENTS BLOOD GLUCOSE REACHED ALMOST MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6881880 | 87310062 | D | A PATIENTS BLOOD GLUCOSE REACHED ALMOST MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6881885 | 87310616 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6881885 | 87310616 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6870723 | 87312655 | D | IT WAS REPORTED THAT A NEEDLE STICK INJURY OCCURRED WITH A BD NEEDLE G X INCH SINCE THE NEEDLE WAS FOUND WITHOUT ITS CAP THE STERILE BARRIER WAS CONSIDERED TO BE BREACHED THERE WAS NO REPORT OF INJURY OR MEDICAL INTERVENTION | Insufficient Information | Device problem code
6870723 | 87312655 | D | IT WAS REPORTED THAT A NEEDLE STICK INJURY OCCURRED WITH A BD NEEDLE G X INCH SINCE THE NEEDLE WAS FOUND WITHOUT ITS CAP THE STERILE BARRIER WAS CONSIDERED TO BE BREACHED THERE WAS NO REPORT OF INJURY OR MEDICAL INTERVENTION | No Known Impact Or Consequence To Patient | Patient problem code
6882029 | 87322001 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6882029 | 87322001 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6881874 | 87322169 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | Positioning Problem | Device problem code
6881874 | 87322169 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6882031 | 87322288 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6882031 | 87322288 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6873265 | 87348048 | D | IT WAS REPORTED THAT THE INTRAAORTIC BALLOON PUMP IABP HAD A LOW BATTERY WARNING ALARM WITH FULLY CHARGED BATTERIES THIS EVENT OCCURRED DURING SERVICETEST PERFORMED BY THE CUSTOMER THERE WAS NO PATIENT INVOLVED THUS NO INJURY OR ADVERSE EVENT WAS REPORTED | No Patient Involvement | Patient problem code
6873265 | 87348048 | D | IT WAS REPORTED THAT THE INTRAAORTIC BALLOON PUMP IABP HAD A LOW BATTERY WARNING ALARM WITH FULLY CHARGED BATTERIES THIS EVENT OCCURRED DURING SERVICETEST PERFORMED BY THE CUSTOMER THERE WAS NO PATIENT INVOLVED THUS NO INJURY OR ADVERSE EVENT WAS REPORTED | Device Alarm System | Device problem code
6892078 | 87374212 | D | A SITE REPRESENTATIVE REPORTED THAT WHILE OUTSIDE OF PROCEDURE THE MOBILE VIEW STATION MVS WOULD NOT HOLD CHARGE AND THE SYSTEM POWERS DOWN WHEN UNPLUGGED FROM AN POWER OUTLET THERE WAS NO PATIENT PRESENT AT THE TIME OF ISSUE | No Patient Involvement | Patient problem code
6892078 | 87374212 | D | A SITE REPRESENTATIVE REPORTED THAT WHILE OUTSIDE OF PROCEDURE THE MOBILE VIEW STATION MVS WOULD NOT HOLD CHARGE AND THE SYSTEM POWERS DOWN WHEN UNPLUGGED FROM AN POWER OUTLET THERE WAS NO PATIENT PRESENT AT THE TIME OF ISSUE | Insufficient Information | Device problem code
6894180 | 87382636 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MMOLL MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS THE DEVICE WAS RETURNED FOR EVALUATION | Insufficient Information | Device problem code
6894180 | 87382636 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE REACHED MMOLL MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS THE DEVICE WAS RETURNED FOR EVALUATION | No Known Impact Or Consequence To Patient | Patient Problem Code
6893842 | 87384386 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893842 | 87384386 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6880380 | 87390431 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY WAS DISTORTED AND CLINICAL INFORMATION COULD NOT BE SEEN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6880380 | 87390431 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY WAS DISTORTED AND CLINICAL INFORMATION COULD NOT BE SEEN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Erratic or Intermittent Display | Device Problem Code
6894864 | 87395243 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6894864 | 87395243 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6896501 | 87441687 | D | IT WAS REPORTED BY THE CUSTOMERS PARENT THAT THE CUSTOMER RECEIVED GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH UNKNOWN BLOOD GLUCOSE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MDL AT THE TIME OF ADMISSION THE CUSTOMER BOLUSED FOR BREAKFAST AND WITHIN AN HOUR WENT INTO A SEIZURE THE CUSTOMER WAS GIVEN GLUCAGON TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS A SEIZURE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED AND THE CALLER DOES NOT BELIEVE THE PUMP OVER DELIVERED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6896501 | 87441687 | D | IT WAS REPORTED BY THE CUSTOMERS PARENT THAT THE CUSTOMER RECEIVED GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH UNKNOWN BLOOD GLUCOSE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MDL AT THE TIME OF ADMISSION THE CUSTOMER BOLUSED FOR BREAKFAST AND WITHIN AN HOUR WENT INTO A SEIZURE THE CUSTOMER WAS GIVEN GLUCAGON TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS A SEIZURE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED AND THE CALLER DOES NOT BELIEVE THE PUMP OVER DELIVERED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Seizures | Patient problem code
6896501 | 87441687 | D | IT WAS REPORTED BY THE CUSTOMERS PARENT THAT THE CUSTOMER RECEIVED GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH UNKNOWN BLOOD GLUCOSE AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL AND MDL AT THE TIME OF ADMISSION THE CUSTOMER BOLUSED FOR BREAKFAST AND WITHIN AN HOUR WENT INTO A SEIZURE THE CUSTOMER WAS GIVEN GLUCAGON TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS A SEIZURE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED AND THE CALLER DOES NOT BELIEVE THE PUMP OVER DELIVERED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Excess Flow or Over-Infusion | Device problem code
6896654 | 87443599 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON HIS LEG PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B ON B AND WAS TREATED WITH MANUAL INJECTION SHOTS OF INSULIN | No Known Impact Or Consequence To Patient | Patient Problem Code
6896654 | 87443599 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON HIS LEG PATIENT WAS SEEN BY A HEALTH CARE PROFESSIONAL AT B ON B AND WAS TREATED WITH MANUAL INJECTION SHOTS OF INSULIN | Insufficient Information | Device problem code
6896569 | 87447542 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH COMPARED TO HER DOCTORS DEVICE BRAND NOT REPORTED THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT PM ON B THE PATIENT CLAIMED OBTAINING BLOOD GLUCOSE READINGS OF MGDL ON THE SUBJECT METER AND MGDL WITH THE OTHER DEVICE PERFORMED WITHIN MINUTES OF EACH OTHER METER TO OTHER METER COMPARISONS DO NOT REASONABLY SUGGEST THAT A MALFUNCTION HAS OCCURRED THERE CAN BE NO PRESUMPTION AS TO WHICH METERS READING IS ERRONEOUS AS THE COMPARISON IS NOT MADE TO A CALIBRATED REFERENCE METHOD THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN NOVOLOG AND LEVEMIR AND CLAIMED THAT IN RESPONSE TO THE ALLEGED ISSUE SHE ADMINISTERED UNITS OF LEVEMIR WHICH SHE ADVISED WAS HER DOSE BASED ON A SLIDING SCALE THE PATIENT REPORTED THAT A COUPLE OF HOURS AFTER THE ALLEGED ISSUE BEGAN SHE DEVELOPED SYMPTOMS OF FEELING SWEATY WITH AN URGENCY TO EAT THERE WAS NO REPORT OF MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE THE PATIENT DENIED TESTING HER BLOOD GLUCOSE ON ANOTHER DEVICE AT THE TIME OF TROUBLESHOOTING THE CSR ESTABLISHED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE CSR CONFIRMED THAT THE TEST STRIP VIAL WAS INTACT THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR ALSO NOTED THAT THE PATIENT DID NOT HAVE TESTING SUPPLIES AVAILABLE TO TEST THE SUBJECT DEVICE REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER ADMINISTERING MEDICATION BASED ON THE ALLEGEDLY ELEVATED READINGS OBTAINED ON THE SUBJECT DEVICE | Sweating | Patient problem code
6896569 | 87447542 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH COMPARED TO HER DOCTORS DEVICE BRAND NOT REPORTED THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT PM ON B THE PATIENT CLAIMED OBTAINING BLOOD GLUCOSE READINGS OF MGDL ON THE SUBJECT METER AND MGDL WITH THE OTHER DEVICE PERFORMED WITHIN MINUTES OF EACH OTHER METER TO OTHER METER COMPARISONS DO NOT REASONABLY SUGGEST THAT A MALFUNCTION HAS OCCURRED THERE CAN BE NO PRESUMPTION AS TO WHICH METERS READING IS ERRONEOUS AS THE COMPARISON IS NOT MADE TO A CALIBRATED REFERENCE METHOD THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN NOVOLOG AND LEVEMIR AND CLAIMED THAT IN RESPONSE TO THE ALLEGED ISSUE SHE ADMINISTERED UNITS OF LEVEMIR WHICH SHE ADVISED WAS HER DOSE BASED ON A SLIDING SCALE THE PATIENT REPORTED THAT A COUPLE OF HOURS AFTER THE ALLEGED ISSUE BEGAN SHE DEVELOPED SYMPTOMS OF FEELING SWEATY WITH AN URGENCY TO EAT THERE WAS NO REPORT OF MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE THE PATIENT DENIED TESTING HER BLOOD GLUCOSE ON ANOTHER DEVICE AT THE TIME OF TROUBLESHOOTING THE CSR ESTABLISHED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE CSR CONFIRMED THAT THE TEST STRIP VIAL WAS INTACT THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR ALSO NOTED THAT THE PATIENT DID NOT HAVE TESTING SUPPLIES AVAILABLE TO TEST THE SUBJECT DEVICE REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER ADMINISTERING MEDICATION BASED ON THE ALLEGEDLY ELEVATED READINGS OBTAINED ON THE SUBJECT DEVICE | Micturation Urgency | Patient problem code
6896569 | 87447542 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH COMPARED TO HER DOCTORS DEVICE BRAND NOT REPORTED THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT PM ON B THE PATIENT CLAIMED OBTAINING BLOOD GLUCOSE READINGS OF MGDL ON THE SUBJECT METER AND MGDL WITH THE OTHER DEVICE PERFORMED WITHIN MINUTES OF EACH OTHER METER TO OTHER METER COMPARISONS DO NOT REASONABLY SUGGEST THAT A MALFUNCTION HAS OCCURRED THERE CAN BE NO PRESUMPTION AS TO WHICH METERS READING IS ERRONEOUS AS THE COMPARISON IS NOT MADE TO A CALIBRATED REFERENCE METHOD THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN NOVOLOG AND LEVEMIR AND CLAIMED THAT IN RESPONSE TO THE ALLEGED ISSUE SHE ADMINISTERED UNITS OF LEVEMIR WHICH SHE ADVISED WAS HER DOSE BASED ON A SLIDING SCALE THE PATIENT REPORTED THAT A COUPLE OF HOURS AFTER THE ALLEGED ISSUE BEGAN SHE DEVELOPED SYMPTOMS OF FEELING SWEATY WITH AN URGENCY TO EAT THERE WAS NO REPORT OF MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE THE PATIENT DENIED TESTING HER BLOOD GLUCOSE ON ANOTHER DEVICE AT THE TIME OF TROUBLESHOOTING THE CSR ESTABLISHED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE CSR CONFIRMED THAT THE TEST STRIP VIAL WAS INTACT THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR ALSO NOTED THAT THE PATIENT DID NOT HAVE TESTING SUPPLIES AVAILABLE TO TEST THE SUBJECT DEVICE REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER ADMINISTERING MEDICATION BASED ON THE ALLEGEDLY ELEVATED READINGS OBTAINED ON THE SUBJECT DEVICE | Injury | Patient problem code
6896569 | 87447542 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH COMPARED TO HER DOCTORS DEVICE BRAND NOT REPORTED THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT PM ON B THE PATIENT CLAIMED OBTAINING BLOOD GLUCOSE READINGS OF MGDL ON THE SUBJECT METER AND MGDL WITH THE OTHER DEVICE PERFORMED WITHIN MINUTES OF EACH OTHER METER TO OTHER METER COMPARISONS DO NOT REASONABLY SUGGEST THAT A MALFUNCTION HAS OCCURRED THERE CAN BE NO PRESUMPTION AS TO WHICH METERS READING IS ERRONEOUS AS THE COMPARISON IS NOT MADE TO A CALIBRATED REFERENCE METHOD THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN NOVOLOG AND LEVEMIR AND CLAIMED THAT IN RESPONSE TO THE ALLEGED ISSUE SHE ADMINISTERED UNITS OF LEVEMIR WHICH SHE ADVISED WAS HER DOSE BASED ON A SLIDING SCALE THE PATIENT REPORTED THAT A COUPLE OF HOURS AFTER THE ALLEGED ISSUE BEGAN SHE DEVELOPED SYMPTOMS OF FEELING SWEATY WITH AN URGENCY TO EAT THERE WAS NO REPORT OF MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE THE PATIENT DENIED TESTING HER BLOOD GLUCOSE ON ANOTHER DEVICE AT THE TIME OF TROUBLESHOOTING THE CSR ESTABLISHED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE CSR CONFIRMED THAT THE TEST STRIP VIAL WAS INTACT THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR ALSO NOTED THAT THE PATIENT DID NOT HAVE TESTING SUPPLIES AVAILABLE TO TEST THE SUBJECT DEVICE REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER ADMINISTERING MEDICATION BASED ON THE ALLEGEDLY ELEVATED READINGS OBTAINED ON THE SUBJECT DEVICE | Hyperglycemia | Patient Problem Code
6896569 | 87447542 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH COMPARED TO HER DOCTORS DEVICE BRAND NOT REPORTED THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT PM ON B THE PATIENT CLAIMED OBTAINING BLOOD GLUCOSE READINGS OF MGDL ON THE SUBJECT METER AND MGDL WITH THE OTHER DEVICE PERFORMED WITHIN MINUTES OF EACH OTHER METER TO OTHER METER COMPARISONS DO NOT REASONABLY SUGGEST THAT A MALFUNCTION HAS OCCURRED THERE CAN BE NO PRESUMPTION AS TO WHICH METERS READING IS ERRONEOUS AS THE COMPARISON IS NOT MADE TO A CALIBRATED REFERENCE METHOD THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN NOVOLOG AND LEVEMIR AND CLAIMED THAT IN RESPONSE TO THE ALLEGED ISSUE SHE ADMINISTERED UNITS OF LEVEMIR WHICH SHE ADVISED WAS HER DOSE BASED ON A SLIDING SCALE THE PATIENT REPORTED THAT A COUPLE OF HOURS AFTER THE ALLEGED ISSUE BEGAN SHE DEVELOPED SYMPTOMS OF FEELING SWEATY WITH AN URGENCY TO EAT THERE WAS NO REPORT OF MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE THE PATIENT DENIED TESTING HER BLOOD GLUCOSE ON ANOTHER DEVICE AT THE TIME OF TROUBLESHOOTING THE CSR ESTABLISHED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT AN APPROVED SAMPLE SITE WAS USED TO OBTAIN THE RESULTS THE CSR CONFIRMED THAT THE TEST STRIP VIAL WAS INTACT THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR ALSO NOTED THAT THE PATIENT DID NOT HAVE TESTING SUPPLIES AVAILABLE TO TEST THE SUBJECT DEVICE REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER ADMINISTERING MEDICATION BASED ON THE ALLEGEDLY ELEVATED READINGS OBTAINED ON THE SUBJECT DEVICE | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6891769 | 87467343 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A BOLUS WAS ADMINISTERED AND THE BASAL RATE WAS ADJUSTED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6891769 | 87467343 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT NOTED THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A BOLUS WAS ADMINISTERED AND THE BASAL RATE WAS ADJUSTED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6892816 | 87470102 | D | IT WAS REPORTED THAT THE CUSTOMER PUMPS HRM TASK DID NOT GRANT MOTOR POWER IN REASONABLE TIME THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE SELFTEST WAS PERFORMED AND TEST PASSED THE CUSTOMER ALSO STATED THAT THE LOAD RESERVOIR AND FILL TUBING PROCESS WAS SUCCESSFUL AND NO OCCLUSION WAS FOUND THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6892816 | 87470102 | D | IT WAS REPORTED THAT THE CUSTOMER PUMPS HRM TASK DID NOT GRANT MOTOR POWER IN REASONABLE TIME THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE SELFTEST WAS PERFORMED AND TEST PASSED THE CUSTOMER ALSO STATED THAT THE LOAD RESERVOIR AND FILL TUBING PROCESS WAS SUCCESSFUL AND NO OCCLUSION WAS FOUND THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Power Problem | Device problem code
6874644 | 87486426 | D | IT WAS REPORTED THAT DURING A ROUTINE VISIT THE PATIENTS BACKUP CONTROLLER HAD A ZEROED OUT DATE AND WOULD NOT HOLD DATE AND TIME THE CONTROLLER WAS EXCHANGED WITH NO REPORTED CONSEQUENCE OR IMPACT TO THE PATIENT NO FURTHER INFORMATION WAS PROVIDED | No Consequences Or Impact To Patient | Patient problem code
6874644 | 87486426 | D | IT WAS REPORTED THAT DURING A ROUTINE VISIT THE PATIENTS BACKUP CONTROLLER HAD A ZEROED OUT DATE AND WOULD NOT HOLD DATE AND TIME THE CONTROLLER WAS EXCHANGED WITH NO REPORTED CONSEQUENCE OR IMPACT TO THE PATIENT NO FURTHER INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
6897405 | 87493874 | N | MEDICAL DEVICES SCITATION INOTROPES ADMINISTERED PATIENT CARDIOVERTED X UNRESPONSIVE TO TREATMENT DIALYSIS CEASED QNB | Component or Accessory Incompatibility | Device problem code
6897405 | 87493874 | N | MEDICAL DEVICES SCITATION INOTROPES ADMINISTERED PATIENT CARDIOVERTED X UNRESPONSIVE TO TREATMENT DIALYSIS CEASED QNB | Incompatible Component/ Accessory | Medical Evaluation result code
6897617 | 87507523 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAS BEEN GOING MORE FREQUENT AND THEY DIDNT SLEEP WELL THE PATIENT SAID THEIR BOWELS ARE BETTER BUT THEY DID HAVE TO TAKE A LAXATIVE THE NEXT DAY THE PATIENTS SPOUSE SAID THEY RECEIVED CONFLICTING INFORMATION ABOUT STIMULATION AND THE SPOUSE WAS CONFUSED THE PATIENTS SPOUSE SAID WHEN STIMULATION WAS INCREASED FREQUENCY INCREASED AND VOID VOLUME DECREASED AS A RESULT OF WHAT WAS REPORTED STIMULATION WAS DECREASED NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6897617 | 87507523 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT HAS BEEN GOING MORE FREQUENT AND THEY DIDNT SLEEP WELL THE PATIENT SAID THEIR BOWELS ARE BETTER BUT THEY DID HAVE TO TAKE A LAXATIVE THE NEXT DAY THE PATIENTS SPOUSE SAID THEY RECEIVED CONFLICTING INFORMATION ABOUT STIMULATION AND THE SPOUSE WAS CONFUSED THE PATIENTS SPOUSE SAID WHEN STIMULATION WAS INCREASED FREQUENCY INCREASED AND VOID VOLUME DECREASED AS A RESULT OF WHAT WAS REPORTED STIMULATION WAS DECREASED NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6897396 | 87508105 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6897396 | 87508105 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6898875 | 87581105 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6898898 | 87581154 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6898897 | 87581481 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6898946 | 87583042 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6898343 | 87583718 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT WAS IN THE HOSPITAL BECAUSE THEY HAVE AN INFECTION AND WERE WAITING ON ANTIBIOTICS FIVE DAYS LATER THE PATIENT WAS RELEASED NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient problem code
6898343 | 87583718 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT WAS IN THE HOSPITAL BECAUSE THEY HAVE AN INFECTION AND WERE WAITING ON ANTIBIOTICS FIVE DAYS LATER THE PATIENT WAS RELEASED NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6898343 | 87583718 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT WAS IN THE HOSPITAL BECAUSE THEY HAVE AN INFECTION AND WERE WAITING ON ANTIBIOTICS FIVE DAYS LATER THE PATIENT WAS RELEASED NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6898139 | 87586818 | D | HILLROM RECEIVED A REPORT FROM THE ACCOUNT STATING THE PATIENT DEVELOPED A NEW STAGE PRESSURE ULCER ON HIS BUTTOCKS THE BED WAS LOCATED AT THE ACCOUNT THIS REPORT WAS FILED IN OUR COMPLAINT HANDLING SYSTEM AS COMPLAINT B | Ulcer | Patient problem code
6898139 | 87586818 | D | HILLROM RECEIVED A REPORT FROM THE ACCOUNT STATING THE PATIENT DEVELOPED A NEW STAGE PRESSURE ULCER ON HIS BUTTOCKS THE BED WAS LOCATED AT THE ACCOUNT THIS REPORT WAS FILED IN OUR COMPLAINT HANDLING SYSTEM AS COMPLAINT B | Insufficient Information | Device problem code
6898139 | 87586818 | D | HILLROM RECEIVED A REPORT FROM THE ACCOUNT STATING THE PATIENT DEVELOPED A NEW STAGE PRESSURE ULCER ON HIS BUTTOCKS THE BED WAS LOCATED AT THE ACCOUNT THIS REPORT WAS FILED IN OUR COMPLAINT HANDLING SYSTEM AS COMPLAINT B | Pressure Sores/Ulcers | Patient problem code
6899841 | 87591047 | D | A USER FACILITY REPORTED THAT ONE MALE PATIENT SUSTAINED FIRST DEGREE BURNS TO THE FACE FOLLOWING HAIR REMOVAL TREATMENT WITH A LUMENIS LIGHTSHEER DESIRE LASER | First-Degree Burn | Patient problem code
6899841 | 87591047 | D | A USER FACILITY REPORTED THAT ONE MALE PATIENT SUSTAINED FIRST DEGREE BURNS TO THE FACE FOLLOWING HAIR REMOVAL TREATMENT WITH A LUMENIS LIGHTSHEER DESIRE LASER | Insufficient Information | Device problem code
6886944 | 87605510 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6886944 | 87605510 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6887010 | 87606294 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6887010 | 87606294 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6886924 | 87606858 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6886924 | 87606858 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6886921 | 87607138 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOUR MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATEB | Component or Accessory Incompatibility | Device Problem Code
6886921 | 87607138 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOUR MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATEB | Incompatible Component/ Accessory | Medical Evaluation result code
6890042 | 87611410 | N | ZYNO MEDICAL HAS RECEIVED THE INVESTIGATION REPORT FROM THE CONTRACT SUPPLIER ON B THE FREEFLOWING ISSUE WITH THE PINCH CLAMP CLOSED WAS CONFIRMED AS AN ISOLATED EVENT WHICH MAY BE DUE TO THE DISLOCATION OF THE PINCH CLAMP THE DETAILS CAN BE FOUND IN THE ATTACHED REPORT UPON RECEIVING THE COMPLAINT IT WAS INITIALLY DETERMINED AS NOT REPORTABLE BY THE MANUFACTURER BY FOLLOWING THE COMPANY PROCEDURES THIS MDR IS FILED RETROSPECTIVELY FILED ACCORDING TO THE CURRENT INTERNAL MDR DECISION GUIDANCE AT ZYNO MEDICAL THE MDR DECISION WAS MADE ON B B | Device Dislodged or Dislocated | Device problem code
6894360 | 87613685 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE FIRST DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED IF YES HOW WAS THE DEVICE REMOVED DID THE JAWS EVENTUALLY OPEN WAS THE SECOND DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED IF YES HOW WAS THE DEVICE REMOVED DID THE JAWS EVENTUALLY OPEN WHAT WAS THE PRODUCT CODE OF THE CARTRIDGE GSTT ECRD ETC | Mechanical Jam | Device problem code
6894360 | 87613685 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE FIRST DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED IF YES HOW WAS THE DEVICE REMOVED DID THE JAWS EVENTUALLY OPEN WAS THE SECOND DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED IF YES HOW WAS THE DEVICE REMOVED DID THE JAWS EVENTUALLY OPEN WHAT WAS THE PRODUCT CODE OF THE CARTRIDGE GSTT ECRD ETC | Mechanical Problem Identified | Medical Evaluation result code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Thrombus | Patient problem code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Loss Of Consciousness | Patient problem code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Hemorrhage | Patient problem code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Ischemia | Patient problem code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Insufficient Information | Device problem code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Cerebrovascular Accident | Patient Problem Code
6900318 | 87616233 | D | IT WAS REPORTED A STROKE AND DEATH OCCURRED ON B A LEFT ATRIAL APPENDAGE LAA CLOSURE PROCEDURE WAS BEING PERFORMED TWO WATCHMAN ACCESS SYSTEMS AND THREE MM WATCHMAN LAA CLOSURE DEVICE AND DELIVERY SYSTEMS WERE USED BUT ALL WERE UNSUCCESSFUL AS THE LAA WAS TOO LARGE TO ACCOMMODATE A CLOSURE DEVICE NO WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED THE PROCEDURE WAS ABORTED BECAUSE OF THIS A TRANSESOPHAGEAL ECHOCARDIOGRAM TEE FROM THE PROCEDURE DID NOT SHOW ANY EVIDENCE OF THROMBUS THE PATIENT WAS SENT TO THE FLOOR FOR MONITORING AND SUBSEQUENTLY SUFFERED AN ISCHEMIC STROKE THE EVENING FOLLOWING THE PROCEDURE THE PATIENT HAD SYMPTOMS OF APHASIA AND AN ALTERED LEVEL OF CONSCIOUSNESS AND BECAME COMBATIVE THE COMPUTED TOMOGRAPHY CT SCAN DID NOT SHOW ANY EVIDENCE OF BLEEDING AND SHOWED NO ACUTE INTRACRANIAL PROCESS A MAGNETIC RESONANCE IMAGING MRIWAS NOT PERFORMED BECAUSE THE PATIENT HAD A PACEMAKER THE PATIENT WAS NOT ON ANY ANTICOAGULANT MEDICATION PRIOR TO THE PROCEDURE BUT WAS GIVEN ONE DOSE OF APIXABAN IMMEDIATELY FOLLOWING THE ABORTED PROCEDURE THE PATIENT HAD BEEN ON ELIQUIS AND MG OF ASPIRIN THE PATIENT WAS DISCHARGED TO A SKILLED NURSING FACILITY ON B THE HOSPITAL WAS NOTIFIED BY THE PATIENTS FAMILY THAT THE PATIENT HAD DIED IN THE NURSING HOME THE EXACT DATE AND CAUSE OF DEATH ARE UNKNOWN | Death | Patient problem code
6899898 | 87622147 | N | B RESULTS OF INVESTIGATION THE VYAIRE FAILURE ANALYSIS LAB RECEIVED THE SUSPECTED DEVICECOMPONENT AND PERFORMED A FAILURE INVESTIGATION THE REPORTED ISSUE WAS DUPLICATED IN THE LABORATORY SETTING THE ROOT CAUSE OF THE REPORTED ISSUE WAS ISOLATED TO A COLD FIRE INTEGRATED CIRCUIT | Fire | Device problem code
6899898 | 87622147 | N | B RESULTS OF INVESTIGATION THE VYAIRE FAILURE ANALYSIS LAB RECEIVED THE SUSPECTED DEVICECOMPONENT AND PERFORMED A FAILURE INVESTIGATION THE REPORTED ISSUE WAS DUPLICATED IN THE LABORATORY SETTING THE ROOT CAUSE OF THE REPORTED ISSUE WAS ISOLATED TO A COLD FIRE INTEGRATED CIRCUIT | Incompatible Component/ Accessory | Medical Evaluation result code
6891649 | 87627333 | N | B PER THE CUSTOMER THERE WERE NO FAULT INDICATOR LIGHTS ILLUMINATED THE WATER WAS NOT CLOUDY OR DISCOLORED AND THERE ARE NO LEAKS PREVENTATIVE MAINTENANCE IS PERFORMED PER THE OPERATORS MANUAL | Material Discolored | Device problem code
6891649 | 87627333 | N | B PER THE CUSTOMER THERE WERE NO FAULT INDICATOR LIGHTS ILLUMINATED THE WATER WAS NOT CLOUDY OR DISCOLORED AND THERE ARE NO LEAKS PREVENTATIVE MAINTENANCE IS PERFORMED PER THE OPERATORS MANUAL | Leak / Splash | Device problem code
6891649 | 87627333 | N | B PER THE CUSTOMER THERE WERE NO FAULT INDICATOR LIGHTS ILLUMINATED THE WATER WAS NOT CLOUDY OR DISCOLORED AND THERE ARE NO LEAKS PREVENTATIVE MAINTENANCE IS PERFORMED PER THE OPERATORS MANUAL | Leakage/Seal | Medical Evaluation result code
6896638 | 87633144 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Inadequacy of Device Shape and/or Size | Device problem code
6896638 | 87633144 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Stress Problem Identified | Medical Evaluation result code
6884502 | 87633815 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE DEVICE STUCK ON TISSUE WHEN IT WOULD NOT RELEASE IF YES HOW WAS THE DEVICE REMOVED FROM TISSUE WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED | No Device Problem Found | Medical Evaluation result code
6890886 | 87634099 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6890886 | 87634099 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6899522 | 87634214 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING B | Component Missing | Device problem code
6899522 | 87634214 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6900862 | 87637052 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6900862 | 87637052 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6900864 | 87637827 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH AND ELEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6900864 | 87637827 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH AND ELEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6900861 | 87637934 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6900861 | 87637934 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6881886 | 87638649 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6881886 | 87638649 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6900925 | 87639269 | D | IT WAS REPORTED THAT A NEEDLE STICK OCCURRED WHILE PERFORMING VENIPUNCTURE ON A PATIENT USING A G X IN NEEDLE X IN TUBING BD SAFETYLOK BLOOD COLLECTION AND INFUSION SET WITH LUER ADAPTER THE EMPLOYEE FOLLOWED BLOOD BORNE PATHOGEN BBP PROTOCOL | Venipuncture | Patient problem code
6900925 | 87639269 | D | IT WAS REPORTED THAT A NEEDLE STICK OCCURRED WHILE PERFORMING VENIPUNCTURE ON A PATIENT USING A G X IN NEEDLE X IN TUBING BD SAFETYLOK BLOOD COLLECTION AND INFUSION SET WITH LUER ADAPTER THE EMPLOYEE FOLLOWED BLOOD BORNE PATHOGEN BBP PROTOCOL | Insufficient Information | Device problem code
6901113 | 87643443 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE EQUIPMENT NOTED THAT THE FLOW SENSOR DIAPHRAGM WAS STUCK TO THE TOP OF THE FLOW SENSOR HOUSING GE HEALTHCARE RECOMMENDED TO THE HOSPITAL TO REPLACE THE FLOW SENSOR | Mechanical Jam | Device problem code
6901113 | 87643443 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE EQUIPMENT NOTED THAT THE FLOW SENSOR DIAPHRAGM WAS STUCK TO THE TOP OF THE FLOW SENSOR HOUSING GE HEALTHCARE RECOMMENDED TO THE HOSPITAL TO REPLACE THE FLOW SENSOR | Mechanical Problem Identified | Medical Evaluation result code
6881864 | 87644664 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | Positioning Problem | Device problem code
6881864 | 87644664 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE THE CANNULA WAS INSERTED CORRECTLY INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6900897 | 87652931 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6900897 | 87652931 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6900870 | 87653023 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6900870 | 87653023 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6896656 | 87653124 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF CAUSING THE PATIENT TO BE UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6896656 | 87653124 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF CAUSING THE PATIENT TO BE UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6900898 | 87657751 | N | NO REPORT OF PATIENT INVOLVEMENT A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT THE POPOFF VALVE AND BREATHING SYSTEM WERE CLEANED | Detachment of Device or device Component | Device problem code
6903277 | 87721784 | D | THE PUBLISHED RESULT OF THE ARTICLE OVERSIRIX A NEW METHOD FOR SIZING AORTIC ENDOGRAFTS IN COMBINATION WITH THE CHIMNEY GRAFTS EARLY EXPERIENCE WITH AORTIC ARCH DISEASE PUBLISHED BY STEFANO FAZZINI ET AL WITHIN THE ANNALS OF VASCULAR SURGERY INDICATED THE FOLLOWING THE STUDY AIMS TO EVALUATE A NEW TAILORED PLANNING NAMED OVERSIRIX BASED ON OSIRIX IMAGING SOFTWARE TO CHOOSE THE CORRECT MAIN GRAFT OVERSIZING IN ORDER TO MINIMIZE TYPE I ENDOLEAK INCIDENCE FROM TO PATIENTS WERE TREATED WITH PARALLEL GRAFTS FOR AORTIC ARCH DISEASES AT THE INSTITUTION THE STUDY INCLUDED PATIENTS WITH SINGLE STENT AND ANTEGRADE FLOW CONFIGURATION THEY WERE DIVIDED IN TWO GROUPS PRE AND POST OVERSIRIX OVERSIRIX WAS CARRIED OUT IN THE RETROSPECTIVE GROUP PRE OVERSIRIX AND IT WAS USED TO PLAN THE ENDOVASCULAR PROCEDURE IN THE PROSPECTIVE GROUP POST OVERSIRIX ON B PATIENT OUT OF THIS STUDY PRESENTED WITH AN THORACIC AORTIC ANEURYSM THAT WAS TREATED WITH GORE TAG THORACIC ENDOPROSTHESIS AND GORE VIABAHN ENDOPROSTHESIS IMPLANTED WITHIN THE LEFT SUBCLAVIAN ARTERY IN A CHIMNEY PROCEDURE DUE TO AN ENDOLEAK TYPE ALONG WITH A REPORTED BIRD BEAK A REINTERVENTION WAS NECESSARY TO PREVENT SAC ENLARGEMENT THE REINTERVENTION WAS UNSPECIFIED ADDITIONALLY IT WAS STATED THAT THE RADIUS OF THE INNER CURVATURE CONTRIBUTED TO THE BIRD BEAK THE PATIENT TOLERATED THE PROCEDURE ADDITIONALLY THE ARTICLE STATED THAT THE REPORTED ENDOLEAKS OCCURRED IN PATIENTS WITH A BIG GUTTERS AREA CAUSED BY A TOO LOW OVERSIZING RATE | Insufficient Information | Device problem code
6903277 | 87721784 | D | THE PUBLISHED RESULT OF THE ARTICLE OVERSIRIX A NEW METHOD FOR SIZING AORTIC ENDOGRAFTS IN COMBINATION WITH THE CHIMNEY GRAFTS EARLY EXPERIENCE WITH AORTIC ARCH DISEASE PUBLISHED BY STEFANO FAZZINI ET AL WITHIN THE ANNALS OF VASCULAR SURGERY INDICATED THE FOLLOWING THE STUDY AIMS TO EVALUATE A NEW TAILORED PLANNING NAMED OVERSIRIX BASED ON OSIRIX IMAGING SOFTWARE TO CHOOSE THE CORRECT MAIN GRAFT OVERSIZING IN ORDER TO MINIMIZE TYPE I ENDOLEAK INCIDENCE FROM TO PATIENTS WERE TREATED WITH PARALLEL GRAFTS FOR AORTIC ARCH DISEASES AT THE INSTITUTION THE STUDY INCLUDED PATIENTS WITH SINGLE STENT AND ANTEGRADE FLOW CONFIGURATION THEY WERE DIVIDED IN TWO GROUPS PRE AND POST OVERSIRIX OVERSIRIX WAS CARRIED OUT IN THE RETROSPECTIVE GROUP PRE OVERSIRIX AND IT WAS USED TO PLAN THE ENDOVASCULAR PROCEDURE IN THE PROSPECTIVE GROUP POST OVERSIRIX ON B PATIENT OUT OF THIS STUDY PRESENTED WITH AN THORACIC AORTIC ANEURYSM THAT WAS TREATED WITH GORE TAG THORACIC ENDOPROSTHESIS AND GORE VIABAHN ENDOPROSTHESIS IMPLANTED WITHIN THE LEFT SUBCLAVIAN ARTERY IN A CHIMNEY PROCEDURE DUE TO AN ENDOLEAK TYPE ALONG WITH A REPORTED BIRD BEAK A REINTERVENTION WAS NECESSARY TO PREVENT SAC ENLARGEMENT THE REINTERVENTION WAS UNSPECIFIED ADDITIONALLY IT WAS STATED THAT THE RADIUS OF THE INNER CURVATURE CONTRIBUTED TO THE BIRD BEAK THE PATIENT TOLERATED THE PROCEDURE ADDITIONALLY THE ARTICLE STATED THAT THE REPORTED ENDOLEAKS OCCURRED IN PATIENTS WITH A BIG GUTTERS AREA CAUSED BY A TOO LOW OVERSIZING RATE | Aneurysm | Patient problem code
6903277 | 87721784 | D | THE PUBLISHED RESULT OF THE ARTICLE OVERSIRIX A NEW METHOD FOR SIZING AORTIC ENDOGRAFTS IN COMBINATION WITH THE CHIMNEY GRAFTS EARLY EXPERIENCE WITH AORTIC ARCH DISEASE PUBLISHED BY STEFANO FAZZINI ET AL WITHIN THE ANNALS OF VASCULAR SURGERY INDICATED THE FOLLOWING THE STUDY AIMS TO EVALUATE A NEW TAILORED PLANNING NAMED OVERSIRIX BASED ON OSIRIX IMAGING SOFTWARE TO CHOOSE THE CORRECT MAIN GRAFT OVERSIZING IN ORDER TO MINIMIZE TYPE I ENDOLEAK INCIDENCE FROM TO PATIENTS WERE TREATED WITH PARALLEL GRAFTS FOR AORTIC ARCH DISEASES AT THE INSTITUTION THE STUDY INCLUDED PATIENTS WITH SINGLE STENT AND ANTEGRADE FLOW CONFIGURATION THEY WERE DIVIDED IN TWO GROUPS PRE AND POST OVERSIRIX OVERSIRIX WAS CARRIED OUT IN THE RETROSPECTIVE GROUP PRE OVERSIRIX AND IT WAS USED TO PLAN THE ENDOVASCULAR PROCEDURE IN THE PROSPECTIVE GROUP POST OVERSIRIX ON B PATIENT OUT OF THIS STUDY PRESENTED WITH AN THORACIC AORTIC ANEURYSM THAT WAS TREATED WITH GORE TAG THORACIC ENDOPROSTHESIS AND GORE VIABAHN ENDOPROSTHESIS IMPLANTED WITHIN THE LEFT SUBCLAVIAN ARTERY IN A CHIMNEY PROCEDURE DUE TO AN ENDOLEAK TYPE ALONG WITH A REPORTED BIRD BEAK A REINTERVENTION WAS NECESSARY TO PREVENT SAC ENLARGEMENT THE REINTERVENTION WAS UNSPECIFIED ADDITIONALLY IT WAS STATED THAT THE RADIUS OF THE INNER CURVATURE CONTRIBUTED TO THE BIRD BEAK THE PATIENT TOLERATED THE PROCEDURE ADDITIONALLY THE ARTICLE STATED THAT THE REPORTED ENDOLEAKS OCCURRED IN PATIENTS WITH A BIG GUTTERS AREA CAUSED BY A TOO LOW OVERSIZING RATE | Hemorrhage | Patient problem code
6901782 | 87735237 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED CONSTANT UPSTREAM OCCLUSION ALARMS DURING TESTING THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | Device Alarm System | Device Problem Code
6901782 | 87735237 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED CONSTANT UPSTREAM OCCLUSION ALARMS DURING TESTING THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Patient Involvement | Patient problem code
6877175 | 87756919 | D | THE TIP OF PEEL AWAY SHEATH WAS FOUND TO BE TORN WHEN THE PHYSICIAN TRIED TO INSERT IT NO ADVERSE EFFECT WAS REPORTED | No Consequences Or Impact To Patient | Patient problem code
6877175 | 87756919 | D | THE TIP OF PEEL AWAY SHEATH WAS FOUND TO BE TORN WHEN THE PHYSICIAN TRIED TO INSERT IT NO ADVERSE EFFECT WAS REPORTED | Insufficient Information | Device problem code
6904540 | 87779976 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6904501 | 87786433 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6904501 | 87786433 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6904504 | 87788319 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6904504 | 87788319 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6904502 | 87788964 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6904502 | 87788964 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6886948 | 87794464 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6886948 | 87794464 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904477 | 87794504 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904477 | 87794504 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904492 | 87794597 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904492 | 87794597 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904476 | 87794644 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904476 | 87794644 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6886917 | 87797645 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6886917 | 87797645 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6905394 | 87804481 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | No Consequences Or Impact To Patient | Patient problem code
6905394 | 87804481 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | Insufficient Information | Device problem code
6886923 | 87805624 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Component or Accessory Incompatibility | Device problem code
6886923 | 87805624 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Incompatible Component/ Accessory | Medical Evaluation result code
6905504 | 87808395 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | Insufficient Information | Device problem code
6905504 | 87808395 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | No Consequences Or Impact To Patient | Patient problem code
6882552 | 87808921 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER LEG FOR BETWEEN AND HOURS SHE STATED THAT THE PINK SLIDE WAS OFF TO THE LEFT OF THE WHITE CIRCLE TREATMENT INCLUDED A BOLUS FROM A NEW POD | No Known Impact Or Consequence To Patient | Patient Problem Code
6882552 | 87808921 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER LEG FOR BETWEEN AND HOURS SHE STATED THAT THE PINK SLIDE WAS OFF TO THE LEFT OF THE WHITE CIRCLE TREATMENT INCLUDED A BOLUS FROM A NEW POD | Insufficient Information | Device problem code
6905503 | 87811206 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | Insufficient Information | Device problem code
6905503 | 87811206 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | No Consequences Or Impact To Patient | Patient problem code
6905550 | 87814014 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | No Consequences Or Impact To Patient | Patient problem code
6905550 | 87814014 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | Insufficient Information | Device problem code
6905647 | 87814075 | N | THE CUSTOMER WAS ABLE TO RESOLVE THE REPORTED ERROR MESSAGE USER ADVISORY UA NOT AT HOME POSITION AFTER POWERONRESTART BY ROTATING THE SHAFT TO HOME POSITION WITH THE INSTRUCTION PROVIDED BY THE ZOLL TECHNICAL SUPPORT IN CANADA THE PLATFORM WILL NOT BE RETURNING TO ZOLL B FOR EVALUATION HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINTS FOR AUTOPULSE SN B | Device Displays Incorrect Message | Device problem code
6905647 | 87814075 | N | THE CUSTOMER WAS ABLE TO RESOLVE THE REPORTED ERROR MESSAGE USER ADVISORY UA NOT AT HOME POSITION AFTER POWERONRESTART BY ROTATING THE SHAFT TO HOME POSITION WITH THE INSTRUCTION PROVIDED BY THE ZOLL TECHNICAL SUPPORT IN CANADA THE PLATFORM WILL NOT BE RETURNING TO ZOLL B FOR EVALUATION HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINTS FOR AUTOPULSE SN B | No Device Problem Found | Medical Evaluation result code
6905512 | 87842269 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | No Consequences Or Impact To Patient | Patient problem code
6905512 | 87842269 | D | RELATED MANUFACTURER REPORT NUMBER FOLLOWING A SYSTEM IMPLANT PROCEDURE DIAGNOSTIC IMAGING REVEALED A LEFT PNEUMOTHORAX FOR WHICH NO TREATMENT WAS REQUIRED THE PATIENT WAS ASYMPTOMATIC AND DISCHARGED THE FOLLOWING DAY | Insufficient Information | Device problem code
6906192 | 87842327 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE PATIENT WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER REPORTED THAT HE ENDED UP IN HOSPITAL DUE TO HIGH BLOOD GLUCOSE AND THE SUGAR WAS IN S MGDL RANGE THE CUSTOMER WAS BOLUSING AND IT DOES NOT SEEM TO BE WORKING NOW AND THOUGHT THAT IT WAS THE SITE AND CHANGED IT THREE TIMES HE BOLUSED FOR HOURS AND SUGARS DID NOT GO BELOW S MGDL AND THE CUSTOMER WAS NOT EATING THE PATIENTS CURRENT BLOOD GLUCOSE WAS MGDL HE IS ON INSULIN AND WAS MGDL TODAY WHILE ON MANUAL INJECTIONS THE CUSTOMER REPORTED THAT ON WEDNESDAY EVENING SUGARS WAS HIGH AND GAVE HIMSELF A BOLUS BEFORE BED AND DID NOT WANT ANYTHING THE BLOOD GLUCOSE WAS MGDL BEFORE BED AND WOKE UP AT MGDL THE CUSTOMER WAS STILL IN S MGDL AND HOURS LATER DOCTOR GOT INVOLVED AND TOLD HIM TO GO TO THE NEAREST EMERGENCY ROOM AND BEEN ON THE OMNI POD FOR FOUR YEARS THEY DID A DISPLACEMENT TEST AND IT PASSED DID A SELFTEST AND IT PASSED THEY DID A BOLUS AND IT DELIVERED SUCCESSFULLY AND IT WAS RECORDED IN THE HISTORY HE WAS CURRENTLY IN THE HOSPITAL AND IS NOT WEARING THE INSULIN PUMP THE CUSTOMER WAS NOT HAVING TUBING CLAMP THE CUSTOMER DECLINED TO CONTINUE TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6906192 | 87842327 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THE PATIENT WAS HOSPITALIZED FOR HIGH BLOOD GLUCOSE THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER REPORTED THAT HE ENDED UP IN HOSPITAL DUE TO HIGH BLOOD GLUCOSE AND THE SUGAR WAS IN S MGDL RANGE THE CUSTOMER WAS BOLUSING AND IT DOES NOT SEEM TO BE WORKING NOW AND THOUGHT THAT IT WAS THE SITE AND CHANGED IT THREE TIMES HE BOLUSED FOR HOURS AND SUGARS DID NOT GO BELOW S MGDL AND THE CUSTOMER WAS NOT EATING THE PATIENTS CURRENT BLOOD GLUCOSE WAS MGDL HE IS ON INSULIN AND WAS MGDL TODAY WHILE ON MANUAL INJECTIONS THE CUSTOMER REPORTED THAT ON WEDNESDAY EVENING SUGARS WAS HIGH AND GAVE HIMSELF A BOLUS BEFORE BED AND DID NOT WANT ANYTHING THE BLOOD GLUCOSE WAS MGDL BEFORE BED AND WOKE UP AT MGDL THE CUSTOMER WAS STILL IN S MGDL AND HOURS LATER DOCTOR GOT INVOLVED AND TOLD HIM TO GO TO THE NEAREST EMERGENCY ROOM AND BEEN ON THE OMNI POD FOR FOUR YEARS THEY DID A DISPLACEMENT TEST AND IT PASSED DID A SELFTEST AND IT PASSED THEY DID A BOLUS AND IT DELIVERED SUCCESSFULLY AND IT WAS RECORDED IN THE HISTORY HE WAS CURRENTLY IN THE HOSPITAL AND IS NOT WEARING THE INSULIN PUMP THE CUSTOMER WAS NOT HAVING TUBING CLAMP THE CUSTOMER DECLINED TO CONTINUE TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6905725 | 87843667 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED SEE A REPORT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT WERE THE INDICATIONS FOR SURGERY COULD YOU PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE HEALTH OF THE TISSUE AT THE ANASTOMOSIS SITE WHAT WAS THE ETIOLOGY OR ORIGIN OF THE INFECTION DURING THE SURGICAL PROCEDURE DURING THE REOPERATION WAS THE EXACT SITE OF THE LEAK IDENTIFIED IF SO WHERE WAS THE LEAK WAS IT ABLE TO BE DETERMINED IF THE LEAK WAS FROM THE TXB OR THE TLC ARE THERE ANY PLANS TO REANASTOMOSE WHAT IS THE CURRENT PATIENT STATUS | Leak / Splash | Device problem code
6905725 | 87843667 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED SEE A REPORT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT WERE THE INDICATIONS FOR SURGERY COULD YOU PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE HEALTH OF THE TISSUE AT THE ANASTOMOSIS SITE WHAT WAS THE ETIOLOGY OR ORIGIN OF THE INFECTION DURING THE SURGICAL PROCEDURE DURING THE REOPERATION WAS THE EXACT SITE OF THE LEAK IDENTIFIED IF SO WHERE WAS THE LEAK WAS IT ABLE TO BE DETERMINED IF THE LEAK WAS FROM THE TXB OR THE TLC ARE THERE ANY PLANS TO REANASTOMOSE WHAT IS THE CURRENT PATIENT STATUS | Leakage/Seal | Medical Evaluation result code
6906319 | 87848523 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING NODULES AND LUMP ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF EDEMA AND SKIN IRRITATION PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES AFTER INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING HEALTH CARE PROFESSIONAL INSTRUCTIONS THE INJECTION TECHNIQUE FOR JUVÉDERM VOLUMA XC WITH REGARD TO THE ANGLE AND ORIENTATION OF THE BEVEL THE DEPTH SUBCUTANEOUS ANDOR SUBMUSCULARSUPRAPERIOSTEAL OF INJECTION AND THE QUANTITY ADMINISTERED MAY VARY DEPENDING ON THE AREA BEING TREATED INJECTION OF JUVÉDERM VOLUMA XC TOO SUPERFICIALLY INTRADERMALLY OR IN LARGE VOLUMES OVER A SMALL AREA MAY RESULT IN VISIBLE AND PERSISTENT LUMPS ANDOR DISCOLORATION JUVÉDERM VOLUMA XC SHOULD BE DISTRIBUTED IN SMALL ALIQUOTS SMALL BOLUSES OF ML TO ML OVER A LARGE AREA TO REDUCE THE RISK OF PERSISTENT LUMPINESS | No Findings Available | Medical Evaluation result code
6906519 | 87849052 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON B MDR | Red Eye(s) | Patient problem code
6906519 | 87849052 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON B MDR | Reaction | Patient problem code
6906519 | 87849052 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON B MDR | Headache | Patient problem code
6906519 | 87849052 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON B MDR | Insufficient Information | Device problem code
6906519 | 87849052 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON B MDR | Injury | Patient problem code
6906561 | 87849797 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW B THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STINGING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON PC MDR | Injury | Patient problem code
6906561 | 87849797 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW B THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STINGING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON PC MDR | Reaction | Patient problem code
6906561 | 87849797 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW B THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STINGING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON PC MDR | Insufficient Information | Device problem code
6906561 | 87849797 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW B THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STINGING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON PC MDR | Red Eye(s) | Patient problem code
6906561 | 87849797 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW B THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION DRY THROAT WHICH LASTED HOURS SKIN REACTION STINGING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE THE ORIGINAL SMOKESMELLCOMPLAINT INVOLVING THE STERRAD S UNIT WAS REPORTED AS A MALFUNCTION ON PC MDR | Headache | Patient problem code
6900789 | 87852540 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT HE CONFIRMED THE BLOOD DETECT ALARM ACTIVATED TWICE IN THE ERROR LOGS UPON INSPECTION OF THE IABP BOTH EXTERNALLY AND INTERNALLY THE FSE NOTICED A SUBSTANTIAL AMOUNT OF MOISTURE INSIDE THE BLOOD DETECT GLASS AND TUBING WHICH WAS SHOWN TO THE BIOMED THE FSE CONDUCTED THE CONDENSATION REMOVAL PROCEDURE TO REMOVE AS MUCH MOISTURE IN THE LINETUBING AS POSSIBLE THE FSE VERIFIED THE CONDENSATION REMOVAL MODULE CRM OPERATION TO BE FULLY FUNCTIONAL THE FSE CHECKED ALL BLOOD DETECT VOLTAGES TO BE WITHIN SPECS AND ALSO CLEANED THE BLOOD DETECT SENSOR TO MAKE SURE THERE WAS NOTHING THAT COULD INCREASE THE PROBABILITY OF A FALSE ALARM THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | False Alarm | Device problem code
6900789 | 87852540 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND REPORTED THAT HE CONFIRMED THE BLOOD DETECT ALARM ACTIVATED TWICE IN THE ERROR LOGS UPON INSPECTION OF THE IABP BOTH EXTERNALLY AND INTERNALLY THE FSE NOTICED A SUBSTANTIAL AMOUNT OF MOISTURE INSIDE THE BLOOD DETECT GLASS AND TUBING WHICH WAS SHOWN TO THE BIOMED THE FSE CONDUCTED THE CONDENSATION REMOVAL PROCEDURE TO REMOVE AS MUCH MOISTURE IN THE LINETUBING AS POSSIBLE THE FSE VERIFIED THE CONDENSATION REMOVAL MODULE CRM OPERATION TO BE FULLY FUNCTIONAL THE FSE CHECKED ALL BLOOD DETECT VOLTAGES TO BE WITHIN SPECS AND ALSO CLEANED THE BLOOD DETECT SENSOR TO MAKE SURE THERE WAS NOTHING THAT COULD INCREASE THE PROBABILITY OF A FALSE ALARM THE IABP PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE | No Device Problem Found | Medical Evaluation result code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | No Audible Alarm | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Loss of Power | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Component or Accessory Incompatibility | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Disconnection | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Electrical Power Problem | Device problem code
6900840 | 87853854 | N | B THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARM NOTED HOWEVER REPLACE BATTERY ALARM AND POWER ERROR CONFIRMED IN THE LONG TRACE DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED REPLACE BATTERY AND THEN ALARMED PUMP ERROR WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J PCB AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Power Source Problem Identified | Medical Evaluation result code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Injury | Patient problem code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Headache | Patient problem code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Red Eye(s) | Patient problem code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Reaction | Patient problem code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Insufficient Information | Device problem code
6906590 | 87854277 | D | A CUSTOMER COMPLAINED OF SMOKESMELL EMITTING FROM THE STERRAD S AND SEVERAL HEALTH CARE WORKERS HCW EXPERIENCED VARIOUS SKIN AND OTHER REACTIONS THIS REPORT IS FOR HCW THE HCW IS A FEMALE WITH INITIALS B AND IT WAS NOTED THE HCW WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT IT WAS ALSO NOTED THE HCW HAD THE FOLLOWING SYMPTOMS RESPIRATORY REACTION STINGING AND BURNING WHICH LASTED HOURS DRY THROAT WHICH LASTED HOURS SKIN REACTION STING ON FACE ARMS AND NECK WHICH LASTED HOURS EYE REACTION RED EYES WHICH LASTED HOURS OTHER REACTION HEADACHE WHICH LASTED HOURS IT WAS NOTED THE HCW RECEIVED MEDICAL ATTENTION AND WAS TREATED WITH INTRAVENOUS CHLORPHENAMINE FOR THE RESPIRATORY AND SKIN REACTION AND PARACETAMOL WHICH IS ALSO KNOWN AS ACETAMINOPHEN AND IS AN OVERTHECOUNTER OTC ORAL MEDICATION TO TREAT HEADACHES IT WAS ALSO NOTED THE TREATMENT RECOMMENDED FOR USE AT HOME WAS CHLORPHENAMINE AND PARACETAMOL WHICH ARE BOTH OTC MEDICATIONS THE HCW HAS NO KNOWN ALLERGIES AND LASTLY THE STATUS OF THE HCW WAS NOTED SHE IS GOOD BASED ON THE NEW INFORMATION RECEIVED THIS EVENT IS NOW DEEMED REPORTABLE AS A SERIOUS INJURY ALTHOUGH THE SYMPTOMS RESOLVED AND THE HCW WAS REPORTED GOOD THE HCW HAD A RESPIRATORY AND SKIN REACTION RECEIVED MEDICAL ATTENTION AND WAS GIVEN AN IV ANTIHISTAMINE | Burning Sensation | Patient problem code
6884701 | 87856428 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Defective Alarm | Device Problem Code
6884701 | 87856428 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | No Device Problem Found | Medical Evaluation result code
6900848 | 87856648 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND HE DID NOT FOUND BLOOD IN THE SYSTEM FSE CALIBRATED BLOOD BACK ALARM CIRCUIT AS PART OF PREVENTIVE MAINTENANCE AND COMPLETE PM WITH FULL CALIBRATION UNIT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | Calibration Problem | Device problem code
6900848 | 87856648 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP AND HE DID NOT FOUND BLOOD IN THE SYSTEM FSE CALIBRATED BLOOD BACK ALARM CIRCUIT AS PART OF PREVENTIVE MAINTENANCE AND COMPLETE PM WITH FULL CALIBRATION UNIT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS RETURNED TO CUSTOMER AND CLEARED FOR CLINICAL USE | No Device Problem Found | Medical Evaluation result code
6884699 | 87856918 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Defective Alarm | Device Problem Code
6884699 | 87856918 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Mechanical Problem Identified | Medical Evaluation result code
6884703 | 87857125 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Defective Alarm | Device Problem Code
6884703 | 87857125 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED A FAULT ALARM THE FREEDOM DRIVER CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Mechanical Problem Identified | Medical Evaluation result code
6901073 | 87858724 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT CAN YOU PLEASE PROVIDE FURTHER DETAIL OF THE EVENT WHAT DEGREE OF HEMORRHOIDS DID THE PATIENT HAVE WHAT CONFIRMATION WAS RECEIVED THAT THE DEVICE WAS FULLY FIRED WHERE WITHIN THE GAP SETTING SCALE WAS THE ORANGE INDICATOR PRIOR TO FIRING DID THE DEVICE CUT IF THE DEVICE CUT WAS THE CUT LINE FULLY CIRCUMFERENTIAL AROUND THE TARGET TISSUE DID THE DEVICE STAPLE WAS THE STAPLE LINE COMPLETE WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE APPEARANCE OF THE DEPLOYED STAPLES BFORMATION IRREGULAR STRAIGHT LEGS WAS THE SAFETY RESET BEFORE REMOVAL ATTEMPTED HOW WAS IT CONFIRMED THAT THE ANVIL WAS FREE FROM THE TISSUE PRIOR TO REMOVING AFTER FIRING WAS THE ADJUSTING KNOB TURNED TO REVOLUTIONS WAS THE DEVICE ROTATED DEGREES IN BOTH DIRECTIONS TO ENSURE THE ANVIL WAS FREE FROM THE TISSUE HOW WAS THE PROCEDURE COMPLETED WAS THERE ANY CHANGE IN THE PROCEDURE OR PATIENT CARE AS A RESULT OF THE EVENT | No Device Problem Found | Medical Evaluation result code
6893788 | 87860277 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893788 | 87860277 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6904987 | 87861686 | N | DYSPHAGIA IS A KNOWN ANTICIPATED ADVERSE EVENT AFTER ENDOSCOPIC ERADICATION THERAPY AND MAY OCCUR AFTER RADIOFREQUENCY ABLATION OR CRYOTHERAPY PROCEDURES DURING THE EGD TO EXAMINE THE SYMPTOMS OF DYSPHAGIA A NARROWING OF THE ESOPHAGUS WAS NOTED HOWEVER THE ENDOSCOPE WAS PASSED THROUGH THE NARROWED AREA AND FURTHER DILATION WAS NOT REQUIRED | Material Deformation | Device problem code
6904987 | 87861686 | N | DYSPHAGIA IS A KNOWN ANTICIPATED ADVERSE EVENT AFTER ENDOSCOPIC ERADICATION THERAPY AND MAY OCCUR AFTER RADIOFREQUENCY ABLATION OR CRYOTHERAPY PROCEDURES DURING THE EGD TO EXAMINE THE SYMPTOMS OF DYSPHAGIA A NARROWING OF THE ESOPHAGUS WAS NOTED HOWEVER THE ENDOSCOPE WAS PASSED THROUGH THE NARROWED AREA AND FURTHER DILATION WAS NOT REQUIRED | Stress Problem Identified | Medical Evaluation result code
6901640 | 87864851 | N | THE EMPLOYEE SUBJECT OF THE REPORTED EVENT SOUGHT AND RECEIVED MEDICAL TREATMENT AT THE USER FACILITY EMERGENCY ROOM AND RETURNED TO WORK THE USER FACILITY STATED THAT THE EMPLOYEE SUBJECT OF THE REPORTED EVENT WAS UNAWARE OF THE PROPER PRACTICE TO SAFELY CLEAR A JAM ON THE SCS AUTOMATION SYSTEM THE SCS AUTOMATION OPERATORS MANUAL STATES ON PAGE IN CASE OF A JAM INVOLVING CONVEYORS PRESS PUSH BUTTON TO DEPRESSURIZE CYLINDER PNEUMATIC SYSTEM BEFORE ATTEMPTING TO REMOVE JAMMED RACK A STERIS SERVICE TECHNICIAN ARRIVED ONSITE INSPECTED THE SCS AUTOMATION SYSTEM AND CONFIRMED IT TO BE OPERATING ACCORDING TO SPECIFICATION A STERIS ACCOUNT MANAGER PERFORMED INSERVICE TRAINING ON WITH THE USER FACILITY EMPLOYEE SUBJECT OF THE REPORTED EVENT ON THE PROPER USAGE OF THE SCS AUTOMATION SYSTEM INCLUDING THE APPROPRIATE WAY TO SAFELY CLEAR A JAM NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Mechanical Jam | Device problem code
6901640 | 87864851 | N | THE EMPLOYEE SUBJECT OF THE REPORTED EVENT SOUGHT AND RECEIVED MEDICAL TREATMENT AT THE USER FACILITY EMERGENCY ROOM AND RETURNED TO WORK THE USER FACILITY STATED THAT THE EMPLOYEE SUBJECT OF THE REPORTED EVENT WAS UNAWARE OF THE PROPER PRACTICE TO SAFELY CLEAR A JAM ON THE SCS AUTOMATION SYSTEM THE SCS AUTOMATION OPERATORS MANUAL STATES ON PAGE IN CASE OF A JAM INVOLVING CONVEYORS PRESS PUSH BUTTON TO DEPRESSURIZE CYLINDER PNEUMATIC SYSTEM BEFORE ATTEMPTING TO REMOVE JAMMED RACK A STERIS SERVICE TECHNICIAN ARRIVED ONSITE INSPECTED THE SCS AUTOMATION SYSTEM AND CONFIRMED IT TO BE OPERATING ACCORDING TO SPECIFICATION A STERIS ACCOUNT MANAGER PERFORMED INSERVICE TRAINING ON WITH THE USER FACILITY EMPLOYEE SUBJECT OF THE REPORTED EVENT ON THE PROPER USAGE OF THE SCS AUTOMATION SYSTEM INCLUDING THE APPROPRIATE WAY TO SAFELY CLEAR A JAM NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Mechanical Problem Identified | Medical Evaluation result code
6878783 | 87866314 | D | IT WAS REPORTED THAT THE BATTERY DISCHARGED FASTER THAN EXPECTED THE BATTERY WAS EXCHANGED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Premature Discharge of Battery | Device problem code
6878783 | 87866314 | D | IT WAS REPORTED THAT THE BATTERY DISCHARGED FASTER THAN EXPECTED THE BATTERY WAS EXCHANGED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Known Impact Or Consequence To Patient | Patient problem code
6882052 | 87868869 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE MECHANISM HAD DEPLOYED PROPERLY HOWEVER PATIENT FELT AS IF THE CANNULA HAD NOT PENETRATED THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6882052 | 87868869 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE MECHANISM HAD DEPLOYED PROPERLY HOWEVER PATIENT FELT AS IF THE CANNULA HAD NOT PENETRATED THE INFUSION SITE | Insufficient Information | Device problem code
6894492 | 87871482 | D | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR NONMALIGNANT PAIN AND OTHER CHRONICINTRACT PAIN TRUNKLIMBS IT WAS REPORTED THAT PATIENT TURNED OFF THEIR DEVICE ON THE DAY OF THE REPORT BECAUSE THEY FELT PAIN AND DISCOMFORT NO FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT ON B THE PATIENT REPORTED THAT SHE STARTED EXPERIENCING PAIN AND DISCOMFORT IN THE FALL OF THE PATIENT ALSO REPORTED THAT SHE STARTED EXPERIENCING THE PAIN AND DISCOMFORT IN B THE PAIN AND DISCOMFORT WAS REPORTEDLY CAUSED BY A DOCTOR THAT WAS GIVING AN INJECTION IN THE BACK AND HIT A NERVE THE DOCTOR DID IT A SECOND TIME AND HIT A NERVE THE PATIENT WAS ON THE TABLE ON HER KNEES WHEN ASKED TO GET UP OFF THE TABLE SHE COULDNT GET UP AND PASSED OUT THEY HAD TO DO CPR SHE WAS IN THE HOSPITAL FOR FIVE DAYS HER LEGS WERE NUMB FOR SIX WEEKS BUT SHE COULD WALK WHEN SHE WAS TAKING A SHOWER THERE WAS A BURNING SENSATION IN HER VAGINA AND RECTUM AND LITTLE BY LITTLE THE BURNING SENSATION GOT WORSE THE DEVICE MADE MORE ELECTRICAL PAINIMPULSES DOWN HER RIGHT LEG AND THEN HAD THE DEVICE TURNED OFF IT IS STILL IMPLANTED NOW BUT IS TURNED OFF AND NOT CAUSING ANY TROUBLE THE PATIENT CURRENTLY TAKES PAIN MEDICATIONS INCLUDING MORPHINE GABAPENTIN AND FENTANYL PATCHES THE PAIN AND DISCOMFORT HAS NOT RESOLVED AND IS NEVER REALLY GONE EXCEPT WHEN LYING IN BED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
6894492 | 87871482 | D | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR NONMALIGNANT PAIN AND OTHER CHRONICINTRACT PAIN TRUNKLIMBS IT WAS REPORTED THAT PATIENT TURNED OFF THEIR DEVICE ON THE DAY OF THE REPORT BECAUSE THEY FELT PAIN AND DISCOMFORT NO FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT ON B THE PATIENT REPORTED THAT SHE STARTED EXPERIENCING PAIN AND DISCOMFORT IN THE FALL OF THE PATIENT ALSO REPORTED THAT SHE STARTED EXPERIENCING THE PAIN AND DISCOMFORT IN B THE PAIN AND DISCOMFORT WAS REPORTEDLY CAUSED BY A DOCTOR THAT WAS GIVING AN INJECTION IN THE BACK AND HIT A NERVE THE DOCTOR DID IT A SECOND TIME AND HIT A NERVE THE PATIENT WAS ON THE TABLE ON HER KNEES WHEN ASKED TO GET UP OFF THE TABLE SHE COULDNT GET UP AND PASSED OUT THEY HAD TO DO CPR SHE WAS IN THE HOSPITAL FOR FIVE DAYS HER LEGS WERE NUMB FOR SIX WEEKS BUT SHE COULD WALK WHEN SHE WAS TAKING A SHOWER THERE WAS A BURNING SENSATION IN HER VAGINA AND RECTUM AND LITTLE BY LITTLE THE BURNING SENSATION GOT WORSE THE DEVICE MADE MORE ELECTRICAL PAINIMPULSES DOWN HER RIGHT LEG AND THEN HAD THE DEVICE TURNED OFF IT IS STILL IMPLANTED NOW BUT IS TURNED OFF AND NOT CAUSING ANY TROUBLE THE PATIENT CURRENTLY TAKES PAIN MEDICATIONS INCLUDING MORPHINE GABAPENTIN AND FENTANYL PATCHES THE PAIN AND DISCOMFORT HAS NOT RESOLVED AND IS NEVER REALLY GONE EXCEPT WHEN LYING IN BED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device problem code
6883212 | 87873165 | N | THE CUSTOMER REPORTED COMPLAINT FOR A LEAK WAS CONFIRMED DURING INITIAL FUNCTIONAL TESTING A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON DURING FUNCTIONAL TESTING THE CATHETER WAS CONNECTED TO A PRESSURIZED INFLATION DEVICE CAPPING THE OUT LUER AND CONNECTING THE IN LUER TO THE PRESSURE INFLATION DEVICE THE CATHETER WAS PRESSURIZED AND IMMEDIATELY A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON A VISUAL INSPECTION OF THE RETURNED CATHETER WAS PERFORMED AND NO ABNORMALITIES WITH THE CATHETER WAS SEEN ALL INFUSION PORTS FLUSHED SUCCESSFULLY DURING MANUFACTURING ALL ICY CATHETERS ARE INSPECTED FOR LEAKS BY SUBJECTING THEM TO PRESSURE TESTING ONLY UNITS THAT PASSED ARE MOVED TO THE NEXT PROCESS PROBABLE CAUSES FOR THE REPORTED COMPLAINT CAN BE A LATENT MATERIAL DEFECT HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO PREVIOUS HISTORY OF RELATED COMPLAINTS FOR ICY CATHETER LOT | Leak / Splash | Device problem code
6883212 | 87873165 | N | THE CUSTOMER REPORTED COMPLAINT FOR A LEAK WAS CONFIRMED DURING INITIAL FUNCTIONAL TESTING A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON DURING FUNCTIONAL TESTING THE CATHETER WAS CONNECTED TO A PRESSURIZED INFLATION DEVICE CAPPING THE OUT LUER AND CONNECTING THE IN LUER TO THE PRESSURE INFLATION DEVICE THE CATHETER WAS PRESSURIZED AND IMMEDIATELY A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON A VISUAL INSPECTION OF THE RETURNED CATHETER WAS PERFORMED AND NO ABNORMALITIES WITH THE CATHETER WAS SEEN ALL INFUSION PORTS FLUSHED SUCCESSFULLY DURING MANUFACTURING ALL ICY CATHETERS ARE INSPECTED FOR LEAKS BY SUBJECTING THEM TO PRESSURE TESTING ONLY UNITS THAT PASSED ARE MOVED TO THE NEXT PROCESS PROBABLE CAUSES FOR THE REPORTED COMPLAINT CAN BE A LATENT MATERIAL DEFECT HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO PREVIOUS HISTORY OF RELATED COMPLAINTS FOR ICY CATHETER LOT | Inflation Problem | Device problem code
6883212 | 87873165 | N | THE CUSTOMER REPORTED COMPLAINT FOR A LEAK WAS CONFIRMED DURING INITIAL FUNCTIONAL TESTING A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON DURING FUNCTIONAL TESTING THE CATHETER WAS CONNECTED TO A PRESSURIZED INFLATION DEVICE CAPPING THE OUT LUER AND CONNECTING THE IN LUER TO THE PRESSURE INFLATION DEVICE THE CATHETER WAS PRESSURIZED AND IMMEDIATELY A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON A VISUAL INSPECTION OF THE RETURNED CATHETER WAS PERFORMED AND NO ABNORMALITIES WITH THE CATHETER WAS SEEN ALL INFUSION PORTS FLUSHED SUCCESSFULLY DURING MANUFACTURING ALL ICY CATHETERS ARE INSPECTED FOR LEAKS BY SUBJECTING THEM TO PRESSURE TESTING ONLY UNITS THAT PASSED ARE MOVED TO THE NEXT PROCESS PROBABLE CAUSES FOR THE REPORTED COMPLAINT CAN BE A LATENT MATERIAL DEFECT HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO PREVIOUS HISTORY OF RELATED COMPLAINTS FOR ICY CATHETER LOT | Defective Device | Device problem code
6883212 | 87873165 | N | THE CUSTOMER REPORTED COMPLAINT FOR A LEAK WAS CONFIRMED DURING INITIAL FUNCTIONAL TESTING A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON DURING FUNCTIONAL TESTING THE CATHETER WAS CONNECTED TO A PRESSURIZED INFLATION DEVICE CAPPING THE OUT LUER AND CONNECTING THE IN LUER TO THE PRESSURE INFLATION DEVICE THE CATHETER WAS PRESSURIZED AND IMMEDIATELY A BOND LEAK WAS NOTED AT THE PROXIMAL END OF THE MEDIAL BALLOON A VISUAL INSPECTION OF THE RETURNED CATHETER WAS PERFORMED AND NO ABNORMALITIES WITH THE CATHETER WAS SEEN ALL INFUSION PORTS FLUSHED SUCCESSFULLY DURING MANUFACTURING ALL ICY CATHETERS ARE INSPECTED FOR LEAKS BY SUBJECTING THEM TO PRESSURE TESTING ONLY UNITS THAT PASSED ARE MOVED TO THE NEXT PROCESS PROBABLE CAUSES FOR THE REPORTED COMPLAINT CAN BE A LATENT MATERIAL DEFECT HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO PREVIOUS HISTORY OF RELATED COMPLAINTS FOR ICY CATHETER LOT | Operational Problem Identified | Medical Evaluation result code
6904491 | 87874075 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6904491 | 87874075 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6886995 | 87874150 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6886995 | 87874150 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6904495 | 87874266 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FOUR DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904495 | 87874266 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FOUR DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6886918 | 87874396 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6886918 | 87874396 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6904486 | 87875255 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904486 | 87875255 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904496 | 87875301 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904496 | 87875301 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904500 | 87875338 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904500 | 87875338 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FOURTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6904499 | 87875363 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904499 | 87875363 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6903086 | 87876635 | N | ZIMMER BIOMET COMPLAINT B COMPLAINT SAMPLE WAS EVALUATED AND THE REPORTED EVENT WAS CONFIRMED THE NAIL DOES NOT STAY IN PLACE AS STATED IN THE COMPLAINT THE SPRING WHICH IS USED TO KEEP THE NAIL WITHIN THE OVERALL ASSEMBLY IS MISSING FROM THE PORTION OF THE ASSEMBLY THAT WAS RETURNED BUT WOULDNT PREVENT THE INSTRUMENT FROM BEING USED CORRECTLY IF ABSENT DHR WAS REVIEWED AND NO DISCREPANCIES WERE FOUND INVESTIGATION RESULTS CONCLUDED THAT THE REPORTED EVENT WAS DUE TO THE MISSING SPRING THERE ARE WARNINGS IN THE PACKAGE INSERT THAT THIS TYPE OF EVENT CAN OCCUR AND RISKS ARE ADDRESSED IN RISK DOCUMENTATION A SUMMARY OF THE INVESTIGATION HAS BEEN SENT TO THE COMPLAINANT IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | Component Missing | Device problem code
6903086 | 87876635 | N | ZIMMER BIOMET COMPLAINT B COMPLAINT SAMPLE WAS EVALUATED AND THE REPORTED EVENT WAS CONFIRMED THE NAIL DOES NOT STAY IN PLACE AS STATED IN THE COMPLAINT THE SPRING WHICH IS USED TO KEEP THE NAIL WITHIN THE OVERALL ASSEMBLY IS MISSING FROM THE PORTION OF THE ASSEMBLY THAT WAS RETURNED BUT WOULDNT PREVENT THE INSTRUMENT FROM BEING USED CORRECTLY IF ABSENT DHR WAS REVIEWED AND NO DISCREPANCIES WERE FOUND INVESTIGATION RESULTS CONCLUDED THAT THE REPORTED EVENT WAS DUE TO THE MISSING SPRING THERE ARE WARNINGS IN THE PACKAGE INSERT THAT THIS TYPE OF EVENT CAN OCCUR AND RISKS ARE ADDRESSED IN RISK DOCUMENTATION A SUMMARY OF THE INVESTIGATION HAS BEEN SENT TO THE COMPLAINANT IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | Missing or Inadequate Safety Measures | Medical Evaluation result code
6880262 | 87879605 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6880262 | 87879605 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6904458 | 87879763 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6904458 | 87879763 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Operational Problem Identified | Medical Evaluation result code
6904475 | 87879802 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904475 | 87879802 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6906863 | 87881147 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6906863 | 87881147 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6906952 | 87886174 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6906952 | 87886174 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6878011 | 87886184 | N | A GALIL SERVICE ENGINEER CONFIRMED THE CUSTOMER COMPLAINT OF THE AUTO VENT VALVE LEAKING VISUAL INSPECTION SHOWED THAT DESICCANT DUST IN THE ARGON GAS DRYER CC HAD CLOGGED THE SCREEN AND OVER PRESSURED THE UNIT OUTLET SCREEN CAUSING DESICCANT TO BE RELEASED INTO THE GAS STREAM CREATING MORE DUST THROUGHOUT THE GAS SYSTEM THE DESICCANT DUST IN THE AREA OF THE AUTO VENT SOLENOID PREVENTED THE SOLENOID NEEDLE FROM SEATING CREATING THE LEAK | Complete Blockage | Device problem code
6878011 | 87886184 | N | A GALIL SERVICE ENGINEER CONFIRMED THE CUSTOMER COMPLAINT OF THE AUTO VENT VALVE LEAKING VISUAL INSPECTION SHOWED THAT DESICCANT DUST IN THE ARGON GAS DRYER CC HAD CLOGGED THE SCREEN AND OVER PRESSURED THE UNIT OUTLET SCREEN CAUSING DESICCANT TO BE RELEASED INTO THE GAS STREAM CREATING MORE DUST THROUGHOUT THE GAS SYSTEM THE DESICCANT DUST IN THE AREA OF THE AUTO VENT SOLENOID PREVENTED THE SOLENOID NEEDLE FROM SEATING CREATING THE LEAK | Environmental Particulates | Device problem code
6878011 | 87886184 | N | A GALIL SERVICE ENGINEER CONFIRMED THE CUSTOMER COMPLAINT OF THE AUTO VENT VALVE LEAKING VISUAL INSPECTION SHOWED THAT DESICCANT DUST IN THE ARGON GAS DRYER CC HAD CLOGGED THE SCREEN AND OVER PRESSURED THE UNIT OUTLET SCREEN CAUSING DESICCANT TO BE RELEASED INTO THE GAS STREAM CREATING MORE DUST THROUGHOUT THE GAS SYSTEM THE DESICCANT DUST IN THE AREA OF THE AUTO VENT SOLENOID PREVENTED THE SOLENOID NEEDLE FROM SEATING CREATING THE LEAK | Leak / Splash | Device problem code
6878011 | 87886184 | N | A GALIL SERVICE ENGINEER CONFIRMED THE CUSTOMER COMPLAINT OF THE AUTO VENT VALVE LEAKING VISUAL INSPECTION SHOWED THAT DESICCANT DUST IN THE ARGON GAS DRYER CC HAD CLOGGED THE SCREEN AND OVER PRESSURED THE UNIT OUTLET SCREEN CAUSING DESICCANT TO BE RELEASED INTO THE GAS STREAM CREATING MORE DUST THROUGHOUT THE GAS SYSTEM THE DESICCANT DUST IN THE AREA OF THE AUTO VENT SOLENOID PREVENTED THE SOLENOID NEEDLE FROM SEATING CREATING THE LEAK | Leakage/Seal | Medical Evaluation result code
6893896 | 87886483 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6893896 | 87886483 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893823 | 87888689 | D | A PATIENTS BLOOD GLUCOSE REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893823 | 87888689 | D | A PATIENTS BLOOD GLUCOSE REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6893861 | 87889541 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893861 | 87889541 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6893897 | 87891035 | D | A PATIENTS BLOOD GLUCOSE REACHED THE LOW S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6893897 | 87891035 | D | A PATIENTS BLOOD GLUCOSE REACHED THE LOW S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Hypoglycemia | Patient Problem Code
6909145 | 87967504 | N | THERE WAS NO INDICATION OF ANY MALFUNCTION OF THE DEVICE | No Device Problem Found | Medical Evaluation result code
6908287 | 87971476 | N | B CATALOG IS UNKNOWN BUT REFERRED TO AS COOK CELECT FILTER NAME AND ADDRESS FOR IMPORTER SITE B SINCE CATALOG IS UNKNOWN K COULD BE EITHER K K OR K B SUMMARY OF INVESTIGATIONAL FINDINGS IT HAS NOT BEEN POSSIBLE TO FURTHER INVESTIGATE OR EVALUATE THIS ALLEGED EVENT BASED ON THE LIMITED INFORMATION PROVIDED TO DATE VIA THE OPERATIVE NOTE STATING EXTREME PAIN THE FILTER FAILING AND UNABLE TO BE REMOVED FILTER RETRIEVAL IS OCCASIONALLY DIFFICULT THIS IS WELLKNOWN FROM PUBLISHED SCIENTIFIC LITERATURE WHERE FILTER RETRIEVALS ARE REFERRED TO AS SIMPLE VS COMPLEX SEVERAL CASE REPORTS PUBLISHED IN SCIENTIFIC LITERATURE DESCRIBE COMPLEX CASES WITH SUCCESSFUL ENDOVASCULAR FILTER RETRIEVALS USING ADDITIONAL ADVANCED TECHNIQUES NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS AND NOTHING INDICATES THAT THE FILTER DID NOT PERFORM AS INTENDED EG INTENDED FOR THE PREVENTION OF RECURRENT PULMONARY EMBOLISM PE VIA PLACEMENT IN THE VENA CAVA COOK WILL REOPEN ITS INVESTIGATION IF FURTHER INFORMATION IS RECEIVED WARRANTING SUPPLEMENTATION IN ACCORDANCE WITH CFR | Difficult to Remove | Device problem code
6908287 | 87971476 | N | B CATALOG IS UNKNOWN BUT REFERRED TO AS COOK CELECT FILTER NAME AND ADDRESS FOR IMPORTER SITE B SINCE CATALOG IS UNKNOWN K COULD BE EITHER K K OR K B SUMMARY OF INVESTIGATIONAL FINDINGS IT HAS NOT BEEN POSSIBLE TO FURTHER INVESTIGATE OR EVALUATE THIS ALLEGED EVENT BASED ON THE LIMITED INFORMATION PROVIDED TO DATE VIA THE OPERATIVE NOTE STATING EXTREME PAIN THE FILTER FAILING AND UNABLE TO BE REMOVED FILTER RETRIEVAL IS OCCASIONALLY DIFFICULT THIS IS WELLKNOWN FROM PUBLISHED SCIENTIFIC LITERATURE WHERE FILTER RETRIEVALS ARE REFERRED TO AS SIMPLE VS COMPLEX SEVERAL CASE REPORTS PUBLISHED IN SCIENTIFIC LITERATURE DESCRIBE COMPLEX CASES WITH SUCCESSFUL ENDOVASCULAR FILTER RETRIEVALS USING ADDITIONAL ADVANCED TECHNIQUES NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS AND NOTHING INDICATES THAT THE FILTER DID NOT PERFORM AS INTENDED EG INTENDED FOR THE PREVENTION OF RECURRENT PULMONARY EMBOLISM PE VIA PLACEMENT IN THE VENA CAVA COOK WILL REOPEN ITS INVESTIGATION IF FURTHER INFORMATION IS RECEIVED WARRANTING SUPPLEMENTATION IN ACCORDANCE WITH CFR | Results Pending Completion of Investigation | Medical Evaluation result code
6909328 | 87971989 | N | THE FIELD REPORT OF GUIDEWIRE COULD NOT BE INSERTED WAS CONFIRMED A TEST GUIDEWIRE WAS INSERTED THROUGH THE CONNECTOR PIN OF THE LEAD THE GUIDEWIRE WAS NOT ABLE TO BE FULLY INSERTED THE GUIDEWIRE WAS BLOCKED AT CM FROM THE CONNECTOR PIN DUE RO INNER COOIL CLOOGED WITH BLOODBODILY FLUIDS THE LEAD WAS OTHERWISE NORMAL | Difficult to Insert | Device problem code
6909328 | 87971989 | N | THE FIELD REPORT OF GUIDEWIRE COULD NOT BE INSERTED WAS CONFIRMED A TEST GUIDEWIRE WAS INSERTED THROUGH THE CONNECTOR PIN OF THE LEAD THE GUIDEWIRE WAS NOT ABLE TO BE FULLY INSERTED THE GUIDEWIRE WAS BLOCKED AT CM FROM THE CONNECTOR PIN DUE RO INNER COOIL CLOOGED WITH BLOODBODILY FLUIDS THE LEAD WAS OTHERWISE NORMAL | Partial Blockage | Device problem code
6909328 | 87971989 | N | THE FIELD REPORT OF GUIDEWIRE COULD NOT BE INSERTED WAS CONFIRMED A TEST GUIDEWIRE WAS INSERTED THROUGH THE CONNECTOR PIN OF THE LEAD THE GUIDEWIRE WAS NOT ABLE TO BE FULLY INSERTED THE GUIDEWIRE WAS BLOCKED AT CM FROM THE CONNECTOR PIN DUE RO INNER COOIL CLOOGED WITH BLOODBODILY FLUIDS THE LEAD WAS OTHERWISE NORMAL | Operational Problem Identified | Medical Evaluation result code
6907716 | 87981843 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJROGENS DISEASE ARE A PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Deflation Problem | Device problem code
6907716 | 87981843 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJROGENS DISEASE ARE A PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Calcified | Device problem code
6907716 | 87981843 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJROGENS DISEASE ARE A PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | No Device Problem Found | Medical Evaluation result code
6907522 | 87982657 | N | STILL IMPLANTED | Insufficient Information | Device Problem Code
6907738 | 87985731 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJOGRENS DISEASE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Deflation Problem | Device problem code
6907738 | 87985731 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJOGRENS DISEASE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Calcified | Device problem code
6907738 | 87985731 | N | THE EVENT OF RHEUMATOID ARTHRITIS AND SJOGRENS DISEASE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | No Device Problem Found | Medical Evaluation result code
6910191 | 87990078 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT STILL HAD THEIR CATH ON AND IS HOPING TO GET IT OFF NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Known Impact Or Consequence To Patient | Patient problem code
6910191 | 87990078 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT STILL HAD THEIR CATH ON AND IS HOPING TO GET IT OFF NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6910032 | 87990570 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WERE THERE ANY ISSUES NOTED IN REGARD TO DEVICE PERFORMANCE IN THE PRIMARY PROCEDURE WHAT HEALTHCARE PROFESSIONAL FIRED THE DEVICE AND WHAT IS HISHER EXPERIENCE DID THE HEALTHCARE PROFESSIONAL CONFIRM A COMPLETE CUTTING WASHER TRANSECTION DID THE HEALTHCARE PROFESSIONAL CONFIRM A COMPLETE DONUT WAS AN INTRA OPERATIVE LEAK TEST PERFORMED IF SO WHAT WERE THE RESULTS WHEN WAS THE BLEEDING IDENTIFIED WERE THERE ANY STAPLES PRESENT IN THE SURGICAL FIELD IF THERE WERE STAPLES PRESENT PLEASE DESCRIBE THE SHAPE OF THE STAPLES ARE THERE ANY PHOTOS AVAILABLE WHAT IS THE CURRENT PATIENT STATUS | Leak / Splash | Device problem code
6910032 | 87990570 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WERE THERE ANY ISSUES NOTED IN REGARD TO DEVICE PERFORMANCE IN THE PRIMARY PROCEDURE WHAT HEALTHCARE PROFESSIONAL FIRED THE DEVICE AND WHAT IS HISHER EXPERIENCE DID THE HEALTHCARE PROFESSIONAL CONFIRM A COMPLETE CUTTING WASHER TRANSECTION DID THE HEALTHCARE PROFESSIONAL CONFIRM A COMPLETE DONUT WAS AN INTRA OPERATIVE LEAK TEST PERFORMED IF SO WHAT WERE THE RESULTS WHEN WAS THE BLEEDING IDENTIFIED WERE THERE ANY STAPLES PRESENT IN THE SURGICAL FIELD IF THERE WERE STAPLES PRESENT PLEASE DESCRIBE THE SHAPE OF THE STAPLES ARE THERE ANY PHOTOS AVAILABLE WHAT IS THE CURRENT PATIENT STATUS | Leakage/Seal | Medical Evaluation result code
6909119 | 87993821 | N | THE PRODUCT WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED WITH BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER THE EXTENDER AND PRESSURE TUBING WERE ALSO RETURNED A VISUAL EXAMINATION OF THE PRODUCT DETECTED THE INNER LUMEN WITHIN THE CATHETER TUBING NEAR THE YFITTING WAS COMPLETELY SEPARATED WITHIN A KINK APPROXIMATELY CM FROM IAB TIP THE INNER LUMEN WAS FOUND TO BE OCCLUDED WITH DRIED BLOOD THE OCCLUSION WAS UNABLE TO BE CLEARED AN UNDERWATER LEAK TEST OF THE BALLOON CATHETER YFITTING EXTRACORPOREAL PRESSURE AND EXTENDER TUBING WAS PERFORMED AND NO OTHER LEAKS WERE DETECTED THE BREAK FOUND IN THE INNER LUMEN APPEARS TO HAVE BEEN THE RESULT OF A SEVERE KINK WHICH EVENTUALLY FAILED AND ALLOWED BLOOD TO LEAK INTO THE MEMBRANE AND CATHETER TUBING CAUSING THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN THE BREAK OCCURRED BUT IT IS POSSIBLY A RESULT OF PATIENT MOVEMENT DURING THE PROCEDURE THE EVALUATION CONFIRMED THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN OR HOW A KINK IN THE INNER LUMEN OCCURS A SEVERE KINK OR CONTINUOUS FLEXING OF A KINK CAN CAUSE THE INNER LUMEN TO BREAK RESULTING IN A LEAK B | Break | Device problem code
6909119 | 87993821 | N | THE PRODUCT WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED WITH BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER THE EXTENDER AND PRESSURE TUBING WERE ALSO RETURNED A VISUAL EXAMINATION OF THE PRODUCT DETECTED THE INNER LUMEN WITHIN THE CATHETER TUBING NEAR THE YFITTING WAS COMPLETELY SEPARATED WITHIN A KINK APPROXIMATELY CM FROM IAB TIP THE INNER LUMEN WAS FOUND TO BE OCCLUDED WITH DRIED BLOOD THE OCCLUSION WAS UNABLE TO BE CLEARED AN UNDERWATER LEAK TEST OF THE BALLOON CATHETER YFITTING EXTRACORPOREAL PRESSURE AND EXTENDER TUBING WAS PERFORMED AND NO OTHER LEAKS WERE DETECTED THE BREAK FOUND IN THE INNER LUMEN APPEARS TO HAVE BEEN THE RESULT OF A SEVERE KINK WHICH EVENTUALLY FAILED AND ALLOWED BLOOD TO LEAK INTO THE MEMBRANE AND CATHETER TUBING CAUSING THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN THE BREAK OCCURRED BUT IT IS POSSIBLY A RESULT OF PATIENT MOVEMENT DURING THE PROCEDURE THE EVALUATION CONFIRMED THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN OR HOW A KINK IN THE INNER LUMEN OCCURS A SEVERE KINK OR CONTINUOUS FLEXING OF A KINK CAN CAUSE THE INNER LUMEN TO BREAK RESULTING IN A LEAK B | Leak / Splash | Device problem code
6909119 | 87993821 | N | THE PRODUCT WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED WITH BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER THE EXTENDER AND PRESSURE TUBING WERE ALSO RETURNED A VISUAL EXAMINATION OF THE PRODUCT DETECTED THE INNER LUMEN WITHIN THE CATHETER TUBING NEAR THE YFITTING WAS COMPLETELY SEPARATED WITHIN A KINK APPROXIMATELY CM FROM IAB TIP THE INNER LUMEN WAS FOUND TO BE OCCLUDED WITH DRIED BLOOD THE OCCLUSION WAS UNABLE TO BE CLEARED AN UNDERWATER LEAK TEST OF THE BALLOON CATHETER YFITTING EXTRACORPOREAL PRESSURE AND EXTENDER TUBING WAS PERFORMED AND NO OTHER LEAKS WERE DETECTED THE BREAK FOUND IN THE INNER LUMEN APPEARS TO HAVE BEEN THE RESULT OF A SEVERE KINK WHICH EVENTUALLY FAILED AND ALLOWED BLOOD TO LEAK INTO THE MEMBRANE AND CATHETER TUBING CAUSING THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN THE BREAK OCCURRED BUT IT IS POSSIBLY A RESULT OF PATIENT MOVEMENT DURING THE PROCEDURE THE EVALUATION CONFIRMED THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN OR HOW A KINK IN THE INNER LUMEN OCCURS A SEVERE KINK OR CONTINUOUS FLEXING OF A KINK CAN CAUSE THE INNER LUMEN TO BREAK RESULTING IN A LEAK B | Material Separation | Device problem code
6909119 | 87993821 | N | THE PRODUCT WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED WITH BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER THE EXTENDER AND PRESSURE TUBING WERE ALSO RETURNED A VISUAL EXAMINATION OF THE PRODUCT DETECTED THE INNER LUMEN WITHIN THE CATHETER TUBING NEAR THE YFITTING WAS COMPLETELY SEPARATED WITHIN A KINK APPROXIMATELY CM FROM IAB TIP THE INNER LUMEN WAS FOUND TO BE OCCLUDED WITH DRIED BLOOD THE OCCLUSION WAS UNABLE TO BE CLEARED AN UNDERWATER LEAK TEST OF THE BALLOON CATHETER YFITTING EXTRACORPOREAL PRESSURE AND EXTENDER TUBING WAS PERFORMED AND NO OTHER LEAKS WERE DETECTED THE BREAK FOUND IN THE INNER LUMEN APPEARS TO HAVE BEEN THE RESULT OF A SEVERE KINK WHICH EVENTUALLY FAILED AND ALLOWED BLOOD TO LEAK INTO THE MEMBRANE AND CATHETER TUBING CAUSING THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN THE BREAK OCCURRED BUT IT IS POSSIBLY A RESULT OF PATIENT MOVEMENT DURING THE PROCEDURE THE EVALUATION CONFIRMED THE REPORTED PROBLEM IT IS DIFFICULT TO DETERMINE WHEN OR HOW A KINK IN THE INNER LUMEN OCCURS A SEVERE KINK OR CONTINUOUS FLEXING OF A KINK CAN CAUSE THE INNER LUMEN TO BREAK RESULTING IN A LEAK B | Fracture Problem | Medical Evaluation result code
6909982 | 87995288 | N | NO INFORMATION ABOUT RETURN OF DEVICE | Insufficient Information | Device problem code
6910009 | 88003938 | D | IT WAS REPORTED THAT PATIENT UNDERWENT TOTAL HIP ARTHROPLASTY ON B SUBSEQUENTLY PATIENT WAS REVISED ON B DUE TO PATIENT FALL AND CUP SPUN OUT THE HEAD AND CUP WERE REPLACED NO FURTHER INFORMATION HAS BEEN PROVIDED | Fall | Patient problem code
6910009 | 88003938 | D | IT WAS REPORTED THAT PATIENT UNDERWENT TOTAL HIP ARTHROPLASTY ON B SUBSEQUENTLY PATIENT WAS REVISED ON B DUE TO PATIENT FALL AND CUP SPUN OUT THE HEAD AND CUP WERE REPLACED NO FURTHER INFORMATION HAS BEEN PROVIDED | Insufficient Information | Device problem code
6879536 | 88013539 | N | THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6879536 | 88013539 | N | THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6908884 | 88015053 | N | IN ADDITION TO THE FLOW RATE ISSUE THE DEVICE HAD A BATTERY OUTSIDE OF WARRANTY ZYNO MEDICAL HAS RECALIBRATED AND TESTED THE DEVICE TO FULL SPECIFICATIONS ON B AND RETURNED THE PUMP TO THE CUSTOMER ON B THE PUMP PERIODIC PREVENTIVE MAINTENANCE PM INTERVAL RECOMMENDED BY ZYNO MEDICAL IN THE INSTRUCTIONS FOR USE IFU IS ONE YEAR OUR COMPANY RECORDS INDICATE THAT THE LAST PM PERFORMED BY ZYNO MEDICAL FOR THIS PUMP WAS ON B | Component or Accessory Incompatibility | Device Problem Code
6908884 | 88015053 | N | IN ADDITION TO THE FLOW RATE ISSUE THE DEVICE HAD A BATTERY OUTSIDE OF WARRANTY ZYNO MEDICAL HAS RECALIBRATED AND TESTED THE DEVICE TO FULL SPECIFICATIONS ON B AND RETURNED THE PUMP TO THE CUSTOMER ON B THE PUMP PERIODIC PREVENTIVE MAINTENANCE PM INTERVAL RECOMMENDED BY ZYNO MEDICAL IN THE INSTRUCTIONS FOR USE IFU IS ONE YEAR OUR COMPANY RECORDS INDICATE THAT THE LAST PM PERFORMED BY ZYNO MEDICAL FOR THIS PUMP WAS ON B | Calibration Problem | Device problem code
6908884 | 88015053 | N | IN ADDITION TO THE FLOW RATE ISSUE THE DEVICE HAD A BATTERY OUTSIDE OF WARRANTY ZYNO MEDICAL HAS RECALIBRATED AND TESTED THE DEVICE TO FULL SPECIFICATIONS ON B AND RETURNED THE PUMP TO THE CUSTOMER ON B THE PUMP PERIODIC PREVENTIVE MAINTENANCE PM INTERVAL RECOMMENDED BY ZYNO MEDICAL IN THE INSTRUCTIONS FOR USE IFU IS ONE YEAR OUR COMPANY RECORDS INDICATE THAT THE LAST PM PERFORMED BY ZYNO MEDICAL FOR THIS PUMP WAS ON B | Failure to Calibrate | Medical Evaluation result code
6907548 | 88016108 | N | NO INFORMATION ABOUT THE RETURN OF DEVICE | Insufficient Information | Device problem code
6907097 | 88018025 | N | NO INFORMATION ABOUT RETURNING DEVICE | Insufficient Information | Device problem code
6906074 | 88018698 | N | B BATCH PAN PA THE ANALYSIS RESULTS SHOWED THAT ONE GSTB RELOAD WAS RECEIVED UNFIRED WITH AN ADDITION PAN ATTACHED THE ADDITIONAL PAN WAS REMOVED FROM THE CARTRIDGE AND LOADED INTO TEST DEVICE THE DEVICE WAS TESTED FOR FUNCTIONALITY IN THE STRAIGHT POSITION AND ACHIEVED ITS COMPLETE STROKE FIRING SEQUENCE WITHOUT ANY DIFFICULTIES THE STAPLE LINE AND CUT LINE WERE COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORM SHAPE WHILE NO CONCLUSION COULD BE REACHED ON WHAT CAUSED THE PAN TO DISLODGE IN ADDITION IT SHOULD BE NOTED THAT AS PART OF OUR QUALITY PROCESS EACH DEVICE IS VISUALLY INSPECTED AND FUNCTIONALLY TESTED DURING MANUFACTURING TO ENSURE THAT THE DEVICE MEETS THE REQUIRED SPECIFICATIONS PRIOR TO SHIPMENT | No Device Problem Found | Medical Evaluation result code
6910739 | 88026404 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE REPORTED COMPLAINT PERFORMING A CALIBRATION OF THE UNIT RESOLVED THE COMPLAINT NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Calibration Problem | Device problem code
6910739 | 88026404 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE REPORTED COMPLAINT PERFORMING A CALIBRATION OF THE UNIT RESOLVED THE COMPLAINT NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Failure to Calibrate | Medical Evaluation result code
6911152 | 88039247 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS HIGH THE PUMP SUPPLIES WERE CHANGED AND A BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL THE CUSTOMER WAS HOSPITALIZED ON B AND DURING THE HOSPITALIZATION THE BG FLUCTUATED EXTREME HIGH AND LOW MGDL THE CAUSE OF THE BG FLUCTUATIONS WAS NOT KNOWN THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP MEDICATION AND INSULIN INJECTIONS DURING A PUMP SYSTEM CHECK INCOMPLETE BOLUS ALERTS AND AN AUTOOFF SAFETY ALARM WERE FOUND TO HAVE OCCURRED THE CUSTOMER ACKNOWLEDGED HAVING NAVIGATED TO THE BOLUS SCREEN AND DID NOT EXIT WITHIN THE ALLOTTED SECONDS THE CUSTOMER ACKNOWLEDGED THE AUTOOFF SAFETY ALARM WAS VALID AS THERE WAS NO PUMP INTERACTION FOR HOURS AND THEN INSULIN WAS NOT RESUMED FOR APPROXIMATELY HOURS THE PUMP TIME WAS FOUND TO BE OFF BY MINUTES THE CUSTOMER DID NOT KNOW THE CAUSE OF THE INCORRECT TIME OR WHEN IT BECAME INCORRECT THE CUSTOMER DID NOT THINK THE TIME CONTRIBUTED TO THE FLUCTUATING BG LEVELS OR HOSPITALIZATION THE CUSTOMER THOUGHT THAT THE HOSPITALIZATION MAY HAVE BEEN TO USER ISSUES REGARDING THE INCOMPLETE BOLUS ALERTS AND AUTOOFF SAFETY ALARM THE CAUSE OF THE LOW BG WAS NOT KNOWN ON B THE CUSTOMERS HEALTHCARE PROVIDER REVERTED THE CUSTOMER TO INSULIN INJECTIONS TO MANAGE BG LEVELS THE CUSTOMER WAS DISCHARGED ON B HOWEVER THE BG LEVELS HAD NOT STABILIZED WHILE USING THE INJECTIONS UPON FOLLOWUP ON B THE CUSTOMER WAS STILL USING INSULIN INJECTIONS AND DECLINED TANDEM TECHNICAL SUPPORTS OFFER TO REACH OUT TO A CLINICAL DIABETES SPECIALIST FOR ASSISTANCE | Hypoglycemia | Patient Problem Code
6911152 | 88039247 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS HIGH THE PUMP SUPPLIES WERE CHANGED AND A BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL THE CUSTOMER WAS HOSPITALIZED ON B AND DURING THE HOSPITALIZATION THE BG FLUCTUATED EXTREME HIGH AND LOW MGDL THE CAUSE OF THE BG FLUCTUATIONS WAS NOT KNOWN THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP MEDICATION AND INSULIN INJECTIONS DURING A PUMP SYSTEM CHECK INCOMPLETE BOLUS ALERTS AND AN AUTOOFF SAFETY ALARM WERE FOUND TO HAVE OCCURRED THE CUSTOMER ACKNOWLEDGED HAVING NAVIGATED TO THE BOLUS SCREEN AND DID NOT EXIT WITHIN THE ALLOTTED SECONDS THE CUSTOMER ACKNOWLEDGED THE AUTOOFF SAFETY ALARM WAS VALID AS THERE WAS NO PUMP INTERACTION FOR HOURS AND THEN INSULIN WAS NOT RESUMED FOR APPROXIMATELY HOURS THE PUMP TIME WAS FOUND TO BE OFF BY MINUTES THE CUSTOMER DID NOT KNOW THE CAUSE OF THE INCORRECT TIME OR WHEN IT BECAME INCORRECT THE CUSTOMER DID NOT THINK THE TIME CONTRIBUTED TO THE FLUCTUATING BG LEVELS OR HOSPITALIZATION THE CUSTOMER THOUGHT THAT THE HOSPITALIZATION MAY HAVE BEEN TO USER ISSUES REGARDING THE INCOMPLETE BOLUS ALERTS AND AUTOOFF SAFETY ALARM THE CAUSE OF THE LOW BG WAS NOT KNOWN ON B THE CUSTOMERS HEALTHCARE PROVIDER REVERTED THE CUSTOMER TO INSULIN INJECTIONS TO MANAGE BG LEVELS THE CUSTOMER WAS DISCHARGED ON B HOWEVER THE BG LEVELS HAD NOT STABILIZED WHILE USING THE INJECTIONS UPON FOLLOWUP ON B THE CUSTOMER WAS STILL USING INSULIN INJECTIONS AND DECLINED TANDEM TECHNICAL SUPPORTS OFFER TO REACH OUT TO A CLINICAL DIABETES SPECIALIST FOR ASSISTANCE | Hyperglycemia | Patient Problem Code
6911152 | 88039247 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS HIGH THE PUMP SUPPLIES WERE CHANGED AND A BOLUS WAS DELIVERED IN AN ATTEMPT TO LOWER THE BG LEVEL THE CUSTOMER WAS HOSPITALIZED ON B AND DURING THE HOSPITALIZATION THE BG FLUCTUATED EXTREME HIGH AND LOW MGDL THE CAUSE OF THE BG FLUCTUATIONS WAS NOT KNOWN THE CUSTOMER WAS TREATED WITH AN INSULIN DRIP MEDICATION AND INSULIN INJECTIONS DURING A PUMP SYSTEM CHECK INCOMPLETE BOLUS ALERTS AND AN AUTOOFF SAFETY ALARM WERE FOUND TO HAVE OCCURRED THE CUSTOMER ACKNOWLEDGED HAVING NAVIGATED TO THE BOLUS SCREEN AND DID NOT EXIT WITHIN THE ALLOTTED SECONDS THE CUSTOMER ACKNOWLEDGED THE AUTOOFF SAFETY ALARM WAS VALID AS THERE WAS NO PUMP INTERACTION FOR HOURS AND THEN INSULIN WAS NOT RESUMED FOR APPROXIMATELY HOURS THE PUMP TIME WAS FOUND TO BE OFF BY MINUTES THE CUSTOMER DID NOT KNOW THE CAUSE OF THE INCORRECT TIME OR WHEN IT BECAME INCORRECT THE CUSTOMER DID NOT THINK THE TIME CONTRIBUTED TO THE FLUCTUATING BG LEVELS OR HOSPITALIZATION THE CUSTOMER THOUGHT THAT THE HOSPITALIZATION MAY HAVE BEEN TO USER ISSUES REGARDING THE INCOMPLETE BOLUS ALERTS AND AUTOOFF SAFETY ALARM THE CAUSE OF THE LOW BG WAS NOT KNOWN ON B THE CUSTOMERS HEALTHCARE PROVIDER REVERTED THE CUSTOMER TO INSULIN INJECTIONS TO MANAGE BG LEVELS THE CUSTOMER WAS DISCHARGED ON B HOWEVER THE BG LEVELS HAD NOT STABILIZED WHILE USING THE INJECTIONS UPON FOLLOWUP ON B THE CUSTOMER WAS STILL USING INSULIN INJECTIONS AND DECLINED TANDEM TECHNICAL SUPPORTS OFFER TO REACH OUT TO A CLINICAL DIABETES SPECIALIST FOR ASSISTANCE | Insufficient Information | Device problem code
6911321 | 88046336 | N | PATIENT INFORMATION NOT PROVIDED DUE TO JAPANESE PATIENT PRIVACY REGULATIONS A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT THE REPRESENTATIVE REPORTED THAT THE ROTATION BELT WAS LOOSE AND TIGHTENING IT RESTORED FUNCTIONALITY TO THE UNIT THE IMAGING SYSTEM THEN PASSED THE SYSTEM CHECKOUT AND WAS FOUND TO BE FULLY FUNCTIONAL | No Device Problem Found | Medical Evaluation result code
6911576 | 88053467 | D | IT WAS REPORTED THAT MULTIPLE TEMPERATURE ALARMS OCCURRED REPORTEDLY THE PUMP WAS NOT EXPOSED TO EXTREME TEMPERATURES THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER CLEARED THE ALARM AND INSULIN DELIVERY WAS RESUMED | No Known Impact Or Consequence To Patient | Patient problem code
6911576 | 88053467 | D | IT WAS REPORTED THAT MULTIPLE TEMPERATURE ALARMS OCCURRED REPORTEDLY THE PUMP WAS NOT EXPOSED TO EXTREME TEMPERATURES THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL THE CUSTOMER CLEARED THE ALARM AND INSULIN DELIVERY WAS RESUMED | Device Alarm System | Device problem code
6910895 | 88054924 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE EQUIPMENT NOTED THE MICROSWITCH CONNECTOR WAS CONTAMINATED WITH DUST THE CONNECTOR WAS CLEANED RESOLVING THE REPORTED COMPLAINT NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Environmental Particulates | Device problem code
6910895 | 88054924 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE EQUIPMENT NOTED THE MICROSWITCH CONNECTOR WAS CONTAMINATED WITH DUST THE CONNECTOR WAS CLEANED RESOLVING THE REPORTED COMPLAINT NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Environment Problem Identified | Medical Evaluation result code
6890302 | 88057481 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6890302 | 88057481 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6911858 | 88062610 | D | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER HCP REGARDING A PATIENTS IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR HCP SAID THE PATIENT IS WANTING TO MEET A LOCAL MANUFACTURE REPRESENTATIVE REP BECAUSE THE PATIENT HAS AIR IN THEIR URINE AND REOC CURRING URINARY TRACT INFECTIONS UTI NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient problem code
6911858 | 88062610 | D | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER HCP REGARDING A PATIENTS IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR HCP SAID THE PATIENT IS WANTING TO MEET A LOCAL MANUFACTURE REPRESENTATIVE REP BECAUSE THE PATIENT HAS AIR IN THEIR URINE AND REOC CURRING URINARY TRACT INFECTIONS UTI NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6911858 | 88062610 | D | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER HCP REGARDING A PATIENTS IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR HCP SAID THE PATIENT IS WANTING TO MEET A LOCAL MANUFACTURE REPRESENTATIVE REP BECAUSE THE PATIENT HAS AIR IN THEIR URINE AND REOC CURRING URINARY TRACT INFECTIONS UTI NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Urinary Tract | Patient problem code
6895121 | 88073564 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
6895121 | 88073564 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Device Dislodged or Dislocated | Device problem code
6893918 | 88077890 | N | RECALIBRATED THE UNIT TO FIX THE REPORTED ISSUE NO REPORT OF PATIENT INVOLVEMENT B THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Calibration Problem | Device problem code
6893918 | 88077890 | N | RECALIBRATED THE UNIT TO FIX THE REPORTED ISSUE NO REPORT OF PATIENT INVOLVEMENT B THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Failure to Calibrate | Medical Evaluation result code
6913148 | 88106716 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF SWELLING IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF EDEMA ADVERSE EVENTS A CLINICAL EVALUATION OF JUVÉDERM ULTRA PLUS XC THE MOST COMMON INJECTION SITE RESPONSES FOR JUVÉDERM ULTRA PLUS XC WERE REDNESS SWELLING TENDERNESS FIRMNESS LUMPSBUMPS DISCOLORATION AND BRUISING | Material Discolored | Device problem code
6913148 | 88106716 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF SWELLING IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF EDEMA ADVERSE EVENTS A CLINICAL EVALUATION OF JUVÉDERM ULTRA PLUS XC THE MOST COMMON INJECTION SITE RESPONSES FOR JUVÉDERM ULTRA PLUS XC WERE REDNESS SWELLING TENDERNESS FIRMNESS LUMPSBUMPS DISCOLORATION AND BRUISING | No Device Problem Found | Medical Evaluation result code
6912819 | 88107227 | N | THE SAMPLE TUBE BD ML DOUBLE WALLED TUBE USED FOR ANALYSIS WAS UNACCEPTABLE FOR USE ON THE ANALYZER THE SAMPLE PROBE WAS HITTING THE SIDE OF THE TUBE WHILE ASPIRATING NO ANALYZER DEFICIENCIES WERE IDENTIFIED PER THE INSTRUCTIONS FOR USE IFU FOR THE ANALYZER THE ONLY DOUBLE WALLED TUBES TO BE USED ON THE ANALYZER ARE THE VACUTAINER PLUS PLASTIC CITRATE TUBE ML OR THE VACUETTE SANDWICH COAGULATION TUBE MLML THE IFU SPECIFICALLY IDENTIFIES THE TUBE USED AS NOT BEING ACCEPTABLE FOR USE ON THE ANALYZER | Coagulation in Device or device Ingredient | Device problem code
6912819 | 88107227 | N | THE SAMPLE TUBE BD ML DOUBLE WALLED TUBE USED FOR ANALYSIS WAS UNACCEPTABLE FOR USE ON THE ANALYZER THE SAMPLE PROBE WAS HITTING THE SIDE OF THE TUBE WHILE ASPIRATING NO ANALYZER DEFICIENCIES WERE IDENTIFIED PER THE INSTRUCTIONS FOR USE IFU FOR THE ANALYZER THE ONLY DOUBLE WALLED TUBES TO BE USED ON THE ANALYZER ARE THE VACUTAINER PLUS PLASTIC CITRATE TUBE ML OR THE VACUETTE SANDWICH COAGULATION TUBE MLML THE IFU SPECIFICALLY IDENTIFIES THE TUBE USED AS NOT BEING ACCEPTABLE FOR USE ON THE ANALYZER | Non-Functional Defect | Medical Evaluation result code
6913241 | 88113251 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE HAD DEPLOYED LATE INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6913241 | 88113251 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE HAD DEPLOYED LATE INDICATING A NEEDLE MECHANISM FAILURE | Activation, Positioning or Separation Problem | Device problem code
6912967 | 88119341 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE UNINTERRUPTIBLE POWER SUPPLY UPS FAILED CAUSING THE CENTRAL NURSES STATION CNS TO LOSE POWER IT ALSO EMITTED A LOUD BEEPING NOISE HE CONNECTED THE CENTRAL NURSES STATION CNS TO A WALL OUTLET TO CONTINUE PATIENT MONITORING NO PATIENT HARM WAS REPORTED THE BIOMEDICAL ENGINEER WAS PROVIDED WITH AN EXCHANGED UNINTERRUPTIBLE POWER SUPPLY UPS | Noise, Audible | Device problem code
6912967 | 88119341 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE UNINTERRUPTIBLE POWER SUPPLY UPS FAILED CAUSING THE CENTRAL NURSES STATION CNS TO LOSE POWER IT ALSO EMITTED A LOUD BEEPING NOISE HE CONNECTED THE CENTRAL NURSES STATION CNS TO A WALL OUTLET TO CONTINUE PATIENT MONITORING NO PATIENT HARM WAS REPORTED THE BIOMEDICAL ENGINEER WAS PROVIDED WITH AN EXCHANGED UNINTERRUPTIBLE POWER SUPPLY UPS | Power Source Problem Identified | Medical Evaluation result code
6897887 | 88120066 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6897887 | 88120066 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6913214 | 88129515 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS B PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6913214 | 88129515 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS B PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6913216 | 88129533 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL AT PM WHILE WEARING THE POD BETWEEN AND HOURS THE POD HAD FELL OFF FROM THE INFUSION SITE BACK WHICH WOULD RESULT IN CANNULA DISLODGING | No Known Impact Or Consequence To Patient | Patient Problem Code
6913216 | 88129533 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL AT PM WHILE WEARING THE POD BETWEEN AND HOURS THE POD HAD FELL OFF FROM THE INFUSION SITE BACK WHICH WOULD RESULT IN CANNULA DISLODGING | Device Dislodged or Dislocated | Device problem code
6913223 | 88129663 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6913223 | 88129663 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6913221 | 88130846 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6913221 | 88130846 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6909925 | 88136423 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY SIX DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6909925 | 88136423 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY SIX DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6909926 | 88137540 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6909926 | 88137540 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6897292 | 88139288 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6897292 | 88139288 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6910776 | 88147022 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEIR INSULIN PUMP HAD MULTIPLE PUMP ERRORS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED FOR PUMP ERRORS AND ATTEMPTED TO CHECK THE HISTORY AND FOUND HARDWARE LOW LEVEL FAILURES THE CUSTOMER STATES PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD DISPLACEMENT TEST WAS SUCCESSFUL AT THIS TIME THE CUSTOMER STATES THEY ARE ABLE TO REWIND THE PUMP THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6910776 | 88147022 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEIR INSULIN PUMP HAD MULTIPLE PUMP ERRORS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL TROUBLESHOOTING WAS PERFORMED FOR PUMP ERRORS AND ATTEMPTED TO CHECK THE HISTORY AND FOUND HARDWARE LOW LEVEL FAILURES THE CUSTOMER STATES PUMP WAS NOT EXPOSED TO A HIGH MAGNETIC FIELD DISPLACEMENT TEST WAS SUCCESSFUL AT THIS TIME THE CUSTOMER STATES THEY ARE ABLE TO REWIND THE PUMP THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Pumping Problem | Device problem code
6896784 | 88149056 | D | A PATIENTS BLOOD GLUCOSE REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6896784 | 88149056 | D | A PATIENTS BLOOD GLUCOSE REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6908019 | 88152198 | N | B AS PER THE SUBSIDIARY TUBING WAS USED DURING CALIBRATION AND THE ERROR OCCURED WHEN THE DEVICE WAS IN THE FULLY OCCLUDE POSITION | Calibration Problem | Device problem code
6908019 | 88152198 | N | B AS PER THE SUBSIDIARY TUBING WAS USED DURING CALIBRATION AND THE ERROR OCCURED WHEN THE DEVICE WAS IN THE FULLY OCCLUDE POSITION | Failure to Calibrate | Medical Evaluation result code
6910645 | 88159034 | N | B AS PER THE SUBSIDIARY CALIBRATION WAS NOT ATTEMPTED IMMEDIATELY PRIOR TO THE CASE BUT MULTIPLE ATTEMPTS TO CALIBRATE WERE UNSUCCESSFUL THEY USED LOCAL CONTROL KNOBS DURING CALIBRATION LOCAL CONTROL KNOBS WERE USED WHEN THE CENTRAL CONTROL MONITOR CCM SLIDER WAS UNAVAILABLE THE SALES ASSOCIATE AND SERVICE ENGINEER CONFIRMED THE GAS CALIBRATION | Calibration Problem | Device problem code
6910645 | 88159034 | N | B AS PER THE SUBSIDIARY CALIBRATION WAS NOT ATTEMPTED IMMEDIATELY PRIOR TO THE CASE BUT MULTIPLE ATTEMPTS TO CALIBRATE WERE UNSUCCESSFUL THEY USED LOCAL CONTROL KNOBS DURING CALIBRATION LOCAL CONTROL KNOBS WERE USED WHEN THE CENTRAL CONTROL MONITOR CCM SLIDER WAS UNAVAILABLE THE SALES ASSOCIATE AND SERVICE ENGINEER CONFIRMED THE GAS CALIBRATION | Failure to Calibrate | Medical Evaluation result code
6893332 | 88160824 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6893332 | 88160824 | D | A PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6881935 | 88160986 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD FOR OVER HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6881935 | 88160986 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD FOR OVER HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE IF THE CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6888522 | 88161181 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT NO LOGS WERE RECORDED AS THIS UNIT WAS THE BACKUP AND WAS NOT BEING USED ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS THE MOST LIKELY ROOT CAUSE OF THE RTCS RESET TO ZERO CAN BE ATTRIBUTED A RUNDOWN OF THE CONTROLLERS RTC COIN CELL DUE TO AN EXTENDED PERIOD OF TIME WITHOUT CONNECTION TO A POWER SOURCE THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6888522 | 88161181 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT NO LOGS WERE RECORDED AS THIS UNIT WAS THE BACKUP AND WAS NOT BEING USED ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER WAS RECEIVED WITH ITS REAL TIME CLOCK RTC RESET TO ZERO HOWEVER WHEN SET TO THE CURRENT DATE AND TIME THE RTC WAS ABLE TO RETAIN THE NEW SETTINGS THE MOST LIKELY ROOT CAUSE OF THE RTCS RESET TO ZERO CAN BE ATTRIBUTED A RUNDOWN OF THE CONTROLLERS RTC COIN CELL DUE TO AN EXTENDED PERIOD OF TIME WITHOUT CONNECTION TO A POWER SOURCE THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6888448 | 88161654 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED MULTIPLE OCCURRENCES OF CONTROLLER FAULT ALARMS WITH FAULT CODES INDICATING SYSUICAPSFAIL THESE ALARMS INDICATE FAILURES OF THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO A CONTROLLER FAULT DUE TO SYSUICAPSFAIL HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device problem code
6888448 | 88161654 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED MULTIPLE OCCURRENCES OF CONTROLLER FAULT ALARMS WITH FAULT CODES INDICATING SYSUICAPSFAIL THESE ALARMS INDICATE FAILURES OF THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO A CONTROLLER FAULT DUE TO SYSUICAPSFAIL HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6911604 | 88169492 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE | Connection Problem | Device problem code
6911604 | 88169492 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE | No Device Problem Found | Medical Evaluation result code
6914343 | 88169760 | D | MEDTRONIC RECEIVED INFORMATION THAT AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE AND ELECTROCARDIOGRAM ECG INDICATED COMPLETE HEART BLOCK CHB ONE DAY AFTER THE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6914343 | 88169760 | D | MEDTRONIC RECEIVED INFORMATION THAT AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE AND ELECTROCARDIOGRAM ECG INDICATED COMPLETE HEART BLOCK CHB ONE DAY AFTER THE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6914929 | 88202998 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6914720 | 88205161 | D | IT WAS REPORTED THAT A MM PERICARDIAL AORTIC VALVE WAS DISABLED VIA A VALVEINVALVE PROCEDURE AFTER AN IMPLANT DURATION OF YEARS MONTHS DUE TO SEVERE AORTIC STENOSIS A MM TRANSCATHETER VALVE WAS SUCCESSFULLY IMPLANTED THE PATIENT WAS DISCHARGED | Aortic Stenosis | Patient problem code
6914720 | 88205161 | D | IT WAS REPORTED THAT A MM PERICARDIAL AORTIC VALVE WAS DISABLED VIA A VALVEINVALVE PROCEDURE AFTER AN IMPLANT DURATION OF YEARS MONTHS DUE TO SEVERE AORTIC STENOSIS A MM TRANSCATHETER VALVE WAS SUCCESSFULLY IMPLANTED THE PATIENT WAS DISCHARGED | Stenosis | Patient problem code
6914720 | 88205161 | D | IT WAS REPORTED THAT A MM PERICARDIAL AORTIC VALVE WAS DISABLED VIA A VALVEINVALVE PROCEDURE AFTER AN IMPLANT DURATION OF YEARS MONTHS DUE TO SEVERE AORTIC STENOSIS A MM TRANSCATHETER VALVE WAS SUCCESSFULLY IMPLANTED THE PATIENT WAS DISCHARGED | Insufficient Information | Device problem code
6914396 | 88215147 | N | THE CUSTOMER WAS ABLE TO RETURN UNITS FROM THE LOT TO NATUS THE FAILURE COULD BE CONFIRMED NATUS INSTRUCTIONS INCLUDE SELECTING THE APPROPRIATE SIZE EAR TIP FOR THE PATIENT AS WELL AS INSTRUCTIONS REGARDING USING LARGER SIZE FOAM TIPS FOR INSERTION INTO THE EAR CANAL NATUS IS INVESTIGATING THIS ISSUE THROUGH CORRECTIVE ACTIONS AT THIS TIME | Inadequacy of Device Shape and/or Size | Device problem code
6914396 | 88215147 | N | THE CUSTOMER WAS ABLE TO RETURN UNITS FROM THE LOT TO NATUS THE FAILURE COULD BE CONFIRMED NATUS INSTRUCTIONS INCLUDE SELECTING THE APPROPRIATE SIZE EAR TIP FOR THE PATIENT AS WELL AS INSTRUCTIONS REGARDING USING LARGER SIZE FOAM TIPS FOR INSERTION INTO THE EAR CANAL NATUS IS INVESTIGATING THIS ISSUE THROUGH CORRECTIVE ACTIONS AT THIS TIME | Stress Problem Identified | Medical Evaluation result code
6907099 | 88222591 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED WAS THE DEVICE STUCK ON TISSUE WHEN THE JAWS WOULD NOT OPEN IF YES HOW WAS THE DEVICE REMOVED WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED NOT OPEN AT LAST USED ECHELON TO RECUT AND REANASTOMOSED THE TISSUE CLOSE THE ATWS JAW TO RECUT THE TISSUE BUT NOT THE CRITICAL PART SO NOT CAUSE PATIENT BLEEDING AND SURGERY DELAYED | Entrapment of Device | Device Problem Code
6907099 | 88222591 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED WAS THE DEVICE STUCK ON TISSUE WHEN THE JAWS WOULD NOT OPEN IF YES HOW WAS THE DEVICE REMOVED WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED NOT OPEN AT LAST USED ECHELON TO RECUT AND REANASTOMOSED THE TISSUE CLOSE THE ATWS JAW TO RECUT THE TISSUE BUT NOT THE CRITICAL PART SO NOT CAUSE PATIENT BLEEDING AND SURGERY DELAYED | Difficult to Open or Close | Device problem code
6907099 | 88222591 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED WAS THE DEVICE STUCK ON TISSUE WHEN THE JAWS WOULD NOT OPEN IF YES HOW WAS THE DEVICE REMOVED WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED NOT OPEN AT LAST USED ECHELON TO RECUT AND REANASTOMOSED THE TISSUE CLOSE THE ATWS JAW TO RECUT THE TISSUE BUT NOT THE CRITICAL PART SO NOT CAUSE PATIENT BLEEDING AND SURGERY DELAYED | Mechanical Jam | Device problem code
6907099 | 88222591 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED WAS THE DEVICE STUCK ON TISSUE WHEN THE JAWS WOULD NOT OPEN IF YES HOW WAS THE DEVICE REMOVED WERE THE DEVICE JAWS EVENTUALLY ABLE TO BE OPENED NOT OPEN AT LAST USED ECHELON TO RECUT AND REANASTOMOSED THE TISSUE CLOSE THE ATWS JAW TO RECUT THE TISSUE BUT NOT THE CRITICAL PART SO NOT CAUSE PATIENT BLEEDING AND SURGERY DELAYED | Mechanical Problem Identified | Medical Evaluation result code
6909929 | 88223710 | N | THE CONCERNED DEVICE LACRIFAST IS NOT DISTRIBUTED IN US UNDER THIS DEVICE NAME BUT IS IDENTICAL TO THE DEVICE LACRIFLOW LACRIMAL STENT DISTRIBUTED IN US UNDER K K THE ACTUAL DEVICE USED WAS RETURNED AND INVESTIGATED THERE WAS A HOLE FROM WHERE THE BOUGIE WAS THOUGHT TO PENETRATE OUT IN THE SIDEWALL OF THE TUBE COLORED IN BLUE WHERE A STAINLESS STEEL RING WAS MOUNTED AND THE RING WAS DISMOUNTED AND MISSING THE RINGS WHEREABOUTS WAS NOT CONFIRMED DURING THE PROCEDURE IN THIS EYE CLINIC AND ACCORDINGLY THE MISSING STAINLESS STEEL RING MIGHT HAVE BEEN REMAINING IN THE PATIENT BODY AS A POSSIBLE CAUSE OF THE PENETRATION OF THE BOUGIE AND DROPPING OFF THE RING DURING THE USE WE SPECULATE AS FOLLOWS WHEN THE DOCTOR TRIED TO INSERT AND ADVANCE THE TUBE WITH THE BOUGIE INSERTED FORCIBLY IN THE OCCLUDED LACRIMAL DUCT EXCESSIVE MECHANICAL FORCE WAS LOADED ON THE RING AND THE DISTAL TIP OF THE TUBE WHERE THE RING WAS MOUNTED AND FURTHER PUSHING THE BOUGIE RESULTED IN DISMOUNTING THE RING AND PENETRATION OF THE BOUGIE WITH THE RING OUT OF THE TUBE | Component Missing | Device problem code
6909929 | 88223710 | N | THE CONCERNED DEVICE LACRIFAST IS NOT DISTRIBUTED IN US UNDER THIS DEVICE NAME BUT IS IDENTICAL TO THE DEVICE LACRIFLOW LACRIMAL STENT DISTRIBUTED IN US UNDER K K THE ACTUAL DEVICE USED WAS RETURNED AND INVESTIGATED THERE WAS A HOLE FROM WHERE THE BOUGIE WAS THOUGHT TO PENETRATE OUT IN THE SIDEWALL OF THE TUBE COLORED IN BLUE WHERE A STAINLESS STEEL RING WAS MOUNTED AND THE RING WAS DISMOUNTED AND MISSING THE RINGS WHEREABOUTS WAS NOT CONFIRMED DURING THE PROCEDURE IN THIS EYE CLINIC AND ACCORDINGLY THE MISSING STAINLESS STEEL RING MIGHT HAVE BEEN REMAINING IN THE PATIENT BODY AS A POSSIBLE CAUSE OF THE PENETRATION OF THE BOUGIE AND DROPPING OFF THE RING DURING THE USE WE SPECULATE AS FOLLOWS WHEN THE DOCTOR TRIED TO INSERT AND ADVANCE THE TUBE WITH THE BOUGIE INSERTED FORCIBLY IN THE OCCLUDED LACRIMAL DUCT EXCESSIVE MECHANICAL FORCE WAS LOADED ON THE RING AND THE DISTAL TIP OF THE TUBE WHERE THE RING WAS MOUNTED AND FURTHER PUSHING THE BOUGIE RESULTED IN DISMOUNTING THE RING AND PENETRATION OF THE BOUGIE WITH THE RING OUT OF THE TUBE | Missing or Inadequate Safety Measures | Medical Evaluation result code
6880294 | 88226358 | N | THE DISTRIBUTOR PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE MAIN MANIFOLD AND CIRCUIT MODULE WERE BLOCKED THE MAIN MANIFOLD AND CIRCUIT MODULE WERE CLEANED THE UNIT WAS RETURNED TO SERVICE NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Partial Blockage | Device problem code
6916005 | 88229838 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING TENDERNESS FEVERISH AND COULD NOT OPEN THEIR MOUTH ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF EDEMA PAIN FEVER AND COMPLAINT ILLDEFINED WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSE POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICEMASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN | Migration or Expulsion of Device | Device problem code
6916005 | 88229838 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING TENDERNESS FEVERISH AND COULD NOT OPEN THEIR MOUTH ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF EDEMA PAIN FEVER AND COMPLAINT ILLDEFINED WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSE POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICEMASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN | Material Discolored | Device problem code
6916005 | 88229838 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING TENDERNESS FEVERISH AND COULD NOT OPEN THEIR MOUTH ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF EDEMA PAIN FEVER AND COMPLAINT ILLDEFINED WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSE POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICEMASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN | Leak / Splash | Device problem code
6916005 | 88229838 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING TENDERNESS FEVERISH AND COULD NOT OPEN THEIR MOUTH ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING FOR THE REPORTED EVENTS OF EDEMA PAIN FEVER AND COMPLAINT ILLDEFINED WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSE POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICEMASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN | Leakage/Seal | Medical Evaluation result code
6893166 | 88230157 | N | A LOG FILE EVALUATION PROVIDED BY THE MANUFACTURER CONFIRMED THE INITIAL DESCRIPTION IN SO FAR THAT ERROR CODES WERE RECORDED INDICATING PROBLEMS WITH THE VACUUM PRESSURE THE VACUUM PRESSURE IS NECESSARY TO MAINTAIN PROPER FUNCTION OF THE VENTILATOR PISTON DIAPHRAGM IE TO AVOID WRINKLING DURING MOVEMENT FOR SAFETY REASONS TO PREVENT FROM SERIOUS DAMAGES THE DEVICE SHUTS DOWN THE AUTOMATIC VENTILATION IT IS HOWEVER NOT PLAUSIBLE THAT THE REPORTED HOLE IN THE ONE PRESSURE READING LINE COULD HAVE CAUSED THE SYMPTOM SINCE THIS WOULD HAVE PRODUCED A DIFFERENT TYPE OF ERROR CODE IN THE LOGS AND WOULD NOT HAVE LED TO A STOP OF VENTILATION IT COULD BE FURTHER DETERMINED THAT THE DEVICE OPERATED ON BATTERY WHICH RAN EMPTY DURING THE PROCEDURE IN QUESTION THIS IS INDEED THE REASON FOR THE REPORTED VENTILATOR FAILURE ALARM WITH A FULLY CHARGED BATTERY THE DEVICE OFFERS A MINIMUM RUN TIME OF MINUTES TYPICALLY THE OPERATION CAN BE CONTINUED FOR TWO HOURS WITHOUT LIMITATIONS THE AUTOMATIC SWITCHOVER TO BATTERY OPERATION IS INDICATED TO THE USER BY A CORRESPONDING ALARM FURTHER ALARM MESSAGES ARE POSTED IF THE BATTERY CAPACITY UNDERRUNS CERTAIN STAGES A TOTAL LOSS OF POWER IS INDICATED BY A POWER FAIL ALARM WHICH IS BUFFERED BY AN INDEPENDENT POWER SOURCE IN THIS CASE ONLY MANUAL VENTILATION REMAINS POSSIBLE DRÄGER FINALLY CONCLUDES THAT THERE WERE TECHNICAL PROBLEMS IN MAINTAINING THE VACUUM PRESSURE BUT INDEED THE ROOT CAUSE FOR THE SHUTDOWN OF AUTOMATIC VENTILATION WAS AN EMPTY BATTERY | Device Displays Incorrect Message | Device problem code
6893166 | 88230157 | N | A LOG FILE EVALUATION PROVIDED BY THE MANUFACTURER CONFIRMED THE INITIAL DESCRIPTION IN SO FAR THAT ERROR CODES WERE RECORDED INDICATING PROBLEMS WITH THE VACUUM PRESSURE THE VACUUM PRESSURE IS NECESSARY TO MAINTAIN PROPER FUNCTION OF THE VENTILATOR PISTON DIAPHRAGM IE TO AVOID WRINKLING DURING MOVEMENT FOR SAFETY REASONS TO PREVENT FROM SERIOUS DAMAGES THE DEVICE SHUTS DOWN THE AUTOMATIC VENTILATION IT IS HOWEVER NOT PLAUSIBLE THAT THE REPORTED HOLE IN THE ONE PRESSURE READING LINE COULD HAVE CAUSED THE SYMPTOM SINCE THIS WOULD HAVE PRODUCED A DIFFERENT TYPE OF ERROR CODE IN THE LOGS AND WOULD NOT HAVE LED TO A STOP OF VENTILATION IT COULD BE FURTHER DETERMINED THAT THE DEVICE OPERATED ON BATTERY WHICH RAN EMPTY DURING THE PROCEDURE IN QUESTION THIS IS INDEED THE REASON FOR THE REPORTED VENTILATOR FAILURE ALARM WITH A FULLY CHARGED BATTERY THE DEVICE OFFERS A MINIMUM RUN TIME OF MINUTES TYPICALLY THE OPERATION CAN BE CONTINUED FOR TWO HOURS WITHOUT LIMITATIONS THE AUTOMATIC SWITCHOVER TO BATTERY OPERATION IS INDICATED TO THE USER BY A CORRESPONDING ALARM FURTHER ALARM MESSAGES ARE POSTED IF THE BATTERY CAPACITY UNDERRUNS CERTAIN STAGES A TOTAL LOSS OF POWER IS INDICATED BY A POWER FAIL ALARM WHICH IS BUFFERED BY AN INDEPENDENT POWER SOURCE IN THIS CASE ONLY MANUAL VENTILATION REMAINS POSSIBLE DRÄGER FINALLY CONCLUDES THAT THERE WERE TECHNICAL PROBLEMS IN MAINTAINING THE VACUUM PRESSURE BUT INDEED THE ROOT CAUSE FOR THE SHUTDOWN OF AUTOMATIC VENTILATION WAS AN EMPTY BATTERY | Loss of Power | Device problem code
6893166 | 88230157 | N | A LOG FILE EVALUATION PROVIDED BY THE MANUFACTURER CONFIRMED THE INITIAL DESCRIPTION IN SO FAR THAT ERROR CODES WERE RECORDED INDICATING PROBLEMS WITH THE VACUUM PRESSURE THE VACUUM PRESSURE IS NECESSARY TO MAINTAIN PROPER FUNCTION OF THE VENTILATOR PISTON DIAPHRAGM IE TO AVOID WRINKLING DURING MOVEMENT FOR SAFETY REASONS TO PREVENT FROM SERIOUS DAMAGES THE DEVICE SHUTS DOWN THE AUTOMATIC VENTILATION IT IS HOWEVER NOT PLAUSIBLE THAT THE REPORTED HOLE IN THE ONE PRESSURE READING LINE COULD HAVE CAUSED THE SYMPTOM SINCE THIS WOULD HAVE PRODUCED A DIFFERENT TYPE OF ERROR CODE IN THE LOGS AND WOULD NOT HAVE LED TO A STOP OF VENTILATION IT COULD BE FURTHER DETERMINED THAT THE DEVICE OPERATED ON BATTERY WHICH RAN EMPTY DURING THE PROCEDURE IN QUESTION THIS IS INDEED THE REASON FOR THE REPORTED VENTILATOR FAILURE ALARM WITH A FULLY CHARGED BATTERY THE DEVICE OFFERS A MINIMUM RUN TIME OF MINUTES TYPICALLY THE OPERATION CAN BE CONTINUED FOR TWO HOURS WITHOUT LIMITATIONS THE AUTOMATIC SWITCHOVER TO BATTERY OPERATION IS INDICATED TO THE USER BY A CORRESPONDING ALARM FURTHER ALARM MESSAGES ARE POSTED IF THE BATTERY CAPACITY UNDERRUNS CERTAIN STAGES A TOTAL LOSS OF POWER IS INDICATED BY A POWER FAIL ALARM WHICH IS BUFFERED BY AN INDEPENDENT POWER SOURCE IN THIS CASE ONLY MANUAL VENTILATION REMAINS POSSIBLE DRÄGER FINALLY CONCLUDES THAT THERE WERE TECHNICAL PROBLEMS IN MAINTAINING THE VACUUM PRESSURE BUT INDEED THE ROOT CAUSE FOR THE SHUTDOWN OF AUTOMATIC VENTILATION WAS AN EMPTY BATTERY | Component or Accessory Incompatibility | Device Problem Code
6893166 | 88230157 | N | A LOG FILE EVALUATION PROVIDED BY THE MANUFACTURER CONFIRMED THE INITIAL DESCRIPTION IN SO FAR THAT ERROR CODES WERE RECORDED INDICATING PROBLEMS WITH THE VACUUM PRESSURE THE VACUUM PRESSURE IS NECESSARY TO MAINTAIN PROPER FUNCTION OF THE VENTILATOR PISTON DIAPHRAGM IE TO AVOID WRINKLING DURING MOVEMENT FOR SAFETY REASONS TO PREVENT FROM SERIOUS DAMAGES THE DEVICE SHUTS DOWN THE AUTOMATIC VENTILATION IT IS HOWEVER NOT PLAUSIBLE THAT THE REPORTED HOLE IN THE ONE PRESSURE READING LINE COULD HAVE CAUSED THE SYMPTOM SINCE THIS WOULD HAVE PRODUCED A DIFFERENT TYPE OF ERROR CODE IN THE LOGS AND WOULD NOT HAVE LED TO A STOP OF VENTILATION IT COULD BE FURTHER DETERMINED THAT THE DEVICE OPERATED ON BATTERY WHICH RAN EMPTY DURING THE PROCEDURE IN QUESTION THIS IS INDEED THE REASON FOR THE REPORTED VENTILATOR FAILURE ALARM WITH A FULLY CHARGED BATTERY THE DEVICE OFFERS A MINIMUM RUN TIME OF MINUTES TYPICALLY THE OPERATION CAN BE CONTINUED FOR TWO HOURS WITHOUT LIMITATIONS THE AUTOMATIC SWITCHOVER TO BATTERY OPERATION IS INDICATED TO THE USER BY A CORRESPONDING ALARM FURTHER ALARM MESSAGES ARE POSTED IF THE BATTERY CAPACITY UNDERRUNS CERTAIN STAGES A TOTAL LOSS OF POWER IS INDICATED BY A POWER FAIL ALARM WHICH IS BUFFERED BY AN INDEPENDENT POWER SOURCE IN THIS CASE ONLY MANUAL VENTILATION REMAINS POSSIBLE DRÄGER FINALLY CONCLUDES THAT THERE WERE TECHNICAL PROBLEMS IN MAINTAINING THE VACUUM PRESSURE BUT INDEED THE ROOT CAUSE FOR THE SHUTDOWN OF AUTOMATIC VENTILATION WAS AN EMPTY BATTERY | Power Source Problem Identified | Medical Evaluation result code
6909152 | 88233192 | N | THE FULL NUMBER FOR THE PRODUCT IN QUESTION IS B IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION AND DETERMINED THAT THE INSTRUMENT TEST WELL IMAGES IN QUESTION SHOWED WELL NUMBERS AND AS SHOWING RED BLOOD CELL ADHERENCE THE PRODUCT HAD ALREADY EXPIRED BEFORE RETENTION TESTING COULD BE DONE BUT THE IMMUCOR LABORATORY HAD ALREADY PERFORMED SPECIFIC LOT RETENTION TESTING IN THE PAST ON SEP WHICH PERFORMED AS EXPECTED AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP AND FOUND THE TESTING INSTRUMENT IN QUESTION TO BE OPERATING AS EXPECTED | Connection Problem | Device problem code
6909152 | 88233192 | N | THE FULL NUMBER FOR THE PRODUCT IN QUESTION IS B IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION AND DETERMINED THAT THE INSTRUMENT TEST WELL IMAGES IN QUESTION SHOWED WELL NUMBERS AND AS SHOWING RED BLOOD CELL ADHERENCE THE PRODUCT HAD ALREADY EXPIRED BEFORE RETENTION TESTING COULD BE DONE BUT THE IMMUCOR LABORATORY HAD ALREADY PERFORMED SPECIFIC LOT RETENTION TESTING IN THE PAST ON SEP WHICH PERFORMED AS EXPECTED AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP AND FOUND THE TESTING INSTRUMENT IN QUESTION TO BE OPERATING AS EXPECTED | No Device Problem Found | Medical Evaluation result code
6897911 | 88233963 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF FROM THE INFUSION SITE ARM | Insufficient Information | Device problem code
6897911 | 88233963 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6888797 | 88234916 | D | INFORMATION WAS RECEIVED FROM A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR SPINAL PAIN AND PERIPHERAL NEUROPATHY VIA A MANUFACTURERS REPRESENTATIVE REP THE REP REPORTED THAT THE PATIENT HAD NOT USED THE DEVICE IN OVER ONE YEAR NOT EXACTLY SURE OF WHEN HE QUIT USING IT THE REP REPORTED THAT THE PATIENT STOPPED USING THE DEVICE BECAUSE DESPITE MANY REPROGRAMMING SESSIONS THE INS WOULD NEVER STAY IN THE PATIENTS FEET WHERE HIS PRIMARY PAIN WAS LOCATED ALONG WITH SOME SECONDARY BACK PAIN THE REP REPORTED THAT THEY ATTEMPTED TO INTERROGATE THE INS BUT NO RESPONSE AND CONCLUDED THE PATIENT WAS IN OVERDISCHARGE THE REP REPORTED THAT THEY USED ANTENNA LOCATE TIMES TO ATTEMPT TO AWAKEN THE DEVICE NO RESPONSE THE REP REPORTED THAT THEY INITIATED A PHYSICIAN MODE RECHARGE PMR AFTER A ONE HOUR TRICKLE CHARGE AND THE DEVICE DID NOT START RECHARGING THE REP REPORTED THAT THE PATIENT ATTEMPTED MORE SESSIONS OF PMR FOR A TOTAL OF HOURS AFTER EACH SESSION WHEN ATTEMPTED TO RECHARGE NO RESPONSE AT ALL NOT EVEN A CALL YOUR DOCTOR SYMBOL FROM THE INS IT WAS CONCLUDED THAT THE INS WOULD NOT RESPOND THE REP REPORTED THAT THERE WAS NO FALL ETC THAT LED TO THE ISSUE JUST THAT THE PATIENT STOPPED USING AND RECHARGING THE DEVICE BECAUSE THE THERAPY FAILED THE REP REPORTED THAT THEY WOULD BE AT THE PATIENTS NEXT FOLLOW UP APPOINTMENT IN OCTOBER TO DISCUSS THE POSSIBILITY OF REVISIONREPLACEMENT THE ISSUE WAS NOT RESOLVED AT THE TIME OF THE REPORT NO FURTHER COMPLICATIONS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
6888797 | 88234916 | D | INFORMATION WAS RECEIVED FROM A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR SPINAL PAIN AND PERIPHERAL NEUROPATHY VIA A MANUFACTURERS REPRESENTATIVE REP THE REP REPORTED THAT THE PATIENT HAD NOT USED THE DEVICE IN OVER ONE YEAR NOT EXACTLY SURE OF WHEN HE QUIT USING IT THE REP REPORTED THAT THE PATIENT STOPPED USING THE DEVICE BECAUSE DESPITE MANY REPROGRAMMING SESSIONS THE INS WOULD NEVER STAY IN THE PATIENTS FEET WHERE HIS PRIMARY PAIN WAS LOCATED ALONG WITH SOME SECONDARY BACK PAIN THE REP REPORTED THAT THEY ATTEMPTED TO INTERROGATE THE INS BUT NO RESPONSE AND CONCLUDED THE PATIENT WAS IN OVERDISCHARGE THE REP REPORTED THAT THEY USED ANTENNA LOCATE TIMES TO ATTEMPT TO AWAKEN THE DEVICE NO RESPONSE THE REP REPORTED THAT THEY INITIATED A PHYSICIAN MODE RECHARGE PMR AFTER A ONE HOUR TRICKLE CHARGE AND THE DEVICE DID NOT START RECHARGING THE REP REPORTED THAT THE PATIENT ATTEMPTED MORE SESSIONS OF PMR FOR A TOTAL OF HOURS AFTER EACH SESSION WHEN ATTEMPTED TO RECHARGE NO RESPONSE AT ALL NOT EVEN A CALL YOUR DOCTOR SYMBOL FROM THE INS IT WAS CONCLUDED THAT THE INS WOULD NOT RESPOND THE REP REPORTED THAT THERE WAS NO FALL ETC THAT LED TO THE ISSUE JUST THAT THE PATIENT STOPPED USING AND RECHARGING THE DEVICE BECAUSE THE THERAPY FAILED THE REP REPORTED THAT THEY WOULD BE AT THE PATIENTS NEXT FOLLOW UP APPOINTMENT IN OCTOBER TO DISCUSS THE POSSIBILITY OF REVISIONREPLACEMENT THE ISSUE WAS NOT RESOLVED AT THE TIME OF THE REPORT NO FURTHER COMPLICATIONS WERE REPORTED | Insufficient Information | Device problem code
6912508 | 88235977 | N | B FAILURE MODE WAS REPLICATED BY SIMULATING SURGEONS TECHNIQUE AND DECOUPLING OF THE ADAPTER FROM THE CANNULA WAS ACHIEVED AND COULD BE REPEATED ROOT CAUSE WAS IDENTIFIED BY PLACING A CLAMP ON THE ADAPTER SLEEVE TO PREVENT INCIDENTAL MOTION ADDITION OF THE CLAMP PREVENTED DECOUPLING IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | Decoupling | Device problem code
6912508 | 88235977 | N | B FAILURE MODE WAS REPLICATED BY SIMULATING SURGEONS TECHNIQUE AND DECOUPLING OF THE ADAPTER FROM THE CANNULA WAS ACHIEVED AND COULD BE REPEATED ROOT CAUSE WAS IDENTIFIED BY PLACING A CLAMP ON THE ADAPTER SLEEVE TO PREVENT INCIDENTAL MOTION ADDITION OF THE CLAMP PREVENTED DECOUPLING IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | No Device Problem Found | Medical Evaluation result code
6892326 | 88241682 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED SEVERAL CONTROLLER FAULT ALARMS WITH FAULT CODES INDICATING SYSUICAPSFAIL ALARMS WITH THIS CODE ARE INDICATIVE OF A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO A CONTROLLER FAULT ALARM WITH FAULT CODES INDICATING SYSUICAPSFAIL WHILE PS WAS ACTIVE DURING BENCH TESTING HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device problem code
6892326 | 88241682 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED SEVERAL CONTROLLER FAULT ALARMS WITH FAULT CODES INDICATING SYSUICAPSFAIL ALARMS WITH THIS CODE ARE INDICATIVE OF A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO A CONTROLLER FAULT ALARM WITH FAULT CODES INDICATING SYSUICAPSFAIL WHILE PS WAS ACTIVE DURING BENCH TESTING HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6892330 | 88242526 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM WITH FAULT CODES INDICATING SYSUICAPSFAIL ALARMS WITH THIS CODE ARE INDICATIVE OF A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA AND ITS RELEASE BY FDA DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OR THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER THEIR EMPLOYEES OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device Problem Code
6892330 | 88242526 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM WITH FAULT CODES INDICATING SYSUICAPSFAIL ALARMS WITH THIS CODE ARE INDICATIVE OF A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH APS CIRCUIT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A LEAKY DIODE ON THE AUTOMATIC POWER SWITCH OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA AND ITS RELEASE BY FDA DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OR THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER THEIR EMPLOYEES OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6907410 | 88245867 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS DUE TO ELEVATED BG LEVELS AND BLOOD IN CANNULA PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6907410 | 88245867 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS DUE TO ELEVATED BG LEVELS AND BLOOD IN CANNULA PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6916031 | 88246026 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM OF TYPE SYSUICAPSFAIL ON B THIS TYPE OF CONTROLLER FAULT ALARM IS INDICATIVE OF A LEAKY DIODE ON THE APS CIRCUIT ADDITIONAL REVIEW REVEALED A VAD STOP ALARM MOST LIKELY DUE TO A DRIVELINE DISCONNECTION AND MULTIPLE POWER SWITCHING EVENTS THESE OBSERVATION ARE CONSIDERED NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE PREMATURE SWITCHING OBSERVED CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES AN INTERNAL INVESTIGATION IS CURRENTLY INVESTIGATING COMMUNICATION ISSUES ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AT THE BENCH THE MANUFACTURER HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE LEAKY DIODES IN THE APS CIRCUIT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device Problem Code
6916031 | 88246026 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM OF TYPE SYSUICAPSFAIL ON B THIS TYPE OF CONTROLLER FAULT ALARM IS INDICATIVE OF A LEAKY DIODE ON THE APS CIRCUIT ADDITIONAL REVIEW REVEALED A VAD STOP ALARM MOST LIKELY DUE TO A DRIVELINE DISCONNECTION AND MULTIPLE POWER SWITCHING EVENTS THESE OBSERVATION ARE CONSIDERED NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE PREMATURE SWITCHING OBSERVED CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES AN INTERNAL INVESTIGATION IS CURRENTLY INVESTIGATING COMMUNICATION ISSUES ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AT THE BENCH THE MANUFACTURER HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE LEAKY DIODES IN THE APS CIRCUIT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communication or Transmission Problem | Device problem code
6916031 | 88246026 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM OF TYPE SYSUICAPSFAIL ON B THIS TYPE OF CONTROLLER FAULT ALARM IS INDICATIVE OF A LEAKY DIODE ON THE APS CIRCUIT ADDITIONAL REVIEW REVEALED A VAD STOP ALARM MOST LIKELY DUE TO A DRIVELINE DISCONNECTION AND MULTIPLE POWER SWITCHING EVENTS THESE OBSERVATION ARE CONSIDERED NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE PREMATURE SWITCHING OBSERVED CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES AN INTERNAL INVESTIGATION IS CURRENTLY INVESTIGATING COMMUNICATION ISSUES ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AT THE BENCH THE MANUFACTURER HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE LEAKY DIODES IN THE APS CIRCUIT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6916031 | 88246026 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM OF TYPE SYSUICAPSFAIL ON B THIS TYPE OF CONTROLLER FAULT ALARM IS INDICATIVE OF A LEAKY DIODE ON THE APS CIRCUIT ADDITIONAL REVIEW REVEALED A VAD STOP ALARM MOST LIKELY DUE TO A DRIVELINE DISCONNECTION AND MULTIPLE POWER SWITCHING EVENTS THESE OBSERVATION ARE CONSIDERED NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE PREMATURE SWITCHING OBSERVED CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES AN INTERNAL INVESTIGATION IS CURRENTLY INVESTIGATING COMMUNICATION ISSUES ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AT THE BENCH THE MANUFACTURER HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE LEAKY DIODES IN THE APS CIRCUIT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6911118 | 88247125 | N | THE FULL NUMBER FOR THE PRODUCT IN QUESTION IS B IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION AND DETERMINED THAT THE INSTRUMENT TEST WELL IMAGE IN QUESTION WAS VISUALLY POSITIVE THE PRODUCT HAD ALREADY EXPIRED BEFORE RETENTION TESTING COULD BE DONE SO THE IMMUCOR LABORATORY PERFORMED A REVIEW OF A DHR WHICH SHOWED THAT ALL INFORMATION WAS ACCEPTABLE DURING THE MANUFACTURING PROCESS FOR THE LOT NUMBER IN QUESTION | Connection Problem | Device problem code
6911118 | 88247125 | N | THE FULL NUMBER FOR THE PRODUCT IN QUESTION IS B IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION AND DETERMINED THAT THE INSTRUMENT TEST WELL IMAGE IN QUESTION WAS VISUALLY POSITIVE THE PRODUCT HAD ALREADY EXPIRED BEFORE RETENTION TESTING COULD BE DONE SO THE IMMUCOR LABORATORY PERFORMED A REVIEW OF A DHR WHICH SHOWED THAT ALL INFORMATION WAS ACCEPTABLE DURING THE MANUFACTURING PROCESS FOR THE LOT NUMBER IN QUESTION | No Device Problem Found | Medical Evaluation result code
6897298 | 88249191 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Apparent Adverse Event | Device problem code
6897298 | 88249191 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Device Problem Found | Medical Evaluation result code
6900901 | 88250851 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6900901 | 88250851 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6909081 | 88251355 | N | THE SAMPLE WAS RECEIVED FROM THE COMPLAINANT AND A SMALL PIECE OF METAL WAS OBSERVED PARTIALLY IMBEDDED IN THE RUBBER SEAL OF THE PLUNGER CAUSING DISCOLORATION OF THE SALINE FURTHER EVALUATION IS BEING CONDUCTED ON THE CONTENTS OF THE SYRINGE | No Device Problem Found | Medical Evaluation result code
6884419 | 88253831 | N | INTUITIVE SURGICAL INC ISI RECEIVED THIS DEVICE INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INSTALLED THIS UNIT INTO THE PRINTED CIRCUIT ASSEMBLY PCA TEST SYSTEM AND IT FAILED TO POWER ON THE CORE OF THE VISION SIDE CART VSC FAILURE ANALYSIS APPLIED POWER TO THE POWER TRAY AT THE TEST BENCH AND FOUND THAT POWER SUPPLY A WAS DEAD THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTIONED WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6884419 | 88253831 | N | INTUITIVE SURGICAL INC ISI RECEIVED THIS DEVICE INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INSTALLED THIS UNIT INTO THE PRINTED CIRCUIT ASSEMBLY PCA TEST SYSTEM AND IT FAILED TO POWER ON THE CORE OF THE VISION SIDE CART VSC FAILURE ANALYSIS APPLIED POWER TO THE POWER TRAY AT THE TEST BENCH AND FOUND THAT POWER SUPPLY A WAS DEAD THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTIONED WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Power Source Problem Identified | Medical Evaluation result code
6916059 | 88253964 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED OCCURRENCES OF PREMATURE POWER SWITCHING EVENTS PRIOR TO BATTERIES REACHING THE B RSOC THRESHOLD ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER SERIAL NUMBER B DID NOT LOG ANY ALARMS FOR THE BATTERY B WITHIN THE ANALYZED PERIOD HOWEVER THE LOG FILES REVEALED OCCURRENCES OF COMMUNICATION ERRORS INVOLVING B AND MOMENTARY DISCONNECTS A MOMENTARY DISCONNECTION MAY CAUSE A BEEP AFTER A POWER SOURCE REGAINS CONNECTION THE FOLLOWING BATTERIES PARTICIPATED DURING SAID EVENTS B HEARTWARE IS INVESTIGATING THE ROOT CAUSE OF DISCONNECTIONS BETWEEN SMR UPGRADED CONTROLLER AND BATTERIES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6916059 | 88253964 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED OCCURRENCES OF PREMATURE POWER SWITCHING EVENTS PRIOR TO BATTERIES REACHING THE B RSOC THRESHOLD ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER SERIAL NUMBER B DID NOT LOG ANY ALARMS FOR THE BATTERY B WITHIN THE ANALYZED PERIOD HOWEVER THE LOG FILES REVEALED OCCURRENCES OF COMMUNICATION ERRORS INVOLVING B AND MOMENTARY DISCONNECTS A MOMENTARY DISCONNECTION MAY CAUSE A BEEP AFTER A POWER SOURCE REGAINS CONNECTION THE FOLLOWING BATTERIES PARTICIPATED DURING SAID EVENTS B HEARTWARE IS INVESTIGATING THE ROOT CAUSE OF DISCONNECTIONS BETWEEN SMR UPGRADED CONTROLLER AND BATTERIES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Material Twisted / Bent | Device problem code
6916059 | 88253964 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED OCCURRENCES OF PREMATURE POWER SWITCHING EVENTS PRIOR TO BATTERIES REACHING THE B RSOC THRESHOLD ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER SERIAL NUMBER B DID NOT LOG ANY ALARMS FOR THE BATTERY B WITHIN THE ANALYZED PERIOD HOWEVER THE LOG FILES REVEALED OCCURRENCES OF COMMUNICATION ERRORS INVOLVING B AND MOMENTARY DISCONNECTS A MOMENTARY DISCONNECTION MAY CAUSE A BEEP AFTER A POWER SOURCE REGAINS CONNECTION THE FOLLOWING BATTERIES PARTICIPATED DURING SAID EVENTS B HEARTWARE IS INVESTIGATING THE ROOT CAUSE OF DISCONNECTIONS BETWEEN SMR UPGRADED CONTROLLER AND BATTERIES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6916059 | 88253964 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED OCCURRENCES OF PREMATURE POWER SWITCHING EVENTS PRIOR TO BATTERIES REACHING THE B RSOC THRESHOLD ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE CONTROLLER SERIAL NUMBER B DID NOT LOG ANY ALARMS FOR THE BATTERY B WITHIN THE ANALYZED PERIOD HOWEVER THE LOG FILES REVEALED OCCURRENCES OF COMMUNICATION ERRORS INVOLVING B AND MOMENTARY DISCONNECTS A MOMENTARY DISCONNECTION MAY CAUSE A BEEP AFTER A POWER SOURCE REGAINS CONNECTION THE FOLLOWING BATTERIES PARTICIPATED DURING SAID EVENTS B HEARTWARE IS INVESTIGATING THE ROOT CAUSE OF DISCONNECTIONS BETWEEN SMR UPGRADED CONTROLLER AND BATTERIES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Non-Functional Defect | Medical Evaluation result code
6896771 | 88254043 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6896771 | 88254043 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6916554 | 88254470 | N | LOW IMPEDANCE ON LEAD MODEL T SERIAL B WAS PREVIOUSLY REPORTED ON MDR | Impedance Problem | Device problem code
6916554 | 88254470 | N | LOW IMPEDANCE ON LEAD MODEL T SERIAL B WAS PREVIOUSLY REPORTED ON MDR | Low impedance | Device Problem Code
6916554 | 88254470 | N | LOW IMPEDANCE ON LEAD MODEL T SERIAL B WAS PREVIOUSLY REPORTED ON MDR | Power Source Problem Identified | Medical Evaluation result code
6882587 | 88257084 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT CAN YOU PLEASE CLARIFY THE EVENT WERE THE STAPLES THAT LOOSENED SPONTANEOUSLY DELIVERED INTO TISSUE UNFORMED MALFORMED ETC OR DID THIS ISSUE OCCUR PRIOR TO THE SURGEON USING THE DEVICE EG STAPLES PROTRUDING FROM DEVICE PRIOR TO USE | Material Deformation | Device Problem Code
6882587 | 88257084 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT CAN YOU PLEASE CLARIFY THE EVENT WERE THE STAPLES THAT LOOSENED SPONTANEOUSLY DELIVERED INTO TISSUE UNFORMED MALFORMED ETC OR DID THIS ISSUE OCCUR PRIOR TO THE SURGEON USING THE DEVICE EG STAPLES PROTRUDING FROM DEVICE PRIOR TO USE | Stress Problem Identified | Medical Evaluation result code
6914297 | 88260879 | N | THE RETURNED DEVICE WAS EVALUATED DURING EVALUATION THE UNIT DID NOT POWER ON USING BATTERY AND WHEN USING AC POWER THE UNIT POWERED ON ONCE FOR A SECOND AND TURNED OFF THE UNIT WAS NOT ABLE TO POWER ON AFTER THE FIRST POWERING ON USING AC THE DEVICE WAS NOT ABLE TO OBTAIN ANY READINGS INTERNAL INSPECTION SHOWED THE BATTERY HAS A LOW VOLTAGE AND THE ACDC MODULE ON THE SYSTEM BOARD IS NOT FIRMLY ATTACHED TO THE SYSTEM CIRCUIT BOARD CAUSING THE UNIT TO NOT CHARGE AND TO POWER ON | Failure to Power Up | Device Problem Code
6914297 | 88260879 | N | THE RETURNED DEVICE WAS EVALUATED DURING EVALUATION THE UNIT DID NOT POWER ON USING BATTERY AND WHEN USING AC POWER THE UNIT POWERED ON ONCE FOR A SECOND AND TURNED OFF THE UNIT WAS NOT ABLE TO POWER ON AFTER THE FIRST POWERING ON USING AC THE DEVICE WAS NOT ABLE TO OBTAIN ANY READINGS INTERNAL INSPECTION SHOWED THE BATTERY HAS A LOW VOLTAGE AND THE ACDC MODULE ON THE SYSTEM BOARD IS NOT FIRMLY ATTACHED TO THE SYSTEM CIRCUIT BOARD CAUSING THE UNIT TO NOT CHARGE AND TO POWER ON | Charging Problem | Device Problem Code
6914297 | 88260879 | N | THE RETURNED DEVICE WAS EVALUATED DURING EVALUATION THE UNIT DID NOT POWER ON USING BATTERY AND WHEN USING AC POWER THE UNIT POWERED ON ONCE FOR A SECOND AND TURNED OFF THE UNIT WAS NOT ABLE TO POWER ON AFTER THE FIRST POWERING ON USING AC THE DEVICE WAS NOT ABLE TO OBTAIN ANY READINGS INTERNAL INSPECTION SHOWED THE BATTERY HAS A LOW VOLTAGE AND THE ACDC MODULE ON THE SYSTEM BOARD IS NOT FIRMLY ATTACHED TO THE SYSTEM CIRCUIT BOARD CAUSING THE UNIT TO NOT CHARGE AND TO POWER ON | Power Source Problem Identified | Medical Evaluation result code
6918967 | 88369661 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6918985 | 88369743 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Material Discolored | Device problem code
6918985 | 88369743 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Non-Functional Defect | Medical Evaluation result code
6918471 | 88369782 | D | FOLLOWING THE BATTERY PERFORMANCE ALERT BPA ADVISORY A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED THE PATIENT WAS ASYMPTOMATIC PREPROCEDURE AND WAS DISCHARGED THE NEXT DAY POSTPROCEDURE | Insufficient Information | Device problem code
6918471 | 88369782 | D | FOLLOWING THE BATTERY PERFORMANCE ALERT BPA ADVISORY A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED THE PATIENT WAS ASYMPTOMATIC PREPROCEDURE AND WAS DISCHARGED THE NEXT DAY POSTPROCEDURE | No Consequences Or Impact To Patient | Patient problem code
6916790 | 88378805 | D | INFORMATION WAS RECEIVED FROM A PATIENT REGARDING THEIR IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR PATIENT SAID THEY ARE HARDLY URINATING AT ALL AND THEY WANTED TO INCREASE STIMULATION BUT COULDNT BECAUSE OF THE PATIENT PROGRAMMER PP ISSUE THEY WERE URINATING REGULARLY THE PATIENT MENTIONED BOTH LEGS WERE BLOWN OUT AND THEY NEED TO WATCH THEIR SALT INTAKE THEY NEED TO EAT MORE HERBS TO MAKE THEIR BODY WORK PROPERLY AS A RESULT OF THE EVENT THE PATIENT WILL CHANGE THEIR BATTERIES IN THEIR PP TO SEE IF THEY CAN INCREASE STIMULATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6916790 | 88378805 | D | INFORMATION WAS RECEIVED FROM A PATIENT REGARDING THEIR IMPLANTABLE NEUROSTIMULATOR INS FOR URINARY DYSFUNCTIONSACRAL NERVE STIMULATION AND GASTROINTESTINALPELVIC FLOOR PATIENT SAID THEY ARE HARDLY URINATING AT ALL AND THEY WANTED TO INCREASE STIMULATION BUT COULDNT BECAUSE OF THE PATIENT PROGRAMMER PP ISSUE THEY WERE URINATING REGULARLY THE PATIENT MENTIONED BOTH LEGS WERE BLOWN OUT AND THEY NEED TO WATCH THEIR SALT INTAKE THEY NEED TO EAT MORE HERBS TO MAKE THEIR BODY WORK PROPERLY AS A RESULT OF THE EVENT THE PATIENT WILL CHANGE THEIR BATTERIES IN THEIR PP TO SEE IF THEY CAN INCREASE STIMULATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6917602 | 88379758 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS SEPSIS IS ASSOCIATED WITH HOMEOSTATIC ABNORMALITIES RANGING FROM ISOLATED THROMBOCYTOPENIA ANDOR SUBCLINICAL ACTIVATION OF BLOOD COAGULATION HYPERCOAGULABILITY THOUGH THE ROOT CAUSE OF THE REPORTED EVENT CANNOT BE CONCLUSIVELY DETERMINED HOWEVER THERE ARE PATIENT PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Coagulation in Device or device Ingredient | Device problem code
6919615 | 88380478 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE STYLET INSERTION TEST RESULT WAS FAILED DUE TO DISTORTION OF DISTAL CONDUCTOR WITHIN IS CONNECTOR PIN VISUAL INSPECTION FOUND THAT THE OUTER INSULATION BREACHED IN VIVO CLAVICLE RIB CRUSHED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6919615 | 88380478 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE STYLET INSERTION TEST RESULT WAS FAILED DUE TO DISTORTION OF DISTAL CONDUCTOR WITHIN IS CONNECTOR PIN VISUAL INSPECTION FOUND THAT THE OUTER INSULATION BREACHED IN VIVO CLAVICLE RIB CRUSHED | Collapse | Device problem code
6919615 | 88380478 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE STYLET INSERTION TEST RESULT WAS FAILED DUE TO DISTORTION OF DISTAL CONDUCTOR WITHIN IS CONNECTOR PIN VISUAL INSPECTION FOUND THAT THE OUTER INSULATION BREACHED IN VIVO CLAVICLE RIB CRUSHED | Break | Device problem code
6919615 | 88380478 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE STYLET INSERTION TEST RESULT WAS FAILED DUE TO DISTORTION OF DISTAL CONDUCTOR WITHIN IS CONNECTOR PIN VISUAL INSPECTION FOUND THAT THE OUTER INSULATION BREACHED IN VIVO CLAVICLE RIB CRUSHED | Non-Functional Defect | Medical Evaluation result code
6918841 | 88381505 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6918033 | 88383321 | N | FACILITY NAME B | Insufficient Information | Device Problem Code
6921475 | 88383564 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED PAIN AND BURNING SENSATION IN THE BACK UP THE SPINE WHEN STIMULATION WAS ON THE PATIENT WILL UNDERGO A REVISION PROCEDURE | Pain | Patient problem code
6921475 | 88383564 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED PAIN AND BURNING SENSATION IN THE BACK UP THE SPINE WHEN STIMULATION WAS ON THE PATIENT WILL UNDERGO A REVISION PROCEDURE | Burning Sensation | Patient problem code
6921475 | 88383564 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED PAIN AND BURNING SENSATION IN THE BACK UP THE SPINE WHEN STIMULATION WAS ON THE PATIENT WILL UNDERGO A REVISION PROCEDURE | Insufficient Information | Device problem code
6917322 | 88384859 | N | NO INFORMATION ABOUT RETURN OF DEVICE | Insufficient Information | Device problem code
6917943 | 88385395 | N | FACILITY NAME B | Insufficient Information | Device Problem Code
6918839 | 88388053 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6917517 | 88390524 | N | THE OF LYMPHOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Calcified | Device problem code
6917517 | 88390524 | N | THE OF LYMPHOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Deflation Problem | Device problem code
6917517 | 88390524 | N | THE OF LYMPHOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Non-Functional Defect | Medical Evaluation result code
6919577 | 88392082 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED A WARM SALINE AND GAUZE TO PROTECT THE PATIENTS SKIN | Leak / Splash | Device problem code
6919577 | 88392082 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED A WARM SALINE AND GAUZE TO PROTECT THE PATIENTS SKIN | Component Incompatible | Device problem code
6919577 | 88392082 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED A WARM SALINE AND GAUZE TO PROTECT THE PATIENTS SKIN | Operational Problem Identified | Medical Evaluation result code
6919244 | 88393252 | D | NO EVIDENCE OF MALFUNCTION BUT THE PATIENT DID EXPERIENCE AN INFECTION I THE IPG POCKET UPDATE WAS RECEIVED PATIENT HAD THEIR SYSTEM EXPLANTED AND PATIENTS INFECTION RESOLVED WITHOUT ISSUE | Infection | Patient problem code
6919244 | 88393252 | D | NO EVIDENCE OF MALFUNCTION BUT THE PATIENT DID EXPERIENCE AN INFECTION I THE IPG POCKET UPDATE WAS RECEIVED PATIENT HAD THEIR SYSTEM EXPLANTED AND PATIENTS INFECTION RESOLVED WITHOUT ISSUE | Insufficient Information | Device Problem Code
6921106 | 88412220 | N | ANALYSIS WAS ABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE STYLUS AND SCREEN WERE NOT CALIBRATED AND THE CONNECTION TO THE XY BOARD WAS RESEATED TO RESOLVE AND THE STYLUS WAS THEN ABLE TO BE CALIBRATED THE SOFTWARE WAS RELOADED AND UPDATED AND THE STYLUS WAS CHECKED TO ENSURE THAT IT STAYED CALIBRATED THE PROGRAMMER PASSED ALL CONSOLE AND SYSTEMS TESTS | Connection Problem | Device problem code
6921106 | 88412220 | N | ANALYSIS WAS ABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE STYLUS AND SCREEN WERE NOT CALIBRATED AND THE CONNECTION TO THE XY BOARD WAS RESEATED TO RESOLVE AND THE STYLUS WAS THEN ABLE TO BE CALIBRATED THE SOFTWARE WAS RELOADED AND UPDATED AND THE STYLUS WAS CHECKED TO ENSURE THAT IT STAYED CALIBRATED THE PROGRAMMER PASSED ALL CONSOLE AND SYSTEMS TESTS | No Device Problem Found | Medical Evaluation result code
6918260 | 88415191 | N | THE CUSTOMER REPORTED THAT THE NEEDLE HAD SEPARATED FROM THE MICROTIP THE COMPANY IS STILL AWAITING ON THE RETURN OF THE DEVICE FOR EVALUATION ONCE THE DEVICE HAS BEEN RECEIVED AND EVALUATED A SUPPLEMENTAL MDR WILL BE FILED WITH THE RESULTS OF THE EVALUATION | No Device Problem Found | Medical Evaluation result code
6917590 | 88417496 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAD BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAD FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6917590 | 88417496 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAD BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAD FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6880626 | 88420935 | N | THE CUSTOMER REPORTED COMPLAINT WAS CONFIRMED DURING THE EVENT LOG REVIEW AND INITIAL FUNCTIONAL TESTING THE PROPYLENE GLYCOL THAT WAS NOTED IN THE SENSOR HOLES WERE CLEANED TO REMEDY THE FAULT THE UNIT HAS PASSED ALL THE FINAL FUNCTIONAL TESTING AND IS WORKING ACCORDING TO SPECIFICATIONS A REVIEW OF THE EVENT LOG WAS PERFORMED AND MULTIPLE ERROR MESSAGE MID AIR TRAP ASSEMBLY MID BC CRIT DATA AND MID SOFTWARE WERE NOTED ON B BUT NOT ON THE REPORTED EVENT DATE OF B HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINT REPORTED FOR THERMOGARD XP IVTM SYSTEM SN B | Device Displays Incorrect Message | Device problem code
6880626 | 88420935 | N | THE CUSTOMER REPORTED COMPLAINT WAS CONFIRMED DURING THE EVENT LOG REVIEW AND INITIAL FUNCTIONAL TESTING THE PROPYLENE GLYCOL THAT WAS NOTED IN THE SENSOR HOLES WERE CLEANED TO REMEDY THE FAULT THE UNIT HAS PASSED ALL THE FINAL FUNCTIONAL TESTING AND IS WORKING ACCORDING TO SPECIFICATIONS A REVIEW OF THE EVENT LOG WAS PERFORMED AND MULTIPLE ERROR MESSAGE MID AIR TRAP ASSEMBLY MID BC CRIT DATA AND MID SOFTWARE WERE NOTED ON B BUT NOT ON THE REPORTED EVENT DATE OF B HISTORICAL COMPLAINTS WERE REVIEWED FOR SERVICE INFORMATION RELATED TO THE REPORTED COMPLAINT AND THERE WAS NO SIMILAR COMPLAINT REPORTED FOR THERMOGARD XP IVTM SYSTEM SN B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6918401 | 88420943 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6918401 | 88420943 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6921666 | 88425028 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ON THEIR BACK AND SHOULDERS AFTER THEY HAD STARTED WEARING THE LIFEVEST THE PATIENTS NURSE REPORTED THAT THE LIFEVEST WAS REMOVED FROM THE PATIENT TO ADDRESS THE RASH AFTER THE LIFEVEST WAS REMOVED IT WAS REPORTED THAT THE RASH HAD IMPROVED THERE WAS NO ALLEGATION THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE PATIENTS REPORTED SKIN IRRITATION | Skin Irritation | Patient problem code
6921666 | 88425028 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ON THEIR BACK AND SHOULDERS AFTER THEY HAD STARTED WEARING THE LIFEVEST THE PATIENTS NURSE REPORTED THAT THE LIFEVEST WAS REMOVED FROM THE PATIENT TO ADDRESS THE RASH AFTER THE LIFEVEST WAS REMOVED IT WAS REPORTED THAT THE RASH HAD IMPROVED THERE WAS NO ALLEGATION THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE PATIENTS REPORTED SKIN IRRITATION | Adverse Event Without Identified Device or Use Problem | Device problem code
6921666 | 88425028 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ON THEIR BACK AND SHOULDERS AFTER THEY HAD STARTED WEARING THE LIFEVEST THE PATIENTS NURSE REPORTED THAT THE LIFEVEST WAS REMOVED FROM THE PATIENT TO ADDRESS THE RASH AFTER THE LIFEVEST WAS REMOVED IT WAS REPORTED THAT THE RASH HAD IMPROVED THERE WAS NO ALLEGATION THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE PATIENTS REPORTED SKIN IRRITATION | Rash | Patient problem code
6921666 | 88425028 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ON THEIR BACK AND SHOULDERS AFTER THEY HAD STARTED WEARING THE LIFEVEST THE PATIENTS NURSE REPORTED THAT THE LIFEVEST WAS REMOVED FROM THE PATIENT TO ADDRESS THE RASH AFTER THE LIFEVEST WAS REMOVED IT WAS REPORTED THAT THE RASH HAD IMPROVED THERE WAS NO ALLEGATION THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE PATIENTS REPORTED SKIN IRRITATION | Itching | Patient problem code
6904463 | 88428976 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B NINE DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6904463 | 88428976 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN B POSITION ON B NINE DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6910473 | 88447907 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH VISUALLY APPEARED TO SHOW A FAINT RING AROUND THE CLEARLY VISIBLE AND SHARP RED BLOOD CELL BUTTON THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED | Connection Problem | Device problem code
6910473 | 88447907 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH VISUALLY APPEARED TO SHOW A FAINT RING AROUND THE CLEARLY VISIBLE AND SHARP RED BLOOD CELL BUTTON THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED | No Device Problem Found | Medical Evaluation result code
6922205 | 88453900 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AFTER WEARING THE LIFEVEST THE RASH WAS DESCRIBED AS HIVES THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO TAKE AN ALLEGRA TABLET FOR THE REPORTED SKIN IRRITATION DURING A FOLLOW UP CALL IT WAS REPORTED THAT THE RASH WAS MUCH BETTER THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION ASSOCIATED WITH THE REPORTED SKIN IRRITATION | Rash | Patient problem code
6922205 | 88453900 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AFTER WEARING THE LIFEVEST THE RASH WAS DESCRIBED AS HIVES THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO TAKE AN ALLEGRA TABLET FOR THE REPORTED SKIN IRRITATION DURING A FOLLOW UP CALL IT WAS REPORTED THAT THE RASH WAS MUCH BETTER THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION ASSOCIATED WITH THE REPORTED SKIN IRRITATION | Skin Irritation | Patient problem code
6922205 | 88453900 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AFTER WEARING THE LIFEVEST THE RASH WAS DESCRIBED AS HIVES THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO TAKE AN ALLEGRA TABLET FOR THE REPORTED SKIN IRRITATION DURING A FOLLOW UP CALL IT WAS REPORTED THAT THE RASH WAS MUCH BETTER THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION ASSOCIATED WITH THE REPORTED SKIN IRRITATION | Urticaria | Patient problem code
6922205 | 88453900 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AFTER WEARING THE LIFEVEST THE RASH WAS DESCRIBED AS HIVES THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO TAKE AN ALLEGRA TABLET FOR THE REPORTED SKIN IRRITATION DURING A FOLLOW UP CALL IT WAS REPORTED THAT THE RASH WAS MUCH BETTER THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION ASSOCIATED WITH THE REPORTED SKIN IRRITATION | Adverse Event Without Identified Device or Use Problem | Device problem code
6922205 | 88453900 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RED ITCHY RASH ALL OVER HIS BODY AFTER WEARING THE LIFEVEST THE RASH WAS DESCRIBED AS HIVES THE PATIENTS PHYSICIAN ADVISED THE PATIENT TO TAKE AN ALLEGRA TABLET FOR THE REPORTED SKIN IRRITATION DURING A FOLLOW UP CALL IT WAS REPORTED THAT THE RASH WAS MUCH BETTER THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION ASSOCIATED WITH THE REPORTED SKIN IRRITATION | Itching | Patient problem code
6920878 | 88707732 | N | THIS EVENT IS RELATED TO ANOTHER EVENT REPORTED UNDER MFR NUMBER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6920878 | 88707732 | N | THIS EVENT IS RELATED TO ANOTHER EVENT REPORTED UNDER MFR NUMBER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6923781 | 88725155 | N | THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO | Break | Device problem code
6923781 | 88725155 | N | THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO | No Device Problem Found | Medical Evaluation result code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Infection | Patient problem code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Infection, Viral | Patient problem code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Inflammation | Patient problem code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Hyperglycemia | Patient Problem Code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Vomiting | Patient problem code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Adverse Event Without Identified Device or Use Problem | Device problem code
6926637 | 88731745 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS MGDL A BOLUS VIA THE PUMP WAS DELIVERED IN AN ATTEMPT TO ADDRESS THE BG LEVEL THE CUSTOMER HAD BEEN VOMITING AND DUE TO DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS ADMITTED TO THE HOSPITAL ON B THE CUSTOMER WAS TREATED WITH AN INSULIN INJECTION AND AN INTRAVENOUS INSULIN DRIP THE DIRECT CAUSE OF THE HIGH BG WAS NOT KNOWN WHILE IN THE HOSPITAL THE CUSTOMER WAS DIAGNOSED WITH A VIRAL INFECTION GALL BLADDER ISSUES AND AN INFLAMED PANCREAS A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER WAS DISCHARGED ON B WITH THE HIGH BG AND DKA RESOLVED THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE | Diabetic Ketoacidosis | Patient problem code
6919309 | 88790735 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DETAILED TRACE ANALYSIS OR PUMP HISTORY CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Electrical Power Problem | Device problem code
6919309 | 88790735 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DETAILED TRACE ANALYSIS OR PUMP HISTORY CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
6919309 | 88790735 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DETAILED TRACE ANALYSIS OR PUMP HISTORY CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Disconnection | Device problem code
6919309 | 88790735 | N | THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST DETAILED TRACE ANALYSIS OR PUMP HISTORY CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Power Source Problem Identified | Medical Evaluation result code
6931330 | 88824930 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED AUDIO ISSUES IS BEING REPORTED UNDER MFR | Output Problem | Device problem code
6931330 | 88824930 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED AUDIO ISSUES IS BEING REPORTED UNDER MFR | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6931340 | 88825585 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED NO AUDIO OUTPUT IS BEING REPORTED UNDER MFR | No Audible Prompt / Feedback | Device problem code
6931340 | 88825585 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED NO AUDIO OUTPUT IS BEING REPORTED UNDER MFR | No Device Problem Found | Medical Evaluation result code
6927134 | 88834225 | N | B PATIENT UNDERWENT RNS SYSTEM IMPLANT ALONG WITH A CONCURRENT RIGHT MEDIAL FRONTAL RESECTION IMMEDIATELY FOLLOWING DAYS OF INVASIVE INTRACRANIAL MONITORING THE PATIENT WAS IMPLANTED WITH THE RNS NEUROSTIMULATOR AND FOUR CORTICAL STRIP LEADS PORT CL SN B RIGHT FRONTAL NOT CONNECTED CL SN B RIGHT FRONTAL NOT CONNECTED CL SN B RIGHT FRONTAL PORT CL SN B RIGHT FRONTAL AND THE PROCEDURE DURATION WAS APPROXIMATELY HOURS | Connection Problem | Device problem code
6927134 | 88834225 | N | B PATIENT UNDERWENT RNS SYSTEM IMPLANT ALONG WITH A CONCURRENT RIGHT MEDIAL FRONTAL RESECTION IMMEDIATELY FOLLOWING DAYS OF INVASIVE INTRACRANIAL MONITORING THE PATIENT WAS IMPLANTED WITH THE RNS NEUROSTIMULATOR AND FOUR CORTICAL STRIP LEADS PORT CL SN B RIGHT FRONTAL NOT CONNECTED CL SN B RIGHT FRONTAL NOT CONNECTED CL SN B RIGHT FRONTAL PORT CL SN B RIGHT FRONTAL AND THE PROCEDURE DURATION WAS APPROXIMATELY HOURS | Non-Functional Defect | Medical Evaluation result code
6932640 | 88834717 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Retraction Problem | Device problem code
6932640 | 88834717 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Device Dislodged or Dislocated | Device problem code
6932640 | 88834717 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Stress Problem Identified | Medical Evaluation result code
6934983 | 88842899 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY DRIED BODY FLUIDBLOOD IN LUMEN OF THE LEAD VISUAL INSPECTION NOTED FOREIGN MATERIAL LIKELY AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION LABORATORY ANALYSIS CONFIRMED THE ALLEGATION BY WIRE INSERTION DIFFICULTY DUE TO FOREIGN MATERIAL IN THE LEAD LUMEN | Device Contamination with Chemical or Other Material | Device problem code
6934983 | 88842899 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY DRIED BODY FLUIDBLOOD IN LUMEN OF THE LEAD VISUAL INSPECTION NOTED FOREIGN MATERIAL LIKELY AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION LABORATORY ANALYSIS CONFIRMED THE ALLEGATION BY WIRE INSERTION DIFFICULTY DUE TO FOREIGN MATERIAL IN THE LEAD LUMEN | Stress Problem Identified | Medical Evaluation result code
6925813 | 88874062 | N | PRODUCT EVENT B | Insufficient Information | Device Problem Code
6928722 | 88919879 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED NAUSEA AND BURNING SENSATION DUE TO AN ALLERGIC REACTION TO THEIR IPG THE PATIENT TESTED POSITIVE FOR ALLERGY TO TITANIUM THE PATIENT UNDERWENT AN IPG EXPLANT PROCEDURE | Nausea | Patient problem code
6928722 | 88919879 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED NAUSEA AND BURNING SENSATION DUE TO AN ALLERGIC REACTION TO THEIR IPG THE PATIENT TESTED POSITIVE FOR ALLERGY TO TITANIUM THE PATIENT UNDERWENT AN IPG EXPLANT PROCEDURE | Hypersensitivity | Patient problem code
6928722 | 88919879 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED NAUSEA AND BURNING SENSATION DUE TO AN ALLERGIC REACTION TO THEIR IPG THE PATIENT TESTED POSITIVE FOR ALLERGY TO TITANIUM THE PATIENT UNDERWENT AN IPG EXPLANT PROCEDURE | Insufficient Information | Device problem code
6928722 | 88919879 | D | A REPORT WAS RECEIVED THAT THE PATIENT EXPERIENCED NAUSEA AND BURNING SENSATION DUE TO AN ALLERGIC REACTION TO THEIR IPG THE PATIENT TESTED POSITIVE FOR ALLERGY TO TITANIUM THE PATIENT UNDERWENT AN IPG EXPLANT PROCEDURE | Burning Sensation | Patient problem code
6930501 | 88920907 | N | THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT MULTIPLE PATIENTS WERE NOTED IN THE ARTICLE HOWEVER A ONE TO ONE CORRELATION COULD NOT BE MADE WITH UNIQUE DEVICE SERIAL NUMBERS POSSIBLE MODELS COULD INCLUDE SINCE NO DEVICE ID WAS PROVIDED IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED REQUEST FOR ADDITIONAL INFORMATION WILL BE MADE AND UPON RECEIPT A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY REFERENCED ARTICLE INCIDENCE AND PREDICTORS OF CARDIAC PERFORATION AFTER PERMANENT PACEMAKER PLACEMENT HEART RHYTHM | No Device Problem Found | Medical Evaluation result code
6921744 | 88922963 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TARGET LIMB AND VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT C EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B PATIENT ID B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Apparent Adverse Event | Device problem code
6921744 | 88922963 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TARGET LIMB AND VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT C EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B PATIENT ID B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Device Problem Found | Medical Evaluation result code
6931050 | 88924050 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED COMPLETE LEAD RETURNED WITH HELIX RETRACTED DETAILED ANALYSIS CONFIRMED THAT LEAD PASSED RELEVANT TESTING STATED THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | Retraction Problem | Device problem code
6931050 | 88924050 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED COMPLETE LEAD RETURNED WITH HELIX RETRACTED DETAILED ANALYSIS CONFIRMED THAT LEAD PASSED RELEVANT TESTING STATED THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | No Device Problem Found | Medical Evaluation result code
6913546 | 88927072 | N | B BATCH UNK ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE CLARIFY HOW WAS THE BLEEDING CONTROLLED ON THE OOZING STAPLE LINES SURGEON USED DIATHERMY TO SPOT COAG WAS THERE ANY ISSUE AT ALL WITH STAPLE FORMATION UNFORMED MALFORMED ETC SOME WERE MALFORMED OR WAS THERE AN ISSUE WITH ANY STAPLES MISSING FROM THE STAPLE LINE INCOMPLETE STAPLE LINE NOT THAT THE SURGEON COULD SEE WAS THERE ANY CHANGE IN PROCEDURE OR PATIENT CARE AS A RESULT OF THE EVENT NO | Failure to Form Staple | Device problem code
6913546 | 88927072 | N | B BATCH UNK ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE CLARIFY HOW WAS THE BLEEDING CONTROLLED ON THE OOZING STAPLE LINES SURGEON USED DIATHERMY TO SPOT COAG WAS THERE ANY ISSUE AT ALL WITH STAPLE FORMATION UNFORMED MALFORMED ETC SOME WERE MALFORMED OR WAS THERE AN ISSUE WITH ANY STAPLES MISSING FROM THE STAPLE LINE INCOMPLETE STAPLE LINE NOT THAT THE SURGEON COULD SEE WAS THERE ANY CHANGE IN PROCEDURE OR PATIENT CARE AS A RESULT OF THE EVENT NO | Material Deformation | Device problem code
6913546 | 88927072 | N | B BATCH UNK ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE CLARIFY HOW WAS THE BLEEDING CONTROLLED ON THE OOZING STAPLE LINES SURGEON USED DIATHERMY TO SPOT COAG WAS THERE ANY ISSUE AT ALL WITH STAPLE FORMATION UNFORMED MALFORMED ETC SOME WERE MALFORMED OR WAS THERE AN ISSUE WITH ANY STAPLES MISSING FROM THE STAPLE LINE INCOMPLETE STAPLE LINE NOT THAT THE SURGEON COULD SEE WAS THERE ANY CHANGE IN PROCEDURE OR PATIENT CARE AS A RESULT OF THE EVENT NO | Component Missing | Device problem code
6913546 | 88927072 | N | B BATCH UNK ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE CLARIFY HOW WAS THE BLEEDING CONTROLLED ON THE OOZING STAPLE LINES SURGEON USED DIATHERMY TO SPOT COAG WAS THERE ANY ISSUE AT ALL WITH STAPLE FORMATION UNFORMED MALFORMED ETC SOME WERE MALFORMED OR WAS THERE AN ISSUE WITH ANY STAPLES MISSING FROM THE STAPLE LINE INCOMPLETE STAPLE LINE NOT THAT THE SURGEON COULD SEE WAS THERE ANY CHANGE IN PROCEDURE OR PATIENT CARE AS A RESULT OF THE EVENT NO | Missing or Inadequate Safety Measures | Medical Evaluation result code
6926831 | 88927842 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF FIRM NODULES SWELLING AND ERYTHEMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS INJECTION OF MORE THAN ML OF JUVÉDERM VOLBELLA XC INJECTABLE GEL HAS NOT BEEN STUDIED FOR LIP AUGMENTATION AND CORRECTION OF PERIORAL RHYTIDS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS MOST SUBJECTS REPORTED AN ISR AFTER TREATMENT WITH THE MOST COMMON BEING SWELLING TENDERNESS AND FIRMNESS THE MAJORITY OF ISRS WERE MILD TO MODERATE IN SEVERITY AND RESOLVED WITHIN DAYS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE CLINICAL STUDY SUBJECTS HAD LUMPSBUMPS OR SWELLING THAT OCCURRED WEEKS TO MONTHS AFTER TREATMENT ALL OF THESE AES WERE MILD OR MODERATE SWELLING WAS TREATED WITH ACETAMINOPHEN OR DOXYCYCLINE AND NO TREATMENT WAS GIVEN FOR THE LUMPSBUMPS ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE | Material Discolored | Device problem code
6926831 | 88927842 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF FIRM NODULES SWELLING AND ERYTHEMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS INJECTION OF MORE THAN ML OF JUVÉDERM VOLBELLA XC INJECTABLE GEL HAS NOT BEEN STUDIED FOR LIP AUGMENTATION AND CORRECTION OF PERIORAL RHYTIDS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS MOST SUBJECTS REPORTED AN ISR AFTER TREATMENT WITH THE MOST COMMON BEING SWELLING TENDERNESS AND FIRMNESS THE MAJORITY OF ISRS WERE MILD TO MODERATE IN SEVERITY AND RESOLVED WITHIN DAYS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE CLINICAL STUDY SUBJECTS HAD LUMPSBUMPS OR SWELLING THAT OCCURRED WEEKS TO MONTHS AFTER TREATMENT ALL OF THESE AES WERE MILD OR MODERATE SWELLING WAS TREATED WITH ACETAMINOPHEN OR DOXYCYCLINE AND NO TREATMENT WAS GIVEN FOR THE LUMPSBUMPS ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE | Non-Functional Defect | Medical Evaluation result code
6902926 | 88932859 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A POWER DISCONNECT ALARM WHEN TWO BATTERIES WERE CONNECTED TO THE CONTROLLER HOWEVER BATTERY WAS NOT RECOGNIZED BY CONTROLLER AND SHOWED A CAPACITY THE BATTERY SAW COLUMN HAD SPURIOUS VALUES WHICH INDICATED A COMMUNICATION PROBLEM BETWEEN THE CONTROLLER AND THE BATTERY HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6902926 | 88932859 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A POWER DISCONNECT ALARM WHEN TWO BATTERIES WERE CONNECTED TO THE CONTROLLER HOWEVER BATTERY WAS NOT RECOGNIZED BY CONTROLLER AND SHOWED A CAPACITY THE BATTERY SAW COLUMN HAD SPURIOUS VALUES WHICH INDICATED A COMMUNICATION PROBLEM BETWEEN THE CONTROLLER AND THE BATTERY HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communication or Transmission Problem | Device problem code
6902926 | 88932859 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A POWER DISCONNECT ALARM WHEN TWO BATTERIES WERE CONNECTED TO THE CONTROLLER HOWEVER BATTERY WAS NOT RECOGNIZED BY CONTROLLER AND SHOWED A CAPACITY THE BATTERY SAW COLUMN HAD SPURIOUS VALUES WHICH INDICATED A COMMUNICATION PROBLEM BETWEEN THE CONTROLLER AND THE BATTERY HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Compatibility Problem | Device problem code
6902926 | 88932859 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A POWER DISCONNECT ALARM WHEN TWO BATTERIES WERE CONNECTED TO THE CONTROLLER HOWEVER BATTERY WAS NOT RECOGNIZED BY CONTROLLER AND SHOWED A CAPACITY THE BATTERY SAW COLUMN HAD SPURIOUS VALUES WHICH INDICATED A COMMUNICATION PROBLEM BETWEEN THE CONTROLLER AND THE BATTERY HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6930519 | 88938200 | N | THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT MULTIPLE PATIENTS WERE NOTED IN THE ARTICLE HOWEVER A ONE TO ONE CORRELATION COULD NOT BE MADE WITH UNIQUE DEVICE SERIAL NUMBERS POSSIBLE MODELS COULD INCLUDE M SINCE NO DEVICE ID WAS PROVIDED IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED REQUEST FOR ADDITIONAL INFORMATION WILL BE MADE AND UPON RECEIPT A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY REFERENCED ARTICLE PREVALENCE OF CENTRAL VENOUS OCCLUSION IN PATIENTS WITH CHRONIC DEFIBRILLATOR LEADS AMERICAN HEART JOURNAL | Communications Problem Identified | Medical Evaluation result code
6930519 | 88938200 | N | THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT MULTIPLE PATIENTS WERE NOTED IN THE ARTICLE HOWEVER A ONE TO ONE CORRELATION COULD NOT BE MADE WITH UNIQUE DEVICE SERIAL NUMBERS POSSIBLE MODELS COULD INCLUDE M SINCE NO DEVICE ID WAS PROVIDED IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED REQUEST FOR ADDITIONAL INFORMATION WILL BE MADE AND UPON RECEIPT A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY REFERENCED ARTICLE PREVALENCE OF CENTRAL VENOUS OCCLUSION IN PATIENTS WITH CHRONIC DEFIBRILLATOR LEADS AMERICAN HEART JOURNAL | No Device Problem Found | Medical Evaluation result code
6932697 | 88940042 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Device Dislodged or Dislocated | Device problem code
6932697 | 88940042 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Retraction Problem | Device problem code
6932697 | 88940042 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY FRACTURES OR ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS RESISTANCE TESTING FOUND THE LEAD WAS ELECTRICALLY CONTINUOUS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED DISLODGMENT AS THE LEAD TIP AND HELIX APPEARED INTACT AND UNDAMAGED | Non-Functional Defect | Medical Evaluation result code
6939510 | 88941023 | N | THE REPORTER CONTACTED ABBOTT DIABETES CARE ALLEGING THE METER WILL NOT TURN ON THIS COMPLAINT IS BEING REPORTED BECAUSE THE REPORTED ISSUE WAS NOT RESOLVED WITH TROUBLESHOOTING THERE WAS NO REPORT OF DEATH SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT | Failure to Power Up | Device problem code
6931575 | 88941392 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS OR HIGH PACING THRESHOLDS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Retraction Problem | Device problem code
6931575 | 88941392 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS OR HIGH PACING THRESHOLDS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Activation, Positioning or Separation Problem | Device Problem Code
6931575 | 88941392 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS OR HIGH PACING THRESHOLDS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Difficult to Remove | Device Problem Code
6931575 | 88941392 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS OR HIGH PACING THRESHOLDS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | High Capture Threshold | Device problem code
6931575 | 88941392 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS OR HIGH PACING THRESHOLDS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Power Source Problem Identified | Medical Evaluation result code
6930128 | 88942933 | N | THIS ACUITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH DRIED BLOODBODY FLUID IN THE TIP REGION THE LEAD TIP APPEARED INTACT AND UNDAMAGED NO PRODUCT ABNORMALITIES WERE IDENTIFIED THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATION OF PERFORATION LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY ESTABLISH A RELATIONSHIP BETWEEN THIS PASSIVE FIXATION LEAD AND THE REPORTED PERFORATION | Material Perforation | Device problem code
6930128 | 88942933 | N | THIS ACUITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH DRIED BLOODBODY FLUID IN THE TIP REGION THE LEAD TIP APPEARED INTACT AND UNDAMAGED NO PRODUCT ABNORMALITIES WERE IDENTIFIED THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATION OF PERFORATION LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY ESTABLISH A RELATIONSHIP BETWEEN THIS PASSIVE FIXATION LEAD AND THE REPORTED PERFORATION | No Device Problem Found | Medical Evaluation result code
6927069 | 88949710 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE JUVÉDERM VOLBELLA XC GROUP AFTER REPEAT TREATMENT WITH JUVÉDERM VOLBELLA XC TREATMENTRELATED AES WERE REPORTED IN OF SUBJECTS INJECTION SITE MASS OCCURRED IN OF SUBJECTS TREATMENTRELATED AES OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS AND INJECTION SITE BRUISING EDEMA INDURATION AND PAIN TREATMENTRELATED AES AFTER REPEAT TREATMENT OCCURRED WITH LOWER INCIDENCE RATES SEVERITY AND DURATION COMPARED TO INITIALTOUCHUP TREATMENT THE FOLLOWING REPORTED ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLBELLA XC FOR LIP AUGMENTATION OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THESE ADVERSE EVENTS WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION INFECTION PARESTHESIA HERPES MIGRATION ANGIOEDEMA AND NECROSIS | Material Discolored | Device problem code
6927069 | 88949710 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE JUVÉDERM VOLBELLA XC GROUP AFTER REPEAT TREATMENT WITH JUVÉDERM VOLBELLA XC TREATMENTRELATED AES WERE REPORTED IN OF SUBJECTS INJECTION SITE MASS OCCURRED IN OF SUBJECTS TREATMENTRELATED AES OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS AND INJECTION SITE BRUISING EDEMA INDURATION AND PAIN TREATMENTRELATED AES AFTER REPEAT TREATMENT OCCURRED WITH LOWER INCIDENCE RATES SEVERITY AND DURATION COMPARED TO INITIALTOUCHUP TREATMENT THE FOLLOWING REPORTED ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLBELLA XC FOR LIP AUGMENTATION OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THESE ADVERSE EVENTS WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION INFECTION PARESTHESIA HERPES MIGRATION ANGIOEDEMA AND NECROSIS | Migration or Expulsion of Device | Device problem code
6927069 | 88949710 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE JUVÉDERM VOLBELLA XC GROUP AFTER REPEAT TREATMENT WITH JUVÉDERM VOLBELLA XC TREATMENTRELATED AES WERE REPORTED IN OF SUBJECTS INJECTION SITE MASS OCCURRED IN OF SUBJECTS TREATMENTRELATED AES OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS AND INJECTION SITE BRUISING EDEMA INDURATION AND PAIN TREATMENTRELATED AES AFTER REPEAT TREATMENT OCCURRED WITH LOWER INCIDENCE RATES SEVERITY AND DURATION COMPARED TO INITIALTOUCHUP TREATMENT THE FOLLOWING REPORTED ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLBELLA XC FOR LIP AUGMENTATION OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THESE ADVERSE EVENTS WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION INFECTION PARESTHESIA HERPES MIGRATION ANGIOEDEMA AND NECROSIS | Device Migration | Medical Evaluation result code
6931904 | 88954072 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY IT WAS CONFIRMED THAT TELEMETRY WAS UNABLE TO BE ESTABLISHED WITH THE DEVICE DETAILED ANALYSIS WAS PERFORMED AND IT WAS DETERMINED THAT THE ONE OF THE RESISTORS WERE MISSING THE RESISTER WAS FOUND WITHIN THE INSIDES OF THE PACEMAKER CASING LABORATORY ANALYSIS CONFIRMED THAT THE CLINICAL OBSERVATION WAS DIRECTLY RELATED TO A COMPONENT ANOMALY | Component Missing | Device problem code
6931904 | 88954072 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY IT WAS CONFIRMED THAT TELEMETRY WAS UNABLE TO BE ESTABLISHED WITH THE DEVICE DETAILED ANALYSIS WAS PERFORMED AND IT WAS DETERMINED THAT THE ONE OF THE RESISTORS WERE MISSING THE RESISTER WAS FOUND WITHIN THE INSIDES OF THE PACEMAKER CASING LABORATORY ANALYSIS CONFIRMED THAT THE CLINICAL OBSERVATION WAS DIRECTLY RELATED TO A COMPONENT ANOMALY | Telemetry Discrepancy | Device problem code
6931904 | 88954072 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY IT WAS CONFIRMED THAT TELEMETRY WAS UNABLE TO BE ESTABLISHED WITH THE DEVICE DETAILED ANALYSIS WAS PERFORMED AND IT WAS DETERMINED THAT THE ONE OF THE RESISTORS WERE MISSING THE RESISTER WAS FOUND WITHIN THE INSIDES OF THE PACEMAKER CASING LABORATORY ANALYSIS CONFIRMED THAT THE CLINICAL OBSERVATION WAS DIRECTLY RELATED TO A COMPONENT ANOMALY | Missing or Inadequate Safety Measures | Medical Evaluation result code
6926827 | 88954994 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF FIRM NODULES SWELLING AND ERYTHEMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS INJECTION OF MORE THAN ML OF JUVÉDERM VOLBELLA XC INJECTABLE GEL HAS NOT BEEN STUDIED FOR LIP AUGMENTATION AND CORRECTION OF PERIORAL RHYTIDS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS MOST SUBJECTS REPORTED AN ISR AFTER TREATMENT WITH THE MOST COMMON BEING SWELLING TENDERNESS AND FIRMNESS THE MAJORITY OF ISRS WERE MILD TO MODERATE IN SEVERITY AND RESOLVED WITHIN DAYS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE CLINICAL STUDY SUBJECTS HAD LUMPSBUMPS OR SWELLING THAT OCCURRED WEEKS TO MONTHS AFTER TREATMENT ALL OF THESE AES WERE MILD OR MODERATE SWELLING WAS TREATED WITH ACETAMINOPHEN OR DOXYCYCLINE AND NO TREATMENT WAS GIVEN FOR THE LUMPSBUMPS ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE | Material Discolored | Device problem code
6926827 | 88954994 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF FIRM NODULES SWELLING AND ERYTHEMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS INJECTION OF MORE THAN ML OF JUVÉDERM VOLBELLA XC INJECTABLE GEL HAS NOT BEEN STUDIED FOR LIP AUGMENTATION AND CORRECTION OF PERIORAL RHYTIDS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS MOST SUBJECTS REPORTED AN ISR AFTER TREATMENT WITH THE MOST COMMON BEING SWELLING TENDERNESS AND FIRMNESS THE MAJORITY OF ISRS WERE MILD TO MODERATE IN SEVERITY AND RESOLVED WITHIN DAYS TREATMENTRELATED AES AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT OCCURRING IN OF SUBJECTS INCLUDED CHAPPED LIPS DIZZINESS DRY LIPS GENERAL PHYSICAL CONDITION ABNORMAL HEADACHE LIP DISORDER LUMPS LIP INJURY ORAL HERPES PRESYNCOPE WOUND AND INJECTION SITE DISCOLORATION DISCOMFORT EDEMA ERYTHEMA EXFOLIATION HYPERAESTHESIA HYPOAESTHESIA LACERATION NODULE PAPULE PARAESTHESIA PRURITUS AND REACTION IN THE CLINICAL STUDY SUBJECTS HAD LUMPSBUMPS OR SWELLING THAT OCCURRED WEEKS TO MONTHS AFTER TREATMENT ALL OF THESE AES WERE MILD OR MODERATE SWELLING WAS TREATED WITH ACETAMINOPHEN OR DOXYCYCLINE AND NO TREATMENT WAS GIVEN FOR THE LUMPSBUMPS ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE | Material and/or Chemical Problem Identified | Medical Evaluation result code
6930089 | 88958548 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT A LEAD REVISION PROCEDURE WAS PERFORMED ONCE THE POCKET WAS OPENED VISUAL INSPECTION OF THE IMPLANTABLE PRODUCTS FOUND EVIDENCE THIS PATIENT HAD TWIDDLERS SYNDROME THE LEAD WAS TESTED WITH THE PACING SYSTEM ANALYZER PSA AND THE MEASUREMENTS REMAINED OUTSIDE OF THE EXPECTED RANGE AND NOISE WAS STILL PRESENT THE RV LEAD WAS SURGICALLY ABANDONED AND A NEW RV LEAD WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Noise, Audible | Device problem code
6930089 | 88958548 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT A LEAD REVISION PROCEDURE WAS PERFORMED ONCE THE POCKET WAS OPENED VISUAL INSPECTION OF THE IMPLANTABLE PRODUCTS FOUND EVIDENCE THIS PATIENT HAD TWIDDLERS SYNDROME THE LEAD WAS TESTED WITH THE PACING SYSTEM ANALYZER PSA AND THE MEASUREMENTS REMAINED OUTSIDE OF THE EXPECTED RANGE AND NOISE WAS STILL PRESENT THE RV LEAD WAS SURGICALLY ABANDONED AND A NEW RV LEAD WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Power Conditioning Problem | Device Problem Code
6930089 | 88958548 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT A LEAD REVISION PROCEDURE WAS PERFORMED ONCE THE POCKET WAS OPENED VISUAL INSPECTION OF THE IMPLANTABLE PRODUCTS FOUND EVIDENCE THIS PATIENT HAD TWIDDLERS SYNDROME THE LEAD WAS TESTED WITH THE PACING SYSTEM ANALYZER PSA AND THE MEASUREMENTS REMAINED OUTSIDE OF THE EXPECTED RANGE AND NOISE WAS STILL PRESENT THE RV LEAD WAS SURGICALLY ABANDONED AND A NEW RV LEAD WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Power Source Problem Identified | Medical Evaluation result code
6927077 | 88967496 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS IN THE CLINICAL STUDY SEVERE TREATMENTRELATED ADVERSE EVENTS OCCURRED IN SUBJECTS THESE ADVERSE EVENTS INCLUDE ANGIOEDEMA AND INJECTION SITE MASS PAIN BRUISING SWELLING ERYTHEMA AND HYPERTROPHY ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE AND ALL EXCEPT THE ANGIOEDEMA REQUIRED NO ACTION ONE SUBJECT EXPERIENCED ANGIOEDEMA IN THE UPPER LIP FOLLOWING TOPICAL ANESTHETIC APPLICATION OF LIDOCAINE TETRACAINE AND INJECTION OF JUVÉDERM ULTRA XC WHICH RESOLVED FOLLOWING ADMINISTRATION OF ORAL ANTIHISTAMINE HYALURONIDASE INJECTION AND ORAL ANTIINFLAMMATORY MEDICATION THE MOST COMMON INJECTION SITE RESPONSES FOR JUVÉDERM ULTRA XC WERE REDNESS SWELLING TENDERNESS FIRMNESS LUMPSBUMPS DISCOLORATION AND BRUISING ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Material Discolored | Device problem code
6927077 | 88967496 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS IN THE CLINICAL STUDY SEVERE TREATMENTRELATED ADVERSE EVENTS OCCURRED IN SUBJECTS THESE ADVERSE EVENTS INCLUDE ANGIOEDEMA AND INJECTION SITE MASS PAIN BRUISING SWELLING ERYTHEMA AND HYPERTROPHY ALL OF THESE EVENTS RESOLVED WITHOUT SEQUELAE AND ALL EXCEPT THE ANGIOEDEMA REQUIRED NO ACTION ONE SUBJECT EXPERIENCED ANGIOEDEMA IN THE UPPER LIP FOLLOWING TOPICAL ANESTHETIC APPLICATION OF LIDOCAINE TETRACAINE AND INJECTION OF JUVÉDERM ULTRA XC WHICH RESOLVED FOLLOWING ADMINISTRATION OF ORAL ANTIHISTAMINE HYALURONIDASE INJECTION AND ORAL ANTIINFLAMMATORY MEDICATION THE MOST COMMON INJECTION SITE RESPONSES FOR JUVÉDERM ULTRA XC WERE REDNESS SWELLING TENDERNESS FIRMNESS LUMPSBUMPS DISCOLORATION AND BRUISING ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Material and/or Chemical Problem Identified | Medical Evaluation result code
6929550 | 88974082 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED VISUAL INSPECTION IDENTIFIED THE STYLET WAS STILL INSERTED IN THE LEADS BODY AND THE HELIX WAS RETRACTED AND THE INSULATION OF THE NECK REGION WAS SLIGHTLY TWISTED ELECTRICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE ELEVATED THRESHOLDS | Retraction Problem | Device problem code
6929550 | 88974082 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED VISUAL INSPECTION IDENTIFIED THE STYLET WAS STILL INSERTED IN THE LEADS BODY AND THE HELIX WAS RETRACTED AND THE INSULATION OF THE NECK REGION WAS SLIGHTLY TWISTED ELECTRICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE ELEVATED THRESHOLDS | Material Twisted / Bent | Device problem code
6929550 | 88974082 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED VISUAL INSPECTION IDENTIFIED THE STYLET WAS STILL INSERTED IN THE LEADS BODY AND THE HELIX WAS RETRACTED AND THE INSULATION OF THE NECK REGION WAS SLIGHTLY TWISTED ELECTRICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE ELEVATED THRESHOLDS | Stress Problem Identified | Medical Evaluation result code
6938747 | 88977560 | D | A REVISION SURGERY WAS PERFORMED DUE TO A PRESSURE SORE THE DEVICE WAS EXPLANTED AND REPLACED THERE WAS NO ALLEGATION OF MALFUNCTION AGAINST THE DEVICE AND THE PATIENTS CONDITION WAS STABLE FOLLOWING THE PROCEDURE | Adverse Event Without Identified Device or Use Problem | Device problem code
6938747 | 88977560 | D | A REVISION SURGERY WAS PERFORMED DUE TO A PRESSURE SORE THE DEVICE WAS EXPLANTED AND REPLACED THERE WAS NO ALLEGATION OF MALFUNCTION AGAINST THE DEVICE AND THE PATIENTS CONDITION WAS STABLE FOLLOWING THE PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
6938032 | 88978170 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS PROPERLY INSERTED INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6938032 | 88978170 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS PROPERLY INSERTED INTO THE INFUSION SITE | Positioning Problem | Device problem code
6936589 | 88978387 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
6936589 | 88978387 | D | A PATIENTS BLOOD GLUCOSE REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6926480 | 88995318 | N | B THE EVENTS OF GUM INFECTION CHEEK BECAME HUGE AREA UNDER EYES HAS BAGS LUMP AND REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6926480 | 88995318 | N | B THE EVENTS OF GUM INFECTION CHEEK BECAME HUGE AREA UNDER EYES HAS BAGS LUMP AND REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Material Discolored | Device problem code
6926480 | 88995318 | N | B THE EVENTS OF GUM INFECTION CHEEK BECAME HUGE AREA UNDER EYES HAS BAGS LUMP AND REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6926480 | 88995318 | N | B THE EVENTS OF GUM INFECTION CHEEK BECAME HUGE AREA UNDER EYES HAS BAGS LUMP AND REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6934985 | 88998526 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION AND STYLET INSERTED IN LEAD SUBSEQUENT ELECTRICAL TESTING AND XRAY DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Activation, Positioning or Separation Problem | Device Problem Code
6934985 | 88998526 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION AND STYLET INSERTED IN LEAD SUBSEQUENT ELECTRICAL TESTING AND XRAY DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Difficult to Remove | Device Problem Code
6934985 | 88998526 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION AND STYLET INSERTED IN LEAD SUBSEQUENT ELECTRICAL TESTING AND XRAY DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Retraction Problem | Device problem code
6934985 | 88998526 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION AND STYLET INSERTED IN LEAD SUBSEQUENT ELECTRICAL TESTING AND XRAY DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED HELIX EXTENSIONRETRACTION PROBLEMS SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | No Findings Available | Medical Evaluation result code
6938931 | 89005367 | D | A REPORT WAS RECEIVED THAT DURING PERMANENT IMPLANT PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED LEAD SPLITTERS THE BSN REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED BUT WAS USED DURING THE PROCEDURE | Defective Component | Device problem code
6938931 | 89005367 | D | A REPORT WAS RECEIVED THAT DURING PERMANENT IMPLANT PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED LEAD SPLITTERS THE BSN REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED BUT WAS USED DURING THE PROCEDURE | No Information | Patient problem code
6933921 | 89012140 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE HELIX REQUIRING MORE TURNS THAT EXPECTED TO EXTEND DURING THE IMPLANT PROCEDURE SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Difficult to Remove | Device Problem Code
6933921 | 89012140 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE HELIX REQUIRING MORE TURNS THAT EXPECTED TO EXTEND DURING THE IMPLANT PROCEDURE SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Activation, Positioning or Separation Problem | Device Problem Code
6933921 | 89012140 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE HELIX REQUIRING MORE TURNS THAT EXPECTED TO EXTEND DURING THE IMPLANT PROCEDURE SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Retraction Problem | Device problem code
6933921 | 89012140 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE HELIX REQUIRING MORE TURNS THAT EXPECTED TO EXTEND DURING THE IMPLANT PROCEDURE SEVERAL IMPLANT TECHNIQUE AND PRODUCT DESIGN FACTORS MAY CONTRIBUTE TO HELIX EXTENSIONRETRACTION DIFFICULTIES INCLUDING TORTUOUS PATIENT ANATOMY OR BENDS IN THE PROXIMAL PORTION OF THE LEAD REMAINING OUTSIDE OF THE BODY SHARP BENDS IN THE STYLET KINKS IN THE LEAD INTRODUCER SHEATH UNDERROTATING OR OVERROTATING THE TERMINAL TOOL FAILURE TO REMOVE THE TERMINAL TOOL BETWEEN EXTENSIONRETRACTION CYCLES TO RELEASE TORQUE STORED IN THE INNER COIL EXCESSIVELY FAST TOOL ROTATION SPEED FULL TURN PER SECOND AND THE PRESENCE OF TISSUE IN THE HELIX WHICH MAY ACCUMULATE WITH MULTIPLE LEAD REPOSITIONING ATTEMPTS IN ADDITION THE INGEVITY LEAD WAS DESIGNED WITH SINGLE FILAR INNER AND OUTER COILS FOR IMPROVED FLEX FATIGUE AND FOR BETTER PERFORMANCE WITHIN AN MRI ENVIRONMENT WITH MULTIPLE LAYERS OF INSULATION BETWEEN THE CONDUCTORS FOR LONGTERM RELIABILITY THIS DESIGN IN CONJUNCTION WITH THE CONTRIBUTING FACTORS MENTIONED ABOVE MAY IMPACT THE RATE AT WHICH TORQUE IS TRANSFERRED FROM THE TERMINAL PIN TO THE HELIX MECHANISM APPROVED INSTRUCTIONS FOR USE PROVIDE BEST PRACTICES TO MITIGATE THESE POTENTIAL CONTRIBUTING FACTORS | Assembly Problem Identified | Medical Evaluation result code
6927074 | 89013019 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL DEVICEINJECTIONRELATED AES AFTER REPEAT TREATMENT WERE MILD TO MODERATE REQUIRED NO ACTION AND RESOLVED WITHOUT SEQUELAE GENERALLY DEVICEINJECTIONRELATED AES WERE LESS SEVERE AFTER REPEAT TREATMENT COMPARED TO INITIALTOUCHUP TREATMENT AND MOST RESOLVED WITHIN MONTHS SIMILAR TO THE INITIALTOUCHUP TREATMENT SUBJECTS EXPERIENCED A DEVICEINJECTIONRELATED AE THAT LASTED MORE THAN DAYS BUT ALL RESOLVED WITHOUT REQUIRING ANY TREATMENT DEVICEINJECTIONRELATED ADVERSE EVENTS OCCURRING IN OF SUBJECTS INCLUDED INJECTION SITE SWELLING INJECTION SITE PAIN AND INJECTION SITE PAPULE ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6927074 | 89013019 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL DEVICEINJECTIONRELATED AES AFTER REPEAT TREATMENT WERE MILD TO MODERATE REQUIRED NO ACTION AND RESOLVED WITHOUT SEQUELAE GENERALLY DEVICEINJECTIONRELATED AES WERE LESS SEVERE AFTER REPEAT TREATMENT COMPARED TO INITIALTOUCHUP TREATMENT AND MOST RESOLVED WITHIN MONTHS SIMILAR TO THE INITIALTOUCHUP TREATMENT SUBJECTS EXPERIENCED A DEVICEINJECTIONRELATED AE THAT LASTED MORE THAN DAYS BUT ALL RESOLVED WITHOUT REQUIRING ANY TREATMENT DEVICEINJECTIONRELATED ADVERSE EVENTS OCCURRING IN OF SUBJECTS INCLUDED INJECTION SITE SWELLING INJECTION SITE PAIN AND INJECTION SITE PAPULE ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6927074 | 89013019 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AND GRANULOMATOUS REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL DEVICEINJECTIONRELATED AES AFTER REPEAT TREATMENT WERE MILD TO MODERATE REQUIRED NO ACTION AND RESOLVED WITHOUT SEQUELAE GENERALLY DEVICEINJECTIONRELATED AES WERE LESS SEVERE AFTER REPEAT TREATMENT COMPARED TO INITIALTOUCHUP TREATMENT AND MOST RESOLVED WITHIN MONTHS SIMILAR TO THE INITIALTOUCHUP TREATMENT SUBJECTS EXPERIENCED A DEVICEINJECTIONRELATED AE THAT LASTED MORE THAN DAYS BUT ALL RESOLVED WITHOUT REQUIRING ANY TREATMENT DEVICEINJECTIONRELATED ADVERSE EVENTS OCCURRING IN OF SUBJECTS INCLUDED INJECTION SITE SWELLING INJECTION SITE PAIN AND INJECTION SITE PAPULE ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6934803 | 89013331 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO MEDICAL DEVICES LEAD IMPLANTED B THE INITIAL REPORTED EVENT OF INFECTION WAS RECEIVED ON B OF NOTE THE SERIOUS INJURY IS NORMALLY SUBMITTED VIA AN ALTERNATIVE SUMMARY REPORT THAT WOULD HAVE BEEN DUE ON B INFORMATION WAS SUBSEQUENTLY RECEIVED ON B AND REVEALED AN OUT OF SPECIFICATION ANALYSIS FINDING AS THIS NEW INFORMATION DOES NOT QUALIFY FOR SUMMARY REPORTING THIS REPORT IS THEREFORE BEING SUBMITTED AS A BIMONTHLY REPORT | Break | Device problem code
6934803 | 89013331 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO MEDICAL DEVICES LEAD IMPLANTED B THE INITIAL REPORTED EVENT OF INFECTION WAS RECEIVED ON B OF NOTE THE SERIOUS INJURY IS NORMALLY SUBMITTED VIA AN ALTERNATIVE SUMMARY REPORT THAT WOULD HAVE BEEN DUE ON B INFORMATION WAS SUBSEQUENTLY RECEIVED ON B AND REVEALED AN OUT OF SPECIFICATION ANALYSIS FINDING AS THIS NEW INFORMATION DOES NOT QUALIFY FOR SUMMARY REPORTING THIS REPORT IS THEREFORE BEING SUBMITTED AS A BIMONTHLY REPORT | Fatigue Problem | Medical Evaluation result code
6937329 | 89015779 | D | PATIENT DEVELOPED HIVES SHORTNESS OF BREATH WHEEZING ITCHING ORAL SWELLING AND ELEVATED BLOOD PRESSURE FOLLOWING INSERTION OF A POWERPICC SOLO CATHETER PATIENT WAS LATER DISCHARGED FROM HOSPITAL | Dyspnea | Patient problem code
6937329 | 89015779 | D | PATIENT DEVELOPED HIVES SHORTNESS OF BREATH WHEEZING ITCHING ORAL SWELLING AND ELEVATED BLOOD PRESSURE FOLLOWING INSERTION OF A POWERPICC SOLO CATHETER PATIENT WAS LATER DISCHARGED FROM HOSPITAL | Urticaria | Patient problem code
6937329 | 89015779 | D | PATIENT DEVELOPED HIVES SHORTNESS OF BREATH WHEEZING ITCHING ORAL SWELLING AND ELEVATED BLOOD PRESSURE FOLLOWING INSERTION OF A POWERPICC SOLO CATHETER PATIENT WAS LATER DISCHARGED FROM HOSPITAL | Insufficient Information | Device problem code
6937329 | 89015779 | D | PATIENT DEVELOPED HIVES SHORTNESS OF BREATH WHEEZING ITCHING ORAL SWELLING AND ELEVATED BLOOD PRESSURE FOLLOWING INSERTION OF A POWERPICC SOLO CATHETER PATIENT WAS LATER DISCHARGED FROM HOSPITAL | Itching | Patient problem code
6937329 | 89015779 | D | PATIENT DEVELOPED HIVES SHORTNESS OF BREATH WHEEZING ITCHING ORAL SWELLING AND ELEVATED BLOOD PRESSURE FOLLOWING INSERTION OF A POWERPICC SOLO CATHETER PATIENT WAS LATER DISCHARGED FROM HOSPITAL | Swelling | Patient problem code
6931616 | 89045784 | N | ONE WEEK LATER THE PATIENT WAS SEEN AND ACCORDING TO A PHYSICIAN FELLOWS DICTATED REPORT THE WOUND DEHISCENCE GREW B THE WEEK AFTER IMPLANTATION THE WOUND LOOKED MUCH IMPROVED NO DRAINAGE OR ERYTHEMA THE PATIENT DENIES ANY FEVER OR CHILLS THE STABLES WERE REMOVED THERE WAS STILL A VERY SMALL OPENING IN THE SKIN WITH MINOR DRAINAGE BLOOD WAS AGAIN DRAWN FOR CULTURES BACTRIM WAS STOPPED AND CLINDAMYCIN AND PROBIOTICS WERE PRESCRIBED ANOTHER INCLINIC WOUND CHECK WAS SCHEDULED FOR THE FOLLOWING WEEK THE PATIENT WAS SEEN URGENTLY BECAUSE BLOOD CULTURES STILL GREW MRSA ONE WEEK LATER THE WOUND LOOKED MUCH IMPROVED THE STERISTRIPS WERE REMOVED AND NO DRAINAGE WAS OBSERVED THE PATIENT WAS SEEN IN B FOR A REGULAR FOLLOWUP VISIT THE PATIENT WHO WAS OFF ANTIBIOTICS FELT WELL AND DENIED WORSENING SHORTENING OF BREATH CHEST PAIN ORTHOPNEA PAROXYSMAL NOCTURNAL DYSPNEA LEG SWELLING AND PALPITATIONS THE PATIENT WAS SEEN IN B FOR NONDEVICE RELATED ISSUES BOSTON SCIENTIFIC RECEIVED INFORMATION THAT A BOSTON SCIENTIFIC REPRESENTATIVE WAS NOT NOTIFIED OF THE INFECTIONPOCKET REVISION THE LOCAL REPRESENTATIVE INVESTIGATED FURTHER AND FOUND THAT TWO SETS OF BLOOD CULTURES WERE DONE ON B AND BOTH WERE NEGATIVE A WOUND CULTURE WAS DONE ON B AND ANOTHER ON B BOTH WOUND CULTURES WERE POSITIVE THE WOUND CULTURE FROM B WAS POSITIVE FOR B | Leak / Splash | Device problem code
6931616 | 89045784 | N | ONE WEEK LATER THE PATIENT WAS SEEN AND ACCORDING TO A PHYSICIAN FELLOWS DICTATED REPORT THE WOUND DEHISCENCE GREW B THE WEEK AFTER IMPLANTATION THE WOUND LOOKED MUCH IMPROVED NO DRAINAGE OR ERYTHEMA THE PATIENT DENIES ANY FEVER OR CHILLS THE STABLES WERE REMOVED THERE WAS STILL A VERY SMALL OPENING IN THE SKIN WITH MINOR DRAINAGE BLOOD WAS AGAIN DRAWN FOR CULTURES BACTRIM WAS STOPPED AND CLINDAMYCIN AND PROBIOTICS WERE PRESCRIBED ANOTHER INCLINIC WOUND CHECK WAS SCHEDULED FOR THE FOLLOWING WEEK THE PATIENT WAS SEEN URGENTLY BECAUSE BLOOD CULTURES STILL GREW MRSA ONE WEEK LATER THE WOUND LOOKED MUCH IMPROVED THE STERISTRIPS WERE REMOVED AND NO DRAINAGE WAS OBSERVED THE PATIENT WAS SEEN IN B FOR A REGULAR FOLLOWUP VISIT THE PATIENT WHO WAS OFF ANTIBIOTICS FELT WELL AND DENIED WORSENING SHORTENING OF BREATH CHEST PAIN ORTHOPNEA PAROXYSMAL NOCTURNAL DYSPNEA LEG SWELLING AND PALPITATIONS THE PATIENT WAS SEEN IN B FOR NONDEVICE RELATED ISSUES BOSTON SCIENTIFIC RECEIVED INFORMATION THAT A BOSTON SCIENTIFIC REPRESENTATIVE WAS NOT NOTIFIED OF THE INFECTIONPOCKET REVISION THE LOCAL REPRESENTATIVE INVESTIGATED FURTHER AND FOUND THAT TWO SETS OF BLOOD CULTURES WERE DONE ON B AND BOTH WERE NEGATIVE A WOUND CULTURE WAS DONE ON B AND ANOTHER ON B BOTH WOUND CULTURES WERE POSITIVE THE WOUND CULTURE FROM B WAS POSITIVE FOR B | Leakage/Seal | Medical Evaluation result code
6931617 | 89046620 | N | ONE WEEK LATER THE PATIENT WAS SEEN AND ACCORDING TO A PHYSICIAN FELLOWS DICTATED REPORT THE WOUND DEHISCENCE GREW B THE WEEK AFTER IMPLANTATION THE WOUND LOOKED MUCH IMPROVED NO DRAINAGE OR ERYTHEMA THE PATIENT DENIES ANY FEVER OR CHILLS THE STABLES WERE REMOVED THERE WAS STILL A VERY SMALL OPENING IN THE SKIN WITH MINOR DRAINAGE BLOOD WAS AGAIN DRAWN FOR CULTURES BACTRIM WAS STOPPED AND CLINDAMYCIN AND PROBIOTICS WERE PRESCRIBED ANOTHER INCLINIC WOUND CHECK WAS SCHEDULED FOR THE FOLLOWING WEEK THE PATIENT WAS SEEN URGENTLY BECAUSE BLOOD CULTURES STILL GREW B ONE WEEK LATER THE WOUND LOOKED MUCH IMPROVED THE STERISTRIPS WERE REMOVED AND NO DRAINAGE WAS OBSERVED THE PATIENT WAS SEEN IN EARLY B FOR A REGULAR FOLLOWUP VISIT THE PATIENT WHO WAS OFF ANTIBIOTICS FELT WELL AND DENIED WORSENING SHORTENING OF BREATH CHEST PAIN ORTHOPNEA PAROXYSMAL NOCTURNAL DYSPNEA LEG SWELLING AND PALPITATIONS THE PATIENT WAS SEEN IN EARLYB FOR NONDEVICE RELATED ISSUES BOSTON SCIENTIFIC RECEIVED INFORMATION THAT A BOSTON SCIENTIFIC REPRESENTATIVE WAS NOT NOTIFIED OF THE INFECTIONPOCKET REVISION THE LOCAL REPRESENTATIVE INVESTIGATED FURTHER AND FOUND THAT TWO SETS OF BLOOD CULTURES WERE DONE ON B AND BOTH WERE NEGATIVE A WOUND CULTURE WAS DONE ON B AND ANOTHER ON B BOTH WOUND CULTURES WERE POSITIVE THE WOUND CULTURE FROM B WAS POSITIVE FOR B | Leak / Splash | Device problem code
6931617 | 89046620 | N | ONE WEEK LATER THE PATIENT WAS SEEN AND ACCORDING TO A PHYSICIAN FELLOWS DICTATED REPORT THE WOUND DEHISCENCE GREW B THE WEEK AFTER IMPLANTATION THE WOUND LOOKED MUCH IMPROVED NO DRAINAGE OR ERYTHEMA THE PATIENT DENIES ANY FEVER OR CHILLS THE STABLES WERE REMOVED THERE WAS STILL A VERY SMALL OPENING IN THE SKIN WITH MINOR DRAINAGE BLOOD WAS AGAIN DRAWN FOR CULTURES BACTRIM WAS STOPPED AND CLINDAMYCIN AND PROBIOTICS WERE PRESCRIBED ANOTHER INCLINIC WOUND CHECK WAS SCHEDULED FOR THE FOLLOWING WEEK THE PATIENT WAS SEEN URGENTLY BECAUSE BLOOD CULTURES STILL GREW B ONE WEEK LATER THE WOUND LOOKED MUCH IMPROVED THE STERISTRIPS WERE REMOVED AND NO DRAINAGE WAS OBSERVED THE PATIENT WAS SEEN IN EARLY B FOR A REGULAR FOLLOWUP VISIT THE PATIENT WHO WAS OFF ANTIBIOTICS FELT WELL AND DENIED WORSENING SHORTENING OF BREATH CHEST PAIN ORTHOPNEA PAROXYSMAL NOCTURNAL DYSPNEA LEG SWELLING AND PALPITATIONS THE PATIENT WAS SEEN IN EARLYB FOR NONDEVICE RELATED ISSUES BOSTON SCIENTIFIC RECEIVED INFORMATION THAT A BOSTON SCIENTIFIC REPRESENTATIVE WAS NOT NOTIFIED OF THE INFECTIONPOCKET REVISION THE LOCAL REPRESENTATIVE INVESTIGATED FURTHER AND FOUND THAT TWO SETS OF BLOOD CULTURES WERE DONE ON B AND BOTH WERE NEGATIVE A WOUND CULTURE WAS DONE ON B AND ANOTHER ON B BOTH WOUND CULTURES WERE POSITIVE THE WOUND CULTURE FROM B WAS POSITIVE FOR B | Leakage/Seal | Medical Evaluation result code
6938803 | 89075649 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS EXTENDED AND DRIED BLOODTISSUE WAS PRESENT AROUND THE NECK REGION RESISTANCE TESTING FOUND THE LEAD WAS NOT ELECTRICALLY CONTINUOUS DETAILED ANALYSIS CONFIRMED THAT THE CATHODE COIL HAD A BREAK AT THE DISTAL END OF THE TERMINAL PIN | Break | Device problem code
6938803 | 89075649 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS EXTENDED AND DRIED BLOODTISSUE WAS PRESENT AROUND THE NECK REGION RESISTANCE TESTING FOUND THE LEAD WAS NOT ELECTRICALLY CONTINUOUS DETAILED ANALYSIS CONFIRMED THAT THE CATHODE COIL HAD A BREAK AT THE DISTAL END OF THE TERMINAL PIN | Fatigue Problem | Medical Evaluation result code
6941548 | 89078240 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Pain | Patient problem code
6941548 | 89078240 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Limited Mobility Of The Implanted Joint | Patient problem code
6941548 | 89078240 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Toxicity | Patient problem code
6941548 | 89078240 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Insufficient Information | Device problem code
6941650 | 89078436 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Toxicity | Patient problem code
6941650 | 89078436 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Limited Mobility Of The Implanted Joint | Patient problem code
6941650 | 89078436 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Pain | Patient problem code
6941650 | 89078436 | D | SEP LITIGATION RECEIVED LITIGATION ALLEGES PAIN BODILY IMPAIRMENT DEBILITATING LACK OF MOBILITY AND HIGH LEVELS OF TOXIC METAL LEVELS | Insufficient Information | Device problem code
6936258 | 89078547 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY DETAILED ANALYSIS INDICATES THAT THE ALLEGATION AGAINST THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS NOT CONFIRMED UPON EXTERNAL VISUAL INSPECTION SCRATCHES AND DENTS WERE NOTED ON THE CASE ALSO THE CASE WAS SWOLLEN MOREOVER INTERNAL VISUAL INSPECTION NOTED NO IRREGULARITIES FURTHER MANUAL ELECTRICAL MEASUREMENTS NOTED NORMAL PACING SENSING AND DEFIBRILLATOR FUNCTIONS THUS THE ICD WAS OPERATING NORMALLY | Scratched Material | Device problem code
6936258 | 89078547 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY DETAILED ANALYSIS INDICATES THAT THE ALLEGATION AGAINST THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS NOT CONFIRMED UPON EXTERNAL VISUAL INSPECTION SCRATCHES AND DENTS WERE NOTED ON THE CASE ALSO THE CASE WAS SWOLLEN MOREOVER INTERNAL VISUAL INSPECTION NOTED NO IRREGULARITIES FURTHER MANUAL ELECTRICAL MEASUREMENTS NOTED NORMAL PACING SENSING AND DEFIBRILLATOR FUNCTIONS THUS THE ICD WAS OPERATING NORMALLY | Stress Problem Identified | Medical Evaluation result code
6934538 | 89094640 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ANALYSIS WAS UNABLE TO CONFIRM REPORTED ALLEGATION THE LEAD PASSED ELECTRICAL TESTING THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | Retraction Problem | Device problem code
6934538 | 89094640 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS RETRACTED ANALYSIS WAS UNABLE TO CONFIRM REPORTED ALLEGATION THE LEAD PASSED ELECTRICAL TESTING THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | No Findings Available | Medical Evaluation result code
6937419 | 89097598 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED IN WHICH THIS RIGHT VENTRICULAR RV LEAD PASSED THE ELECTRICAL TEST FURTHER IT WAS NOTED ON VISUAL OBSERVATION THAT THE EXTRACTING STYLET WAS STUCK INSIDE THE DISTAL SEGMENT | Mechanical Jam | Device problem code
6937419 | 89097598 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED IN WHICH THIS RIGHT VENTRICULAR RV LEAD PASSED THE ELECTRICAL TEST FURTHER IT WAS NOTED ON VISUAL OBSERVATION THAT THE EXTRACTING STYLET WAS STUCK INSIDE THE DISTAL SEGMENT | Mechanical Problem Identified | Medical Evaluation result code
6930268 | 89103970 | N | B LOT ML WORK ORDER B WAS REVIEWED THIS LOT WAS MANUFACTURED ON AND THE EXPIRATION DATE IS ONE PROBE IN THIS LOT SN B FAILED TO PROGRAM THERE WERE NO OTHER ISSUES WITH THE MANUFACTURING OR TESTING OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT WERE THE INDICATIONS FOR ABLATION THERAPY IN THE SURGEONS OPINION WHAT WERE THE CONTRIBUTING FACTORS THAT LED TO THE PROBE BREAKING WHAT IS THE CURRENT PATIENT STATUS | Break | Device problem code
6930268 | 89103970 | N | B LOT ML WORK ORDER B WAS REVIEWED THIS LOT WAS MANUFACTURED ON AND THE EXPIRATION DATE IS ONE PROBE IN THIS LOT SN B FAILED TO PROGRAM THERE WERE NO OTHER ISSUES WITH THE MANUFACTURING OR TESTING OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT WERE THE INDICATIONS FOR ABLATION THERAPY IN THE SURGEONS OPINION WHAT WERE THE CONTRIBUTING FACTORS THAT LED TO THE PROBE BREAKING WHAT IS THE CURRENT PATIENT STATUS | Fatigue Problem | Medical Evaluation result code
6938750 | 89106981 | N | B THE EVENTS OF LATE SEROMA SILICON GRANULOMATA CYST AND LYMPHOMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Deflation Problem | Device problem code
6938750 | 89106981 | N | B THE EVENTS OF LATE SEROMA SILICON GRANULOMATA CYST AND LYMPHOMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Calcified | Device problem code
6938750 | 89106981 | N | B THE EVENTS OF LATE SEROMA SILICON GRANULOMATA CYST AND LYMPHOMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY BASED ON INFORMATION REPORTED TO FDA AND FOUND IN MEDICAL LITERATURE A POSSIBLE ASSOCIATION HAS BEEN IDENTIFIED BETWEEN BREAST IMPLANTS AND THE RARE DEVELOPMENT OF ANAPLASTIC LARGE CELL LYMPHOMA ALCL A TYPE OF NONHODGKINS LYMPHOMA WOMEN WITH BREAST IMPLANTS MAY HAVE A VERY SMALL BUT INCREASED RISK OF DEVELOPING ALCL IN THE FLUID OR SCAR CAPSULE ADJACENT TO THE IMPLANT ALCL HAS BEEN REPORTED GLOBALLY IN PATIENTS WITH AN IMPLANT HISTORY THAT INCLUDES ALLERGANS AND OTHER MANUFACTURERS BREAST IMPLANTS YOU SHOULD CONSIDER THE POSSIBILITY OF ALCL WHEN YOU HAVE A PATIENT WITH LATE ONSET PERSISTENT PERIIMPLANT SEROMA IN SOME CASES PATIENTS PRESENTED WITH CAPSULAR CONTRACTURE OR MASSES ADJACENT TO THE BREAST IMPLANT WHEN TESTING FOR ALCL COLLECT FRESH SEROMA FLUID AND REPRESENTATIVE PORTIONS OF THE CAPSULE AND SEND FOR PATHOLOGY TESTS TO RULE OUT ALCL IF YOUR PATIENT IS DIAGNOSED WITH PERIIMPLANT ALCL DEVELOP AN INDIVIDUALIZED TREATMENT PLAN IN COORDINATION WITH A MULTIDISCIPLINARY CARE TEAM BECAUSE OF THE SMALL NUMBER OF CASES WORLDWIDE THERE IS NO DEFINED CONSENSUS TREATMENT REGIMEN FOR PERIIMPLANT ALCL | Non-Functional Defect | Medical Evaluation result code
6933942 | 89111495 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS PATIENT DEVELOPED A HEMATOMA FOLLOWING AN IMPLANT PROCEDURE A PRESSURE DRESSING WAS USED AS TREATMENT AND THE PATIENT WAS DISCHARGED THE PACEMAKER REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6933942 | 89111495 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS PATIENT DEVELOPED A HEMATOMA FOLLOWING AN IMPLANT PROCEDURE A PRESSURE DRESSING WAS USED AS TREATMENT AND THE PATIENT WAS DISCHARGED THE PACEMAKER REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
6943462 | 89112157 | D | IT WAS REPORTED THAT THE CANNULA WAS NOT IN THE SKIN IT WAS NOT SPECIFICALLY STATED WHETHER THE CANNULA WAS VISIBLE OUTSIDE OF THE POD HOWEVER THE PINK SLIDE HAD MOVED FORWARD THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS | No Known Impact Or Consequence To Patient | Patient Problem Code
6943462 | 89112157 | D | IT WAS REPORTED THAT THE CANNULA WAS NOT IN THE SKIN IT WAS NOT SPECIFICALLY STATED WHETHER THE CANNULA WAS VISIBLE OUTSIDE OF THE POD HOWEVER THE PINK SLIDE HAD MOVED FORWARD THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS | Insufficient Information | Device problem code
6934388 | 89118202 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS PACING THERAPY WAS NOT DELIVERED WHEN IT WAS REQUIRED ON THE ATRIAL CHANNEL THE PHYSICIAN SUSPECTED THE CAUSE WAS MOST LIKELY DUE TO A TYPE OF EXIT BLOCK NO DEVICE ISSUES WERE IDENTIFIED A REVISION PROCEDURE WAS PERFORMED AND THE ATRIAL LEAD WAS REMOVED FROM SERVICE AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
6934388 | 89118202 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS PACING THERAPY WAS NOT DELIVERED WHEN IT WAS REQUIRED ON THE ATRIAL CHANNEL THE PHYSICIAN SUSPECTED THE CAUSE WAS MOST LIKELY DUE TO A TYPE OF EXIT BLOCK NO DEVICE ISSUES WERE IDENTIFIED A REVISION PROCEDURE WAS PERFORMED AND THE ATRIAL LEAD WAS REMOVED FROM SERVICE AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Adverse Event Without Identified Device or Use Problem | Device problem code
6934997 | 89124825 | N | THE FACILITY REPORTEDLY OBSERVED THAT THE METRICIDE OPA PLUS HIGH LEVEL DISINFECTANT USED IN THEIR CER OPTIMA AUTOMATED ENDOSCOPE REPROCESSOR HLD RESERVOIR WAS CLOUDY IN APPEARANCE THE FACILITY CONFIRMED THAT THE OPA SOLUTION WAS NOT EXPIRED AND THE MINIMUM REQUIRED CONCENTRATION MRC PASSED FOR THE REPROCESSING CYCLE THE INITIAL REPORTER THOUGHT SOMEONE MAY HAVE ACCIDENTALLY TOPPED OFF THE DISINFECTANT RESERVOIR WITH METRIZYME DETERGENT WHICH THEY USE AS PART OF THEIR PRECLEANING STEP PRIOR TO HIGHLEVEL DISINFECTING SCOPES THEREFORE IT IS UNKNOWN IF SCOPES WERE PROPERLY HIGHLEVEL DISINFECTED IT IS ALSO UNKNOWN IF SCOPES WERE USED ON PATIENTS THUS POTENTIAL RISK OF PATIENT ADVERSE EFFECTS FROM IMPROPER HLD OF SCOPES MEDIVATORS TECHNICAL SERVICES ADVISED THE CUSTOMER TO DUMP THE DISINFECTANT FROM THE RESERVOIR RINSE THOROUGHLY WITH WATER THEN REFILL WITH THE OPA SOLUTION AND TEST THE MRC USING THE SOLUTION TEST STRIPS THE CUSTOMER FOLLOWED THESE DIRECTIONS AND CONFIRMED THAT THE MRC PASSED TO DATE THERE HAVE BEEN NO REPORTS OF PATIENT HARM MEDIVATORS WILL CONTINUE MONITORING THIS COMPLAINT IN THE MEDIVATORS COMPLAINT HANDLING SYSTEM | Device Handling Problem | Device problem code
6934997 | 89124825 | N | THE FACILITY REPORTEDLY OBSERVED THAT THE METRICIDE OPA PLUS HIGH LEVEL DISINFECTANT USED IN THEIR CER OPTIMA AUTOMATED ENDOSCOPE REPROCESSOR HLD RESERVOIR WAS CLOUDY IN APPEARANCE THE FACILITY CONFIRMED THAT THE OPA SOLUTION WAS NOT EXPIRED AND THE MINIMUM REQUIRED CONCENTRATION MRC PASSED FOR THE REPROCESSING CYCLE THE INITIAL REPORTER THOUGHT SOMEONE MAY HAVE ACCIDENTALLY TOPPED OFF THE DISINFECTANT RESERVOIR WITH METRIZYME DETERGENT WHICH THEY USE AS PART OF THEIR PRECLEANING STEP PRIOR TO HIGHLEVEL DISINFECTING SCOPES THEREFORE IT IS UNKNOWN IF SCOPES WERE PROPERLY HIGHLEVEL DISINFECTED IT IS ALSO UNKNOWN IF SCOPES WERE USED ON PATIENTS THUS POTENTIAL RISK OF PATIENT ADVERSE EFFECTS FROM IMPROPER HLD OF SCOPES MEDIVATORS TECHNICAL SERVICES ADVISED THE CUSTOMER TO DUMP THE DISINFECTANT FROM THE RESERVOIR RINSE THOROUGHLY WITH WATER THEN REFILL WITH THE OPA SOLUTION AND TEST THE MRC USING THE SOLUTION TEST STRIPS THE CUSTOMER FOLLOWED THESE DIRECTIONS AND CONFIRMED THAT THE MRC PASSED TO DATE THERE HAVE BEEN NO REPORTS OF PATIENT HARM MEDIVATORS WILL CONTINUE MONITORING THIS COMPLAINT IN THE MEDIVATORS COMPLAINT HANDLING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6938208 | 89133009 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A HEAT RASH UNDER THE THERAPY ELECTRODES AND FRONT RIGHT ECG ELECTRODE THE PATIENTS DOCTOR RECOMMENDED THE USE OF A CORTISONE CREAM ON THE AREA THE FINAL MEDICAL OUTCOME OF THE IRRITATION IS UNKNOWN | Irritation | Patient problem code
6938208 | 89133009 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A HEAT RASH UNDER THE THERAPY ELECTRODES AND FRONT RIGHT ECG ELECTRODE THE PATIENTS DOCTOR RECOMMENDED THE USE OF A CORTISONE CREAM ON THE AREA THE FINAL MEDICAL OUTCOME OF THE IRRITATION IS UNKNOWN | Insufficient Information | Device problem code
6938208 | 89133009 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A HEAT RASH UNDER THE THERAPY ELECTRODES AND FRONT RIGHT ECG ELECTRODE THE PATIENTS DOCTOR RECOMMENDED THE USE OF A CORTISONE CREAM ON THE AREA THE FINAL MEDICAL OUTCOME OF THE IRRITATION IS UNKNOWN | Rash | Patient problem code
6933132 | 89154268 | N | B THE PUMP WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Component Missing | Device problem code
6933132 | 89154268 | N | B THE PUMP WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
6933132 | 89154268 | N | B THE PUMP WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6933132 | 89154268 | N | B THE PUMP WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | No Device Problem Found | Medical Evaluation result code
6933700 | 89155003 | N | LOTSERIAL NUMBERS WERE NOT PROVIDED THEREFORE A REVIEW OF THE MANUFACTURING RECORDS COULD NOT BE CONDUCTED PER THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE IFU ADVERSE EVENTS THAT MAY OCCUR AND OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK AND ANEURYSM ENLARGEMENT B | No Device Problem Found | Medical Evaluation result code
6932661 | 89156668 | N | THE LOT HISTORY RECORD FOR THE INSPAN PLATE PART LOT BH WAS REVIEWED FOLLOWING THE INCIDENT NO NCMR DEVIATIONS OR CAPAS RESULTED FROM THE MANUFACTURING OF LOT BH THE INSPAN INTERSPINOUS PLATE SYSTEM WAS SOLD NONSTERILE THE ENDUSER IS RESPONSIBLE FOR ENSURING THAT THE DEVICE IS STERILIZED PRIOR TO USE ADDITIONALLY SPINEFRONTIER PROVIDES A CLEANING AND STERILIZATION IFU WHICH ADEQUATELY DESCRIBES THE CLEANING AND STERILIZATION PROCESSES THE USER SITE COULD NOT PRODUCE CLEANING AND STERILIZATION RECORD FOR THE SET USED IN THIS SURGERY CONSEQUENTLY SPINEFRONTIER COULD NOT ESTABLISH IF THE SET WAS STERILIZED PRIOR TO SURGERY | Problem with Sterilization | Device problem code
6932661 | 89156668 | N | THE LOT HISTORY RECORD FOR THE INSPAN PLATE PART LOT BH WAS REVIEWED FOLLOWING THE INCIDENT NO NCMR DEVIATIONS OR CAPAS RESULTED FROM THE MANUFACTURING OF LOT BH THE INSPAN INTERSPINOUS PLATE SYSTEM WAS SOLD NONSTERILE THE ENDUSER IS RESPONSIBLE FOR ENSURING THAT THE DEVICE IS STERILIZED PRIOR TO USE ADDITIONALLY SPINEFRONTIER PROVIDES A CLEANING AND STERILIZATION IFU WHICH ADEQUATELY DESCRIBES THE CLEANING AND STERILIZATION PROCESSES THE USER SITE COULD NOT PRODUCE CLEANING AND STERILIZATION RECORD FOR THE SET USED IN THIS SURGERY CONSEQUENTLY SPINEFRONTIER COULD NOT ESTABLISH IF THE SET WAS STERILIZED PRIOR TO SURGERY | Device Incorrectly Disinfected/Sterilised During Reprocessing | Medical Evaluation result code
6941844 | 89160279 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6931249 | 89173792 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS WHEN THE POD WAS REMOVED IT WAS NOTICED THAT THE CANNULA DID NOT GO INTO THE SKIN | Activation, Positioning or Separation Problem | Device problem code
6931249 | 89173792 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR BETWEEN AND HOURS WHEN THE POD WAS REMOVED IT WAS NOTICED THAT THE CANNULA DID NOT GO INTO THE SKIN | No Known Impact Or Consequence To Patient | Patient Problem Code
6936944 | 89174445 | D | IT WAS REPORTED THAT AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM IT WAS NOTICED THAT THE PATIENTS SITE WAS RED SWOLLEN HARD AND FULL OF PUS PATIENT WENT TO THE HOSPITAL AND WAS PRESCRIBED FLUCLOXACILLIN MG FOR DAYS FOR HER SITE INFECTION | Insufficient Information | Device problem code
6936944 | 89174445 | D | IT WAS REPORTED THAT AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM IT WAS NOTICED THAT THE PATIENTS SITE WAS RED SWOLLEN HARD AND FULL OF PUS PATIENT WENT TO THE HOSPITAL AND WAS PRESCRIBED FLUCLOXACILLIN MG FOR DAYS FOR HER SITE INFECTION | Abscess | Patient problem code
6936944 | 89174445 | D | IT WAS REPORTED THAT AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM IT WAS NOTICED THAT THE PATIENTS SITE WAS RED SWOLLEN HARD AND FULL OF PUS PATIENT WENT TO THE HOSPITAL AND WAS PRESCRIBED FLUCLOXACILLIN MG FOR DAYS FOR HER SITE INFECTION | Infection | Patient problem code
6936944 | 89174445 | D | IT WAS REPORTED THAT AFTER WEARING THE POD BETWEEN AND HOURS ON THE ARM IT WAS NOTICED THAT THE PATIENTS SITE WAS RED SWOLLEN HARD AND FULL OF PUS PATIENT WENT TO THE HOSPITAL AND WAS PRESCRIBED FLUCLOXACILLIN MG FOR DAYS FOR HER SITE INFECTION | Swelling | Patient problem code
6939906 | 89180966 | N | B CONVERSATION WITH B PATIENT SAFETY MANAGER CONDUCTED ON B B REPORTED THE FOLLOWING PATIENT INFORMATION THE PATIENT HAD A LONG HISTORY OF CHRONIC HEALTH ISSUES INCLUDING CANCER AT A YOUNGER AGE CHRONIC CARDIAC ISSUES LEADING TO HEART FAILURE PREGNANCY WHICH ACCELERATED HEART ISSUE PATIENT DEATH OCCURRED APPROX WEEKS TO A MONTH AFTER THE REPORTED INCIDENT WITH ARTERIAL LINE NO STATEMENT WOULD BE PROVIDED BY A DOCTOR THE CUSTOMER ALSO INDICATED VIA PHONE CONVERSATION ON B THAT THERE WAS NO HARM OR INJURY TO THE PATIENT AS A RESULT OF THE ALLEGED DEFECT THE PATIENT DIED AT A LATER DATE AND THE DEATH HAD NOTHING TO DO WITH THE REPORTED ISSUE PER COMPLAINT REPORTER | Defective Device | Device problem code
6939906 | 89180966 | N | B CONVERSATION WITH B PATIENT SAFETY MANAGER CONDUCTED ON B B REPORTED THE FOLLOWING PATIENT INFORMATION THE PATIENT HAD A LONG HISTORY OF CHRONIC HEALTH ISSUES INCLUDING CANCER AT A YOUNGER AGE CHRONIC CARDIAC ISSUES LEADING TO HEART FAILURE PREGNANCY WHICH ACCELERATED HEART ISSUE PATIENT DEATH OCCURRED APPROX WEEKS TO A MONTH AFTER THE REPORTED INCIDENT WITH ARTERIAL LINE NO STATEMENT WOULD BE PROVIDED BY A DOCTOR THE CUSTOMER ALSO INDICATED VIA PHONE CONVERSATION ON B THAT THERE WAS NO HARM OR INJURY TO THE PATIENT AS A RESULT OF THE ALLEGED DEFECT THE PATIENT DIED AT A LATER DATE AND THE DEATH HAD NOTHING TO DO WITH THE REPORTED ISSUE PER COMPLAINT REPORTER | No Device Problem Found | Medical Evaluation result code
6938939 | 89185620 | N | THE EVENT OF ALLERGIC REACTION AUTOIMMUNE REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY ADDITIONAL COMPLICATIONS AFTER BREAST IMPLANT SURGERY THE FOLLOWING MAY OCCUR ANDOR PERSIST WITH VARYING INTENSITY ANDOR FOR A VARYING LENGTH OF TIME HEMATOMASEROMA IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING AND BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS CAN FORM IN THE TISSUE CAPSULE SURROUNDING THE IMPLANT WITH SYMPTOMS THAT MAY INCLUDE PAIN AND FIRMNESS LYMPHADENOPATHY HAS ALSO BEEN REPORTED IN SOME WOMEN WITH IMPLANTS CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Calcified | Device problem code
6938939 | 89185620 | N | THE EVENT OF ALLERGIC REACTION AUTOIMMUNE REACTION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY ADDITIONAL COMPLICATIONS AFTER BREAST IMPLANT SURGERY THE FOLLOWING MAY OCCUR ANDOR PERSIST WITH VARYING INTENSITY ANDOR FOR A VARYING LENGTH OF TIME HEMATOMASEROMA IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING AND BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS CAN FORM IN THE TISSUE CAPSULE SURROUNDING THE IMPLANT WITH SYMPTOMS THAT MAY INCLUDE PAIN AND FIRMNESS LYMPHADENOPATHY HAS ALSO BEEN REPORTED IN SOME WOMEN WITH IMPLANTS CONNECTIVE TISSUE DISEASE CTD CONCERN OVER THE ASSOCIATION OF BREAST IMPLANTS TO THE DEVELOPMENT OF AUTOIMMUNE OR CONNECTIVE TISSUE DISEASES SUCH AS LUPUS SCLERODERMA OR RHEUMATOID ARTHRITIS WAS RAISED BECAUSE OF CASES REPORTED IN THE LITERATURE WITH SMALL NUMBERS OF WOMEN WITH IMPLANTS A REVIEW OF SEVERAL LARGE EPIDEMIOLOGICAL STUDIES OF WOMEN WITH AND WITHOUT IMPLANTS INDICATES THAT THESE DISEASES ARE NO MORE COMMON IN WOMEN WITH IMPLANTS THAN THOSE IN WOMEN WITHOUT IMPLANTS | Deflation Problem | Device problem code
6940693 | 89223610 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT IS THE CURRENT PATIENT STATUS THE PATIENT RECOVERED DOES THE SURGEON ROUTINELY LOAD AND INSPECT EACH CLIP WITHIN THE JAWS OF THE DEVICE OF THE STRUCTURE BEFORE CLOSING ON THE VESSEL NO INFORMATION HOW MANY CLIPS DOES THE SURGEON ROUTINELY PLACE ON THE PATIENTS SIDE NO INFORMATION WHERE THERE ANY ISSUES NOTED WITH CLIP FORMATION IN THE INITIAL PROCEDURE OR IN THE REOPERATION THE BLEEDING OCCURRED FROM THE LEFT GASTROEPIPLOIC ARTERY IN THE INITIAL PROCEDURE ON DAYS AFTER THE INITIAL PROCEDURE THE PATIENT HAD A FEVER THE BLEEDING OCCURRED FROM AROUND THE LEFT GASTROEPIPLOIC ARTERY POSTOPERATIVELY DR SAID THAT THE BLEEDING SITE WAS NOT IDENTIFIED WHETHER FROM THE LEFT GASTROEPIPLOIC ARTERY OR OTHER SITE NEAR WERE THERE ANY ISSUES WITH THE DEVICE IN THE INITIAL PROCEDURE NO HOW WAS THE BLEED DIAGNOSED NO INFORMATION ADDITIONAL INFORMATION FROM THE SALES REP IN THE REOPERATION HAND SUTURE AND ELECTRICAL SCALPEL WERE USED TO STOP BLEEDING | Insufficient Information | Device problem code
6940693 | 89223610 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT IS THE CURRENT PATIENT STATUS THE PATIENT RECOVERED DOES THE SURGEON ROUTINELY LOAD AND INSPECT EACH CLIP WITHIN THE JAWS OF THE DEVICE OF THE STRUCTURE BEFORE CLOSING ON THE VESSEL NO INFORMATION HOW MANY CLIPS DOES THE SURGEON ROUTINELY PLACE ON THE PATIENTS SIDE NO INFORMATION WHERE THERE ANY ISSUES NOTED WITH CLIP FORMATION IN THE INITIAL PROCEDURE OR IN THE REOPERATION THE BLEEDING OCCURRED FROM THE LEFT GASTROEPIPLOIC ARTERY IN THE INITIAL PROCEDURE ON DAYS AFTER THE INITIAL PROCEDURE THE PATIENT HAD A FEVER THE BLEEDING OCCURRED FROM AROUND THE LEFT GASTROEPIPLOIC ARTERY POSTOPERATIVELY DR SAID THAT THE BLEEDING SITE WAS NOT IDENTIFIED WHETHER FROM THE LEFT GASTROEPIPLOIC ARTERY OR OTHER SITE NEAR WERE THERE ANY ISSUES WITH THE DEVICE IN THE INITIAL PROCEDURE NO HOW WAS THE BLEED DIAGNOSED NO INFORMATION ADDITIONAL INFORMATION FROM THE SALES REP IN THE REOPERATION HAND SUTURE AND ELECTRICAL SCALPEL WERE USED TO STOP BLEEDING | Operational Problem Identified | Medical Evaluation result code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Disability | Patient problem code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Ulcer | Patient problem code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Claudication | Patient problem code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Stenosis | Patient problem code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Insufficient Information | Device problem code
6941019 | 89226469 | D | AN ENDURANT ILIAC LIMB WAS MODIFIED AND IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT IN PREPARATION FOR OPEN REPAIR OF AN ACUTE ISCHEMIA DUE TO AN AORTIC ULCERLIKE LESION A PATIENT PREVIOUSLY OBSERVED AT THE CLINIC FOR STABLE BILATERAL BUTTOCK CLAUDICATION WAS ADMITTED WITH BILATERAL ACUTE ISCHEMIA A COMPUTED TOMOGRAPHY SCAN IDENTIFIED AN INFRARENAL ULCER LIKE AORTIC LESION WHICH CONDITIONED A SIGNIFICANT STENOSIS OF THE AORTA AT THAT LEVEL IN THE CONTEXT OF A HYPOPLASTIC AORTOILIAC AXIS AFTER THREEDIMENSIONAL RECONSTRUCTION AND CASE PLANNING NONDEFERRABLE ENDOVASCULAR TREATMENT WAS COMPLETED BY TRIMMING AN MM MEDTRONIC ENDURANT ILIAC LIMB MM PROXIMAL AND MMDISTAL DIAMETERS THE CENTERLINE RECONSTRUCTION AND THE SUBSEQUENT ANGIOGRAPHIC CONFIRMATION INDICATED A DISTANCE OF MM FROM THE LOWEST RENAL ARTERY TO THE AORTIC BIFURCATION ACCORDINGLY THE TWO PROXIMAL STENTS WERE REMOVED IN A WAY SIMILAR TO THE PREVIOUSLY DESCRIBED CASE FOR AN IMPLANT LENGTH OF MM THE PATIENT RECOVERED PEDAL PULSES IN THE OPERATING ROOM AND WAS DISCHARGED UNDER ANTIPLATELET THERAPY THE DAY AFTER THREE MONTHS AFTER THE PROCEDURE SOME DEGREE OF BUTTOCK CLAUDICATION WAS REPORTED AT THE FOLLOWUP VISIT WHICH CORRELATED WITH BILATERAL COMMON ILIAC STENOSIS IN THE IMAGING STUDIES PERFORMED TO OFFER A LONGLASTING FIX THE PREVIOUS REPAIR WAS EXPLANTED AND AN AORTOBIFEMORAL DACRON BYPASS WAS PERFORMED ALLOWING THE PATIENT TO HAVE FULL RECOVERY OF BOTH PEDAL PULSES AND NO WALKING IMPAIRMENT MONTHS AFTER THE SURGERY | Ischemia | Patient problem code
6940453 | 89229268 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FEELS LIKE THEY RETAINED MORE CC IN THE PAST HOURS THAN THE NIGHT PRIOR CC THEY SLEPT THE WHOLE NIGHT LAST NIGHT WHEREAS THE NIGHT PRIOR THEY WOKE UP TO USE THE RESTROOM MULTIPLE TIMES THE NEXT DAY PATIENT HAD DRIED BLOOD ON THEIR BANDAGE SO THEY HAD THE HOME HEALTH NURSE CHANGE IT FOR THEM NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6940453 | 89229268 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FEELS LIKE THEY RETAINED MORE CC IN THE PAST HOURS THAN THE NIGHT PRIOR CC THEY SLEPT THE WHOLE NIGHT LAST NIGHT WHEREAS THE NIGHT PRIOR THEY WOKE UP TO USE THE RESTROOM MULTIPLE TIMES THE NEXT DAY PATIENT HAD DRIED BLOOD ON THEIR BANDAGE SO THEY HAD THE HOME HEALTH NURSE CHANGE IT FOR THEM NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6941805 | 89233151 | N | THE CUSTOMER RESPONDED TO FOLLOW UP QUESTIONS VIA EMAIL ON B AND INDICATED THAT SHE WAS DIAGNOSED WITH MASTITIS AND PRESCRIBED AN ANTIBIOTIC SHE STATED THAT THE MASTITIS WAS UNDER CONTROL AND CONFIRMED THAT THOUGH SHE CONTINUED TO HAVE ISSUES WITH THE PUMP SHE HAD NOT CALLED BACK TO CONDUCT TROUBLESHOOTING SHE WAS ENCOURAGED TO CALL BACK SO THAT TROUBLESHOOTING COULD BE CONDUCTED TO HELP RESOLVE THE ISSUE SHE WAS EXPERIENCING BASED ON THE RESULTS OF B IT CANNOT BE DEFINITIVELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS MASTITIS IS USUALLY A BENIGN SELFLIMITING INFECTION WITH FEW CONSEQUENCES FOR THE SUCKLING INFANT THE RISK OF MASTITIS IS HIGHER AMONG WOMEN WHO HAVE BREASTFED PREVIOUSLY ESPECIALLY THOSE WITH A HISTORY OR MASTITIS RIORDAN WAMBACH TH ED P BREASTFEEDING AND HUMAN LACTATION MASTITIS REQUIRES PROMPT MEDICAL ATTENTION FOR THE MOTHER FOR PAIN RELIEF AND PRESCRIPTION ANTIBIOTICS TO AVOID PROGRESSION TO OVERWHELMING SEPSIS | Component or Accessory Incompatibility | Device Problem Code
6941805 | 89233151 | N | THE CUSTOMER RESPONDED TO FOLLOW UP QUESTIONS VIA EMAIL ON B AND INDICATED THAT SHE WAS DIAGNOSED WITH MASTITIS AND PRESCRIBED AN ANTIBIOTIC SHE STATED THAT THE MASTITIS WAS UNDER CONTROL AND CONFIRMED THAT THOUGH SHE CONTINUED TO HAVE ISSUES WITH THE PUMP SHE HAD NOT CALLED BACK TO CONDUCT TROUBLESHOOTING SHE WAS ENCOURAGED TO CALL BACK SO THAT TROUBLESHOOTING COULD BE CONDUCTED TO HELP RESOLVE THE ISSUE SHE WAS EXPERIENCING BASED ON THE RESULTS OF B IT CANNOT BE DEFINITIVELY CONCLUDED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE CUSTOMERS MASTITIS MASTITIS IS USUALLY A BENIGN SELFLIMITING INFECTION WITH FEW CONSEQUENCES FOR THE SUCKLING INFANT THE RISK OF MASTITIS IS HIGHER AMONG WOMEN WHO HAVE BREASTFED PREVIOUSLY ESPECIALLY THOSE WITH A HISTORY OR MASTITIS RIORDAN WAMBACH TH ED P BREASTFEEDING AND HUMAN LACTATION MASTITIS REQUIRES PROMPT MEDICAL ATTENTION FOR THE MOTHER FOR PAIN RELIEF AND PRESCRIPTION ANTIBIOTICS TO AVOID PROGRESSION TO OVERWHELMING SEPSIS | Device Difficult to Operate | Medical Evaluation result code
6940036 | 89236142 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WERE THERE ANY ISSUES NOTED WITH DEVICE PERFORMANCE IN INITIAL PROCEDURE THERE ARE IDENTIFIABLE ABNORMAL SHAPED STAPLES ONLY WITH A FEW MINUTE STAPLES AT THE ENDS HE MADE THE MANUAL SUTURE WERE THERE ANY MALFORMED STAPLES IDENTIFIED AT ANY TIME IF SO PLEASE DESCRIBE THEIR SHAPE I DID NOT IDENTIFY ABNORMAL SHAPED STAPLES ONLY HAD A FEW STAPLES ON THE ENDS MADE THE MANUAL SUTURE HOW AND WHEN WAS THE LEAK IDENTIFIED AFTER SURGERY HOW WAS THE LEAK ADDRESSED MANUAL SUTURE WHAT IS THE CURRENT PATIENT STATUS THE PATIENT IS ALREADY AT HOME NORMAL SITUATION WERE THE STAPLES MALFORMED WHAT IS MEANT BY ONLY A FEW STAPLES AT THE ENDS THEY ARE CLIPS THAT ARE AT THE END OF THE CHARGE CHANNEL | Leak / Splash | Device problem code
6940036 | 89236142 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WERE THERE ANY ISSUES NOTED WITH DEVICE PERFORMANCE IN INITIAL PROCEDURE THERE ARE IDENTIFIABLE ABNORMAL SHAPED STAPLES ONLY WITH A FEW MINUTE STAPLES AT THE ENDS HE MADE THE MANUAL SUTURE WERE THERE ANY MALFORMED STAPLES IDENTIFIED AT ANY TIME IF SO PLEASE DESCRIBE THEIR SHAPE I DID NOT IDENTIFY ABNORMAL SHAPED STAPLES ONLY HAD A FEW STAPLES ON THE ENDS MADE THE MANUAL SUTURE HOW AND WHEN WAS THE LEAK IDENTIFIED AFTER SURGERY HOW WAS THE LEAK ADDRESSED MANUAL SUTURE WHAT IS THE CURRENT PATIENT STATUS THE PATIENT IS ALREADY AT HOME NORMAL SITUATION WERE THE STAPLES MALFORMED WHAT IS MEANT BY ONLY A FEW STAPLES AT THE ENDS THEY ARE CLIPS THAT ARE AT THE END OF THE CHARGE CHANNEL | Material Deformation | Device problem code
6940036 | 89236142 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WERE THERE ANY ISSUES NOTED WITH DEVICE PERFORMANCE IN INITIAL PROCEDURE THERE ARE IDENTIFIABLE ABNORMAL SHAPED STAPLES ONLY WITH A FEW MINUTE STAPLES AT THE ENDS HE MADE THE MANUAL SUTURE WERE THERE ANY MALFORMED STAPLES IDENTIFIED AT ANY TIME IF SO PLEASE DESCRIBE THEIR SHAPE I DID NOT IDENTIFY ABNORMAL SHAPED STAPLES ONLY HAD A FEW STAPLES ON THE ENDS MADE THE MANUAL SUTURE HOW AND WHEN WAS THE LEAK IDENTIFIED AFTER SURGERY HOW WAS THE LEAK ADDRESSED MANUAL SUTURE WHAT IS THE CURRENT PATIENT STATUS THE PATIENT IS ALREADY AT HOME NORMAL SITUATION WERE THE STAPLES MALFORMED WHAT IS MEANT BY ONLY A FEW STAPLES AT THE ENDS THEY ARE CLIPS THAT ARE AT THE END OF THE CHARGE CHANNEL | Leakage/Seal | Medical Evaluation result code
6940897 | 89237097 | D | AN ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM X MM ABDOMINAL AORTIC ANEURYSM ON AN UNKNOWN DATE CURRENTLY ON CT THE AORTIC NECK DIAMETER WAS MM AND THE LEFT ILIAC LIMB DIAMETER MEASURED MM MM MM AND MM FROM PROXIMAL TO DISTAL IT WAS REPORTED THAT THE PATIENT PRESENTED EMERGENTLY WITH ABDOMINAL PAIN CT IMAGES REVEALED A TYPE B ENDOLEAK FROM LEFT COMMON ILIAC ARTERY ON THE MAINBODY SIDE THE CT ALSO SHOWED THAT THE PROXIMAL EDGE OF THE STENT GRAFT ONLY HAD SEAL FOR MM BEFORE THE ABDOMINAL AORTIC ANEURYSM THE LOWEST RIGHT RENAL ARTERY WAS MM PROXIMAL TO THE ANEURYSM SAC A MM DIAMETER ENDURANT CUFF WAS PLACED AT THE LEVEL OF THE RENAL ARTERIES AND DIRECTLY POST CUFF PLACEMENT ANOTHER ANGIOGRAM WAS PREFORMED TO PROVE THE PATENCY OF THE RENAL ARTERIES THIS WAS FOLLOWED BY THE IMPLANT OF ENDOANCHORS USING ANGIOGRAM THE ORIGIN OF THE LEFT HYPOGASTRIC WAS IDENTIFIED AND MARKED AND THE VESSEL WAS EMBOLISED A STRAIGHT SEGMENT OF THE LEFT EXTERNAL ILIAC WAS IDENTIFIED AND A XX ENDURANT LIMB WAS DEPLOYED AND BALLOONED WITH A RELIANT A COMPLETION ANGIOGRAM WAS PERFORMED SHOWING THAT THE RENAL ARTERIES WERE PATENT THE LEFT TYPE B ENDOLEAK WAS RESOLVED AND THE RIGHT LIMB WAS WIDELY PATENT WITH GOOD SEAL AND APPOSITION ACCORDING TO THE PHYSICIAN THE CAUSE OF THE EVENT WAS DISEASE PROGRESSION NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | No Consequences Or Impact To Patient | Patient problem code
6940897 | 89237097 | D | AN ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM X MM ABDOMINAL AORTIC ANEURYSM ON AN UNKNOWN DATE CURRENTLY ON CT THE AORTIC NECK DIAMETER WAS MM AND THE LEFT ILIAC LIMB DIAMETER MEASURED MM MM MM AND MM FROM PROXIMAL TO DISTAL IT WAS REPORTED THAT THE PATIENT PRESENTED EMERGENTLY WITH ABDOMINAL PAIN CT IMAGES REVEALED A TYPE B ENDOLEAK FROM LEFT COMMON ILIAC ARTERY ON THE MAINBODY SIDE THE CT ALSO SHOWED THAT THE PROXIMAL EDGE OF THE STENT GRAFT ONLY HAD SEAL FOR MM BEFORE THE ABDOMINAL AORTIC ANEURYSM THE LOWEST RIGHT RENAL ARTERY WAS MM PROXIMAL TO THE ANEURYSM SAC A MM DIAMETER ENDURANT CUFF WAS PLACED AT THE LEVEL OF THE RENAL ARTERIES AND DIRECTLY POST CUFF PLACEMENT ANOTHER ANGIOGRAM WAS PREFORMED TO PROVE THE PATENCY OF THE RENAL ARTERIES THIS WAS FOLLOWED BY THE IMPLANT OF ENDOANCHORS USING ANGIOGRAM THE ORIGIN OF THE LEFT HYPOGASTRIC WAS IDENTIFIED AND MARKED AND THE VESSEL WAS EMBOLISED A STRAIGHT SEGMENT OF THE LEFT EXTERNAL ILIAC WAS IDENTIFIED AND A XX ENDURANT LIMB WAS DEPLOYED AND BALLOONED WITH A RELIANT A COMPLETION ANGIOGRAM WAS PERFORMED SHOWING THAT THE RENAL ARTERIES WERE PATENT THE LEFT TYPE B ENDOLEAK WAS RESOLVED AND THE RIGHT LIMB WAS WIDELY PATENT WITH GOOD SEAL AND APPOSITION ACCORDING TO THE PHYSICIAN THE CAUSE OF THE EVENT WAS DISEASE PROGRESSION NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Insufficient Information | Device problem code
6947254 | 89237450 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER BACK FOR BETWEEN AND HOURS SHE WAS UNSURE IF THE CANNULA WAS PROPERLY INSERTED AND FELT INSULIN ON THE SKIN | No Known Impact Or Consequence To Patient | Patient Problem Code
6947254 | 89237450 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER BACK FOR BETWEEN AND HOURS SHE WAS UNSURE IF THE CANNULA WAS PROPERLY INSERTED AND FELT INSULIN ON THE SKIN | Positioning Problem | Device problem code
6939508 | 89239622 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH VERIO IQ METER WAS READING BLOOD SAMPLES AS CONTROL SOLUTION THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT AM ON B THE PATIENT CLAIMED THAT HE TESTED HIS BLOOD GLUCOSE WITH THE SUBJECT METER AND OBTAINED A RESULT OF MGDL WHICH THE SUBJECT METER FLAGGED AS A CONTROL SOLUTION TEST RESULT THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN TYPE AND DOSE OF MEDICATION WAS NOT REPORTED AND SHE DENIED MAKING ANY CHANGES TO HER USUAL DIABETES MANAGEMENT REGIMEN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT IMMEDIATELY AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF FEELING THIRST TIREDNESS EXHAUSTION AND EXCESSIVE PEEING THE PATIENT DENIED RECEIVING MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE AT THE TIME OF TROUBLESHOOTING THE CSR CONFIRMED THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR WALKED THE PATIENT THROUGH A RETEST HOWEVER THE ISSUE REMAINED UNRESOLVED REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED INACCURACY ISSUE BEGAN | Injury | Patient problem code
6939508 | 89239622 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT HER ONETOUCH VERIO IQ METER WAS READING BLOOD SAMPLES AS CONTROL SOLUTION THE COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER SERVICE REPRESENTATIVE CSR DOCUMENTATION THE PATIENT REPORTED THAT THE ALLEGED ISSUE BEGAN AT AM ON B THE PATIENT CLAIMED THAT HE TESTED HIS BLOOD GLUCOSE WITH THE SUBJECT METER AND OBTAINED A RESULT OF MGDL WHICH THE SUBJECT METER FLAGGED AS A CONTROL SOLUTION TEST RESULT THE PATIENT MANAGES HER DIABETES WITH SELFADJUSTED INSULIN TYPE AND DOSE OF MEDICATION WAS NOT REPORTED AND SHE DENIED MAKING ANY CHANGES TO HER USUAL DIABETES MANAGEMENT REGIMEN IN RESPONSE TO THE ALLEGED ISSUE THE PATIENT REPORTED THAT IMMEDIATELY AFTER THE ALLEGED ISSUE OCCURRED SHE DEVELOPED SYMPTOMS OF FEELING THIRST TIREDNESS EXHAUSTION AND EXCESSIVE PEEING THE PATIENT DENIED RECEIVING MEDICAL TREATMENT AS A RESULT OF THE ALLEGED ISSUE AT THE TIME OF TROUBLESHOOTING THE CSR CONFIRMED THAT THE PATIENT WAS FOLLOWING THE CORRECT TESTING PROCEDURE AND THAT THE TEST STRIPS HAD BEEN STORED PROPERLY WERE NOT OPEN BEYOND THEIR DISCARD DATE AND HAD NOT EXPIRED THE CSR WALKED THE PATIENT THROUGH A RETEST HOWEVER THE ISSUE REMAINED UNRESOLVED REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED INACCURACY ISSUE BEGAN | Insufficient Information | Device problem code
6882036 | 89246320 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6882036 | 89246320 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6941835 | 89246417 | N | THE EVENT OF VARIED INJURIES ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Deflation Problem | Device problem code
6941835 | 89246417 | N | THE EVENT OF VARIED INJURIES ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Calcified | Device problem code
6941835 | 89246417 | N | THE EVENT OF VARIED INJURIES ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Results Pending Completion of Investigation | Medical Evaluation result code
6884091 | 89251210 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE ILLUMINATOR INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE A VISUAL INSPECTION WAS PERFORMED AND THE FRONT OF THE UNIT WAS FOUND TO BE CLOGGED WITH DUST THE UNIT WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH AN ERROR THE COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6884091 | 89251210 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE ILLUMINATOR INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE A VISUAL INSPECTION WAS PERFORMED AND THE FRONT OF THE UNIT WAS FOUND TO BE CLOGGED WITH DUST THE UNIT WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH AN ERROR THE COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Environmental Particulates | Device problem code
6884091 | 89251210 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE ILLUMINATOR INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE A VISUAL INSPECTION WAS PERFORMED AND THE FRONT OF THE UNIT WAS FOUND TO BE CLOGGED WITH DUST THE UNIT WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH AN ERROR THE COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6884091 | 89251210 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE ILLUMINATOR INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE A VISUAL INSPECTION WAS PERFORMED AND THE FRONT OF THE UNIT WAS FOUND TO BE CLOGGED WITH DUST THE UNIT WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH AN ERROR THE COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Complete Blockage | Device problem code
6884091 | 89251210 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE ILLUMINATOR INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE A VISUAL INSPECTION WAS PERFORMED AND THE FRONT OF THE UNIT WAS FOUND TO BE CLOGGED WITH DUST THE UNIT WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH AN ERROR THE COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Environment Problem Identified | Medical Evaluation result code
6885024 | 89253166 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED THROUGH LOG FILE ANALYSIS WHICH REVEALED A CONTROLLER FAULT ALARM ON THE REPORTED EVENT DATE WITH ERROR CODES INDICATING SYSUICDATAFLASHFAIL CONTROLLER FAULTS OF TYPE DATAFLASHFAIL ARE SUSPECTED TO BE RELATED TO THE UIC FAILURE TO UPDATE A PARAMETER SETTING IN THE DATA FLASH MEMORY THIS WOULD PREVENT THE NEW PARAMETER SETTINGS FROM BEING SAVED AND MIGHT CAUSE THE SETTINGS TO BE LOST ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN INABILITY TO ACCEPT ANY PARAMETER ADJUSTMENTS THAT INVOLVE THE UIC FROM THE MONITOR THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO THE UIC FAILURE TO UPDATE PARAMETER SETTINGS IN DATA FLASH MEMORY PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device Problem Code
6885024 | 89253166 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED THROUGH LOG FILE ANALYSIS WHICH REVEALED A CONTROLLER FAULT ALARM ON THE REPORTED EVENT DATE WITH ERROR CODES INDICATING SYSUICDATAFLASHFAIL CONTROLLER FAULTS OF TYPE DATAFLASHFAIL ARE SUSPECTED TO BE RELATED TO THE UIC FAILURE TO UPDATE A PARAMETER SETTING IN THE DATA FLASH MEMORY THIS WOULD PREVENT THE NEW PARAMETER SETTINGS FROM BEING SAVED AND MIGHT CAUSE THE SETTINGS TO BE LOST ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN INABILITY TO ACCEPT ANY PARAMETER ADJUSTMENTS THAT INVOLVE THE UIC FROM THE MONITOR THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO THE UIC FAILURE TO UPDATE PARAMETER SETTINGS IN DATA FLASH MEMORY PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Flare or Flash | Device problem code
6885024 | 89253166 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE THE REPORTED EVENT WAS CONFIRMED THROUGH LOG FILE ANALYSIS WHICH REVEALED A CONTROLLER FAULT ALARM ON THE REPORTED EVENT DATE WITH ERROR CODES INDICATING SYSUICDATAFLASHFAIL CONTROLLER FAULTS OF TYPE DATAFLASHFAIL ARE SUSPECTED TO BE RELATED TO THE UIC FAILURE TO UPDATE A PARAMETER SETTING IN THE DATA FLASH MEMORY THIS WOULD PREVENT THE NEW PARAMETER SETTINGS FROM BEING SAVED AND MIGHT CAUSE THE SETTINGS TO BE LOST ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE FAILED TO MEET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE TO AN INABILITY TO ACCEPT ANY PARAMETER ADJUSTMENTS THAT INVOLVE THE UIC FROM THE MONITOR THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO THE UIC FAILURE TO UPDATE PARAMETER SETTINGS IN DATA FLASH MEMORY PER THE INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Results Pending Completion of Investigation | Medical Evaluation result code
6943322 | 89256067 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6943322 | 89256067 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6943735 | 89256087 | N | INFORMATION RECEIVED SUGGESTS THAT THIS IS NOT AN ALLEGED DEFECT OF THE DEVICE THE PATIENT FELL REQUIRING REVISION SURGERY | Defective Device | Device problem code
6943317 | 89257077 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component or Accessory Incompatibility | Device problem code
6943317 | 89257077 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Incompatible Component/ Accessory | Medical Evaluation result code
6943423 | 89257769 | D | THE PATIENT REPORTED GOING TO THE EMERGENCY ROOM AND BEING DIAGNOSED WITH HYPERGLYCEMIA HER BLOOD GLUCOSE HAD REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR BETWEEN AND HOURS TREATMENT INCLUDED TWO BAGS OF IV | Hyperglycemia | Patient problem code
6943423 | 89257769 | D | THE PATIENT REPORTED GOING TO THE EMERGENCY ROOM AND BEING DIAGNOSED WITH HYPERGLYCEMIA HER BLOOD GLUCOSE HAD REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR BETWEEN AND HOURS TREATMENT INCLUDED TWO BAGS OF IV | Insufficient Information | Device problem code
6943501 | 89259326 | N | B PER THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE IFU ADVERSE EVENTS THAT MAY OCCUR AND OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK AND ANEURYSM ENLARGEMENT | No Device Problem Found | Medical Evaluation result code
6946433 | 89261074 | N | B RESULTS OF INVESTIGATION A VYAIRE MEDICAL FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR DID NOT FIND ANY LEAKS ON THE VENTILATOR THE VENTILATOR PASSED THE PATIENT CIRCUIT CALIBRATION AND PERFORMANCE CHECK CORRECTLY THE FSR PERFORMED THE DRIVER DISPLACEMENT INDICATOR DDI CALIBRATION AND RAN THE VENTILATOR TO VERIFY THAT THE DEVICE MAINTAINED THE SETTINGS CORRECTLY THE DEVICE MEETS MANUFACTURERS SPECIFICATIONS | Calibration Problem | Device problem code
6946433 | 89261074 | N | B RESULTS OF INVESTIGATION A VYAIRE MEDICAL FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR DID NOT FIND ANY LEAKS ON THE VENTILATOR THE VENTILATOR PASSED THE PATIENT CIRCUIT CALIBRATION AND PERFORMANCE CHECK CORRECTLY THE FSR PERFORMED THE DRIVER DISPLACEMENT INDICATOR DDI CALIBRATION AND RAN THE VENTILATOR TO VERIFY THAT THE DEVICE MAINTAINED THE SETTINGS CORRECTLY THE DEVICE MEETS MANUFACTURERS SPECIFICATIONS | Leak / Splash | Device problem code
6946433 | 89261074 | N | B RESULTS OF INVESTIGATION A VYAIRE MEDICAL FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR DID NOT FIND ANY LEAKS ON THE VENTILATOR THE VENTILATOR PASSED THE PATIENT CIRCUIT CALIBRATION AND PERFORMANCE CHECK CORRECTLY THE FSR PERFORMED THE DRIVER DISPLACEMENT INDICATOR DDI CALIBRATION AND RAN THE VENTILATOR TO VERIFY THAT THE DEVICE MAINTAINED THE SETTINGS CORRECTLY THE DEVICE MEETS MANUFACTURERS SPECIFICATIONS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6946433 | 89261074 | N | B RESULTS OF INVESTIGATION A VYAIRE MEDICAL FIELD SERVICE REPRESENTATIVE FSR EVALUATED THE DEVICE ONSITE THE FSR DID NOT FIND ANY LEAKS ON THE VENTILATOR THE VENTILATOR PASSED THE PATIENT CIRCUIT CALIBRATION AND PERFORMANCE CHECK CORRECTLY THE FSR PERFORMED THE DRIVER DISPLACEMENT INDICATOR DDI CALIBRATION AND RAN THE VENTILATOR TO VERIFY THAT THE DEVICE MAINTAINED THE SETTINGS CORRECTLY THE DEVICE MEETS MANUFACTURERS SPECIFICATIONS | Leakage/Seal | Medical Evaluation result code
6945028 | 89261099 | N | THIS RIGHT VENTRICULAR RV LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATION OF PERFORATION LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED ISSUE AS THE TIP AND HELIX OF THE LEAD WERE INTACT AND APPEARED UNDAMAGED RESISTANCE TESTING DETERMINED THE LEAD WAS ELECTRICALLY CONTINUOUS | Retraction Problem | Device problem code
6945028 | 89261099 | N | THIS RIGHT VENTRICULAR RV LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATION OF PERFORATION LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED ISSUE AS THE TIP AND HELIX OF THE LEAD WERE INTACT AND APPEARED UNDAMAGED RESISTANCE TESTING DETERMINED THE LEAD WAS ELECTRICALLY CONTINUOUS | Material Perforation | Device Problem Code
6945028 | 89261099 | N | THIS RIGHT VENTRICULAR RV LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATION OF PERFORATION LABORATORY ANALYSIS WAS UNABLE TO CONCLUSIVELY DETERMINE THE ROOT CAUSE OF THE REPORTED ISSUE AS THE TIP AND HELIX OF THE LEAD WERE INTACT AND APPEARED UNDAMAGED RESISTANCE TESTING DETERMINED THE LEAD WAS ELECTRICALLY CONTINUOUS | Non-Functional Defect | Medical Evaluation result code
6943843 | 89263644 | D | MALFUNCTIONING RIGHT VENTRICULAR PACING LEAD PRESENTING WITH HISTORY OF COMPLETE HEART BLOCK INTERMITTENTLY AND PACER MALFUNCTION IN THE FORM OF NONCAPTURE OF THE RIGHT VENTRICULAR LEAD WITH DIZZINESS AND SYNCOPE | Dizziness | Patient problem code
6943843 | 89263644 | D | MALFUNCTIONING RIGHT VENTRICULAR PACING LEAD PRESENTING WITH HISTORY OF COMPLETE HEART BLOCK INTERMITTENTLY AND PACER MALFUNCTION IN THE FORM OF NONCAPTURE OF THE RIGHT VENTRICULAR LEAD WITH DIZZINESS AND SYNCOPE | Syncope | Patient problem code
6943843 | 89263644 | D | MALFUNCTIONING RIGHT VENTRICULAR PACING LEAD PRESENTING WITH HISTORY OF COMPLETE HEART BLOCK INTERMITTENTLY AND PACER MALFUNCTION IN THE FORM OF NONCAPTURE OF THE RIGHT VENTRICULAR LEAD WITH DIZZINESS AND SYNCOPE | Insufficient Information | Device problem code
6943843 | 89263644 | D | MALFUNCTIONING RIGHT VENTRICULAR PACING LEAD PRESENTING WITH HISTORY OF COMPLETE HEART BLOCK INTERMITTENTLY AND PACER MALFUNCTION IN THE FORM OF NONCAPTURE OF THE RIGHT VENTRICULAR LEAD WITH DIZZINESS AND SYNCOPE | Complete Heart Block | Patient problem code
6944129 | 89265207 | D | MEDTRONIC RECEIVED INFORMATION THAT TEN DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6944129 | 89265207 | D | MEDTRONIC RECEIVED INFORMATION THAT TEN DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6943739 | 89270461 | N | INFORMATION RECEIVED SUGGESTS THAT THIS IS NOT AN ALLEGED DEFECT OF THE DEVICE THE PATIENT FELL REQUIRING REVISION SURGERY | Defective Device | Device problem code
6943739 | 89270461 | N | INFORMATION RECEIVED SUGGESTS THAT THIS IS NOT AN ALLEGED DEFECT OF THE DEVICE THE PATIENT FELL REQUIRING REVISION SURGERY | No Findings Available | Medical Evaluation result code
6943919 | 89271930 | N | INFORMATION RECEIVED SUGGESTS THAT THIS IS NOT AN ALLEGED DEFECT OF THE DEVICE THE PATIENT FELL REQUIRING REVISION SURGERY | Defective Device | Device problem code
6943919 | 89271930 | N | INFORMATION RECEIVED SUGGESTS THAT THIS IS NOT AN ALLEGED DEFECT OF THE DEVICE THE PATIENT FELL REQUIRING REVISION SURGERY | No Findings Available | Medical Evaluation result code
6943405 | 89272220 | N | B JOURNAL REFERENCE OPEN CARDIOVASCULAR MEDICINE JOURNAL CORONARY PERFORATION OF DISTAL DIAGONAL BRANCH FOLLOWED BY PROLONGED RECURRENT CARDIAC TAMPONADE FINALLY RESOLVED WITH PERICARDIOTOMY THE POTENTIAL RISK OF HYDROPHILIC GUIDEWIRES DOI | Material Perforation | Device problem code
6943405 | 89272220 | N | B JOURNAL REFERENCE OPEN CARDIOVASCULAR MEDICINE JOURNAL CORONARY PERFORATION OF DISTAL DIAGONAL BRANCH FOLLOWED BY PROLONGED RECURRENT CARDIAC TAMPONADE FINALLY RESOLVED WITH PERICARDIOTOMY THE POTENTIAL RISK OF HYDROPHILIC GUIDEWIRES DOI | No Device Problem Found | Medical Evaluation result code
6888974 | 89288939 | N | THE DEVICE WAS IN THE ANTERIOR MEDIASTINUM THE CLEARANCE LOOP SEPARATION WAS NOTICED ABOUT HOURS AFTER SURGERY IN THE CVICU AN XRAY WAS TAKEN AND SHOWED THAT THE CLEARANCE LOOP WAS INSIDE THE CHEST TUBE THE CLEARANCE APPARATUS WAS REMOVED AND THE CHEST TUBE WAS CONNECTED TO THE DRAINAGE CANISTER AT END OF USE A CT SCAN CONFIRMED THAT THE LOOP WAS NOT IN THE CHEST | Leak / Splash | Device problem code
6888974 | 89288939 | N | THE DEVICE WAS IN THE ANTERIOR MEDIASTINUM THE CLEARANCE LOOP SEPARATION WAS NOTICED ABOUT HOURS AFTER SURGERY IN THE CVICU AN XRAY WAS TAKEN AND SHOWED THAT THE CLEARANCE LOOP WAS INSIDE THE CHEST TUBE THE CLEARANCE APPARATUS WAS REMOVED AND THE CHEST TUBE WAS CONNECTED TO THE DRAINAGE CANISTER AT END OF USE A CT SCAN CONFIRMED THAT THE LOOP WAS NOT IN THE CHEST | Material Separation | Device problem code
6888974 | 89288939 | N | THE DEVICE WAS IN THE ANTERIOR MEDIASTINUM THE CLEARANCE LOOP SEPARATION WAS NOTICED ABOUT HOURS AFTER SURGERY IN THE CVICU AN XRAY WAS TAKEN AND SHOWED THAT THE CLEARANCE LOOP WAS INSIDE THE CHEST TUBE THE CLEARANCE APPARATUS WAS REMOVED AND THE CHEST TUBE WAS CONNECTED TO THE DRAINAGE CANISTER AT END OF USE A CT SCAN CONFIRMED THAT THE LOOP WAS NOT IN THE CHEST | Fracture Problem | Medical Evaluation result code
6948955 | 89289898 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Component Missing | Device problem code
6948955 | 89289898 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Missing or Inadequate Safety Measures | Medical Evaluation result code
6948952 | 89291270 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Component Missing | Device problem code
6948952 | 89291270 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Missing or Inadequate Safety Measures | Medical Evaluation result code
6948932 | 89293859 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Component Missing | Device problem code
6948932 | 89293859 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Missing or Inadequate Safety Measures | Medical Evaluation result code
6949009 | 89294292 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Component Missing | Device problem code
6949009 | 89294292 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Missing or Inadequate Safety Measures | Medical Evaluation result code
6949031 | 89295016 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Component Missing | Device problem code
6949031 | 89295016 | N | IT WAS REPORTED THAT FIVE DIFFERENT INFUSION SETS WERE ATTEMPTED TO BE USED AND NOTED TO HAVE ADHESIVE ROLLED UP ON ITSELF AND ANOTHER WITH ADHESIVE MISSING SEE MFR | Missing or Inadequate Safety Measures | Medical Evaluation result code
6945519 | 89339959 | N | PERFORMANCE VARIES IN AND OUTOFSPECIFICATION WE HAVE EXTENSIVELY TESTED VITRECTOMY OUR R D DEPARTMENT HAS MADE A CHANGE AS WAS ALREADY INDICATED IN THE B THE NEW VITRECTOMY WILL SOON BE AVAILABLE OUR CSC DEPARTMENT CAN TELL WHEN THE NEW SERIES IS AVAILABLE NO FURTHER INFORMATION AVAILABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER B THERE IS LIMITED INFORMATION AND NO INFORMATION ON THE OUTCOME OF THIS PATIENT | Insufficient Information | Device problem code
6945519 | 89339959 | N | PERFORMANCE VARIES IN AND OUTOFSPECIFICATION WE HAVE EXTENSIVELY TESTED VITRECTOMY OUR R D DEPARTMENT HAS MADE A CHANGE AS WAS ALREADY INDICATED IN THE B THE NEW VITRECTOMY WILL SOON BE AVAILABLE OUR CSC DEPARTMENT CAN TELL WHEN THE NEW SERIES IS AVAILABLE NO FURTHER INFORMATION AVAILABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER B THERE IS LIMITED INFORMATION AND NO INFORMATION ON THE OUTCOME OF THIS PATIENT | Results Pending Completion of Investigation | Medical Evaluation result code
6950675 | 89342147 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950675 | 89342147 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | No Findings Available | Medical Evaluation result code
6950680 | 89343221 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950680 | 89343221 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6950697 | 89343246 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950697 | 89343246 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6950674 | 89343290 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950674 | 89343290 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6950676 | 89343381 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950676 | 89343381 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6950722 | 89343427 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950722 | 89343427 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6950678 | 89343431 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6950678 | 89343431 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Operational Problem Identified | Medical Evaluation result code
6948433 | 89349717 | N | THE DEVICE WAS NOT EXPLANTED THE MANUFACTURING RECORDS WERE REVIEWED AND NO NONCONFORMITIES WERE FOUND A REVIEW OF THE IPG DIAGNOSTIC DATA FOUND NO EVIDENCE OF A DEVICE ISSUE RELATED TO THE NATURE OF THE COMPLAINT | No Device Problem Found | Medical Evaluation result code
6948104 | 89358087 | N | DEVICE ISSUE WITH INOMAX DSIR B WAS CREATED AS B THE DEVICE INVESTIGATION WAS COMPLETED ON B THE REGIONAL SERVICE CENTER RSC SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO APPARENT IPL FAILURE FOLLOWED BY A SYSTEM SHUTDOWN ALARM DUE TO AN INTER PROCESSOR LINK COMMUNICATIONS FAILURE IN SEQUENCE B CONSISTENT WITH THE REPORTED EVENT TIMING THE SYSTEM SHUTDOWN ALARM IS CONSISTENT WITH THE REPORTED CONTINUOUS AUDIBLE ALARM AND SERVICE REQUIRED SCREEN DISPLAY NOT PREVIOUSLY SEEN BY THE USER A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE DEVICE SERVICE POOL THE ROOT CAUSE FOR THE PATIENT OXYGEN DESATUATION WAS DUE TO THE DELIVERY FAILURE ALARM FOLLOWED BY A SYSTEM SHUTDOWN WHICH STOPPED THE DEVICE FROM DELIVERING INOMAX THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM | Failure to Deliver | Device problem code
6948104 | 89358087 | N | DEVICE ISSUE WITH INOMAX DSIR B WAS CREATED AS B THE DEVICE INVESTIGATION WAS COMPLETED ON B THE REGIONAL SERVICE CENTER RSC SERVICE LOG REVIEW REVEALED A DELIVERY FAILURE ALARM DUE TO APPARENT IPL FAILURE FOLLOWED BY A SYSTEM SHUTDOWN ALARM DUE TO AN INTER PROCESSOR LINK COMMUNICATIONS FAILURE IN SEQUENCE B CONSISTENT WITH THE REPORTED EVENT TIMING THE SYSTEM SHUTDOWN ALARM IS CONSISTENT WITH THE REPORTED CONTINUOUS AUDIBLE ALARM AND SERVICE REQUIRED SCREEN DISPLAY NOT PREVIOUSLY SEEN BY THE USER A FULL FUNCTIONAL TEST WAS PERFORMED AND THE DEVICE OPERATED ACCORDING TO SPECIFICATIONS SO IT WAS RETURNED TO THE DEVICE SERVICE POOL THE ROOT CAUSE FOR THE PATIENT OXYGEN DESATUATION WAS DUE TO THE DELIVERY FAILURE ALARM FOLLOWED BY A SYSTEM SHUTDOWN WHICH STOPPED THE DEVICE FROM DELIVERING INOMAX THIS CONDITION WILL BE TRACKED AND TRENDED UNDER THE COMPANYS QUALITY SYSTEM | Alarm System Problem Identified | Medical Evaluation result code
6950597 | 89358533 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Material Discolored | Device problem code
6950597 | 89358533 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Non-Functional Defect | Medical Evaluation result code
6950596 | 89360720 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6882711 | 89371743 | N | TRACEABILITY REVIEWED AND INSPECTION RECORD DOES NOT REVEAL ANY NON COMPLIANCE ON THIS LOT NO INFORMATION ABOUT THE RETURN | Insufficient Information | Device problem code
6917677 | 89380795 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Positioning Failure | Device Problem Code
6917677 | 89380795 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Assembly Problem Identified | Medical Evaluation result code
6917588 | 89380902 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917588 | 89380902 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6917690 | 89381088 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917690 | 89381088 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6917616 | 89382100 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917616 | 89382100 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6917628 | 89382106 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917628 | 89382106 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6917599 | 89386735 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917599 | 89386735 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6885715 | 89388903 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6885715 | 89388903 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Incompatible Component/ Accessory | Medical Evaluation result code
6924070 | 89390355 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS NOTED THAT THE DEVICE EXPERIENCED INTERMITTENT CONNECTION WITH THE PROGRAMMER HOWEVER PASSED FUNCTIONAL TESTING | Connection Problem | Device problem code
6924070 | 89390355 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS NOTED THAT THE DEVICE EXPERIENCED INTERMITTENT CONNECTION WITH THE PROGRAMMER HOWEVER PASSED FUNCTIONAL TESTING | No Device Problem Found | Medical Evaluation result code
6917726 | 89406953 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6917726 | 89406953 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6952288 | 89408696 | D | IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL AFTER UNDERGOING A TRIAL IMPLANT PROCEDURE DUE TO FOOT PAIN AND WEAKNESS IN HIS LOWER LEG AND FOOT PATIENT WAS OBSERVED IN THE HOSPITAL AND THE ISSUE RESOLVED ON ITS OWN PATIENT WAS LATER DISCHARGED TO HOME | Insufficient Information | Device problem code
6952288 | 89408696 | D | IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL AFTER UNDERGOING A TRIAL IMPLANT PROCEDURE DUE TO FOOT PAIN AND WEAKNESS IN HIS LOWER LEG AND FOOT PATIENT WAS OBSERVED IN THE HOSPITAL AND THE ISSUE RESOLVED ON ITS OWN PATIENT WAS LATER DISCHARGED TO HOME | Weakness | Patient problem code
6952288 | 89408696 | D | IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL AFTER UNDERGOING A TRIAL IMPLANT PROCEDURE DUE TO FOOT PAIN AND WEAKNESS IN HIS LOWER LEG AND FOOT PATIENT WAS OBSERVED IN THE HOSPITAL AND THE ISSUE RESOLVED ON ITS OWN PATIENT WAS LATER DISCHARGED TO HOME | Pain | Patient problem code
6897908 | 89411739 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA AND ITS RELEASE BY FDA DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OR THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER THEIR EMPLOYEES OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6897908 | 89411739 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA AND ITS RELEASE BY FDA DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OR THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER THEIR EMPLOYEES OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6921126 | 89414246 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS NOTED THAT THE RETURNED PATIENT CABLE FAILED A CONTINUITY TEST | No Device Problem Found | Medical Evaluation result code
6891839 | 89414826 | N | THE INVESTIGATION DETERMINED THAT HIGHER AND LOWER THAN EXPECTED RESULTS WERE OBTAINED FROM TWO DIFFERENT LEVELS OF NONVITROS BIORAD CONTROLS AND A SINGLE LEVEL OF VITROS TDM PV FLUID USING TWO DIFFERENT VITROS VALP REAGENT LOTS ON A VITROS FS SYSTEM THE MOST LIKELY ASSIGNABLE CAUSE IS AN INSTRUMENT RELATED EVENT A VITROS GENT WITHINRUN PRECISION TEST USED TO ASSESS INSTRUMENT PERFORMANCE WAS OUTSIDE OF ACCEPTABLE GUIDELINES INDICATING THE VITROS FS SYSTEM WAS NOT PERFORMING AS INTENDED AN ORTHO FIELD ENGINEER PERFORMED A RESET OF THE BELT TENSION FOR SECONDARY METERING PERFORMED ALL SECONDARY METERING ADJUSTMENTS AND PERFORMED THE XARM AT READ ADJUSTMENT FOR READING THE CUVETTE IN THE PHOTOMETER A POST SERVICE VITROS GENT WITHINRUN PRECISION TEST WAS WITHIN ACCEPTABLE GUIDELINES HOWEVER ALL VITROS VALP RESULTS WERE LOWER THAN EXPECTED AFTER SERVICE ACCEPTABLE PERFORMANCE WAS OBTAINED AFTER THE CUSTOMER PERFORMED THE ROUTINE MAINTENANCE PROCEDURES OF REPLACING THE PHOTOMETER LAMP PERFORMING A WATER BLANK AND RECALIBRATING VITROS VALP REAGENT LOT | Low Test Results | Device problem code
6891839 | 89414826 | N | THE INVESTIGATION DETERMINED THAT HIGHER AND LOWER THAN EXPECTED RESULTS WERE OBTAINED FROM TWO DIFFERENT LEVELS OF NONVITROS BIORAD CONTROLS AND A SINGLE LEVEL OF VITROS TDM PV FLUID USING TWO DIFFERENT VITROS VALP REAGENT LOTS ON A VITROS FS SYSTEM THE MOST LIKELY ASSIGNABLE CAUSE IS AN INSTRUMENT RELATED EVENT A VITROS GENT WITHINRUN PRECISION TEST USED TO ASSESS INSTRUMENT PERFORMANCE WAS OUTSIDE OF ACCEPTABLE GUIDELINES INDICATING THE VITROS FS SYSTEM WAS NOT PERFORMING AS INTENDED AN ORTHO FIELD ENGINEER PERFORMED A RESET OF THE BELT TENSION FOR SECONDARY METERING PERFORMED ALL SECONDARY METERING ADJUSTMENTS AND PERFORMED THE XARM AT READ ADJUSTMENT FOR READING THE CUVETTE IN THE PHOTOMETER A POST SERVICE VITROS GENT WITHINRUN PRECISION TEST WAS WITHIN ACCEPTABLE GUIDELINES HOWEVER ALL VITROS VALP RESULTS WERE LOWER THAN EXPECTED AFTER SERVICE ACCEPTABLE PERFORMANCE WAS OBTAINED AFTER THE CUSTOMER PERFORMED THE ROUTINE MAINTENANCE PROCEDURES OF REPLACING THE PHOTOMETER LAMP PERFORMING A WATER BLANK AND RECALIBRATING VITROS VALP REAGENT LOT | Results Pending Completion of Investigation | Medical Evaluation result code
6952233 | 89414834 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6911149 | 89428554 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6911149 | 89428554 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6902913 | 89430257 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT B DISCHARGED FROM TO OVER THE COURSE OF HOURS NEAR THE REPORTED EVENT DATE NO ANOMALIES WERE DETECTED ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT COULD NOT BE DUPLICATED AT THE BENCH LEVEL THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY ALWAYS WAIT UNTIL THE READY TURNS ON TO DISCONNECT THE BATTERY FROM THE BATTERY CHARGER HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Disconnection | Device problem code
6902913 | 89430257 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT B DISCHARGED FROM TO OVER THE COURSE OF HOURS NEAR THE REPORTED EVENT DATE NO ANOMALIES WERE DETECTED ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT COULD NOT BE DUPLICATED AT THE BENCH LEVEL THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY ALWAYS WAIT UNTIL THE READY TURNS ON TO DISCONNECT THE BATTERY FROM THE BATTERY CHARGER HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Material Twisted / Bent | Device problem code
6902913 | 89430257 | N | ONE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT B DISCHARGED FROM TO OVER THE COURSE OF HOURS NEAR THE REPORTED EVENT DATE NO ANOMALIES WERE DETECTED ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT COULD NOT BE DUPLICATED AT THE BENCH LEVEL THE INSTRUCTIONS FOR USE AND PATIENT MANUAL OUTLINE THE STEPS FOR HANDLING AND PROPERLY CONNECTING POWER SOURCES INCLUDING ENSURING PROPER ALIGNMENT AS WELL AS INSTRUCTIONS NOT TO TWIST OR FORCE CONNECTIONS TOGETHER IN ORDER TO PREVENT DAMAGES ADDITIONALLY A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE AND THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIMES IT ALSO PROVIDES INFORMATION ABOUT PROPER CARE OF THE SYSTEM AND WHAT TO DO IN CASE OF AN EMERGENCY ALWAYS WAIT UNTIL THE READY TURNS ON TO DISCONNECT THE BATTERY FROM THE BATTERY CHARGER HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Results Pending Completion of Investigation | Medical Evaluation result code
6885684 | 89431117 | N | CLINICAL EVALUATION OF B B MD PHD CHIEF PHYSICIAN B HOSPITAL MAIL FROM ACCORDING TO THE AVAILABLE DATA THE PATIENT HAD MEASUREMENTS PERFORMED BY AN ABL ANALYZER DURING THE PERIOD B THROUGH B NINETEEN OF MEASUREMENTS SHOW POTASSIUM VALUES IN THE RANGE MMOLL TWO POTASSIUM MEASUREMENTS ARE CLEARLY OUTSIDE THIS RANGE THE FIRST ABNORMAL MEASUREMENT ON B IS MMOLL AND ACCORDING TO THE AVAILABLE DATA THIS MEASUREMENT PROMPTED TREATMENT OF THE PATIENT WITH GLUCOSE AND INSULIN NINE MINUTES LATER POTASSIUM WAS MEASURED AT MMOLL THE FOLLOWING HOURS AND MINUTES EIGHT POTASSIUM MEASUREMENTS WERE ALL STABLE WITHIN THE RANGE MMOLL ON B POTASSIUM WAS MEASURED AT MMOLL AND THREE MINUTES LATER POTASSIUM MEASURED AT MMOLL SODIUM VALUES WERE STABLE DURING THE MEASUREMENT PERIOD WITHIN THE RANGE MMOLL EXCEPT FOR B WHERE THE VALUE WAS MEASURED AT MMOLL THREE MINUTES LATER THE VALUE WAS MMOLL IT IS ESTIMATED THAT THE LARGE VARIATIONS IN POTASSIUM VALUES WITHIN NINE MINUTES ON B AND THE LARGE VARIATION IN POTASSIUM AND SODIUM VALUES WITHIN THREE MINUTES ON B IS UNLIKELY TO BE CAUSED BY PHYSIOLOGICAL EVENTS WHETHER THE VARIATIONS ARE CAUSED BY SAMPLING ERRORS OR MEASUREMENT ERRORS IS DIFFICULT TO ESTABLISH BASED UPON CURRENT AVAILABLE INFORMATION DESPITE THAT CALIBRATION QC MEASUREMENTS HAVE BEEN REPORTED WITHOUT ERRORS IN THE PERIOD PRIOR THE CURRENT EVENT A MALFUNCTION OF THE SENSORS OR ANALYZER CANNOT BE RULED OUT BASED UPON THE CURRENT AVAILABLE INFORMATION SINCE THE POTENTIAL ERRONEOUS GLUCOSE AND INSULIN TREATMENT WAS REPORTED TO BE THE CONSEQUENCE OF A VERY HIGH K IT IS CONSIDERED LIKELY THAT THIS MUST HAVE BEEN INITIATED SHORTLY AFTER THE MEASUREMENT ON B RESULTING IN A POTASSIUM VALUE OF MMOLL AS THE FOLLOWING EIGHT POTASSIUM MEASUREMENTS DURING THE FOLLOWING PERIOD OF HOURS AND MINUTES WERE ALL STABLE WITHIN THE RANGE MMOLL IT IS NOT CONSIDERED LIKELY THAT THE POTENTIAL ERRONEOUS POTASSIUM MEASUREMENT LEADING TO A POTENTIAL WRONGFUL TREATMENT OF THE PATIENT CONTRIBUTED TO THE DEATH OF THE PATIENT RADIOMETER HAS RECEIVED DATA LOGS FOR INVESTIGATION RESULTS OF THE INVESTIGATION WILL BE REPORTED IN FOLLOWUP REPORTS | Calibration Problem | Device problem code
6885684 | 89431117 | N | CLINICAL EVALUATION OF B B MD PHD CHIEF PHYSICIAN B HOSPITAL MAIL FROM ACCORDING TO THE AVAILABLE DATA THE PATIENT HAD MEASUREMENTS PERFORMED BY AN ABL ANALYZER DURING THE PERIOD B THROUGH B NINETEEN OF MEASUREMENTS SHOW POTASSIUM VALUES IN THE RANGE MMOLL TWO POTASSIUM MEASUREMENTS ARE CLEARLY OUTSIDE THIS RANGE THE FIRST ABNORMAL MEASUREMENT ON B IS MMOLL AND ACCORDING TO THE AVAILABLE DATA THIS MEASUREMENT PROMPTED TREATMENT OF THE PATIENT WITH GLUCOSE AND INSULIN NINE MINUTES LATER POTASSIUM WAS MEASURED AT MMOLL THE FOLLOWING HOURS AND MINUTES EIGHT POTASSIUM MEASUREMENTS WERE ALL STABLE WITHIN THE RANGE MMOLL ON B POTASSIUM WAS MEASURED AT MMOLL AND THREE MINUTES LATER POTASSIUM MEASURED AT MMOLL SODIUM VALUES WERE STABLE DURING THE MEASUREMENT PERIOD WITHIN THE RANGE MMOLL EXCEPT FOR B WHERE THE VALUE WAS MEASURED AT MMOLL THREE MINUTES LATER THE VALUE WAS MMOLL IT IS ESTIMATED THAT THE LARGE VARIATIONS IN POTASSIUM VALUES WITHIN NINE MINUTES ON B AND THE LARGE VARIATION IN POTASSIUM AND SODIUM VALUES WITHIN THREE MINUTES ON B IS UNLIKELY TO BE CAUSED BY PHYSIOLOGICAL EVENTS WHETHER THE VARIATIONS ARE CAUSED BY SAMPLING ERRORS OR MEASUREMENT ERRORS IS DIFFICULT TO ESTABLISH BASED UPON CURRENT AVAILABLE INFORMATION DESPITE THAT CALIBRATION QC MEASUREMENTS HAVE BEEN REPORTED WITHOUT ERRORS IN THE PERIOD PRIOR THE CURRENT EVENT A MALFUNCTION OF THE SENSORS OR ANALYZER CANNOT BE RULED OUT BASED UPON THE CURRENT AVAILABLE INFORMATION SINCE THE POTENTIAL ERRONEOUS GLUCOSE AND INSULIN TREATMENT WAS REPORTED TO BE THE CONSEQUENCE OF A VERY HIGH K IT IS CONSIDERED LIKELY THAT THIS MUST HAVE BEEN INITIATED SHORTLY AFTER THE MEASUREMENT ON B RESULTING IN A POTASSIUM VALUE OF MMOLL AS THE FOLLOWING EIGHT POTASSIUM MEASUREMENTS DURING THE FOLLOWING PERIOD OF HOURS AND MINUTES WERE ALL STABLE WITHIN THE RANGE MMOLL IT IS NOT CONSIDERED LIKELY THAT THE POTENTIAL ERRONEOUS POTASSIUM MEASUREMENT LEADING TO A POTENTIAL WRONGFUL TREATMENT OF THE PATIENT CONTRIBUTED TO THE DEATH OF THE PATIENT RADIOMETER HAS RECEIVED DATA LOGS FOR INVESTIGATION RESULTS OF THE INVESTIGATION WILL BE REPORTED IN FOLLOWUP REPORTS | Results Pending Completion of Investigation | Medical Evaluation result code
6918375 | 89432556 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918375 | 89432556 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6902725 | 89433341 | D | IT WAS REPORTED THAT THE BATTERY DISCHARGED FASTER THAN EXPECTED THE BATTERY WAS REPLACED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Premature Discharge of Battery | Device problem code
6902725 | 89433341 | D | IT WAS REPORTED THAT THE BATTERY DISCHARGED FASTER THAN EXPECTED THE BATTERY WAS REPLACED NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6952859 | 89455754 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF SWELLING AND PUFFINESS IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS JUVÉDERM VOLUMA XC IS TO BE USED AS SUPPLIED MODIFICATION OR USE OF THE PRODUCT OUTSIDE THE DIRECTIONS FOR USE MAY ADVERSELY IMPACT THE STERILITY HOMOGENEITY AND PERFORMANCE OF THE PRODUCT ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING HEALTH CARE PROFESSIONAL INSTRUCTIONS PATIENTS MAY EXPERIENCE TREATMENT SITE RESPONSES WHICH TYPICALLY RESOLVE WITHIN TO WEEKS ICE MAY BE APPLIED FOR A BRIEF PERIOD FOLLOWING TREATMENT TO MINIMIZE SWELLING AND REDUCE PAIN | No Device Problem Found | Medical Evaluation result code
6953954 | 89458330 | N | SO FAR WE DO NOT KNOW IF A PATIENT GOT HURT WITH THIS DEVICE THE INFORMATION IS STILL MISSING BUT AS THE AWARENESS DATE WAS ON SEPTEMBER TH WE DECIDED TO REPORT THE INCIDENT ANYWAY TO MAKE A FOLLOWUP POSSIBLE IF THERE ARE FURTHER INFORMATION AVAILABLE THE PIN PRESSURE HAD A MINOR DEVIATION ON OF MEASURING POINTS CONSIDERING THE MINOR DEVIATION OF THE DEVICE WE ASSUME THAT THE DEVICE DID NOT CONTRIBUTE TO THE DESCRIPTION OF THE POSSIBLE EVENT | Component Missing | Device problem code
6953954 | 89458330 | N | SO FAR WE DO NOT KNOW IF A PATIENT GOT HURT WITH THIS DEVICE THE INFORMATION IS STILL MISSING BUT AS THE AWARENESS DATE WAS ON SEPTEMBER TH WE DECIDED TO REPORT THE INCIDENT ANYWAY TO MAKE A FOLLOWUP POSSIBLE IF THERE ARE FURTHER INFORMATION AVAILABLE THE PIN PRESSURE HAD A MINOR DEVIATION ON OF MEASURING POINTS CONSIDERING THE MINOR DEVIATION OF THE DEVICE WE ASSUME THAT THE DEVICE DID NOT CONTRIBUTE TO THE DESCRIPTION OF THE POSSIBLE EVENT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6920683 | 89463675 | N | PRODUCT EVENT SUMMARY THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO | Break | Device problem code
6920683 | 89463675 | N | PRODUCT EVENT SUMMARY THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH AT THE FIRST RIBCLAVICLE LOCATION WHILE IN VIVO | Fatigue Problem | Medical Evaluation result code
6920677 | 89463797 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD INTERMITTENT LOSS OF COMMUNICATION WHEN USING A GOOD PROGRAMMER | Communication or Transmission Problem | Device problem code
6920677 | 89463797 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD INTERMITTENT LOSS OF COMMUNICATION WHEN USING A GOOD PROGRAMMER | Communications Problem Identified | Medical Evaluation result code
6920726 | 89464743 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD AN INTERMITTENT LOSS OF COMMUNICATION WITH THE PROGRAMMER | Communication or Transmission Problem | Device problem code
6920726 | 89464743 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD AN INTERMITTENT LOSS OF COMMUNICATION WITH THE PROGRAMMER | Communications Problem Identified | Medical Evaluation result code
6905218 | 89465017 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS UNABLE TO BE REVIEWED AS WE WERE NOT ABLE TO OBTAIN THE SERIAL NUMBER FOR THE IABP ASSOCIATED WITH THIS COMPLAINT THE COMPANY REPRESENTATIVE REPORTED THAT THE ESIS ECG LEADWIRES WERE EXCHANGED FOR ANOTHER SET AND THE PROBLEM WAS RESOLVED IT WAS REPORTED THAT THE CUSTOMER LATER CONNECTED THE FAULTY LEADWIRES TO ANOTHER IABP AND HAD THE SAME ISSUE OF NO ECG SIGNAL INDICATING THAT THE ISSUE WAS WITH THE WIRES AND NOT THE IABP | Component or Accessory Incompatibility | Device Problem Code
6905218 | 89465017 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS UNABLE TO BE REVIEWED AS WE WERE NOT ABLE TO OBTAIN THE SERIAL NUMBER FOR THE IABP ASSOCIATED WITH THIS COMPLAINT THE COMPANY REPRESENTATIVE REPORTED THAT THE ESIS ECG LEADWIRES WERE EXCHANGED FOR ANOTHER SET AND THE PROBLEM WAS RESOLVED IT WAS REPORTED THAT THE CUSTOMER LATER CONNECTED THE FAULTY LEADWIRES TO ANOTHER IABP AND HAD THE SAME ISSUE OF NO ECG SIGNAL INDICATING THAT THE ISSUE WAS WITH THE WIRES AND NOT THE IABP | Communications Problem Identified | Medical Evaluation result code
6896913 | 89466859 | N | WHILE THERE IS NO INDICATION THAT A SERIOUS INJURY RESULTED IN THIS EVENT THERE HAS BEEN A PREVIOUS REPORT RECEIVED WHERE THIS MALFUNCTION RESULTED IN A SERIOUS INJURY THEREFORE IT MUST BE PRESUMED THAT RECURRENCE OF THIS MALFUNCTION COULD POSSIBLY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR REQUIRE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE SUCH AS SUCH THIS EVENT IS REPORTABLE PER CFR PART THE INSERT WAS INSPECTED AND FOUND TO BE CLOGGED THE INSERT WAS NOT TESTED FOR TEMPERATURE DUE TO BEING CLOGGED AFTER TESTING THE INSERT WAS FOUND TO HAVE PLASTIC FLASH CLOGGING THE INSERT THE INSERT WAS TESTED USING A G UNIT AT CC OF WATER FLOW THE TEST UNIT WAS G THERMOMETER ID CAL DATE CAL DUE IN CONCLUSION THE COMPLAINT FOR THE INSERT GETTING HOT HAS FAILED FOR HAVING NO WATER FLOW CAUSED BY FLASH DEBRIS MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO OBTAIN THE PATIENT OUTCOME | Flare or Flash | Device problem code
6896913 | 89466859 | N | WHILE THERE IS NO INDICATION THAT A SERIOUS INJURY RESULTED IN THIS EVENT THERE HAS BEEN A PREVIOUS REPORT RECEIVED WHERE THIS MALFUNCTION RESULTED IN A SERIOUS INJURY THEREFORE IT MUST BE PRESUMED THAT RECURRENCE OF THIS MALFUNCTION COULD POSSIBLY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR REQUIRE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE SUCH AS SUCH THIS EVENT IS REPORTABLE PER CFR PART THE INSERT WAS INSPECTED AND FOUND TO BE CLOGGED THE INSERT WAS NOT TESTED FOR TEMPERATURE DUE TO BEING CLOGGED AFTER TESTING THE INSERT WAS FOUND TO HAVE PLASTIC FLASH CLOGGING THE INSERT THE INSERT WAS TESTED USING A G UNIT AT CC OF WATER FLOW THE TEST UNIT WAS G THERMOMETER ID CAL DATE CAL DUE IN CONCLUSION THE COMPLAINT FOR THE INSERT GETTING HOT HAS FAILED FOR HAVING NO WATER FLOW CAUSED BY FLASH DEBRIS MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO OBTAIN THE PATIENT OUTCOME | Complete Blockage | Device problem code
6896913 | 89466859 | N | WHILE THERE IS NO INDICATION THAT A SERIOUS INJURY RESULTED IN THIS EVENT THERE HAS BEEN A PREVIOUS REPORT RECEIVED WHERE THIS MALFUNCTION RESULTED IN A SERIOUS INJURY THEREFORE IT MUST BE PRESUMED THAT RECURRENCE OF THIS MALFUNCTION COULD POSSIBLY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR REQUIRE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE SUCH AS SUCH THIS EVENT IS REPORTABLE PER CFR PART THE INSERT WAS INSPECTED AND FOUND TO BE CLOGGED THE INSERT WAS NOT TESTED FOR TEMPERATURE DUE TO BEING CLOGGED AFTER TESTING THE INSERT WAS FOUND TO HAVE PLASTIC FLASH CLOGGING THE INSERT THE INSERT WAS TESTED USING A G UNIT AT CC OF WATER FLOW THE TEST UNIT WAS G THERMOMETER ID CAL DATE CAL DUE IN CONCLUSION THE COMPLAINT FOR THE INSERT GETTING HOT HAS FAILED FOR HAVING NO WATER FLOW CAUSED BY FLASH DEBRIS MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO OBTAIN THE PATIENT OUTCOME | Overheating of Device | Device Problem Code
6920788 | 89468589 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT STYLUSPOINTER PROBLEM THE STYLUS WAS FOUND OUT OF CALIBRATION | Calibration Problem | Device problem code
6920788 | 89468589 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT STYLUSPOINTER PROBLEM THE STYLUS WAS FOUND OUT OF CALIBRATION | Results Pending Completion of Investigation | Medical Evaluation result code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Material Separation | Device problem code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Component Missing | Device problem code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Detachment of Device or device Component | Device problem code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6918091 | 89471004 | N | IT WAS REPORTED THAT THE LIGHT FELL AND HIT THE GROUND A STRYKER FIELD SERVICE TECHNICIAN SFST WAS DISPATCHED TO THE CUSTOMER SITE WHILE INTERVIEWING THE USERS IT WAS REPORTED THAT THE LIGHT SWITCH FROM THE WALL WAS NOT WORKING AND THE NURSE WENT TO THE LIGHT TO TURN IT ON AT THE SOURCE AND IT DISENGAGED THE NURSE CAUGHT THE LIGHT AND IT WAS MOVED OUT OF THE SURGICAL AREA DURING INVESTIGATION IT WAS FOUND THAT THE SEPARATION WAS A RESULT OF A MISSING CCLIP BEING THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THE LIGHT WAS RETURNED TO FULL USE THERE WAS NO IMPACT TO THE PATIENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Missing or Inadequate Safety Measures | Medical Evaluation result code
6918350 | 89471164 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6918350 | 89471164 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6924100 | 89471302 | N | PRODUCT EVENT SUMMARY AT ANALYSIS THE ANALYZER DEMONSTRATED A LOSS OF COMMUNICATION WITH THE PROGRAMMER AND FAILED ITS TESTS IT IS TO BE SENT ON FOR REPAIR AND RESTOCK | Communication or Transmission Problem | Device problem code
6924100 | 89471302 | N | PRODUCT EVENT SUMMARY AT ANALYSIS THE ANALYZER DEMONSTRATED A LOSS OF COMMUNICATION WITH THE PROGRAMMER AND FAILED ITS TESTS IT IS TO BE SENT ON FOR REPAIR AND RESTOCK | Communications Problem Identified | Medical Evaluation result code
6918190 | 89471967 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918190 | 89471967 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6918251 | 89472024 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918251 | 89472024 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6923738 | 89473274 | N | AT ANALYSIS THE INSTRUMENT WAS CHECKED AND CLEANED THE STYLUS CALIBRATION WAS COMPLETED AND THE SOFTWARE WAS UPDATED IT WAS CONFIRMED THAT THE PROGRAMMER WAS WORKING OK | Calibration Problem | Device problem code
6923738 | 89473274 | N | AT ANALYSIS THE INSTRUMENT WAS CHECKED AND CLEANED THE STYLUS CALIBRATION WAS COMPLETED AND THE SOFTWARE WAS UPDATED IT WAS CONFIRMED THAT THE PROGRAMMER WAS WORKING OK | Results Pending Completion of Investigation | Medical Evaluation result code
6918363 | 89474806 | N | B ACCORDING TO THE B THE PROBLEM WAS SOLVED BY THE BIOMED AT THE HOSPITAL THE BIOMED REBOOTED THE PUMP AND IT WORKED FINE THE CUSTOMER BELIEVES THE PROBLEM WAS CAUSED BECAUSE THEY ACCIDENTALLY HIT THE ONOFF BUTTON THUS THE FAILURE COULD BE CONFIRMED THE DATA IS ALSO BEING HANDLED THROUGH A DESIGNATED MAQUET CARDIOPULMONARY TRENDING AND APPLICABLE INVESTIGATION PROCESS DUE TO THIS NO FURTHER INVESTIGATION INITIATIONS WILL BE COMPLETED AT THIS TIME SINCE THE REPORTED FAILURE DID NOT CONTRIBUTE TO A DEATH OR SERIOUS INJURY NO CORRECTIVE ACTION IS NEEDED IN ADDITION AT THIS TIME IT CANNOT BE CONCLUDED THAT THIS IS A SYSTEMIC ERROR NO CORRECTIVE ACTION IS NEEDED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED AFTER NEW INFORMATION HAS BEEN RECEIVED | Computer Software Problem | Device problem code
6918363 | 89474806 | N | B ACCORDING TO THE B THE PROBLEM WAS SOLVED BY THE BIOMED AT THE HOSPITAL THE BIOMED REBOOTED THE PUMP AND IT WORKED FINE THE CUSTOMER BELIEVES THE PROBLEM WAS CAUSED BECAUSE THEY ACCIDENTALLY HIT THE ONOFF BUTTON THUS THE FAILURE COULD BE CONFIRMED THE DATA IS ALSO BEING HANDLED THROUGH A DESIGNATED MAQUET CARDIOPULMONARY TRENDING AND APPLICABLE INVESTIGATION PROCESS DUE TO THIS NO FURTHER INVESTIGATION INITIATIONS WILL BE COMPLETED AT THIS TIME SINCE THE REPORTED FAILURE DID NOT CONTRIBUTE TO A DEATH OR SERIOUS INJURY NO CORRECTIVE ACTION IS NEEDED IN ADDITION AT THIS TIME IT CANNOT BE CONCLUDED THAT THIS IS A SYSTEMIC ERROR NO CORRECTIVE ACTION IS NEEDED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED AFTER NEW INFORMATION HAS BEEN RECEIVED | Software Problem Identified | Medical Evaluation result code
6920899 | 89474967 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE BATTERY CONTACTS WERE CONTAMINATED ANALYSIS ALSO FOUND THE HEART WIRE CONTACTS WERE STUCK STICKY | Mechanical Jam | Device problem code
6920899 | 89474967 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE BATTERY CONTACTS WERE CONTAMINATED ANALYSIS ALSO FOUND THE HEART WIRE CONTACTS WERE STUCK STICKY | Mechanical Problem Identified | Medical Evaluation result code
6923698 | 89475963 | N | MANUFACTURERS ANALYSIS NOTED THAT THE DEVICE FAILED INCOMING TESTING IT DID NOT COMMUNICATE WITH THE PROGRAMMER | Communication or Transmission Problem | Device problem code
6923698 | 89475963 | N | MANUFACTURERS ANALYSIS NOTED THAT THE DEVICE FAILED INCOMING TESTING IT DID NOT COMMUNICATE WITH THE PROGRAMMER | Communications Problem Identified | Medical Evaluation result code
6917643 | 89476426 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B FOUR MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THESE TWO IMPLANTS FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6917643 | 89476426 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B FOUR MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THESE TWO IMPLANTS FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6908938 | 89476440 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER CCC TO REPORT THE DISCORDANT TOTAL HUMAN CHORIONIC GONADOTROPIN THCG RESULT QUALITY CONTROL QC AND CALIBRATION WERE ACCEPTABLE SIEMENS IS INVESTIGATING THE ISSUE | Calibration Problem | Device problem code
6908938 | 89476440 | N | THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER CCC TO REPORT THE DISCORDANT TOTAL HUMAN CHORIONIC GONADOTROPIN THCG RESULT QUALITY CONTROL QC AND CALIBRATION WERE ACCEPTABLE SIEMENS IS INVESTIGATING THE ISSUE | Results Pending Completion of Investigation | Medical Evaluation result code
6923764 | 89477331 | N | THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Break | Device problem code
6923764 | 89477331 | N | THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Fracture Problem | Medical Evaluation result code
6918720 | 89478843 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6918720 | 89478843 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6923774 | 89480657 | N | AT ANALYSIS THE DEVICE WOULD NOT COMMUNICATE WITH A KNOWN GOOD PROGRAMMER AND WAS TO BE SENT ON FOR REPAIR AND RESTOCK | Communication or Transmission Problem | Device problem code
6923774 | 89480657 | N | AT ANALYSIS THE DEVICE WOULD NOT COMMUNICATE WITH A KNOWN GOOD PROGRAMMER AND WAS TO BE SENT ON FOR REPAIR AND RESTOCK | Communications Problem Identified | Medical Evaluation result code
6919974 | 89480659 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT OBSTRUCTION FOUND IN THE VENTRICULAR PORT SMALL PIECE OF PLASTIC INSIDE VENTICULAR PORT | Obstruction of Flow | Device problem code
6919974 | 89480659 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT OBSTRUCTION FOUND IN THE VENTRICULAR PORT SMALL PIECE OF PLASTIC INSIDE VENTICULAR PORT | Operational Problem Identified | Medical Evaluation result code
6925840 | 89480975 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO POWER ON COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Insufficient Information | Device Problem Code
6925840 | 89480975 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE WAS UNABLE TO POWER ON COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6921065 | 89483060 | N | OUR EVALUATION FOUND AN EXTERNAL CUT AT ANGLE COVER APPROXIMATELY MM FROM TIP WITH TOOL MARKS AND CUT MARKS AT THREAD WRAP EPOXY CAUSING A LEAK THERE IS A MISSING PIECE OF ANGLE COVER IT MEASURES APPROXIMATELY MMXMM THERE ARE SCRATCHES AT THE DISTAL HEAD DEBRIS ON THE VERTEBRAE AND SCRAPE MARKS INSIDE WORKING CHANNEL IT APPEARS THAT A SHARP INSTRUMENT OF SOME TYPE HAS CUT INTO THE DISTAL HEAD MATERIAL AND CUT OUT A ROUND SECTION OF THE ANGLE COVER MATERIAL | Scratched Material | Device problem code
6921065 | 89483060 | N | OUR EVALUATION FOUND AN EXTERNAL CUT AT ANGLE COVER APPROXIMATELY MM FROM TIP WITH TOOL MARKS AND CUT MARKS AT THREAD WRAP EPOXY CAUSING A LEAK THERE IS A MISSING PIECE OF ANGLE COVER IT MEASURES APPROXIMATELY MMXMM THERE ARE SCRATCHES AT THE DISTAL HEAD DEBRIS ON THE VERTEBRAE AND SCRAPE MARKS INSIDE WORKING CHANNEL IT APPEARS THAT A SHARP INSTRUMENT OF SOME TYPE HAS CUT INTO THE DISTAL HEAD MATERIAL AND CUT OUT A ROUND SECTION OF THE ANGLE COVER MATERIAL | Leak / Splash | Device problem code
6921065 | 89483060 | N | OUR EVALUATION FOUND AN EXTERNAL CUT AT ANGLE COVER APPROXIMATELY MM FROM TIP WITH TOOL MARKS AND CUT MARKS AT THREAD WRAP EPOXY CAUSING A LEAK THERE IS A MISSING PIECE OF ANGLE COVER IT MEASURES APPROXIMATELY MMXMM THERE ARE SCRATCHES AT THE DISTAL HEAD DEBRIS ON THE VERTEBRAE AND SCRAPE MARKS INSIDE WORKING CHANNEL IT APPEARS THAT A SHARP INSTRUMENT OF SOME TYPE HAS CUT INTO THE DISTAL HEAD MATERIAL AND CUT OUT A ROUND SECTION OF THE ANGLE COVER MATERIAL | Component Missing | Device problem code
6921065 | 89483060 | N | OUR EVALUATION FOUND AN EXTERNAL CUT AT ANGLE COVER APPROXIMATELY MM FROM TIP WITH TOOL MARKS AND CUT MARKS AT THREAD WRAP EPOXY CAUSING A LEAK THERE IS A MISSING PIECE OF ANGLE COVER IT MEASURES APPROXIMATELY MMXMM THERE ARE SCRATCHES AT THE DISTAL HEAD DEBRIS ON THE VERTEBRAE AND SCRAPE MARKS INSIDE WORKING CHANNEL IT APPEARS THAT A SHARP INSTRUMENT OF SOME TYPE HAS CUT INTO THE DISTAL HEAD MATERIAL AND CUT OUT A ROUND SECTION OF THE ANGLE COVER MATERIAL | Missing or Inadequate Safety Measures | Medical Evaluation result code
6954565 | 89483240 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6921104 | 89485632 | N | THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE ANALYST NOTED THAT THE HELIX WOULD NOT BE EXTENDED DUE TO DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP OVERRETRACTED | Mechanical Jam | Device problem code
6921104 | 89485632 | N | THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE ANALYST NOTED THAT THE HELIX WOULD NOT BE EXTENDED DUE TO DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP OVERRETRACTED | Mechanical Problem Identified | Medical Evaluation result code
6918281 | 89485759 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918281 | 89485759 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6920023 | 89487759 | N | MANUFACTURERS ANALYSIS INDICATED THAT THE PROGRAMMERS DISPLAY HAD A PORTION THAT WAS UNRESPONSIVE THE PROGRAMMER WAS TO BE SCRAPPED | Component or Accessory Incompatibility | Device problem code
6920023 | 89487759 | N | MANUFACTURERS ANALYSIS INDICATED THAT THE PROGRAMMERS DISPLAY HAD A PORTION THAT WAS UNRESPONSIVE THE PROGRAMMER WAS TO BE SCRAPPED | Operational Problem Identified | Medical Evaluation result code
6919600 | 89488699 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Break | Device problem code
6919600 | 89488699 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Fracture Problem | Medical Evaluation result code
6918290 | 89493325 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918290 | 89493325 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6918396 | 89493340 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918396 | 89493340 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6920758 | 89495226 | N | MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER HAD NOISE ON SENSING THE ANALYZER PASSING INCOMING TESTING | Noise, Audible | Device problem code
6920758 | 89495226 | N | MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER HAD NOISE ON SENSING THE ANALYZER PASSING INCOMING TESTING | No Device Problem Found | Medical Evaluation result code
6920316 | 89496986 | N | ANALYSIS CONFIRMED THE STYLUS DID NOT WORK CORRECTLY IT WAS FARTHER OFF ON THE TOP LEFT THE CONNECTION FROM THE XY BOARD TO THE OVERLAY WAS RESEATED TO RESOLVE THE ISSUE ANALYSIS ALSO FOUND THE PROGRAMMER FAILED LEFT ANTENNA CONNECTION THE ANTENNA CONNECTION WAS RESEATED TO RESOLVE THE ISSUE | Connection Problem | Device problem code
6920316 | 89496986 | N | ANALYSIS CONFIRMED THE STYLUS DID NOT WORK CORRECTLY IT WAS FARTHER OFF ON THE TOP LEFT THE CONNECTION FROM THE XY BOARD TO THE OVERLAY WAS RESEATED TO RESOLVE THE ISSUE ANALYSIS ALSO FOUND THE PROGRAMMER FAILED LEFT ANTENNA CONNECTION THE ANTENNA CONNECTION WAS RESEATED TO RESOLVE THE ISSUE | Interoperability Problem Identified | Medical Evaluation result code
6920682 | 89504550 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE PATIENT CONNECTOR HAD DISTURBED PINS NOT ABLE TO PLUG INTO THE PATIENT CONNECTION | Connection Problem | Device problem code
6920682 | 89504550 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE PATIENT CONNECTOR HAD DISTURBED PINS NOT ABLE TO PLUG INTO THE PATIENT CONNECTION | Interoperability Problem Identified | Medical Evaluation result code
6951888 | 89505559 | D | IT WAS REPORTED THAT THE INSULIN PUMP HAD AN ERROR ALARM IT WAS NOTED THAT THE PUMP NOT EXPOSED TO ANY MAGNETIC FIELDS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6951888 | 89505559 | D | IT WAS REPORTED THAT THE INSULIN PUMP HAD AN ERROR ALARM IT WAS NOTED THAT THE PUMP NOT EXPOSED TO ANY MAGNETIC FIELDS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MMOLL THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Device Alarm System | Device problem code
6929027 | 89506379 | N | B PUMP ERROR FOUND IN PUMP HISTORY A DETAILED TRACE ANALYSIS CONFIRMED ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A MISSING RESERVOIR TUBE RETAINER AND RESERVOIR TUBE ORING THERE WAS ALSO A SCRATCHED CASE AND MINOR SCRATCHED LCD WINDOW | Component or Accessory Incompatibility | Device problem code
6929027 | 89506379 | N | B PUMP ERROR FOUND IN PUMP HISTORY A DETAILED TRACE ANALYSIS CONFIRMED ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A MISSING RESERVOIR TUBE RETAINER AND RESERVOIR TUBE ORING THERE WAS ALSO A SCRATCHED CASE AND MINOR SCRATCHED LCD WINDOW | Computer Software Problem | Device problem code
6929027 | 89506379 | N | B PUMP ERROR FOUND IN PUMP HISTORY A DETAILED TRACE ANALYSIS CONFIRMED ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A MISSING RESERVOIR TUBE RETAINER AND RESERVOIR TUBE ORING THERE WAS ALSO A SCRATCHED CASE AND MINOR SCRATCHED LCD WINDOW | Component Missing | Device problem code
6929027 | 89506379 | N | B PUMP ERROR FOUND IN PUMP HISTORY A DETAILED TRACE ANALYSIS CONFIRMED ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A MISSING RESERVOIR TUBE RETAINER AND RESERVOIR TUBE ORING THERE WAS ALSO A SCRATCHED CASE AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6929027 | 89506379 | N | B PUMP ERROR FOUND IN PUMP HISTORY A DETAILED TRACE ANALYSIS CONFIRMED ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAT V FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR THE PUMP WAS RECEIVED WITH A MISSING RESERVOIR TUBE RETAINER AND RESERVOIR TUBE ORING THERE WAS ALSO A SCRATCHED CASE AND MINOR SCRATCHED LCD WINDOW | Software Problem Identified | Medical Evaluation result code
6955320 | 89512597 | N | THE REPORTER HAS DECLINED TO PROVIDE ALLERGAN FURTHER INFORMATION REGARDING EVENT PRODUCT OR PATIENT DETAILS THE EVENTS OF ALLERGIC REACTION SWELLING AND FLUID IN CHEST ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS HEALTH CARE PROFESSIONAL INSTRUCTIONS THE HEALTH CARE PROFESSIONAL SHOULD INSTRUCT THE PATIENT TO PROMPTLY REPORT TO HERHIM ANY EVIDENCE OF PROBLEMS POSSIBLY ASSOCIATED WITH THE USE OF JUVÉDERM ULTRA INJECTABLE GEL | Migration or Expulsion of Device | Device problem code
6955320 | 89512597 | N | THE REPORTER HAS DECLINED TO PROVIDE ALLERGAN FURTHER INFORMATION REGARDING EVENT PRODUCT OR PATIENT DETAILS THE EVENTS OF ALLERGIC REACTION SWELLING AND FLUID IN CHEST ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS HEALTH CARE PROFESSIONAL INSTRUCTIONS THE HEALTH CARE PROFESSIONAL SHOULD INSTRUCT THE PATIENT TO PROMPTLY REPORT TO HERHIM ANY EVIDENCE OF PROBLEMS POSSIBLY ASSOCIATED WITH THE USE OF JUVÉDERM ULTRA INJECTABLE GEL | Leak / Splash | Device problem code
6955320 | 89512597 | N | THE REPORTER HAS DECLINED TO PROVIDE ALLERGAN FURTHER INFORMATION REGARDING EVENT PRODUCT OR PATIENT DETAILS THE EVENTS OF ALLERGIC REACTION SWELLING AND FLUID IN CHEST ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS HEALTH CARE PROFESSIONAL INSTRUCTIONS THE HEALTH CARE PROFESSIONAL SHOULD INSTRUCT THE PATIENT TO PROMPTLY REPORT TO HERHIM ANY EVIDENCE OF PROBLEMS POSSIBLY ASSOCIATED WITH THE USE OF JUVÉDERM ULTRA INJECTABLE GEL | Material Discolored | Device problem code
6955320 | 89512597 | N | THE REPORTER HAS DECLINED TO PROVIDE ALLERGAN FURTHER INFORMATION REGARDING EVENT PRODUCT OR PATIENT DETAILS THE EVENTS OF ALLERGIC REACTION SWELLING AND FLUID IN CHEST ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS HEALTH CARE PROFESSIONAL INSTRUCTIONS THE HEALTH CARE PROFESSIONAL SHOULD INSTRUCT THE PATIENT TO PROMPTLY REPORT TO HERHIM ANY EVIDENCE OF PROBLEMS POSSIBLY ASSOCIATED WITH THE USE OF JUVÉDERM ULTRA INJECTABLE GEL | Device Migration | Medical Evaluation result code
6923096 | 89514369 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER WOULD NOT COMMUNICATE WITH THE PROGRAMMER | Communication or Transmission Problem | Device problem code
6923096 | 89514369 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER WOULD NOT COMMUNICATE WITH THE PROGRAMMER | Communications Problem Identified | Medical Evaluation result code
6922821 | 89515262 | N | PRODUCT EVENT SUMMARY THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE INNER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO | Break | Device problem code
6922821 | 89515262 | N | PRODUCT EVENT SUMMARY THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE INNER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO | Fracture Problem | Medical Evaluation result code
6955569 | 89517744 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING PATIENT CODE B NO CODE AVAILABLE SECONDARY SURGERY LENS EXPLANT LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT B | Component Missing | Device problem code
6955569 | 89517744 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING PATIENT CODE B NO CODE AVAILABLE SECONDARY SURGERY LENS EXPLANT LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6910854 | 89524223 | N | B MANUFACTURING SITE EVALUATION SAMPLES RECEIVED UNOPENED RACEPACKS AND OPENED ANALYSIS AND RESULTS THERE ARE NO PREVIOUS COMPLAINTS OF THE SAME CODEBATCH MANUFACTURED AND DISTRIBUTED IN THE MARKET B UNITS OF THIS CODE BATCH THERE ARE NO UNITS IN STOCK RECEIVED CLOSED SAMPLES AND OPEN AND UNUSED SAMPLES WITH THE NEEDLE DETACHED FROM THE THREAD AND THE THREAD IS STILL WOUND ON THE PACK TESTED THE NEEDLE ATTACHMENT OF THE CLOSED SAMPLES RECEIVED AND THE RESULTS DO NOT FULFILL THE OEM REQUIREMENTS REVIEWED THE BATCH MANUFACTURING RECORD THIS PRODUCT HAD A NORMAL PROCESS AND THE RESULTS DURING THE PROCESS FULFILLED THE OEM REQUIREMENTS FINAL CONCLUSION COMPLAINT IS JUSTIFIED TAKING INTO ACCOUNT THAT THE RESULTS OF CLOSED SAMPLES RECEIVED DO NOT FULFILL THE OEM SPECIFICATIONS CORRECTIVEPREVENTIVE ACTIONS THIS COMPLAINT IS RECORDED FOR TRENDING ANALYSIS TO ASSESS IF ACTIONS ARE NEEDED IN THE FUTURE | Detachment of Device or device Component | Device problem code
6910854 | 89524223 | N | B MANUFACTURING SITE EVALUATION SAMPLES RECEIVED UNOPENED RACEPACKS AND OPENED ANALYSIS AND RESULTS THERE ARE NO PREVIOUS COMPLAINTS OF THE SAME CODEBATCH MANUFACTURED AND DISTRIBUTED IN THE MARKET B UNITS OF THIS CODE BATCH THERE ARE NO UNITS IN STOCK RECEIVED CLOSED SAMPLES AND OPEN AND UNUSED SAMPLES WITH THE NEEDLE DETACHED FROM THE THREAD AND THE THREAD IS STILL WOUND ON THE PACK TESTED THE NEEDLE ATTACHMENT OF THE CLOSED SAMPLES RECEIVED AND THE RESULTS DO NOT FULFILL THE OEM REQUIREMENTS REVIEWED THE BATCH MANUFACTURING RECORD THIS PRODUCT HAD A NORMAL PROCESS AND THE RESULTS DURING THE PROCESS FULFILLED THE OEM REQUIREMENTS FINAL CONCLUSION COMPLAINT IS JUSTIFIED TAKING INTO ACCOUNT THAT THE RESULTS OF CLOSED SAMPLES RECEIVED DO NOT FULFILL THE OEM SPECIFICATIONS CORRECTIVEPREVENTIVE ACTIONS THIS COMPLAINT IS RECORDED FOR TRENDING ANALYSIS TO ASSESS IF ACTIONS ARE NEEDED IN THE FUTURE | Interoperability Problem Identified | Medical Evaluation result code
6922579 | 89525821 | N | THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Break | Device problem code
6922579 | 89525821 | N | THE PARTIAL LEAD WAS RETURNED IN SEGMENTS ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Fracture Problem | Medical Evaluation result code
6955572 | 89526074 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING B LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT B | Component Missing | Device problem code
6955572 | 89526074 | N | DEVICE EVALUATION LENS WAS RETURNED IN LIQUID IN THE LENS VIAL VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING B LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6907102 | 89528990 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6907102 | 89528990 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6929251 | 89531114 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6929251 | 89531114 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6918285 | 89538049 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Jam | Device problem code
6918285 | 89538049 | N | IT WAS REPORTED THAT A SURGICAL LIGHT FELL AND THAT A NURSE CAUGHT THE LIGHT WHEN THE STRYKER FIELD SERVICE TECHNICIAN SFST ARRIVED ONSITE TO REPLACE THE LIGHT THE CUSTOMER REQUESTED THAT THE SFST INSPECT THE OTHER LIGHTS THAT WERE CONVERTED TO LOW CEILING AT THE SAME TIME UPON INSPECTION THE SFST FOUND THAT THE CCLIP WAS NOT FOUND TO BE INSTALLED IN THE OTHER LIGHTS THE SFST INSTALLED THE CCLIP AND THE SFST CHECKED FOR FUNCTIONALITY THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE EVENT REPORTED TO HAVE OCCURRED | Mechanical Problem Identified | Medical Evaluation result code
6910815 | 89540779 | N | THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AN ATTEMPT WAS MADE TO REPLICATE THE ISSUE BY MANIPULATING THE BATTERYS CONNECTION TO A CONTROLLER DURING THE ANALYSIS OF THE UNIT RESULTS REVEALED THAT THE ELECTRICAL CONNECTION BETWEEN THE BATTERY AND CONTROLLER WAS STABLE THE LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER B DID NOT HAVE THE SMR SOFTWARE THE CONTROLLER LOG ALARM FILE COVERING THE REPORTED EVENT DATE WAS NOT AVAILABLE FOR ANALYSIS ANALYSIS OF THE DATA LOG FILES REVEALED SEVERAL PREMATURE POWER SWITCHING EVENTS THAT WERE DUE TO MOMENTARY DISCONNECTIONS ANDOR COMMUNICATION ERRORS BETWEEN THE CONTROLLER AND BATTERY HOWEVER THE REPORTED BEEPS ARE MOST LIKELY ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERY THE MOST LIKELY ROOT CAUSE OF THE REPORTED BEEPS CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN A CONTROLLER AND THE BATTERY HEARTWARE IS INVESTIGATING COMMUNICATION ERRORS ON NONSMR CONTROLLERS AND MOMENTARY DISCONNECTIONS PER INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Connection Problem | Device problem code
6910815 | 89540779 | N | THE BATTERY B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING AN ATTEMPT WAS MADE TO REPLICATE THE ISSUE BY MANIPULATING THE BATTERYS CONNECTION TO A CONTROLLER DURING THE ANALYSIS OF THE UNIT RESULTS REVEALED THAT THE ELECTRICAL CONNECTION BETWEEN THE BATTERY AND CONTROLLER WAS STABLE THE LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER B DID NOT HAVE THE SMR SOFTWARE THE CONTROLLER LOG ALARM FILE COVERING THE REPORTED EVENT DATE WAS NOT AVAILABLE FOR ANALYSIS ANALYSIS OF THE DATA LOG FILES REVEALED SEVERAL PREMATURE POWER SWITCHING EVENTS THAT WERE DUE TO MOMENTARY DISCONNECTIONS ANDOR COMMUNICATION ERRORS BETWEEN THE CONTROLLER AND BATTERY HOWEVER THE REPORTED BEEPS ARE MOST LIKELY ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERY THE MOST LIKELY ROOT CAUSE OF THE REPORTED BEEPS CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN A CONTROLLER AND THE BATTERY HEARTWARE IS INVESTIGATING COMMUNICATION ERRORS ON NONSMR CONTROLLERS AND MOMENTARY DISCONNECTIONS PER INSTRUCTIONS FOR USE IFU THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Interoperability Problem Identified | Medical Evaluation result code
6956014 | 89560637 | N | DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAVE BEEN COMPLETED THE BELT PASSED INCOMING EVALUATION THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES AND INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING THE ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WERE VERIFIED THE MONITOR FAILED INCOMING TESTING UPON EVALUATION THE DRIVEN GROUND CIRCUITRY WAS OPEN THE CAUSE OF THE INCOMING TEST FAILURE IS THE OPEN DRIVEN GROUND THE CAUSE OF THE OPEN DRIVEN GROUND IS AN OPEN R RESISTOR THE OPEN RESISTOR IS CONSISTENT WITH THE OBSERVED NONLIFEVEST DEFIBRILLATIONS THERE IS NO INDICATION THAT THE OPEN RESISTOR CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH | No Display / Image | Device problem code
6956014 | 89560637 | N | DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAVE BEEN COMPLETED THE BELT PASSED INCOMING EVALUATION THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES AND INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING THE ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WERE VERIFIED THE MONITOR FAILED INCOMING TESTING UPON EVALUATION THE DRIVEN GROUND CIRCUITRY WAS OPEN THE CAUSE OF THE INCOMING TEST FAILURE IS THE OPEN DRIVEN GROUND THE CAUSE OF THE OPEN DRIVEN GROUND IS AN OPEN R RESISTOR THE OPEN RESISTOR IS CONSISTENT WITH THE OBSERVED NONLIFEVEST DEFIBRILLATIONS THERE IS NO INDICATION THAT THE OPEN RESISTOR CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH | Non-Functional Defect | Medical Evaluation result code
6955904 | 89561228 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF UNDER MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS GIVEN INTRAVENOUS GLUCOSE TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS INCOHERENCE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE CUSTOMER WAS ON A DIET DURING THE INCIDENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6955904 | 89561228 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF UNDER MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS GIVEN INTRAVENOUS GLUCOSE TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS INCOHERENCE THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE CUSTOMER WAS ON A DIET DURING THE INCIDENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
6955704 | 89561426 | N | PRODUCT EVENT SUMMARY THE PROXIMAL SEGMENT OF THE LEAD WAS RETURNED ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO THE INITIAL REPORTED EVENT WAS RECEIVED ON JUL OF NOTE THE REPORTABLE SERIOUS INJURY OF INFECTION IS NORMALLY SUBMITTED VIA AN ALTERNATIVE SUMMARY REPORT SUBMISSION THAT WOULD HAVE BEEN DUE ON B INFORMATION WAS SUBSEQUENTLY RECEIVED ON OCT AND THE LEAD TESTED OUT OF SPECIFICATION AS THERE IS NEW INFORMATION THAT REASONABLY SUGGESTS THE DEVICE HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A MALFUNCTION THIS EVENT NO LONGER QUALIFIES FOR SUMMARY REPORTING AND IS THEREFORE BEING SUBMITTED ON A BIMONTHLY TIMEFRAME | Break | Device problem code
6955704 | 89561426 | N | PRODUCT EVENT SUMMARY THE PROXIMAL SEGMENT OF THE LEAD WAS RETURNED ANALYZED AND THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO THE INITIAL REPORTED EVENT WAS RECEIVED ON JUL OF NOTE THE REPORTABLE SERIOUS INJURY OF INFECTION IS NORMALLY SUBMITTED VIA AN ALTERNATIVE SUMMARY REPORT SUBMISSION THAT WOULD HAVE BEEN DUE ON B INFORMATION WAS SUBSEQUENTLY RECEIVED ON OCT AND THE LEAD TESTED OUT OF SPECIFICATION AS THERE IS NEW INFORMATION THAT REASONABLY SUGGESTS THE DEVICE HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A MALFUNCTION THIS EVENT NO LONGER QUALIFIES FOR SUMMARY REPORTING AND IS THEREFORE BEING SUBMITTED ON A BIMONTHLY TIMEFRAME | Fracture Problem | Medical Evaluation result code
6955916 | 89565898 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6955916 | 89565898 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6957311 | 89566820 | N | NO INFORMATION ABOUT RETURN OF DEVICE | Insufficient Information | Device problem code
6956979 | 89568923 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6956979 | 89568923 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6956979 | 89568923 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6956979 | 89568923 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Device Migration | Medical Evaluation result code
6924683 | 89573435 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT THE STYLUS AND THE CURSOR WERE NOT ALIGNED DUE TO THE CONNECTION BETWEEN THE OVERLAY AND THE XY PRINTED CIRCUIT BOARD ASSEMBLY | Failure to Align | Device problem code
6924683 | 89573435 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT THE STYLUS AND THE CURSOR WERE NOT ALIGNED DUE TO THE CONNECTION BETWEEN THE OVERLAY AND THE XY PRINTED CIRCUIT BOARD ASSEMBLY | Connection Problem | Device problem code
6924683 | 89573435 | N | PRODUCT EVENT SUMMARY ANALYSIS CONFIRMED THE REPORTED EVENT THE STYLUS AND THE CURSOR WERE NOT ALIGNED DUE TO THE CONNECTION BETWEEN THE OVERLAY AND THE XY PRINTED CIRCUIT BOARD ASSEMBLY | Interoperability Problem Identified | Medical Evaluation result code
6957673 | 89582660 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT L EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US STUDY NAME B PATIENT ID BCOMBINATION PRODUCT IS APPLICABLE DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Apparent Adverse Event | Device problem code
6957673 | 89582660 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT L EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US STUDY NAME B PATIENT ID BCOMBINATION PRODUCT IS APPLICABLE DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Device Problem Found | Medical Evaluation result code
6957601 | 89582670 | N | IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH A FOLLOWUP MEDWATCH WILL BE FILED AS APPROPRIATE THE DEVICE WAS RECEIVED AND EVALUATED THE NEEDLES OF THE DEVICE WERE INSPECTED UNDER MAGNIFICATION AND THERE WERE NO MISSING PARTS OF THE NEEDLE THE SLEEVE WAS REMOVED FROM THE BARREL OF THE GUN AND BOTH IMPLANTS WERE MISSING THERE WAS A MEETING WITH THE NEW PRODUCT QUALITY ENGINEER AND THE NEW PRODUCT DEVELOPMENT ENGINEER AND THE GUNS WERE EXAMINED THERE WERE NO ANOMALIES NOTED DURING THE INSPECTION OF THE TWO GUNS THE SALES REP WAS CONTACTED AND QUESTIONS WERE ASKED ABOUT WHAT HAPPENED TO THE IMPLANTS THE FIRST IMPLANT WAS PLACED PROPERLY WHEN THE GUN WAS PULLED BACK TO GET READY TO PLACE THE SECOND IMPLANT IT FELL OFF THE NEEDLE ON THE GUN THE SALES REP THINKS THAT THE SURGEON PULLED THE GUN OUT TOO QUICKLY AND THIS CAUSED THE SECOND IMPLANT TO COME OUT OF THE GUN THE SALES REP EXPLAINED THAT THE SECOND IMPLANT WAS REMOVED FROM THE PATIENT BUT THE FIRST IMPLANT REMAINED IN THE PATIENT FOR BOTH DEVICES THERE WERE NO PATIENT CONSEQUENCES OR DELAYS THIS COMPLAINT CANNOT BE CONFIRMED A DHR REVIEW FOR THE LOT L HAS BEEN CONDUCTED TO DETERMINE IF THERE WERE ANY INTERNAL PROCESSING ISSUES WHICH WOULD HAVE CONTRIBUTED TO THE NATURE OF THE PRODUCT COMPLAINT OUR RESULTS INDICATE THAT THIS BATCH OF PRODUCT WAS PROCESSED WITHOUT INCIDENT AND THEREFORE THERE IS NO EVIDENCE OF MANUFACTURING ANOMALIES ON THE PAPERWORK REVIEWED FURTHER A REVIEW INTO THE DEPUY SYNTHES MITEK COMPLAINTS SYSTEM REVEALED NO OTHER COMPLAINTS OF ANY TYPE FOR THIS LOT OF B DEVICES THAT WERE RELEASED TO DISTRIBUTION AT THIS POINT IN TIME NO CORRECTIVE ACTION IS REQUIRED AND NO FURTHER ACTION IS WARRANTED HOWEVER DEPUY SYNTHES MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD B SEE ASSOCIATED MEDWATCH | Material Separation | Device problem code
6957601 | 89582670 | N | IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH A FOLLOWUP MEDWATCH WILL BE FILED AS APPROPRIATE THE DEVICE WAS RECEIVED AND EVALUATED THE NEEDLES OF THE DEVICE WERE INSPECTED UNDER MAGNIFICATION AND THERE WERE NO MISSING PARTS OF THE NEEDLE THE SLEEVE WAS REMOVED FROM THE BARREL OF THE GUN AND BOTH IMPLANTS WERE MISSING THERE WAS A MEETING WITH THE NEW PRODUCT QUALITY ENGINEER AND THE NEW PRODUCT DEVELOPMENT ENGINEER AND THE GUNS WERE EXAMINED THERE WERE NO ANOMALIES NOTED DURING THE INSPECTION OF THE TWO GUNS THE SALES REP WAS CONTACTED AND QUESTIONS WERE ASKED ABOUT WHAT HAPPENED TO THE IMPLANTS THE FIRST IMPLANT WAS PLACED PROPERLY WHEN THE GUN WAS PULLED BACK TO GET READY TO PLACE THE SECOND IMPLANT IT FELL OFF THE NEEDLE ON THE GUN THE SALES REP THINKS THAT THE SURGEON PULLED THE GUN OUT TOO QUICKLY AND THIS CAUSED THE SECOND IMPLANT TO COME OUT OF THE GUN THE SALES REP EXPLAINED THAT THE SECOND IMPLANT WAS REMOVED FROM THE PATIENT BUT THE FIRST IMPLANT REMAINED IN THE PATIENT FOR BOTH DEVICES THERE WERE NO PATIENT CONSEQUENCES OR DELAYS THIS COMPLAINT CANNOT BE CONFIRMED A DHR REVIEW FOR THE LOT L HAS BEEN CONDUCTED TO DETERMINE IF THERE WERE ANY INTERNAL PROCESSING ISSUES WHICH WOULD HAVE CONTRIBUTED TO THE NATURE OF THE PRODUCT COMPLAINT OUR RESULTS INDICATE THAT THIS BATCH OF PRODUCT WAS PROCESSED WITHOUT INCIDENT AND THEREFORE THERE IS NO EVIDENCE OF MANUFACTURING ANOMALIES ON THE PAPERWORK REVIEWED FURTHER A REVIEW INTO THE DEPUY SYNTHES MITEK COMPLAINTS SYSTEM REVEALED NO OTHER COMPLAINTS OF ANY TYPE FOR THIS LOT OF B DEVICES THAT WERE RELEASED TO DISTRIBUTION AT THIS POINT IN TIME NO CORRECTIVE ACTION IS REQUIRED AND NO FURTHER ACTION IS WARRANTED HOWEVER DEPUY SYNTHES MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD B SEE ASSOCIATED MEDWATCH | Component Missing | Device problem code
6957601 | 89582670 | N | IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH A FOLLOWUP MEDWATCH WILL BE FILED AS APPROPRIATE THE DEVICE WAS RECEIVED AND EVALUATED THE NEEDLES OF THE DEVICE WERE INSPECTED UNDER MAGNIFICATION AND THERE WERE NO MISSING PARTS OF THE NEEDLE THE SLEEVE WAS REMOVED FROM THE BARREL OF THE GUN AND BOTH IMPLANTS WERE MISSING THERE WAS A MEETING WITH THE NEW PRODUCT QUALITY ENGINEER AND THE NEW PRODUCT DEVELOPMENT ENGINEER AND THE GUNS WERE EXAMINED THERE WERE NO ANOMALIES NOTED DURING THE INSPECTION OF THE TWO GUNS THE SALES REP WAS CONTACTED AND QUESTIONS WERE ASKED ABOUT WHAT HAPPENED TO THE IMPLANTS THE FIRST IMPLANT WAS PLACED PROPERLY WHEN THE GUN WAS PULLED BACK TO GET READY TO PLACE THE SECOND IMPLANT IT FELL OFF THE NEEDLE ON THE GUN THE SALES REP THINKS THAT THE SURGEON PULLED THE GUN OUT TOO QUICKLY AND THIS CAUSED THE SECOND IMPLANT TO COME OUT OF THE GUN THE SALES REP EXPLAINED THAT THE SECOND IMPLANT WAS REMOVED FROM THE PATIENT BUT THE FIRST IMPLANT REMAINED IN THE PATIENT FOR BOTH DEVICES THERE WERE NO PATIENT CONSEQUENCES OR DELAYS THIS COMPLAINT CANNOT BE CONFIRMED A DHR REVIEW FOR THE LOT L HAS BEEN CONDUCTED TO DETERMINE IF THERE WERE ANY INTERNAL PROCESSING ISSUES WHICH WOULD HAVE CONTRIBUTED TO THE NATURE OF THE PRODUCT COMPLAINT OUR RESULTS INDICATE THAT THIS BATCH OF PRODUCT WAS PROCESSED WITHOUT INCIDENT AND THEREFORE THERE IS NO EVIDENCE OF MANUFACTURING ANOMALIES ON THE PAPERWORK REVIEWED FURTHER A REVIEW INTO THE DEPUY SYNTHES MITEK COMPLAINTS SYSTEM REVEALED NO OTHER COMPLAINTS OF ANY TYPE FOR THIS LOT OF B DEVICES THAT WERE RELEASED TO DISTRIBUTION AT THIS POINT IN TIME NO CORRECTIVE ACTION IS REQUIRED AND NO FURTHER ACTION IS WARRANTED HOWEVER DEPUY SYNTHES MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD B SEE ASSOCIATED MEDWATCH | Missing or Inadequate Safety Measures | Medical Evaluation result code
6923290 | 89584808 | N | PRODUCT EVENT SUMMARY AT ANALYSIS THE STATED REASON FOR RETURN OF STYLUS AND CURSOR NOT ALIGNING ON THE SCREEN WAS CONFIRMED THE XY PRINTED CIRCUIT BOARDOVERLAY CABLECONNECTOR WAS CLEANED AND RESEATED AND THE STYLUS AND OVERLAY WERE ABLE TO BE CALIBRATED USING THE POINT METHOD THE PROGRAMMER FAILED AN INCOMING TEST RELATED TO THE LEFT ANTENNA CONNECTION AND THE LEFT ANTENNA WAS THEREFORE CONNECTED THE CIRCUIT BOARDS AND MECHANICAL PARTS WERE INSPECTED THE HARD DRIVE RECONFIGURED AND THE SOFTWARE UPDATED AND RELOADED | Connection Problem | Device problem code
6923290 | 89584808 | N | PRODUCT EVENT SUMMARY AT ANALYSIS THE STATED REASON FOR RETURN OF STYLUS AND CURSOR NOT ALIGNING ON THE SCREEN WAS CONFIRMED THE XY PRINTED CIRCUIT BOARDOVERLAY CABLECONNECTOR WAS CLEANED AND RESEATED AND THE STYLUS AND OVERLAY WERE ABLE TO BE CALIBRATED USING THE POINT METHOD THE PROGRAMMER FAILED AN INCOMING TEST RELATED TO THE LEFT ANTENNA CONNECTION AND THE LEFT ANTENNA WAS THEREFORE CONNECTED THE CIRCUIT BOARDS AND MECHANICAL PARTS WERE INSPECTED THE HARD DRIVE RECONFIGURED AND THE SOFTWARE UPDATED AND RELOADED | Failure to Align | Device problem code
6923290 | 89584808 | N | PRODUCT EVENT SUMMARY AT ANALYSIS THE STATED REASON FOR RETURN OF STYLUS AND CURSOR NOT ALIGNING ON THE SCREEN WAS CONFIRMED THE XY PRINTED CIRCUIT BOARDOVERLAY CABLECONNECTOR WAS CLEANED AND RESEATED AND THE STYLUS AND OVERLAY WERE ABLE TO BE CALIBRATED USING THE POINT METHOD THE PROGRAMMER FAILED AN INCOMING TEST RELATED TO THE LEFT ANTENNA CONNECTION AND THE LEFT ANTENNA WAS THEREFORE CONNECTED THE CIRCUIT BOARDS AND MECHANICAL PARTS WERE INSPECTED THE HARD DRIVE RECONFIGURED AND THE SOFTWARE UPDATED AND RELOADED | Interoperability Problem Identified | Medical Evaluation result code
6957260 | 89586508 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6957260 | 89586508 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6936509 | 89588853 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS B | Insufficient Information | Device problem code
6936509 | 89588853 | D | A PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS B | No Known Impact Or Consequence To Patient | Patient Problem Code
6957259 | 89592137 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6957259 | 89592137 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6957258 | 89592464 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6957258 | 89592464 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Incompatible Component/ Accessory | Medical Evaluation result code
6903213 | 89595100 | N | AFTER THE PROCEDURE THERE WAS NO ABNORMALITY ON THE ENDOSCOPIC IMAGE USING THE SUBJECT OTVS IN THE FACILITY THE SUBJECT OTVS AND OTVSHDLE WERE RETURNED TO OLYMPUS MEDICAL SYSTEMS CORPOMSC FOR THE EVALUATION OMSC PERFORMED THE INVESTIGATION USING THE FOLLOWING THE SUBJECT OTVS OTVSHDLE COMBINED WITH THE SUBJECT OTVS OMSC INVESTIGATED THESE DEVICES HOWEVER THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY FOUND OMSC CHECKED THE ENDOSCOPIC IMAGE AFTER THE SUBJECT OTVS WAS TURNED ON IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE THESE DEVICES WERE OPERATED FOR HOURS IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE AFTER THESE DEVICES WERE OPERATED FOR HOURS THE IMPACT WAS APPLIED TO THE SUBJECT OTVS OMSC CHECKED THESE DEVICES AND FOUND THE FOLLOWING THERE WAS NO FOREIGN MATERIAL ADHERED TO THE INSIDE OF THE SCOPE CONNECTER IN THE SUBJECT OTVS THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THERE WAS A SUSPECTED DISCONNECTION IN THE CABLE OF THE OTVSHDLE OMSC COULD NOT IDENTIFY THE ROOT CAUSE BECAUSE THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY WAS FOUND IN THE SUBJECT OTVS BASED ON THESE INVESTIGATIONS OMSC SURMISED THIS PHENOMENON MIGHT HAVE BEEN OCCURRED BY THE FOLLOWING THE ABNORMALITY IN THE COMMUNICATION BETWEEN THE SUBJECT OTVS AND THE OTVSHDLE OCCURRED TEMPORARILY BECAUSE THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED THE ABNORMALITY IN THE SIGNAL OF THE OTVSHDLE OCCURRED TEMPORARILY DUE TO DISCONNECTION IN THE CABLE OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED DP CIRCUIT BOARD WHICH HAD WHITE BALANCE SETTING AND THE COLOR ADJUSTMENT WAS NOT FUNCTIONED NORMALLY TEMPORARILY DUE TO UNSPECIFIED NOISES BY THE USAGE ENVIRONMENT THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED OTVS INSTRUCTION MANUAL STATES THE CORRESPONDING METHOD IN CASE OF AN ABNORMALITY THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED | Device Contamination with Chemical or Other Material | Device problem code
6903213 | 89595100 | N | AFTER THE PROCEDURE THERE WAS NO ABNORMALITY ON THE ENDOSCOPIC IMAGE USING THE SUBJECT OTVS IN THE FACILITY THE SUBJECT OTVS AND OTVSHDLE WERE RETURNED TO OLYMPUS MEDICAL SYSTEMS CORPOMSC FOR THE EVALUATION OMSC PERFORMED THE INVESTIGATION USING THE FOLLOWING THE SUBJECT OTVS OTVSHDLE COMBINED WITH THE SUBJECT OTVS OMSC INVESTIGATED THESE DEVICES HOWEVER THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY FOUND OMSC CHECKED THE ENDOSCOPIC IMAGE AFTER THE SUBJECT OTVS WAS TURNED ON IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE THESE DEVICES WERE OPERATED FOR HOURS IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE AFTER THESE DEVICES WERE OPERATED FOR HOURS THE IMPACT WAS APPLIED TO THE SUBJECT OTVS OMSC CHECKED THESE DEVICES AND FOUND THE FOLLOWING THERE WAS NO FOREIGN MATERIAL ADHERED TO THE INSIDE OF THE SCOPE CONNECTER IN THE SUBJECT OTVS THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THERE WAS A SUSPECTED DISCONNECTION IN THE CABLE OF THE OTVSHDLE OMSC COULD NOT IDENTIFY THE ROOT CAUSE BECAUSE THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY WAS FOUND IN THE SUBJECT OTVS BASED ON THESE INVESTIGATIONS OMSC SURMISED THIS PHENOMENON MIGHT HAVE BEEN OCCURRED BY THE FOLLOWING THE ABNORMALITY IN THE COMMUNICATION BETWEEN THE SUBJECT OTVS AND THE OTVSHDLE OCCURRED TEMPORARILY BECAUSE THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED THE ABNORMALITY IN THE SIGNAL OF THE OTVSHDLE OCCURRED TEMPORARILY DUE TO DISCONNECTION IN THE CABLE OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED DP CIRCUIT BOARD WHICH HAD WHITE BALANCE SETTING AND THE COLOR ADJUSTMENT WAS NOT FUNCTIONED NORMALLY TEMPORARILY DUE TO UNSPECIFIED NOISES BY THE USAGE ENVIRONMENT THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED OTVS INSTRUCTION MANUAL STATES THE CORRESPONDING METHOD IN CASE OF AN ABNORMALITY THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED | Disconnection | Device problem code
6903213 | 89595100 | N | AFTER THE PROCEDURE THERE WAS NO ABNORMALITY ON THE ENDOSCOPIC IMAGE USING THE SUBJECT OTVS IN THE FACILITY THE SUBJECT OTVS AND OTVSHDLE WERE RETURNED TO OLYMPUS MEDICAL SYSTEMS CORPOMSC FOR THE EVALUATION OMSC PERFORMED THE INVESTIGATION USING THE FOLLOWING THE SUBJECT OTVS OTVSHDLE COMBINED WITH THE SUBJECT OTVS OMSC INVESTIGATED THESE DEVICES HOWEVER THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY FOUND OMSC CHECKED THE ENDOSCOPIC IMAGE AFTER THE SUBJECT OTVS WAS TURNED ON IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE THESE DEVICES WERE OPERATED FOR HOURS IT WAS REPEATED TIMES THAT OMSC SET THE WHITE BALANCE AFTER THESE DEVICES WERE OPERATED FOR HOURS THE IMPACT WAS APPLIED TO THE SUBJECT OTVS OMSC CHECKED THESE DEVICES AND FOUND THE FOLLOWING THERE WAS NO FOREIGN MATERIAL ADHERED TO THE INSIDE OF THE SCOPE CONNECTER IN THE SUBJECT OTVS THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THERE WAS A SUSPECTED DISCONNECTION IN THE CABLE OF THE OTVSHDLE OMSC COULD NOT IDENTIFY THE ROOT CAUSE BECAUSE THIS PHENOMENON WAS NOT REPRODUCED AND NO ABNORMALITY WAS FOUND IN THE SUBJECT OTVS BASED ON THESE INVESTIGATIONS OMSC SURMISED THIS PHENOMENON MIGHT HAVE BEEN OCCURRED BY THE FOLLOWING THE ABNORMALITY IN THE COMMUNICATION BETWEEN THE SUBJECT OTVS AND THE OTVSHDLE OCCURRED TEMPORARILY BECAUSE THERE WAS THE FOREIGN MATERIAL ADHERED TO THE CONNECTER OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED THE ABNORMALITY IN THE SIGNAL OF THE OTVSHDLE OCCURRED TEMPORARILY DUE TO DISCONNECTION IN THE CABLE OF THE OTVSHDLE THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED DP CIRCUIT BOARD WHICH HAD WHITE BALANCE SETTING AND THE COLOR ADJUSTMENT WAS NOT FUNCTIONED NORMALLY TEMPORARILY DUE TO UNSPECIFIED NOISES BY THE USAGE ENVIRONMENT THEREFORE THE ABNORMALITY IN THE ENDOSCOPIC IMAGE OCCURRED OTVS INSTRUCTION MANUAL STATES THE CORRESPONDING METHOD IN CASE OF AN ABNORMALITY THERE WERE NO FURTHER DETAILS PROVIDED IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE SUPPLEMENTED | Interoperability Problem Identified | Medical Evaluation result code
6957960 | 89595522 | N | MEDTRONIC COMPLAINT REPORT B ARTICLE HERNIA SUPPL SS IPOM COMBINED WITH ANTERIOR COMPONENT SEPARATION TECHNIQUE IN GIANT INCISIONAL VENTRAL HERNIA REPAIRS SHUPYK NATIONAL MEDICAL ACADEMY OF POSTGRADUATE EDUCATION KYIV UKRAINE AUTHORS Y P FELESHTYNSKY V V SMISHCHUK V F VATAMANIUK | No Device Problem Found | Medical Evaluation result code
6948221 | 89600709 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948221 | 89600709 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6948333 | 89602275 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948333 | 89602275 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | No Audible Alarm | Device problem code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Disconnection | Device problem code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Loss of Power | Device problem code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Stress Problem Identified | Medical Evaluation result code
6939294 | 89603790 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST POWER LOSS ALARM AND REPLACE BATTERY NOW CONFIRMED IN THE LONG TRACE HISTORY DETAIL TRACE ANALYSIS CONFIRMED THE PUMP ALARMED POWER LOSS ALARM AND REPLACE BATTERY NOW WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR AT CONNECTOR BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Power Source Problem Identified | Medical Evaluation result code
6948342 | 89605424 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948342 | 89605424 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6948596 | 89605677 | N | B THE DEVICE SAMPLE WAS RETURNED FOR EVALUATION INITIAL REVIEW OF THE RETURNED SAMPLE INDICATES A JUNCTURE HUB LEAK DURING USE | Leak / Splash | Device problem code
6948596 | 89605677 | N | B THE DEVICE SAMPLE WAS RETURNED FOR EVALUATION INITIAL REVIEW OF THE RETURNED SAMPLE INDICATES A JUNCTURE HUB LEAK DURING USE | Leakage/Seal | Medical Evaluation result code
6948256 | 89605881 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948256 | 89605881 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6924083 | 89614832 | N | AT ANALYSIS THE ANALYZER GENERATED A WARNING DURING THE BENCH TEST LOSS OF COMMUNICATION THOUGH THE ANALYZER WAS RESET THREE TIMES THE WARNING KEPT RECURRING THE ANALYZER WAS BEING SENT ON FOR REPAIR AND REALLOCATION | Communication or Transmission Problem | Device problem code
6924083 | 89614832 | N | AT ANALYSIS THE ANALYZER GENERATED A WARNING DURING THE BENCH TEST LOSS OF COMMUNICATION THOUGH THE ANALYZER WAS RESET THREE TIMES THE WARNING KEPT RECURRING THE ANALYZER WAS BEING SENT ON FOR REPAIR AND REALLOCATION | Communications Problem Identified | Medical Evaluation result code
6923730 | 89616256 | N | THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Break | Device problem code
6923730 | 89616256 | N | THE DISTAL PORTION OF THE LEAD WAS RETURNED ANALYZED ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO A DEPRESSION WHILE IN VIVO | Fracture Problem | Medical Evaluation result code
6938280 | 89617282 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MINOR SCRATCHED LCD WINDOW B | Scratched Material | Device problem code
6938280 | 89617282 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER AND MINOR SCRATCHED LCD WINDOW B | Component Missing | Device problem code
6958474 | 89617453 | D | IT WAS REPORTED B THE PATIENT WAS EXPERIENCING A HEATING SENSATION WHEN RECHARGING THE SCS IPG A REPLACEMENT CHARGER WAS GIVEN TO THE PATIENT BUT THE ISSUE PERSISTED SUBSEQUENTLY THE PATIENTS SCS IPG WAS REPLACED WITH A NEW ONE | Insufficient Information | Device problem code
6958474 | 89617453 | D | IT WAS REPORTED B THE PATIENT WAS EXPERIENCING A HEATING SENSATION WHEN RECHARGING THE SCS IPG A REPLACEMENT CHARGER WAS GIVEN TO THE PATIENT BUT THE ISSUE PERSISTED SUBSEQUENTLY THE PATIENTS SCS IPG WAS REPLACED WITH A NEW ONE | Burning Sensation | Patient problem code
6958625 | 89621454 | D | A REPORT WAS RECEIVED THAT DURING PERMANENT IMPLANT PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED LEAD THE BSN SALES REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED THE PATIENT WAS DOING WELL POSTOPERATIVELY | No Consequences Or Impact To Patient | Patient problem code
6958625 | 89621454 | D | A REPORT WAS RECEIVED THAT DURING PERMANENT IMPLANT PROCEDURE THE PATIENT WAS IMPLANTED WITH AN EXPIRED LEAD THE BSN SALES REPRESENTATIVE CONFIRMED THAT IT WAS NOT EXPIRED THE PATIENT WAS DOING WELL POSTOPERATIVELY | Defective Device | Device problem code
6923240 | 89621632 | N | AT ANALYSIS THE ANALYZER FAILED ITS LIGHTEMITTING DIODE OPERATIONAL TEST INDICATING AN OUT OF SPECIFICATION FINDING IN AN ELECTRICAL MANNER THE ANALYZER WILL BE SENT ON FOR REPAIR AND RESTOCK | No Device Problem Found | Medical Evaluation result code
6923240 | 89621632 | N | AT ANALYSIS THE ANALYZER FAILED ITS LIGHTEMITTING DIODE OPERATIONAL TEST INDICATING AN OUT OF SPECIFICATION FINDING IN AN ELECTRICAL MANNER THE ANALYZER WILL BE SENT ON FOR REPAIR AND RESTOCK | Missing or Inadequate Safety Measures | Medical Evaluation result code
6922899 | 89621783 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE MONTH AND EIGHTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6922899 | 89621783 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE MONTH AND EIGHTEEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6958247 | 89622955 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6958247 | 89622955 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Incompatible Component/ Accessory | Medical Evaluation result code
6923814 | 89626966 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED THERE WAS BLOOD ON THE DISTAL CONDUCTOR OF THE LEAD AND IT WAS OBSTRUCTED THE ANALYST NOTED A TEST STYLET STOPPED INSERTION AT CM DUE TO DRIED BLOOD OBSTRUCTING THE DISTAL CONDUCTOR LUMEN | Obstruction of Flow | Device problem code
6923814 | 89626966 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED THERE WAS BLOOD ON THE DISTAL CONDUCTOR OF THE LEAD AND IT WAS OBSTRUCTED THE ANALYST NOTED A TEST STYLET STOPPED INSERTION AT CM DUE TO DRIED BLOOD OBSTRUCTING THE DISTAL CONDUCTOR LUMEN | Operational Problem Identified | Medical Evaluation result code
6923743 | 89630389 | N | THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE HELIX COULD NOT BE EXTENDED DUE TO THE DISTAL CONDUCTOR DISTORTED WITHIN THE DF CONNECTOR PIN AND THE DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP VISUALLY OVERRETRACTED | Mechanical Jam | Device problem code
6923743 | 89630389 | N | THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE DISTAL CONDUCTOR OF THE LEAD WAS EXTRINSICALLY OVERROTATED THE ANALYST NOTED THAT THE HELIX COULD NOT BE EXTENDED DUE TO THE DISTAL CONDUCTOR DISTORTED WITHIN THE DF CONNECTOR PIN AND THE DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP VISUALLY OVERRETRACTED | No Device Problem Found | Medical Evaluation result code
6925355 | 89630411 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER WAS HAVING SENSING ISSUES THE ANALYZER PASSED INCOMING TESTING | Device Sensing Problem | Device problem code
6925355 | 89630411 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER WAS HAVING SENSING ISSUES THE ANALYZER PASSED INCOMING TESTING | No Device Problem Found | Medical Evaluation result code
6952139 | 89630687 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | No Device Problem Found | Medical Evaluation result code
6924133 | 89633182 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE PROGRAMMER STYLUS DID NOT ALIGN WITH THE DISPLAY CURSOR IT WAS INDICATED THAT THE CONNECTION BETWEEN THE OVERLAY AND PRINTED CIRCUIT BOARD PCB REQUIRED A RESEAT THE PROGRAMMER WAS RETURNED TO SERVICE | Failure to Align | Device problem code
6924133 | 89633182 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE PROGRAMMER STYLUS DID NOT ALIGN WITH THE DISPLAY CURSOR IT WAS INDICATED THAT THE CONNECTION BETWEEN THE OVERLAY AND PRINTED CIRCUIT BOARD PCB REQUIRED A RESEAT THE PROGRAMMER WAS RETURNED TO SERVICE | Connection Problem | Device problem code
6924133 | 89633182 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE PROGRAMMER STYLUS DID NOT ALIGN WITH THE DISPLAY CURSOR IT WAS INDICATED THAT THE CONNECTION BETWEEN THE OVERLAY AND PRINTED CIRCUIT BOARD PCB REQUIRED A RESEAT THE PROGRAMMER WAS RETURNED TO SERVICE | Interoperability Problem Identified | Medical Evaluation result code
6924594 | 89634920 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE ANALYST NOTED THAT THE HELIX WOULD NOT BE EXTENDED DUE TO DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP OVERRETRACTED PHOTO TAKEN AND SAVED | Mechanical Jam | Device problem code
6924594 | 89634920 | N | PRODUCT EVENT SUMMARY THE FULL LEAD WAS RETURNED AND ANALYZED ANALYSIS INDICATED THAT THE HELIX EXCEEDED SPECIFICATION THE NUMBER OF TURNS TO EXTEND AND RETRACT THE ANALYST NOTED THAT THE HELIX WOULD NOT BE EXTENDED DUE TO DRIVE SHAFT LUG STUCK BEHIND THE RETRACTION STOP OVERRETRACTED PHOTO TAKEN AND SAVED | Mechanical Problem Identified | Medical Evaluation result code
6913147 | 89642287 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE | Partial Blockage | Device Problem Code
6913147 | 89642287 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE | Material Twisted / Bent | Device problem code
6913147 | 89642287 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE | Mechanical Jam | Device problem code
6913147 | 89642287 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE | Mechanical Problem Identified | Medical Evaluation result code
6959402 | 89647338 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED INACCURATE VALUES IS BEING REPORTED UNDER MFR NUMBER | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6959402 | 89647338 | N | B DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA THE REPORTED INACCURATE VALUES IS BEING REPORTED UNDER MFR NUMBER | Results Pending Completion of Investigation | Medical Evaluation result code
6952099 | 89652600 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Problem | Device problem code
6952099 | 89652600 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Source Problem Identified | Medical Evaluation result code
6952198 | 89652700 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Problem | Device problem code
6952198 | 89652700 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Source Problem Identified | Medical Evaluation result code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Battery Problem | Device Problem Code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Computer Software Problem | Device problem code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Component or Accessory Incompatibility | Device problem code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Interoperability Problem Identified | Medical Evaluation result code
6960290 | 89672958 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES POWER MANAGEMENT TOOL SHOWS THAT THE BACKUP BATTERY LOADED VOLTAGE LOADED VLITH AND UNLOADED VOLTAGE UL VLITH WERE WITHIN SPEC RANGE THE INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY OR FAILED BATTERY ALARMS DURING TESTING NO UNEXPECTED PUMP ERROR ALARMS DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR ALARM AND ERROR DUE TO A SOFTWARE ERROR THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY B | Software Problem Identified | Medical Evaluation result code
6960261 | 89675938 | D | EDWARDS LIFESCIENCES MAINTAINS AN IMPLANT PATIENT REGISTRY THIS REGISTRY IS A PATIENT TRACKING MECHANISM FOR SERIALIZED EDWARDS IMPLANTABLE DEVICES BIOPROSTHETIC HEART VALVES AND ANNULOPLASTY RINGS RATHER THAN A TRUE POSTMARKET SURVEILLANCE REGISTRY THROUGH THE REGISTRY EDWARDS IS NOTIFIED WHEN THESE DEVICES ARE IMPLANTED IN ADDITION PATIENT ANDOR DEVICE STATUS MAY BE REPORTED TO THE REGISTRY VIA THE IMPLANTATION DATA CARDS THE INFORMATION IS RECEIVED FROM VARIOUS SOURCES EG SURGEON HOSPITAL AND PATIENT FAMILY MEMBERS AND IS NOT RECEIVED IN THE FORM OF A CONVENTIONAL CUSTOMER COMPLAINT THE INFORMATION REPORTED MAY OR MAY NOT BE RELATED TO THE EDWARDS DEVICE IN THIS CASE IT WAS REPORTED THAT A MITRAL ANNULOPLASTY RING IMPLANTED FOR THREE YEARS AND THREE MONTHS WAS EXPLANTED DUE TO STENOSIS AND INSUFFICIENCY SECONDARY TO HOST TISSUE GROWTH THROUGH THE OPENING OF THE VALVE THE DEVICE WAS EXPLANTED AND THE PATIENTS MITRAL VALVE WAS REPLACED WITH AN EDWARDS BIOPROSTHETIC VALVE INTRAOPERATIVE ECHOCARDIOGRAM REVEALED THAT THE VALVE WAS IN GOOD POSITION WITH NO PVL PATIENT TOLERATED PROCEDURE WELL AND WAS DISCHARGED TO CVICU IN STABLE CONDITION SHE EXPERIENCED POSTOP ATRIAL FIBRILLATION AND WAS TREATED WITH MEDICATIONS AND THEN CARDIOVERSION ON POD SHE WAS DISCHARGED IN GOOD CONDITION ON REHAB ON POD | Atrial Fibrillation | Patient problem code
6960261 | 89675938 | D | EDWARDS LIFESCIENCES MAINTAINS AN IMPLANT PATIENT REGISTRY THIS REGISTRY IS A PATIENT TRACKING MECHANISM FOR SERIALIZED EDWARDS IMPLANTABLE DEVICES BIOPROSTHETIC HEART VALVES AND ANNULOPLASTY RINGS RATHER THAN A TRUE POSTMARKET SURVEILLANCE REGISTRY THROUGH THE REGISTRY EDWARDS IS NOTIFIED WHEN THESE DEVICES ARE IMPLANTED IN ADDITION PATIENT ANDOR DEVICE STATUS MAY BE REPORTED TO THE REGISTRY VIA THE IMPLANTATION DATA CARDS THE INFORMATION IS RECEIVED FROM VARIOUS SOURCES EG SURGEON HOSPITAL AND PATIENT FAMILY MEMBERS AND IS NOT RECEIVED IN THE FORM OF A CONVENTIONAL CUSTOMER COMPLAINT THE INFORMATION REPORTED MAY OR MAY NOT BE RELATED TO THE EDWARDS DEVICE IN THIS CASE IT WAS REPORTED THAT A MITRAL ANNULOPLASTY RING IMPLANTED FOR THREE YEARS AND THREE MONTHS WAS EXPLANTED DUE TO STENOSIS AND INSUFFICIENCY SECONDARY TO HOST TISSUE GROWTH THROUGH THE OPENING OF THE VALVE THE DEVICE WAS EXPLANTED AND THE PATIENTS MITRAL VALVE WAS REPLACED WITH AN EDWARDS BIOPROSTHETIC VALVE INTRAOPERATIVE ECHOCARDIOGRAM REVEALED THAT THE VALVE WAS IN GOOD POSITION WITH NO PVL PATIENT TOLERATED PROCEDURE WELL AND WAS DISCHARGED TO CVICU IN STABLE CONDITION SHE EXPERIENCED POSTOP ATRIAL FIBRILLATION AND WAS TREATED WITH MEDICATIONS AND THEN CARDIOVERSION ON POD SHE WAS DISCHARGED IN GOOD CONDITION ON REHAB ON POD | Stenosis | Patient problem code
6960261 | 89675938 | D | EDWARDS LIFESCIENCES MAINTAINS AN IMPLANT PATIENT REGISTRY THIS REGISTRY IS A PATIENT TRACKING MECHANISM FOR SERIALIZED EDWARDS IMPLANTABLE DEVICES BIOPROSTHETIC HEART VALVES AND ANNULOPLASTY RINGS RATHER THAN A TRUE POSTMARKET SURVEILLANCE REGISTRY THROUGH THE REGISTRY EDWARDS IS NOTIFIED WHEN THESE DEVICES ARE IMPLANTED IN ADDITION PATIENT ANDOR DEVICE STATUS MAY BE REPORTED TO THE REGISTRY VIA THE IMPLANTATION DATA CARDS THE INFORMATION IS RECEIVED FROM VARIOUS SOURCES EG SURGEON HOSPITAL AND PATIENT FAMILY MEMBERS AND IS NOT RECEIVED IN THE FORM OF A CONVENTIONAL CUSTOMER COMPLAINT THE INFORMATION REPORTED MAY OR MAY NOT BE RELATED TO THE EDWARDS DEVICE IN THIS CASE IT WAS REPORTED THAT A MITRAL ANNULOPLASTY RING IMPLANTED FOR THREE YEARS AND THREE MONTHS WAS EXPLANTED DUE TO STENOSIS AND INSUFFICIENCY SECONDARY TO HOST TISSUE GROWTH THROUGH THE OPENING OF THE VALVE THE DEVICE WAS EXPLANTED AND THE PATIENTS MITRAL VALVE WAS REPLACED WITH AN EDWARDS BIOPROSTHETIC VALVE INTRAOPERATIVE ECHOCARDIOGRAM REVEALED THAT THE VALVE WAS IN GOOD POSITION WITH NO PVL PATIENT TOLERATED PROCEDURE WELL AND WAS DISCHARGED TO CVICU IN STABLE CONDITION SHE EXPERIENCED POSTOP ATRIAL FIBRILLATION AND WAS TREATED WITH MEDICATIONS AND THEN CARDIOVERSION ON POD SHE WAS DISCHARGED IN GOOD CONDITION ON REHAB ON POD | Insufficient Information | Device problem code
6960313 | 89677748 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF ABSCESS INFECTION SWELLING AND WHEN SENT FOR AN MRI THE PATIENT HAD TO BE ADMITTED TO THE EMERGENCY ROOM AND REQUIRED SEDATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Material Discolored | Device problem code
6960313 | 89677748 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF ABSCESS INFECTION SWELLING AND WHEN SENT FOR AN MRI THE PATIENT HAD TO BE ADMITTED TO THE EMERGENCY ROOM AND REQUIRED SEDATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Migration or Expulsion of Device | Device problem code
6960313 | 89677748 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF ABSCESS INFECTION SWELLING AND WHEN SENT FOR AN MRI THE PATIENT HAD TO BE ADMITTED TO THE EMERGENCY ROOM AND REQUIRED SEDATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Leak / Splash | Device problem code
6960313 | 89677748 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF ABSCESS INFECTION SWELLING AND WHEN SENT FOR AN MRI THE PATIENT HAD TO BE ADMITTED TO THE EMERGENCY ROOM AND REQUIRED SEDATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Leakage/Seal | Medical Evaluation result code
6960313 | 89677748 | N | B FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF ABSCESS INFECTION SWELLING AND WHEN SENT FOR AN MRI THE PATIENT HAD TO BE ADMITTED TO THE EMERGENCY ROOM AND REQUIRED SEDATION ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Device Migration | Medical Evaluation result code
6960780 | 89678425 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6960780 | 89678425 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
6960763 | 89678535 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD SUBSEQUENTLY PATIENT WAS TAKEN TO THE HOSPITAL DIAGNOSED WITH DIABETIC KETOACIDOSIS AND TREATED WITH AN INSULIN DRIP | Insufficient Information | Device problem code
6960763 | 89678535 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD SUBSEQUENTLY PATIENT WAS TAKEN TO THE HOSPITAL DIAGNOSED WITH DIABETIC KETOACIDOSIS AND TREATED WITH AN INSULIN DRIP | No Known Impact Or Consequence To Patient | Patient problem code
6960767 | 89678894 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6960767 | 89678894 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
6941434 | 89703575 | N | B ONE OPENUSED RESERVOIR WAS EVALUATED A VISUAL INSPECTION WAS PERFORMED AND THE PLUNGER STOPPERPLUNGER AND O RINGS WERE FOUND TO BE MISSING ANALYSIS WAS UNABLE TO CONFIRM CUSTOMER COMPLAINT | Component Missing | Device problem code
6941434 | 89703575 | N | B ONE OPENUSED RESERVOIR WAS EVALUATED A VISUAL INSPECTION WAS PERFORMED AND THE PLUNGER STOPPERPLUNGER AND O RINGS WERE FOUND TO BE MISSING ANALYSIS WAS UNABLE TO CONFIRM CUSTOMER COMPLAINT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6939952 | 89705674 | N | THE ANESTHESIOLOGIST ADJUSTED THE POPOFF VALVE TO RESOLVE THE REPORTED FAULT PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Detachment of Device or device Component | Device problem code
6939952 | 89705674 | N | THE ANESTHESIOLOGIST ADJUSTED THE POPOFF VALVE TO RESOLVE THE REPORTED FAULT PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS | Interoperability Problem Identified | Medical Evaluation result code
6961076 | 89705980 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED LESION THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT E EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US FOREIGN B STUDY NAME B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Apparent Adverse Event | Device problem code
6961076 | 89705980 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED LESION THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT E EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US FOREIGN B STUDY NAME B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONSADVERSE EVENTS | No Findings Available | Medical Evaluation result code
6961247 | 89711367 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT K EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONS ADVERSE EVENTS | No Apparent Adverse Event | Device problem code
6961247 | 89711367 | N | THE PATIENT REQUIRED REVASCULARIZATION OF THE TREATED VESSEL THIS IS BEING REPORTED AS A FOLLOWUP TO THE CLINICAL STUDY PATIENT INFORMATION REGARDING RELEVANT TESTS OR LABORATORY DATA IS UNKNOWN THIS INFORMATION WAS NOT AVAILABLE FROM THE FACILITY THE DRUG INFORMATION IS NOTED BELOW STELLAREX OTW DRUGCOATED ANGIOPLASTY BALLOON PACLITAXEL DRUG MG THERAPY DATE B ADJUNCT TO PTA IN THE TREATMENT OF DENOVO OR POST PTA OCCLUDED STENOTIC OR RESTENOTIC LESIONS EXCLUDING INSTENT IN FEMPOP ARTERIES LOT K EXPIRATION DATE UDI AND PMA NUMBERS ARE NOT APPLICABLE THIS DEVICE WAS USED IN CLINICAL APPLICATION PRIOR TO BEING AVAILABLE IN THE US B DURING THE INDEX PROCEDURE THE PRODUCT WORKED AS INTENDED THUS NO PRODUCT EVALUATION WAS REQUIRED PER THE IFU RESTENOSIS IS LISTED AS A POTENTIAL COMPLICATIONS ADVERSE EVENTS | No Findings Available | Medical Evaluation result code
6925123 | 89717277 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE ANALYZER WOULD NOT COMMUNICATE WITH A PROGRAMMER THE ANALYZER WAS TO BE SENT FOR REPAIR | Communication or Transmission Problem | Device problem code
6925123 | 89717277 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS CONFIRMED THE CUSTOMER COMMENT THAT THE ANALYZER WOULD NOT COMMUNICATE WITH A PROGRAMMER THE ANALYZER WAS TO BE SENT FOR REPAIR | Communications Problem Identified | Medical Evaluation result code
6910764 | 89717693 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED TO HAVE A CLEAR HALO AROUND THE RED BLOOD CELL BUTTON AND THEN A PINKISH OUTER CIRCLE ALSO | Connection Problem | Device problem code
6910764 | 89717693 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED TO HAVE A CLEAR HALO AROUND THE RED BLOOD CELL BUTTON AND THEN A PINKISH OUTER CIRCLE ALSO | Interoperability Problem Identified | Medical Evaluation result code
6961438 | 89717695 | N | BATCH REVIEW PERFORMED ON SEPTEMBER B ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS STEM REVISION AT YEARS FOLLOWING CEMENTLESS PRIMARY THA THE STEM LOOKS SLIGHTLY UNDERSIZED THE CUP EVIDENTLY PROUD FROM ACETABULUM IT IS CONCEIVABLE THAT THE STEM HAS NOT BEEN ABLE TO FIND ADEQUATE PRIMARY STABILITY AND THEREFORE EFFECTIVE BONE INGROWTH WAS NEVER ACHIEVED THERE IS NO INDICATION THAT A FAULTY IMPLANT CAUSED THE ADVERSE EVENT ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE TISSUE AND CULTURES HAVE BEEN FOUND TO HAVE NO GROWTH OF ANY FORM OF BACTERIA THE SURGERY WAS COMPLETED SUCCESSFULLY ON B | Unstable | Device Problem Code
6961438 | 89717695 | N | BATCH REVIEW PERFORMED ON SEPTEMBER B ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS STEM REVISION AT YEARS FOLLOWING CEMENTLESS PRIMARY THA THE STEM LOOKS SLIGHTLY UNDERSIZED THE CUP EVIDENTLY PROUD FROM ACETABULUM IT IS CONCEIVABLE THAT THE STEM HAS NOT BEEN ABLE TO FIND ADEQUATE PRIMARY STABILITY AND THEREFORE EFFECTIVE BONE INGROWTH WAS NEVER ACHIEVED THERE IS NO INDICATION THAT A FAULTY IMPLANT CAUSED THE ADVERSE EVENT ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE TISSUE AND CULTURES HAVE BEEN FOUND TO HAVE NO GROWTH OF ANY FORM OF BACTERIA THE SURGERY WAS COMPLETED SUCCESSFULLY ON B | Inadequacy of Device Shape and/or Size | Device problem code
6961438 | 89717695 | N | BATCH REVIEW PERFORMED ON SEPTEMBER B ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS STEM REVISION AT YEARS FOLLOWING CEMENTLESS PRIMARY THA THE STEM LOOKS SLIGHTLY UNDERSIZED THE CUP EVIDENTLY PROUD FROM ACETABULUM IT IS CONCEIVABLE THAT THE STEM HAS NOT BEEN ABLE TO FIND ADEQUATE PRIMARY STABILITY AND THEREFORE EFFECTIVE BONE INGROWTH WAS NEVER ACHIEVED THERE IS NO INDICATION THAT A FAULTY IMPLANT CAUSED THE ADVERSE EVENT ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE TISSUE AND CULTURES HAVE BEEN FOUND TO HAVE NO GROWTH OF ANY FORM OF BACTERIA THE SURGERY WAS COMPLETED SUCCESSFULLY ON B | Stress Problem Identified | Medical Evaluation result code
6925566 | 89718095 | N | PRODUCT ANALYSIS THE FULL LEAD WAS RETURNED IN SEGMENTS AND ANALYZED ANALYSIS INDICATED THAT THE INNER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO | Break | Device problem code
6921177 | 89718595 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM ON B OF TYPE UICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT IN ADDITION REVIEW OF THE CONTROLLER LOG FILES REVEALED A CRITICAL BATTERY ALARM ON B THIS ALARM WAS MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THIS IS AN INCIDENTAL FINDING AND NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device Problem Code
6921177 | 89718595 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM ON B OF TYPE UICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT IN ADDITION REVIEW OF THE CONTROLLER LOG FILES REVEALED A CRITICAL BATTERY ALARM ON B THIS ALARM WAS MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THIS IS AN INCIDENTAL FINDING AND NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communication or Transmission Problem | Device problem code
6921177 | 89718595 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM ON B OF TYPE UICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT IN ADDITION REVIEW OF THE CONTROLLER LOG FILES REVEALED A CRITICAL BATTERY ALARM ON B THIS ALARM WAS MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THIS IS AN INCIDENTAL FINDING AND NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Communications Problem Identified | Medical Evaluation result code
6921177 | 89718595 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED A CONTROLLER FAULT ALARM ON B OF TYPE UICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT IN ADDITION REVIEW OF THE CONTROLLER LOG FILES REVEALED A CRITICAL BATTERY ALARM ON B THIS ALARM WAS MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THIS IS AN INCIDENTAL FINDING AND NOT RELATED TO THE REPORTED EVENT THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Fracture Problem | Medical Evaluation result code
6927804 | 89719003 | N | THE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT SYSTEM TESTING DID NOT INDICATE ANY ABNORMAL OPERATION OF THE CONTROLLER CONTROLLER WAS RECEIVED UPGRADED WITH SOFTWARE MAINTENANCE RELEASE SMR VERSIONS FOR UIC AND FOR PMC THE ALARMS MENTIONED IN THE EVENT DETAILS COULD BE OBSERVED DURING THE SMR UPDATE HOWEVER THEY ARE NOT ASSOCIATED WITH A CONTROLLER MALFUNCTION THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | No Device Problem Found | Medical Evaluation result code
6924092 | 89719460 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER WAS HAVING PROBLEMS WITH THE ATRIAL CHANNEL THE ANALYZER PASSED INCOMING TESTING IT WAS INDICATED THAT THE FRONT CONNECTOR WAS LOOSE | Loose or Intermittent Connection | Device problem code
6924092 | 89719460 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER WAS HAVING PROBLEMS WITH THE ATRIAL CHANNEL THE ANALYZER PASSED INCOMING TESTING IT WAS INDICATED THAT THE FRONT CONNECTOR WAS LOOSE | Interoperability Problem Identified | Medical Evaluation result code
6924092 | 89719460 | N | PRODUCT EVENT SUMMARY MANUFACTURERS ANALYSIS WAS UNABLE TO CONFIRM THE CUSTOMER COMMENT THAT THE ANALYZER WAS HAVING PROBLEMS WITH THE ATRIAL CHANNEL THE ANALYZER PASSED INCOMING TESTING IT WAS INDICATED THAT THE FRONT CONNECTOR WAS LOOSE | Detachment of Device or device Component | Device problem code
6961582 | 89722839 | D | IT WAS REPORTED THAT PATIENT HAD BEEN DIAGNOSED WITH DIABETIC KETOACIDOSIS AFTER WEARING THE POD PATIENT WAS SUBSEQUENTLY HOSPITALIZED HOWEVER NO FURTHER DETAILS COULD BE OBTAINED | Diabetic Ketoacidosis | Patient problem code
6961582 | 89722839 | D | IT WAS REPORTED THAT PATIENT HAD BEEN DIAGNOSED WITH DIABETIC KETOACIDOSIS AFTER WEARING THE POD PATIENT WAS SUBSEQUENTLY HOSPITALIZED HOWEVER NO FURTHER DETAILS COULD BE OBTAINED | Insufficient Information | Device problem code
6945380 | 89724382 | N | B INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6945380 | 89724382 | N | B INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6944817 | 89726236 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE MONTH AND TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944817 | 89726236 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE MONTH AND TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6945273 | 89729992 | N | B INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6945273 | 89729992 | N | B INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6944744 | 89733175 | N | B THE PUMP WAS RECEIVED WITH A CONSTANT FLASHING WHITE LIGHT ON THE DISPLAY AND BEEPING DUE TO A CRACKEDBLEEDING ON LCD GLASS THE DEVICE WAS UNABLE TO PERFORM FUNCTIONAL TESTS DUE TO MALFUNCTIONING DISPLAY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW | Noise, Audible | Device problem code
6944744 | 89733175 | N | B THE PUMP WAS RECEIVED WITH A CONSTANT FLASHING WHITE LIGHT ON THE DISPLAY AND BEEPING DUE TO A CRACKEDBLEEDING ON LCD GLASS THE DEVICE WAS UNABLE TO PERFORM FUNCTIONAL TESTS DUE TO MALFUNCTIONING DISPLAY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW | No Device Problem Found | Medical Evaluation result code
6944744 | 89733175 | N | B THE PUMP WAS RECEIVED WITH A CONSTANT FLASHING WHITE LIGHT ON THE DISPLAY AND BEEPING DUE TO A CRACKEDBLEEDING ON LCD GLASS THE DEVICE WAS UNABLE TO PERFORM FUNCTIONAL TESTS DUE TO MALFUNCTIONING DISPLAY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6944744 | 89733175 | N | B THE PUMP WAS RECEIVED WITH A CONSTANT FLASHING WHITE LIGHT ON THE DISPLAY AND BEEPING DUE TO A CRACKEDBLEEDING ON LCD GLASS THE DEVICE WAS UNABLE TO PERFORM FUNCTIONAL TESTS DUE TO MALFUNCTIONING DISPLAY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
6958914 | 89736127 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Injury | Patient problem code
6958914 | 89736127 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Insufficient Information | Device problem code
6962105 | 89737286 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Injury | Patient problem code
6962105 | 89737286 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Insufficient Information | Device problem code
6962193 | 89738784 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Insufficient Information | Device problem code
6962193 | 89738784 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Injury | Patient problem code
6939107 | 89749118 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED ONE CONTROLLER FAULT ALARMS LOGGED ON B THIS ALARMS WAS OF TYPE SYSUICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device problem code
6939107 | 89749118 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED ONE CONTROLLER FAULT ALARMS LOGGED ON B THIS ALARMS WAS OF TYPE SYSUICAPSFAIL INDICATING A MALFUNCTION OF THE AUTOMATIC POWER SWITCH APS CIRCUIT HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A LEAKY DIODE ON THE APS CIRCUIT THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Operational Problem Identified | Medical Evaluation result code
6962511 | 89752730 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE PROVIDE ADDITIONAL DETAIL OF THE EVENT WHAT PROCEDURE WAS BEING PERFORMED WHEN WAS THE PROCEDURE DATE DID THE GASTRIC LEAK OCCUR DURING THE ORIGINAL PROCEDURE IF YES HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK CONTROLLED WAS THERE ANY CHANGE IN THE PROCEDURE AS A RESULT OF THE LEAK WAS THERE ANY PATIENT CONSEQUENCE OR DID THE GASTRIC LEAK OCCUR POSTOPERATIVELY IF YES HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK CONTROLLED PLEASE PROVIDE FURTHER DETAIL OF THE POSTOPERATIVE EVENT WHAT IS THE CURRENT PATIENT STATUS PROCEDURE PRIMARY LAP SLEEVE GASTRECTOMY DATE B GASTRIC LEAK HAPPENED POST OP APPROX WEEKS PATIENT STATUS GOOD ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY ISSUES NOTED WITH STAPLER PERFORMANCE IN PRIMARY PROCEDURE WHY DOES THE SURGEON BELIEVE THAT THE LEAK WEEKS POSTOP IS RELATED TO THE ETHICON DEVICE WHAT COLOR RELOADS WERE USED WHAT IS THE GENDER BMI OR PATIENT WHAT IS THE CURRENT PATIENT STATUS | Leak / Splash | Device problem code
6962511 | 89752730 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED CAN YOU PLEASE PROVIDE ADDITIONAL DETAIL OF THE EVENT WHAT PROCEDURE WAS BEING PERFORMED WHEN WAS THE PROCEDURE DATE DID THE GASTRIC LEAK OCCUR DURING THE ORIGINAL PROCEDURE IF YES HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK CONTROLLED WAS THERE ANY CHANGE IN THE PROCEDURE AS A RESULT OF THE LEAK WAS THERE ANY PATIENT CONSEQUENCE OR DID THE GASTRIC LEAK OCCUR POSTOPERATIVELY IF YES HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK CONTROLLED PLEASE PROVIDE FURTHER DETAIL OF THE POSTOPERATIVE EVENT WHAT IS THE CURRENT PATIENT STATUS PROCEDURE PRIMARY LAP SLEEVE GASTRECTOMY DATE B GASTRIC LEAK HAPPENED POST OP APPROX WEEKS PATIENT STATUS GOOD ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY ISSUES NOTED WITH STAPLER PERFORMANCE IN PRIMARY PROCEDURE WHY DOES THE SURGEON BELIEVE THAT THE LEAK WEEKS POSTOP IS RELATED TO THE ETHICON DEVICE WHAT COLOR RELOADS WERE USED WHAT IS THE GENDER BMI OR PATIENT WHAT IS THE CURRENT PATIENT STATUS | Leakage/Seal | Medical Evaluation result code
6962185 | 89754282 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Insufficient Information | Device problem code
6962185 | 89754282 | D | IT WAS REPORTED THAT THE NURSE RECEIVED A NEEDLE STICK INJURY WHEN ACTIVATING THE SAFETY MECHANISM ON A ML BD SAFETYGLIDE INSULIN SYRINGE G X MM MEDICAL INTERVENTION IS UNKNOWN | Injury | Patient problem code
6932589 | 89761216 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY THE LEAD WAS INSPECTED AND ANALYZED THE LEAD HAD BEEN SEVERED AND WAS RETURNED IN TWO SEGMENTS EACH RETURNED SEGMENT WAS DETERMINED TO BE ELECTRICALLY CONTINUOUS A XRAY WAS PERFORMED WHICH SHOWED NO FRACTURES THE CLINICAL OBSERVATION WAS NOT ABLE TO BE CONFIRMED | No Device Problem Found | Medical Evaluation result code
6928312 | 89761346 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL EXAMINATION NOTED DRIED BLOOD BODY FLUID IN THE LEAD LUMEN A LABORATORY TEST WIRE WAS INSERTED IN THE LEAD BUT COULD BE ADVANCED NO FURTHER THAN APPROXIMATELY MM FROM THE TERMINAL PIN ANALYSIS CONFIRMED THE PRESENCE OF FOREIGN MATERIAL WITHIN THE LEAD LUMEN IN THIS AREA TESTING OF THE FOREIGN MATERIAL CONFIRMED IT WAS A COMBINATION OF COAGULATED BODILY FLUID AND POLYMER AS A RESULT OF THE GUIDE WIRELEAD INTERACTION THIS WAS CONFIRMED TO BE THE ROOT CAUSE OF THE DIFFICULTY EXPERIENCED ADVANCING THE WHISPER WIRE IN THE FIELD | Device Contamination with Chemical or Other Material | Device problem code
6928312 | 89761346 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL EXAMINATION NOTED DRIED BLOOD BODY FLUID IN THE LEAD LUMEN A LABORATORY TEST WIRE WAS INSERTED IN THE LEAD BUT COULD BE ADVANCED NO FURTHER THAN APPROXIMATELY MM FROM THE TERMINAL PIN ANALYSIS CONFIRMED THE PRESENCE OF FOREIGN MATERIAL WITHIN THE LEAD LUMEN IN THIS AREA TESTING OF THE FOREIGN MATERIAL CONFIRMED IT WAS A COMBINATION OF COAGULATED BODILY FLUID AND POLYMER AS A RESULT OF THE GUIDE WIRELEAD INTERACTION THIS WAS CONFIRMED TO BE THE ROOT CAUSE OF THE DIFFICULTY EXPERIENCED ADVANCING THE WHISPER WIRE IN THE FIELD | Contamination of Environment by Device | Medical Evaluation result code
6962944 | 89775288 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6963008 | 89775641 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6963495 | 89796464 | N | ADDITIONAL BREAST SHIELDS AND INSERTS WERE SENT TO THE CUSTOMER IN FOLLOW UP WITH THE CUSTOMER SHE STATED THAT SHE WAS EXPERIENCING PAIN USING THE MM BREAST SHIELDS THAT CAME WITH THE PUMP SHE TRIED THE MM BECAUSE SHE THOUGHT SHE NEEDED A LARGER SIZE BUT THEY CAUSED EVEN MORE PAIN SHE SWITCHED TO THE MM BREAST SHIELDS AND DID NOT EXPERIENCE PAIN BUT SHE WASNT EXPRESSING MUCH MILK SO SHE SWITCHED BACK TO THE MM SHE STATED THAT SHE WAS EXPERIENCING REDNESS AROUND THE AREOLA AND WAS PRESCRIBED JACK NEWMANS NIPPLE CREAM BY HER PHYSICIAN SHE STATED THAT SHE STILL HAS A LITTLE PAIN DUE TO SENSITIVITY ISSUES THE INSTRUCTIONS FOR USE ITEM NUMBER B REV E AND PAGE PROVIDES INSTRUCTIONS FOR BREAST SHIELD SIZING INCLUDING REFERENCE TO MEDELABREASTSHIELDSCOM AND REFERRAL TO A LACTATION CONSULTANT FOR ASSISTANCE IN CHOOSING THE CORRECT SIZE BREAST SHIELD | Inadequacy of Device Shape and/or Size | Device problem code
6963495 | 89796464 | N | ADDITIONAL BREAST SHIELDS AND INSERTS WERE SENT TO THE CUSTOMER IN FOLLOW UP WITH THE CUSTOMER SHE STATED THAT SHE WAS EXPERIENCING PAIN USING THE MM BREAST SHIELDS THAT CAME WITH THE PUMP SHE TRIED THE MM BECAUSE SHE THOUGHT SHE NEEDED A LARGER SIZE BUT THEY CAUSED EVEN MORE PAIN SHE SWITCHED TO THE MM BREAST SHIELDS AND DID NOT EXPERIENCE PAIN BUT SHE WASNT EXPRESSING MUCH MILK SO SHE SWITCHED BACK TO THE MM SHE STATED THAT SHE WAS EXPERIENCING REDNESS AROUND THE AREOLA AND WAS PRESCRIBED JACK NEWMANS NIPPLE CREAM BY HER PHYSICIAN SHE STATED THAT SHE STILL HAS A LITTLE PAIN DUE TO SENSITIVITY ISSUES THE INSTRUCTIONS FOR USE ITEM NUMBER B REV E AND PAGE PROVIDES INSTRUCTIONS FOR BREAST SHIELD SIZING INCLUDING REFERENCE TO MEDELABREASTSHIELDSCOM AND REFERRAL TO A LACTATION CONSULTANT FOR ASSISTANCE IN CHOOSING THE CORRECT SIZE BREAST SHIELD | Stress Problem Identified | Medical Evaluation result code
6962946 | 89798096 | N | B WHAT COLOR CARTRIDGES WERE USED WAS BUTTRESSING MATERIAL USED WAS THE DEVICE DIFFICULT TO CLOSE WERE ANY UNUSUAL SOUNDS NOTED WHAT DOES THE SURGEON MEAN BY PLICATION PLEASE PROVIDE ADDITIONAL INFORMATION ABOUT THE CLIPS THAT WERE OPEN DID THE SURGERY START AS LAPAROSCOPIC OR OPEN WHAT IS THE CURRENT PATIENT STATUS | Difficult to Open or Close | Device problem code
6962946 | 89798096 | N | B WHAT COLOR CARTRIDGES WERE USED WAS BUTTRESSING MATERIAL USED WAS THE DEVICE DIFFICULT TO CLOSE WERE ANY UNUSUAL SOUNDS NOTED WHAT DOES THE SURGEON MEAN BY PLICATION PLEASE PROVIDE ADDITIONAL INFORMATION ABOUT THE CLIPS THAT WERE OPEN DID THE SURGERY START AS LAPAROSCOPIC OR OPEN WHAT IS THE CURRENT PATIENT STATUS | Operational Problem Identified | Medical Evaluation result code
6964109 | 89805951 | N | AS SOON AS ELLEX MEDICAL BECAME AWARE OF THE INCIDENT THE CALIBRATION OF THE LASER SYSTEM AND ALIGNMENT OF THE OPTICAL PATHS WAS VERIFIED BY ELLEX SERVICE THROUGH A PAF PRODUCT ACCEPTANCE FAULT RECORD ATTACHED ALL THE PARAMETERS IN ALIGNMENTS CHECKS INCLUDING OFFSET IN THE PAF ARE WITHIN THE ACCEPTANCE CRITERIA WHICH PROVES THAT THE DEVICE IS NOT AT FAULT AND IS PERFORMING WITHIN THE NORMAL SPECIFICATIONS AND THERE WAS NO ISSUE WITH ALIGNMENT MECHANICAL OPERATION OR CALIBRATION THE GENERAL DEVICE ASPECTS WITH RESPECT TO USABILITY INSTRUCTIONS FOR USE AND THE RISK CONTROLS WERE ALSO EVALUATED ADEQUACY OF DEVICE DESIGN INDICATIONS OF PARAMETERS RELEVANT TO SAFETY ANTERIOR AND POSTERIOR OFFSET LOCATED ON THE LASER HEAD CONTROL KNOB ARE LETTER AND COLOUR CODED ACCURACY OF CONTROLS OFFSET RANGE IS WELL DEFINED WITH THE REQUIRED TOLERANCE PAF POST THE EVENT ALSO CONFIRMED THE OFFSET RESULTS WERE NORMAL ADEQUACY OF LABELLING OPERATOR MANUAL INCLUDED IN ATTACHMENT INCLUDES SUFFICIENT INSTRUCTIONS WARNINGS PRECAUTIONS FOR THE SAFE USE OF THE DEVICE AND THE LABELLING IS CONSIDERED ADEQUATE RISK MANAGEMENT REVIEW INDICATED THAT THE RISK CONTROL MEASURES RISK ASSESSMENT RA TANGO IN ATTACHMENT ARE ADEQUATE AND WITHIN THE NORMAL USE RISK LEVELS ARE ACCEPTABLE CUSTOMER COMPLAINTS THE DEVICE HAS BEEN IN USE FOR LESS THAN A YEAR NO OTHER UNTOWARD INCIDENTS RELATED TO DEVICE USE WERE REPORTED | Calibration Problem | Device problem code
6964109 | 89805951 | N | AS SOON AS ELLEX MEDICAL BECAME AWARE OF THE INCIDENT THE CALIBRATION OF THE LASER SYSTEM AND ALIGNMENT OF THE OPTICAL PATHS WAS VERIFIED BY ELLEX SERVICE THROUGH A PAF PRODUCT ACCEPTANCE FAULT RECORD ATTACHED ALL THE PARAMETERS IN ALIGNMENTS CHECKS INCLUDING OFFSET IN THE PAF ARE WITHIN THE ACCEPTANCE CRITERIA WHICH PROVES THAT THE DEVICE IS NOT AT FAULT AND IS PERFORMING WITHIN THE NORMAL SPECIFICATIONS AND THERE WAS NO ISSUE WITH ALIGNMENT MECHANICAL OPERATION OR CALIBRATION THE GENERAL DEVICE ASPECTS WITH RESPECT TO USABILITY INSTRUCTIONS FOR USE AND THE RISK CONTROLS WERE ALSO EVALUATED ADEQUACY OF DEVICE DESIGN INDICATIONS OF PARAMETERS RELEVANT TO SAFETY ANTERIOR AND POSTERIOR OFFSET LOCATED ON THE LASER HEAD CONTROL KNOB ARE LETTER AND COLOUR CODED ACCURACY OF CONTROLS OFFSET RANGE IS WELL DEFINED WITH THE REQUIRED TOLERANCE PAF POST THE EVENT ALSO CONFIRMED THE OFFSET RESULTS WERE NORMAL ADEQUACY OF LABELLING OPERATOR MANUAL INCLUDED IN ATTACHMENT INCLUDES SUFFICIENT INSTRUCTIONS WARNINGS PRECAUTIONS FOR THE SAFE USE OF THE DEVICE AND THE LABELLING IS CONSIDERED ADEQUATE RISK MANAGEMENT REVIEW INDICATED THAT THE RISK CONTROL MEASURES RISK ASSESSMENT RA TANGO IN ATTACHMENT ARE ADEQUATE AND WITHIN THE NORMAL USE RISK LEVELS ARE ACCEPTABLE CUSTOMER COMPLAINTS THE DEVICE HAS BEEN IN USE FOR LESS THAN A YEAR NO OTHER UNTOWARD INCIDENTS RELATED TO DEVICE USE WERE REPORTED | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6963189 | 89809645 | N | B DATE SENT BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DOES THE SURGEON BELIEVE THAT THE POSTOPERATIVE LEAK NOTED IN ARTICLE WAS RELATED TO THE ETHICON STAPLER IF YES WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED PLEASE EXPLAIN | Leak / Splash | Device problem code
6963189 | 89809645 | N | B DATE SENT BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DOES THE SURGEON BELIEVE THAT THE POSTOPERATIVE LEAK NOTED IN ARTICLE WAS RELATED TO THE ETHICON STAPLER IF YES WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED PLEASE EXPLAIN | Leakage/Seal | Medical Evaluation result code
6963050 | 89811159 | N | B DATE SENT BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DOES THE SURGEON BELIEVE THAT THE PATIENT LEAK AND BLEEDING WAS RELATED TO THE ETHICON STAPLER | Leak / Splash | Device problem code
6963050 | 89811159 | N | B DATE SENT BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT DOES THE SURGEON BELIEVE THAT THE PATIENT LEAK AND BLEEDING WAS RELATED TO THE ETHICON STAPLER | Leakage/Seal | Medical Evaluation result code
6951961 | 89821469 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Problem | Device problem code
6951961 | 89821469 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON THE PCB OF THE AC POWER ADAPTER WHICH CONTRIBUTED TO THE FIELD COMPLAINT OF POWER UP ANOMALY | Power Source Problem Identified | Medical Evaluation result code
6947916 | 89822122 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY B THIS B ANTIBODY SCREENING TEST WELL WAS KNOWN TO BE B FOR BOTH E ANTIGEN AND K ANTIGEN | Connection Problem | Device problem code
6947916 | 89822122 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION WHICH APPEARED VISUALLY B THIS B ANTIBODY SCREENING TEST WELL WAS KNOWN TO BE B FOR BOTH E ANTIGEN AND K ANTIGEN | Interoperability Problem Identified | Medical Evaluation result code
6933450 | 89824934 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD INTERMITTENT LOSS OF COMMUNICATION WHEN USED WITH KNOWN GOOD PROGRAMMER | Communication or Transmission Problem | Device problem code
6933450 | 89824934 | N | PRODUCT EVENT SUMMARY ANALYSIS FOUND THE ANALYZER HAD INTERMITTENT LOSS OF COMMUNICATION WHEN USED WITH KNOWN GOOD PROGRAMMER | Communications Problem Identified | Medical Evaluation result code
6964959 | 89828308 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6964959 | 89828308 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6964959 | 89828308 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF GRANULOMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING FOR THE REPORTED EVENT OF GRANULOMA ADVERSE EVENTS C OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6960991 | 89837316 | D | PATIENT WITH CHRONIC PELVIC AND ABDOMINAL PAIN NAUSEA WEIGHT LOSS FOLLOWING PLACEMENT OF ESSURE B BILATERAL FALLOPIAN TUBE ESSURE REMOVAL PERFORMED B | Nausea | Patient problem code
6960991 | 89837316 | D | PATIENT WITH CHRONIC PELVIC AND ABDOMINAL PAIN NAUSEA WEIGHT LOSS FOLLOWING PLACEMENT OF ESSURE B BILATERAL FALLOPIAN TUBE ESSURE REMOVAL PERFORMED B | Insufficient Information | Device problem code
6960991 | 89837316 | D | PATIENT WITH CHRONIC PELVIC AND ABDOMINAL PAIN NAUSEA WEIGHT LOSS FOLLOWING PLACEMENT OF ESSURE B BILATERAL FALLOPIAN TUBE ESSURE REMOVAL PERFORMED B | Weight Fluctuations | Patient problem code
6960991 | 89837316 | D | PATIENT WITH CHRONIC PELVIC AND ABDOMINAL PAIN NAUSEA WEIGHT LOSS FOLLOWING PLACEMENT OF ESSURE B BILATERAL FALLOPIAN TUBE ESSURE REMOVAL PERFORMED B | Pain, Abdominal | Patient problem code
6960249 | 89842390 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6960249 | 89842390 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6965347 | 89844764 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6947141 | 89849734 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6947141 | 89849734 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
6951477 | 89853246 | D | IT WAS REPORTED THAT A SPECTRUM PUMP DISPLAYED FALSE DOWNSTREAM OCCLUSION MESSAGES DURING TESTING IN THE BIOMEDICAL SERVICE AREA THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Patient Involvement | Patient problem code
6951477 | 89853246 | D | IT WAS REPORTED THAT A SPECTRUM PUMP DISPLAYED FALSE DOWNSTREAM OCCLUSION MESSAGES DURING TESTING IN THE BIOMEDICAL SERVICE AREA THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | Device Displays Incorrect Message | Device Problem Code
6921145 | 89856133 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED A SERIES OF POWER UP EVENTS TIMESTAMPED WITH AT WHICH IMPLIES THAT THE RTC LOST BATTERY POWER WHILE THE CONTROLLER WAS POWERED OFF ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING TO VERIFY IF THE REAL TIME CLOCK WAS WORKING THE CONTROLLER DATE AND TIME WAS SET TO THE CORRECT TIME AND LEFT TO RUN FOR OVER AN HOUR IT WAS ABLE TO KEEP AN ACCURATE DATE AND TIME EVEN AFTER A DAY HAD ELAPSED THE TEMPORARY RESET WAS MOST LIKELY DUE TO A DISCHARGE OF THE BATTERY DURING STORAGE THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO COIN CELL RUN DOWN HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Leak / Splash | Device problem code
6921145 | 89856133 | N | ONE CONTROLLER WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED A SERIES OF POWER UP EVENTS TIMESTAMPED WITH AT WHICH IMPLIES THAT THE RTC LOST BATTERY POWER WHILE THE CONTROLLER WAS POWERED OFF ANALYSIS OF THE DEVICE REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING TO VERIFY IF THE REAL TIME CLOCK WAS WORKING THE CONTROLLER DATE AND TIME WAS SET TO THE CORRECT TIME AND LEFT TO RUN FOR OVER AN HOUR IT WAS ABLE TO KEEP AN ACCURATE DATE AND TIME EVEN AFTER A DAY HAD ELAPSED THE TEMPORARY RESET WAS MOST LIKELY DUE TO A DISCHARGE OF THE BATTERY DURING STORAGE THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO COIN CELL RUN DOWN HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Leakage/Seal | Medical Evaluation result code
6965913 | 89860104 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS IN THE S MGDL WITH A TRACE OF KETONES THE CAUSE WAS NOT KNOWN A CORRECTION BOLUS WAS DELIVERED THE CUSTOMER WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG AND DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS TREATED WITH FLUIDS AND HUMALOG THE HIGH BG AND DKA WERE RESOLVED THE CUSTOMER WAS DISCHARGED ON B THE CUSTOMER REVERTED TO USING LANTUS FOR INSULIN THERAPY A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER ACKNOWLEDGED THE RESULTS OF THE SYSTEM CHECK AND WAS SATISFIED | Adverse Event Without Identified Device or Use Problem | Device problem code
6965913 | 89860104 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS IN THE S MGDL WITH A TRACE OF KETONES THE CAUSE WAS NOT KNOWN A CORRECTION BOLUS WAS DELIVERED THE CUSTOMER WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG AND DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS TREATED WITH FLUIDS AND HUMALOG THE HIGH BG AND DKA WERE RESOLVED THE CUSTOMER WAS DISCHARGED ON B THE CUSTOMER REVERTED TO USING LANTUS FOR INSULIN THERAPY A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER ACKNOWLEDGED THE RESULTS OF THE SYSTEM CHECK AND WAS SATISFIED | Hyperglycemia | Patient problem code
6965913 | 89860104 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG LEVEL WAS IN THE S MGDL WITH A TRACE OF KETONES THE CAUSE WAS NOT KNOWN A CORRECTION BOLUS WAS DELIVERED THE CUSTOMER WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL FOR THE HIGH BG AND DIABETIC KETOACIDOSIS DKA THE CUSTOMER WAS TREATED WITH FLUIDS AND HUMALOG THE HIGH BG AND DKA WERE RESOLVED THE CUSTOMER WAS DISCHARGED ON B THE CUSTOMER REVERTED TO USING LANTUS FOR INSULIN THERAPY A PUMP SYSTEM CHECK WAS PERFORMED AND NO DEVICE ISSUES WERE IDENTIFIED THE CUSTOMER ACKNOWLEDGED THE RESULTS OF THE SYSTEM CHECK AND WAS SATISFIED | Diabetic Ketoacidosis | Patient problem code
6950761 | 89867095 | N | B PHYSIOCONTROL EVALUATED THE DEVICE AND VERIFIED THE REPORTED ISSUE IT WAS OBSERVED THAT PCBMOUNTED JACK CONNECTOR DESIGNATOR J WAS NOT CORRECTLY SEATED IN ITS CONNECTOR RECEPTACLE ON THE POWER PCB ASSEMBLY PHYSIO PROPERLY SEATED PCBMOUNTED JACK CONNECTOR DESIGNATOR J INTO ITS CONNECTOR ON THE POWER PCB ASSEMBLY PROPER DEVICE OPERATION WAS THEN OBSERVED THROUGH FUNCTIONAL AND PERFORMANCE TESTING THE DEVICE WAS SUBSEQUENTLY RETURNED TO THE CUSTOMER B | No Device Problem Found | Medical Evaluation result code
6967149 | 89915624 | N | B THE LOT NUMBER WAS NOT PROVIDED THEREFORE THE UDI IS NOT AVAILABLE B THE REVIEW OF THE MANUFACTURING PAPERWORK COULD NOT BE PERFORMED AS THE LOT NUMBERS WERE NOT AVAILABLE B ACCORDING TO THE GORE EXCLUDER AAA ENDOPROSTHESES INSTRUCTIONS FOR USE ADVERSE EVENTS WHICH MAY OCCUR AND REQUIRE INTERVENTION INCLUDE ENDOLEAK AND ANEURYSM ENLARGMENT | No Device Problem Found | Medical Evaluation result code
6967040 | 89917454 | N | B LOTSERIAL NUMBERS WERE NOT PROVIDED THEREFORE A REVIEW OF THE MANUFACTURING RECORDS COULD NOT BE CONDUCTED THEREFORE THE PUBLISHED DATE OF THE ARTICLE WAS USEDPER THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE IFU ADVERSE EVENTS THAT MAY OCCUR AND OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK AND ANEURYSM ENLARGEMENT | No Device Problem Found | Medical Evaluation result code
6965949 | 89923187 | D | IT WAS REPORTED THAT ON B THE PATIENT WAS AWAKENED POSSIBLY WITH A PAIN OR NUMBNESS IN THE CHEEK ONCE AWAKE BOTH THE PATIENT AND SPOUSE NOTICED SOME SLURRING OF THE PATIENTS SPEECH THE PATIENT ALSO NOTED NUMBNESS AND PARENTHESIS OF THE RIGHT HAND THE PATIENT WENT BACK TO SLEEP AND WOKE UP IN THE MORNING FEELING IMPROVED BUT NOT FULLY BACK TO NORMAL THE PATIENT WENT TO THE EMERGENCY ROOM ER WHERE A HEAD COMPUTERIZED TOMOGRAPHY CT SCAN WAS PERFORMED THERE WERE NO ACUTE CHANGES BUT THERE WAS EVIDENCE OF AN OLD OCCIPITAL INFARCT UPON ARRIVAL TO THE ER PROTHROMBIN TIME PT WAS INTERNATIONAL NORMALIZED RATIO INR WAS AND PARTIAL THROMBOPLASTIN TIME PTT WAS THE NEUROLOGIST SAW THE PATIENT AND NOTED MILD DYSARTHRIA LEFT VISUAL FIELD CUT POSSIBLY FROM OLD STROKE AND DECREASED HEARING SUBJECTS NORMAL BASELINE A COMPUTERIZED TOMOGRAPHY ANGIOGRAM CTA WAS UNREMARKABLE AND NO PROXIMAL FLOW LIMITING STENOSIS WAS SEEN THE PATIENTS SPELL RESOLVED AND NO FURTHER WORKUP WAS INDICATED THE NEUROLOGIST INDICATED THAT A SMALL STROKE COULD NOT BE RULED OUT BUT SINCE THE PATIENT IS ALREADY ON THE APPROPRIATE MEDICAL THERAPY NO FURTHER TREATMENT WAS WARRANTED THE PATIENT CONTINUED TO IMPROVE AND WAS DISCHARGED ON B WITH ONLY MINIMAL DYSARTHRIA AND NO OTHER RESIDUAL DEFICITS THE EVENT WAS DEEMED RESOLVED THE PRINCIPAL INVESTIGATOR INDICATED THE EVENT WAS POSSIBLY RELATED TO THE VENTRICULAR ASSIST DEVICE VAD SYSTEM AND PATIENT CONDITION THE VAD REMAINS IN USE NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Numbness | Patient problem code
6965949 | 89923187 | D | IT WAS REPORTED THAT ON B THE PATIENT WAS AWAKENED POSSIBLY WITH A PAIN OR NUMBNESS IN THE CHEEK ONCE AWAKE BOTH THE PATIENT AND SPOUSE NOTICED SOME SLURRING OF THE PATIENTS SPEECH THE PATIENT ALSO NOTED NUMBNESS AND PARENTHESIS OF THE RIGHT HAND THE PATIENT WENT BACK TO SLEEP AND WOKE UP IN THE MORNING FEELING IMPROVED BUT NOT FULLY BACK TO NORMAL THE PATIENT WENT TO THE EMERGENCY ROOM ER WHERE A HEAD COMPUTERIZED TOMOGRAPHY CT SCAN WAS PERFORMED THERE WERE NO ACUTE CHANGES BUT THERE WAS EVIDENCE OF AN OLD OCCIPITAL INFARCT UPON ARRIVAL TO THE ER PROTHROMBIN TIME PT WAS INTERNATIONAL NORMALIZED RATIO INR WAS AND PARTIAL THROMBOPLASTIN TIME PTT WAS THE NEUROLOGIST SAW THE PATIENT AND NOTED MILD DYSARTHRIA LEFT VISUAL FIELD CUT POSSIBLY FROM OLD STROKE AND DECREASED HEARING SUBJECTS NORMAL BASELINE A COMPUTERIZED TOMOGRAPHY ANGIOGRAM CTA WAS UNREMARKABLE AND NO PROXIMAL FLOW LIMITING STENOSIS WAS SEEN THE PATIENTS SPELL RESOLVED AND NO FURTHER WORKUP WAS INDICATED THE NEUROLOGIST INDICATED THAT A SMALL STROKE COULD NOT BE RULED OUT BUT SINCE THE PATIENT IS ALREADY ON THE APPROPRIATE MEDICAL THERAPY NO FURTHER TREATMENT WAS WARRANTED THE PATIENT CONTINUED TO IMPROVE AND WAS DISCHARGED ON B WITH ONLY MINIMAL DYSARTHRIA AND NO OTHER RESIDUAL DEFICITS THE EVENT WAS DEEMED RESOLVED THE PRINCIPAL INVESTIGATOR INDICATED THE EVENT WAS POSSIBLY RELATED TO THE VENTRICULAR ASSIST DEVICE VAD SYSTEM AND PATIENT CONDITION THE VAD REMAINS IN USE NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Stenosis | Patient problem code
6965949 | 89923187 | D | IT WAS REPORTED THAT ON B THE PATIENT WAS AWAKENED POSSIBLY WITH A PAIN OR NUMBNESS IN THE CHEEK ONCE AWAKE BOTH THE PATIENT AND SPOUSE NOTICED SOME SLURRING OF THE PATIENTS SPEECH THE PATIENT ALSO NOTED NUMBNESS AND PARENTHESIS OF THE RIGHT HAND THE PATIENT WENT BACK TO SLEEP AND WOKE UP IN THE MORNING FEELING IMPROVED BUT NOT FULLY BACK TO NORMAL THE PATIENT WENT TO THE EMERGENCY ROOM ER WHERE A HEAD COMPUTERIZED TOMOGRAPHY CT SCAN WAS PERFORMED THERE WERE NO ACUTE CHANGES BUT THERE WAS EVIDENCE OF AN OLD OCCIPITAL INFARCT UPON ARRIVAL TO THE ER PROTHROMBIN TIME PT WAS INTERNATIONAL NORMALIZED RATIO INR WAS AND PARTIAL THROMBOPLASTIN TIME PTT WAS THE NEUROLOGIST SAW THE PATIENT AND NOTED MILD DYSARTHRIA LEFT VISUAL FIELD CUT POSSIBLY FROM OLD STROKE AND DECREASED HEARING SUBJECTS NORMAL BASELINE A COMPUTERIZED TOMOGRAPHY ANGIOGRAM CTA WAS UNREMARKABLE AND NO PROXIMAL FLOW LIMITING STENOSIS WAS SEEN THE PATIENTS SPELL RESOLVED AND NO FURTHER WORKUP WAS INDICATED THE NEUROLOGIST INDICATED THAT A SMALL STROKE COULD NOT BE RULED OUT BUT SINCE THE PATIENT IS ALREADY ON THE APPROPRIATE MEDICAL THERAPY NO FURTHER TREATMENT WAS WARRANTED THE PATIENT CONTINUED TO IMPROVE AND WAS DISCHARGED ON B WITH ONLY MINIMAL DYSARTHRIA AND NO OTHER RESIDUAL DEFICITS THE EVENT WAS DEEMED RESOLVED THE PRINCIPAL INVESTIGATOR INDICATED THE EVENT WAS POSSIBLY RELATED TO THE VENTRICULAR ASSIST DEVICE VAD SYSTEM AND PATIENT CONDITION THE VAD REMAINS IN USE NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Cerebrovascular Accident | Patient Problem Code
6965949 | 89923187 | D | IT WAS REPORTED THAT ON B THE PATIENT WAS AWAKENED POSSIBLY WITH A PAIN OR NUMBNESS IN THE CHEEK ONCE AWAKE BOTH THE PATIENT AND SPOUSE NOTICED SOME SLURRING OF THE PATIENTS SPEECH THE PATIENT ALSO NOTED NUMBNESS AND PARENTHESIS OF THE RIGHT HAND THE PATIENT WENT BACK TO SLEEP AND WOKE UP IN THE MORNING FEELING IMPROVED BUT NOT FULLY BACK TO NORMAL THE PATIENT WENT TO THE EMERGENCY ROOM ER WHERE A HEAD COMPUTERIZED TOMOGRAPHY CT SCAN WAS PERFORMED THERE WERE NO ACUTE CHANGES BUT THERE WAS EVIDENCE OF AN OLD OCCIPITAL INFARCT UPON ARRIVAL TO THE ER PROTHROMBIN TIME PT WAS INTERNATIONAL NORMALIZED RATIO INR WAS AND PARTIAL THROMBOPLASTIN TIME PTT WAS THE NEUROLOGIST SAW THE PATIENT AND NOTED MILD DYSARTHRIA LEFT VISUAL FIELD CUT POSSIBLY FROM OLD STROKE AND DECREASED HEARING SUBJECTS NORMAL BASELINE A COMPUTERIZED TOMOGRAPHY ANGIOGRAM CTA WAS UNREMARKABLE AND NO PROXIMAL FLOW LIMITING STENOSIS WAS SEEN THE PATIENTS SPELL RESOLVED AND NO FURTHER WORKUP WAS INDICATED THE NEUROLOGIST INDICATED THAT A SMALL STROKE COULD NOT BE RULED OUT BUT SINCE THE PATIENT IS ALREADY ON THE APPROPRIATE MEDICAL THERAPY NO FURTHER TREATMENT WAS WARRANTED THE PATIENT CONTINUED TO IMPROVE AND WAS DISCHARGED ON B WITH ONLY MINIMAL DYSARTHRIA AND NO OTHER RESIDUAL DEFICITS THE EVENT WAS DEEMED RESOLVED THE PRINCIPAL INVESTIGATOR INDICATED THE EVENT WAS POSSIBLY RELATED TO THE VENTRICULAR ASSIST DEVICE VAD SYSTEM AND PATIENT CONDITION THE VAD REMAINS IN USE NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6965949 | 89923187 | D | IT WAS REPORTED THAT ON B THE PATIENT WAS AWAKENED POSSIBLY WITH A PAIN OR NUMBNESS IN THE CHEEK ONCE AWAKE BOTH THE PATIENT AND SPOUSE NOTICED SOME SLURRING OF THE PATIENTS SPEECH THE PATIENT ALSO NOTED NUMBNESS AND PARENTHESIS OF THE RIGHT HAND THE PATIENT WENT BACK TO SLEEP AND WOKE UP IN THE MORNING FEELING IMPROVED BUT NOT FULLY BACK TO NORMAL THE PATIENT WENT TO THE EMERGENCY ROOM ER WHERE A HEAD COMPUTERIZED TOMOGRAPHY CT SCAN WAS PERFORMED THERE WERE NO ACUTE CHANGES BUT THERE WAS EVIDENCE OF AN OLD OCCIPITAL INFARCT UPON ARRIVAL TO THE ER PROTHROMBIN TIME PT WAS INTERNATIONAL NORMALIZED RATIO INR WAS AND PARTIAL THROMBOPLASTIN TIME PTT WAS THE NEUROLOGIST SAW THE PATIENT AND NOTED MILD DYSARTHRIA LEFT VISUAL FIELD CUT POSSIBLY FROM OLD STROKE AND DECREASED HEARING SUBJECTS NORMAL BASELINE A COMPUTERIZED TOMOGRAPHY ANGIOGRAM CTA WAS UNREMARKABLE AND NO PROXIMAL FLOW LIMITING STENOSIS WAS SEEN THE PATIENTS SPELL RESOLVED AND NO FURTHER WORKUP WAS INDICATED THE NEUROLOGIST INDICATED THAT A SMALL STROKE COULD NOT BE RULED OUT BUT SINCE THE PATIENT IS ALREADY ON THE APPROPRIATE MEDICAL THERAPY NO FURTHER TREATMENT WAS WARRANTED THE PATIENT CONTINUED TO IMPROVE AND WAS DISCHARGED ON B WITH ONLY MINIMAL DYSARTHRIA AND NO OTHER RESIDUAL DEFICITS THE EVENT WAS DEEMED RESOLVED THE PRINCIPAL INVESTIGATOR INDICATED THE EVENT WAS POSSIBLY RELATED TO THE VENTRICULAR ASSIST DEVICE VAD SYSTEM AND PATIENT CONDITION THE VAD REMAINS IN USE NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Insufficient Information | Device problem code
6944801 | 89926311 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6944801 | 89926311 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944834 | 89927204 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6944834 | 89927204 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6964936 | 89930111 | N | DISTRIBUTOR NARRATIVE THE MANUFACTURER UNIMAX MEDICAL SYSTEMS INC IS RESPONSIBLE FOR PERFORMING EVALUATION INVESTIGATION AND ANY REMEDIAL ACTIONS RELATED TO THIS REPORTED DEVICE ISSUE PER AGREEMENT WITH B | No Device Problem Found | Medical Evaluation result code
6922937 | 89930264 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6922937 | 89930264 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6922916 | 89930598 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6922916 | 89930598 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944794 | 89935772 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6944794 | 89935772 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6967047 | 89937151 | N | PATIENTS DATE OF BIRTH AND AGE UNAVAILABLE PATIENTS WEIGHT UNAVAILABLE DEVICE EVALUATION FIBROUS MATERIAL AND BURNT BIOLOGICAL MATERIAL SEEN AT THE TIP ONE CONFIRMED DEAD FIBER AND ONE QUADRANT COULD NOT BE ASSESSED BECAUSE OF BURNT ON BIOLOGICAL MATERIAL SEEN ON THE TIP FTIR SPECTRA EVALUATION REVEALED THE COMPOSITION OF MATERIAL ON THE DISTAL TIP WAS POLYTETRAFLUOROETHYLENE TEFLON | Power Problem | Device problem code
6967047 | 89937151 | N | PATIENTS DATE OF BIRTH AND AGE UNAVAILABLE PATIENTS WEIGHT UNAVAILABLE DEVICE EVALUATION FIBROUS MATERIAL AND BURNT BIOLOGICAL MATERIAL SEEN AT THE TIP ONE CONFIRMED DEAD FIBER AND ONE QUADRANT COULD NOT BE ASSESSED BECAUSE OF BURNT ON BIOLOGICAL MATERIAL SEEN ON THE TIP FTIR SPECTRA EVALUATION REVEALED THE COMPOSITION OF MATERIAL ON THE DISTAL TIP WAS POLYTETRAFLUOROETHYLENE TEFLON | Power Source Problem Identified | Medical Evaluation result code
6916694 | 89937294 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR | Component Incompatible | Device problem code
6916694 | 89937294 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR | Interoperability Problem Identified | Medical Evaluation result code
6913522 | 89938539 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT IS THE CORRECT PROCEDURE DATE AS IT WAS LISTED AS B WAS THE PROCEDURE PERFORMED ON B HOW DID THE DEVICE MALFUNCTIONDID NOT FIRE CORRECTLY WHERE WITHIN THE GAP SETTING SCALE WAS THE ORANGE INDICATOR PRIOR TO FIRING WHAT CONFIRMATION WAS RECEIVED THAT THE DEVICE WAS FULLY FIRED DID THE DEVICE STAPLE WAS THE STAPLE LINE COMPLETE WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT WERE THE STAPLES DEPLOYED INTO THE TISSUE WERE THE STAPLES SEEN IN THE SURGICAL FIELD DID THE DEVICE CUT WAS THE CUT LINE FULLY CIRCUMFERENTIAL AROUND THE TARGET TISSUE IF THE DEVICE FULLY CUT WERE BOTH TISSUE DONUTS PRESENT IF THE DEVICE FULLY CUT WERE BOTH DONUTS COMPLETE HOW MANY COUNTERCLOCKWISE REVOLUTIONS WERE USED TO OPEN THE DEVICE HOW WAS IT CONFIRMED THAT THE ANVIL WAS FREE FROM THE TISSUE PRIOR TO REMOVING | Misfire | Device Problem Code
6913522 | 89938539 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT IS THE CORRECT PROCEDURE DATE AS IT WAS LISTED AS B WAS THE PROCEDURE PERFORMED ON B HOW DID THE DEVICE MALFUNCTIONDID NOT FIRE CORRECTLY WHERE WITHIN THE GAP SETTING SCALE WAS THE ORANGE INDICATOR PRIOR TO FIRING WHAT CONFIRMATION WAS RECEIVED THAT THE DEVICE WAS FULLY FIRED DID THE DEVICE STAPLE WAS THE STAPLE LINE COMPLETE WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT WERE THE STAPLES DEPLOYED INTO THE TISSUE WERE THE STAPLES SEEN IN THE SURGICAL FIELD DID THE DEVICE CUT WAS THE CUT LINE FULLY CIRCUMFERENTIAL AROUND THE TARGET TISSUE IF THE DEVICE FULLY CUT WERE BOTH TISSUE DONUTS PRESENT IF THE DEVICE FULLY CUT WERE BOTH DONUTS COMPLETE HOW MANY COUNTERCLOCKWISE REVOLUTIONS WERE USED TO OPEN THE DEVICE HOW WAS IT CONFIRMED THAT THE ANVIL WAS FREE FROM THE TISSUE PRIOR TO REMOVING | Component Missing | Device problem code
6913522 | 89938539 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT IS THE CORRECT PROCEDURE DATE AS IT WAS LISTED AS B WAS THE PROCEDURE PERFORMED ON B HOW DID THE DEVICE MALFUNCTIONDID NOT FIRE CORRECTLY WHERE WITHIN THE GAP SETTING SCALE WAS THE ORANGE INDICATOR PRIOR TO FIRING WHAT CONFIRMATION WAS RECEIVED THAT THE DEVICE WAS FULLY FIRED DID THE DEVICE STAPLE WAS THE STAPLE LINE COMPLETE WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT WERE THE STAPLES DEPLOYED INTO THE TISSUE WERE THE STAPLES SEEN IN THE SURGICAL FIELD DID THE DEVICE CUT WAS THE CUT LINE FULLY CIRCUMFERENTIAL AROUND THE TARGET TISSUE IF THE DEVICE FULLY CUT WERE BOTH TISSUE DONUTS PRESENT IF THE DEVICE FULLY CUT WERE BOTH DONUTS COMPLETE HOW MANY COUNTERCLOCKWISE REVOLUTIONS WERE USED TO OPEN THE DEVICE HOW WAS IT CONFIRMED THAT THE ANVIL WAS FREE FROM THE TISSUE PRIOR TO REMOVING | Fire | Device problem code
6913522 | 89938539 | N | B ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT IS THE CORRECT PROCEDURE DATE AS IT WAS LISTED AS B WAS THE PROCEDURE PERFORMED ON B HOW DID THE DEVICE MALFUNCTIONDID NOT FIRE CORRECTLY WHERE WITHIN THE GAP SETTING SCALE WAS THE ORANGE INDICATOR PRIOR TO FIRING WHAT CONFIRMATION WAS RECEIVED THAT THE DEVICE WAS FULLY FIRED DID THE DEVICE STAPLE WAS THE STAPLE LINE COMPLETE WERE THERE ANY STAPLES MISSING FROM THE STAPLE LINE WHAT WAS THE SHAPE OF THE STAPLES BFORMATION IRREGULAR LEGS STRAIGHT WERE THE STAPLES DEPLOYED INTO THE TISSUE WERE THE STAPLES SEEN IN THE SURGICAL FIELD DID THE DEVICE CUT WAS THE CUT LINE FULLY CIRCUMFERENTIAL AROUND THE TARGET TISSUE IF THE DEVICE FULLY CUT WERE BOTH TISSUE DONUTS PRESENT IF THE DEVICE FULLY CUT WERE BOTH DONUTS COMPLETE HOW MANY COUNTERCLOCKWISE REVOLUTIONS WERE USED TO OPEN THE DEVICE HOW WAS IT CONFIRMED THAT THE ANVIL WAS FREE FROM THE TISSUE PRIOR TO REMOVING | Missing or Inadequate Safety Measures | Medical Evaluation result code
6918952 | 89939488 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND THE INSTRUMENT TO BE OPERATING AS EXPECTED | Connection Problem | Device problem code
6918952 | 89939488 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND THE INSTRUMENT TO BE OPERATING AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
6925980 | 89939862 | N | IT WAS REPORTED THAT THE LIGHT HEAD CAME OFF THE BIOMED CALLED STATING THEY WILL USE A THIRD PARTY TO REPAIR THE COMPLAINT IF STRYKER RECEIVES ANY FURTHER INFORMATION ABOUT THIS COMPLAINT A SUPPLEMENTAL WILL BE FILED THERE WAS NO PATIENT INVOLVEMENT OR ADVERSE CONSEQUENCE REPORTED | Detachment of Device or device Component | Device problem code
6927344 | 89942262 | N | THE MTML WAS RETURNED TO ISI AND EVALUATED FAILURE ANALYSIS WAS UNABLE TO BE REPRODUCED THE CUSTOMER REPORTED FAILURE MODE HOWEVER FAILURE ANALYSIS CONFIRMED THAT RELATED SYSTEM ERROR CODE AS HAVING OCCURRED VIA THE SYSTEM LOGS A VISUAL INSPECTION WAS PERFORMED SINE CYCLE WAS PERFORMED WITHOUT ANY ISSUES TESTS PERFORMED VIA DTE AND MATLAB PASSED NO TROUBLE WAS FOUND WITH THE EVALUATION OF THE MTML THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Device Displays Incorrect Message | Device problem code
6927344 | 89942262 | N | THE MTML WAS RETURNED TO ISI AND EVALUATED FAILURE ANALYSIS WAS UNABLE TO BE REPRODUCED THE CUSTOMER REPORTED FAILURE MODE HOWEVER FAILURE ANALYSIS CONFIRMED THAT RELATED SYSTEM ERROR CODE AS HAVING OCCURRED VIA THE SYSTEM LOGS A VISUAL INSPECTION WAS PERFORMED SINE CYCLE WAS PERFORMED WITHOUT ANY ISSUES TESTS PERFORMED VIA DTE AND MATLAB PASSED NO TROUBLE WAS FOUND WITH THE EVALUATION OF THE MTML THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | No Device Problem Found | Medical Evaluation result code
6927797 | 89944290 | N | THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER IN USE DURING THE EVENT WAS B LOG ANALYSIS REVEALED THAT THE BATTERY WOULD DISCHARGE AT AN EXPECTED RATE THE REPORTED EVENT NOT LASTING LONG ENOUGH COULD NOT BE CONFIRMED OR DUPLICATED DURING FUNCTIONAL TESTING THE RETURNED BATTERY PASSED FUNCTIONAL TESTING THE BATTERY DISCHARGED AT AN EXPECTED RATE WHEN TESTED THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Leak / Splash | Device problem code
6927797 | 89944290 | N | THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER IN USE DURING THE EVENT WAS B LOG ANALYSIS REVEALED THAT THE BATTERY WOULD DISCHARGE AT AN EXPECTED RATE THE REPORTED EVENT NOT LASTING LONG ENOUGH COULD NOT BE CONFIRMED OR DUPLICATED DURING FUNCTIONAL TESTING THE RETURNED BATTERY PASSED FUNCTIONAL TESTING THE BATTERY DISCHARGED AT AN EXPECTED RATE WHEN TESTED THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Failure to Discharge | Device Problem Code
6927797 | 89944290 | N | THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER IN USE DURING THE EVENT WAS B LOG ANALYSIS REVEALED THAT THE BATTERY WOULD DISCHARGE AT AN EXPECTED RATE THE REPORTED EVENT NOT LASTING LONG ENOUGH COULD NOT BE CONFIRMED OR DUPLICATED DURING FUNCTIONAL TESTING THE RETURNED BATTERY PASSED FUNCTIONAL TESTING THE BATTERY DISCHARGED AT AN EXPECTED RATE WHEN TESTED THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Leakage/Seal | Medical Evaluation result code
6926650 | 89944358 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE | Connection Problem | Device problem code
6926650 | 89944358 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY WEAK POSITIVE | Interoperability Problem Identified | Medical Evaluation result code
6927857 | 89945236 | N | CONTROLLER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE CONTROLLER IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE FUNCTIONAL ANALYSIS OF THE CONTROLLER CONFIRMED THAT THE PS POWER CONNECTOR DID NOT WORK PROPERLY SYSTEM TESTING CONFIRMED A CONTROLLER FAULT ALARM WHILE PS WAS ACTIVE REVIEW OF THE LOG FILES ALSO REVEALED CONTROLLER FAULT ALARM WHILE PS WAS ACTIVE THE CONTROLLER WAS IN USE WHEN THE REPORTED EVENT OCCURRED THIS MALFUNCTION IS MOST LIKELY DUE TO A LEAKY DIODE D ON THE AUTOMATIC POWER SWITCH CIRCUIT OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Defective Alarm | Device Problem Code
6927857 | 89945236 | N | CONTROLLER B WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE CONTROLLER IN RELATION TO THE REPORTED EVENT REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE FUNCTIONAL ANALYSIS OF THE CONTROLLER CONFIRMED THAT THE PS POWER CONNECTOR DID NOT WORK PROPERLY SYSTEM TESTING CONFIRMED A CONTROLLER FAULT ALARM WHILE PS WAS ACTIVE REVIEW OF THE LOG FILES ALSO REVEALED CONTROLLER FAULT ALARM WHILE PS WAS ACTIVE THE CONTROLLER WAS IN USE WHEN THE REPORTED EVENT OCCURRED THIS MALFUNCTION IS MOST LIKELY DUE TO A LEAKY DIODE D ON THE AUTOMATIC POWER SWITCH CIRCUIT OF THE CONTROLLER THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA | Mechanical Problem Identified | Medical Evaluation result code
6931905 | 89945744 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT NO NOISE WAS ABLE TO BE CREATED WITH ADDITIONAL TROUBLESHOOTING THE LEAD CONFIGURATION WAS REPROGRAMMED TO BIPOLAR SENSING AND UNIPOLAR PACING NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME IF ADDITIONAL INFORMATION BECOMES AVAILABLE THIS REPORT WILL BE UPDATED | Noise, Audible | Device problem code
6931905 | 89945744 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT NO NOISE WAS ABLE TO BE CREATED WITH ADDITIONAL TROUBLESHOOTING THE LEAD CONFIGURATION WAS REPROGRAMMED TO BIPOLAR SENSING AND UNIPOLAR PACING NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME IF ADDITIONAL INFORMATION BECOMES AVAILABLE THIS REPORT WILL BE UPDATED | Power Conditioning Problem | Device Problem Code
6931905 | 89945744 | N | ADDITIONAL INFORMATION RECEIVED FROM THE LOCAL FIELD REPRESENTATIVE INDICATED THAT NO NOISE WAS ABLE TO BE CREATED WITH ADDITIONAL TROUBLESHOOTING THE LEAD CONFIGURATION WAS REPROGRAMMED TO BIPOLAR SENSING AND UNIPOLAR PACING NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME IF ADDITIONAL INFORMATION BECOMES AVAILABLE THIS REPORT WILL BE UPDATED | Power Source Problem Identified | Medical Evaluation result code
6928004 | 89946620 | N | EVALUATION SUMMARY THE STENT WAS NOT POSITIONED ON THE BALLOON BETWEEN THE MARKER BANDS STENT HAD MOVE PROXIMAL OVER THE MARKER BAND A KINK ALONG THE DISTAL SHAFT WAS APPARENT CM PROXIMAL TO THE DISTAL TIP DEFORMATION WAS EVIDENT TO THE DISTAL TIP WITH TIP BEING JAGGED AND UNEVEN THE INNER LUMEN PATENCY WAS NOT VERIFIED WITH A INCH MANDREL DUE TO A KINK ALONG THE DISTAL SHAFT IMAGE REVIEW TWO PHOTOGRAPHS WERE PROVIDED BY THE ACCOUNT ONE PHOTO CAPTURED THE PROXIMAL PILLOW AND STENT WRAPS THERE WAS NO DEFORMATION EVIDENT TO THE STENT STRUTS THE OTHER PHOTO CAPTURED THE STENT THERE WAS NO DEFORMATION EVIDENT TO THE STENT WRAPS AND STRUTS | Break | Device problem code
6928004 | 89946620 | N | EVALUATION SUMMARY THE STENT WAS NOT POSITIONED ON THE BALLOON BETWEEN THE MARKER BANDS STENT HAD MOVE PROXIMAL OVER THE MARKER BAND A KINK ALONG THE DISTAL SHAFT WAS APPARENT CM PROXIMAL TO THE DISTAL TIP DEFORMATION WAS EVIDENT TO THE DISTAL TIP WITH TIP BEING JAGGED AND UNEVEN THE INNER LUMEN PATENCY WAS NOT VERIFIED WITH A INCH MANDREL DUE TO A KINK ALONG THE DISTAL SHAFT IMAGE REVIEW TWO PHOTOGRAPHS WERE PROVIDED BY THE ACCOUNT ONE PHOTO CAPTURED THE PROXIMAL PILLOW AND STENT WRAPS THERE WAS NO DEFORMATION EVIDENT TO THE STENT STRUTS THE OTHER PHOTO CAPTURED THE STENT THERE WAS NO DEFORMATION EVIDENT TO THE STENT WRAPS AND STRUTS | Material Deformation | Device problem code
6928004 | 89946620 | N | EVALUATION SUMMARY THE STENT WAS NOT POSITIONED ON THE BALLOON BETWEEN THE MARKER BANDS STENT HAD MOVE PROXIMAL OVER THE MARKER BAND A KINK ALONG THE DISTAL SHAFT WAS APPARENT CM PROXIMAL TO THE DISTAL TIP DEFORMATION WAS EVIDENT TO THE DISTAL TIP WITH TIP BEING JAGGED AND UNEVEN THE INNER LUMEN PATENCY WAS NOT VERIFIED WITH A INCH MANDREL DUE TO A KINK ALONG THE DISTAL SHAFT IMAGE REVIEW TWO PHOTOGRAPHS WERE PROVIDED BY THE ACCOUNT ONE PHOTO CAPTURED THE PROXIMAL PILLOW AND STENT WRAPS THERE WAS NO DEFORMATION EVIDENT TO THE STENT STRUTS THE OTHER PHOTO CAPTURED THE STENT THERE WAS NO DEFORMATION EVIDENT TO THE STENT WRAPS AND STRUTS | Stress Problem Identified | Medical Evaluation result code
6937277 | 89948598 | N | EVALUATION SUMMARY POST MARKET VIGILANCE PMV LED AN EVALUATION OF THE DEVICE INITIAL VISUAL INSPECTION OF THE INSTRUMENT NOTED NO ABNORMALITIES FUNCTIONALLY THE INSTRUMENT WAS SINGLE USE LOADING UNIT DURING TESTING OF THE INSTRUMENTS A SKIP WAS NOTED IN THE FIRING STROKE AFTER CLOSURE OF THE LOADING UNIT JAWS AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK THIS EXAMINATION NOTED SHEARED TEETH ON THE FIRING RACK RECORDS FROM EACH MANUFACTURING LOT ARE THOROUGHLY REVIEWED TO ENSURE THAT PRODUCTS ARE RELEASED MEETING ALL QUALITY RELEASE SPECIFICATIONS AT THE TIME OF MANUFACTURE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR IN ANY OF THE FOLLOWING CIRCUMSTANCES FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS IN EITHER OF THESE CIRCUMSTANCES IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT IN ADDITION STAPLES MAY NOT FORM PROPERLY AND TISSUE MAY NOT BE FULLY TRANSECTED SHOULD NEW INFORMATION BECOME AVAILABLE THE FILE WILL BE REOPENED AND THE INVESTIGATION SUMMARY AMENDED AS APPROPRIATE | Break | Device problem code
6937277 | 89948598 | N | EVALUATION SUMMARY POST MARKET VIGILANCE PMV LED AN EVALUATION OF THE DEVICE INITIAL VISUAL INSPECTION OF THE INSTRUMENT NOTED NO ABNORMALITIES FUNCTIONALLY THE INSTRUMENT WAS SINGLE USE LOADING UNIT DURING TESTING OF THE INSTRUMENTS A SKIP WAS NOTED IN THE FIRING STROKE AFTER CLOSURE OF THE LOADING UNIT JAWS AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK THIS EXAMINATION NOTED SHEARED TEETH ON THE FIRING RACK RECORDS FROM EACH MANUFACTURING LOT ARE THOROUGHLY REVIEWED TO ENSURE THAT PRODUCTS ARE RELEASED MEETING ALL QUALITY RELEASE SPECIFICATIONS AT THE TIME OF MANUFACTURE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR IN ANY OF THE FOLLOWING CIRCUMSTANCES FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS IN EITHER OF THESE CIRCUMSTANCES IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT IN ADDITION STAPLES MAY NOT FORM PROPERLY AND TISSUE MAY NOT BE FULLY TRANSECTED SHOULD NEW INFORMATION BECOME AVAILABLE THE FILE WILL BE REOPENED AND THE INVESTIGATION SUMMARY AMENDED AS APPROPRIATE | Retraction Problem | Device problem code
6937277 | 89948598 | N | EVALUATION SUMMARY POST MARKET VIGILANCE PMV LED AN EVALUATION OF THE DEVICE INITIAL VISUAL INSPECTION OF THE INSTRUMENT NOTED NO ABNORMALITIES FUNCTIONALLY THE INSTRUMENT WAS SINGLE USE LOADING UNIT DURING TESTING OF THE INSTRUMENTS A SKIP WAS NOTED IN THE FIRING STROKE AFTER CLOSURE OF THE LOADING UNIT JAWS AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK THIS EXAMINATION NOTED SHEARED TEETH ON THE FIRING RACK RECORDS FROM EACH MANUFACTURING LOT ARE THOROUGHLY REVIEWED TO ENSURE THAT PRODUCTS ARE RELEASED MEETING ALL QUALITY RELEASE SPECIFICATIONS AT THE TIME OF MANUFACTURE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR IN ANY OF THE FOLLOWING CIRCUMSTANCES FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS IN EITHER OF THESE CIRCUMSTANCES IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT IN ADDITION STAPLES MAY NOT FORM PROPERLY AND TISSUE MAY NOT BE FULLY TRANSECTED SHOULD NEW INFORMATION BECOME AVAILABLE THE FILE WILL BE REOPENED AND THE INVESTIGATION SUMMARY AMENDED AS APPROPRIATE | Non-Functional Defect | Medical Evaluation result code
6928837 | 89948605 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6928837 | 89948605 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6941433 | 89952000 | D | COMPLAINANT ALLEGED THAT THE DEVICES PACER RATE WAS NOT ADJUSTABLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Pacing Problem | Device Problem Code
6941433 | 89952000 | D | COMPLAINANT ALLEGED THAT THE DEVICES PACER RATE WAS NOT ADJUSTABLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6930788 | 89952324 | N | VISUAL RECEIPT INSPECTION RECEIVED ONE USED STAND ALONE CL CHEMOLOCK BAG SPIKE LOT IS UNKNOWN ONE NEW CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT ONE USED EMPTY BAG HOSPIRA ML NACL ONE USED DRY SPIKE CHEMOLOCK CONNECTOR ONE USED CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT UNKNOWN ONE USED IV SET MODEL UNKNOWN CONNECTED TO ONE USED SPINNING SPIROS ONE USED EMPTY BAXTER ML BAG THE ONE USED CL HAD NO VISIBLE ANOMALIES PRESENT THE ONE USED CL ATTACHED TO THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT IT WAS NOTICED THAT THE ML BAG WAS NOT SECURELY ATTACHED TO THE CL DEVICE WHEN RECEIVED THE PORT WAS NOT ATTACHED ALL THE WAY ONTO THE SHOULDER OF THE SPIKE FLUSH WITH THE BASE OF THE SPIKE THE ONE PACKAGED CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT FUNCTIONAL TESTING THE CL WAS REMOVED FROM THE ML BAG THE BAG WAS FILLED WITH WATER THE RETURNED CL DEVICE WAS THEN ATTACHED TO THE BAG BY ATTACHING THE BAG PORT ALL THE WAY FLUSH OVER THE SHOULDER OF THE CL SPIKE THE PACKAGED CL WITH THE ROLLER CLAMP IN THE OFF POSITION WAS THEN ATTACHED TO THE CL THE ROLLER CLAMP WAS THEN OPENED AND WATER FLOWED THROUGH THE CLCL UNTIL EMPTY NO LEAKS OR BAG SEPARATION OCCURRED FINAL ANALYSIS SUMMARY THE REPORTED PRODUCT PROBLEM FOR CL SEPARATINGLEAKING FROM THE ML BAG PORT COULD NOT BE REPLICATEDCONFIRMED THERE WAS NO SEPARATIONLEAKAGE FOUND BETWEEN THE SPIKEBAG PORT CONNECTION AS LONG AS THE PORT WAS SECURELY ATTACHED OVER THE SHOULDER FLUSH WITH THE BASE OF THE SPIKE OF THE CL ICU MEDICAL WILL MONITOR AND TREND | Connection Problem | Device problem code
6930788 | 89952324 | N | VISUAL RECEIPT INSPECTION RECEIVED ONE USED STAND ALONE CL CHEMOLOCK BAG SPIKE LOT IS UNKNOWN ONE NEW CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT ONE USED EMPTY BAG HOSPIRA ML NACL ONE USED DRY SPIKE CHEMOLOCK CONNECTOR ONE USED CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT UNKNOWN ONE USED IV SET MODEL UNKNOWN CONNECTED TO ONE USED SPINNING SPIROS ONE USED EMPTY BAXTER ML BAG THE ONE USED CL HAD NO VISIBLE ANOMALIES PRESENT THE ONE USED CL ATTACHED TO THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT IT WAS NOTICED THAT THE ML BAG WAS NOT SECURELY ATTACHED TO THE CL DEVICE WHEN RECEIVED THE PORT WAS NOT ATTACHED ALL THE WAY ONTO THE SHOULDER OF THE SPIKE FLUSH WITH THE BASE OF THE SPIKE THE ONE PACKAGED CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT FUNCTIONAL TESTING THE CL WAS REMOVED FROM THE ML BAG THE BAG WAS FILLED WITH WATER THE RETURNED CL DEVICE WAS THEN ATTACHED TO THE BAG BY ATTACHING THE BAG PORT ALL THE WAY FLUSH OVER THE SHOULDER OF THE CL SPIKE THE PACKAGED CL WITH THE ROLLER CLAMP IN THE OFF POSITION WAS THEN ATTACHED TO THE CL THE ROLLER CLAMP WAS THEN OPENED AND WATER FLOWED THROUGH THE CLCL UNTIL EMPTY NO LEAKS OR BAG SEPARATION OCCURRED FINAL ANALYSIS SUMMARY THE REPORTED PRODUCT PROBLEM FOR CL SEPARATINGLEAKING FROM THE ML BAG PORT COULD NOT BE REPLICATEDCONFIRMED THERE WAS NO SEPARATIONLEAKAGE FOUND BETWEEN THE SPIKEBAG PORT CONNECTION AS LONG AS THE PORT WAS SECURELY ATTACHED OVER THE SHOULDER FLUSH WITH THE BASE OF THE SPIKE OF THE CL ICU MEDICAL WILL MONITOR AND TREND | Material Separation | Device problem code
6930788 | 89952324 | N | VISUAL RECEIPT INSPECTION RECEIVED ONE USED STAND ALONE CL CHEMOLOCK BAG SPIKE LOT IS UNKNOWN ONE NEW CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT ONE USED EMPTY BAG HOSPIRA ML NACL ONE USED DRY SPIKE CHEMOLOCK CONNECTOR ONE USED CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT UNKNOWN ONE USED IV SET MODEL UNKNOWN CONNECTED TO ONE USED SPINNING SPIROS ONE USED EMPTY BAXTER ML BAG THE ONE USED CL HAD NO VISIBLE ANOMALIES PRESENT THE ONE USED CL ATTACHED TO THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT IT WAS NOTICED THAT THE ML BAG WAS NOT SECURELY ATTACHED TO THE CL DEVICE WHEN RECEIVED THE PORT WAS NOT ATTACHED ALL THE WAY ONTO THE SHOULDER OF THE SPIKE FLUSH WITH THE BASE OF THE SPIKE THE ONE PACKAGED CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT FUNCTIONAL TESTING THE CL WAS REMOVED FROM THE ML BAG THE BAG WAS FILLED WITH WATER THE RETURNED CL DEVICE WAS THEN ATTACHED TO THE BAG BY ATTACHING THE BAG PORT ALL THE WAY FLUSH OVER THE SHOULDER OF THE CL SPIKE THE PACKAGED CL WITH THE ROLLER CLAMP IN THE OFF POSITION WAS THEN ATTACHED TO THE CL THE ROLLER CLAMP WAS THEN OPENED AND WATER FLOWED THROUGH THE CLCL UNTIL EMPTY NO LEAKS OR BAG SEPARATION OCCURRED FINAL ANALYSIS SUMMARY THE REPORTED PRODUCT PROBLEM FOR CL SEPARATINGLEAKING FROM THE ML BAG PORT COULD NOT BE REPLICATEDCONFIRMED THERE WAS NO SEPARATIONLEAKAGE FOUND BETWEEN THE SPIKEBAG PORT CONNECTION AS LONG AS THE PORT WAS SECURELY ATTACHED OVER THE SHOULDER FLUSH WITH THE BASE OF THE SPIKE OF THE CL ICU MEDICAL WILL MONITOR AND TREND | Leak / Splash | Device problem code
6930788 | 89952324 | N | VISUAL RECEIPT INSPECTION RECEIVED ONE USED STAND ALONE CL CHEMOLOCK BAG SPIKE LOT IS UNKNOWN ONE NEW CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT ONE USED EMPTY BAG HOSPIRA ML NACL ONE USED DRY SPIKE CHEMOLOCK CONNECTOR ONE USED CL CM APPX ML ADMIN SET W DROP INLINE CHAMBER CHEMOLOCK SPIROS BAG HANGER DROPIN RED CAP LOT UNKNOWN ONE USED IV SET MODEL UNKNOWN CONNECTED TO ONE USED SPINNING SPIROS ONE USED EMPTY BAXTER ML BAG THE ONE USED CL HAD NO VISIBLE ANOMALIES PRESENT THE ONE USED CL ATTACHED TO THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT THE CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT IT WAS NOTICED THAT THE ML BAG WAS NOT SECURELY ATTACHED TO THE CL DEVICE WHEN RECEIVED THE PORT WAS NOT ATTACHED ALL THE WAY ONTO THE SHOULDER OF THE SPIKE FLUSH WITH THE BASE OF THE SPIKE THE ONE PACKAGED CL DEVICE HAD NO VISIBLE ANOMALIES PRESENT FUNCTIONAL TESTING THE CL WAS REMOVED FROM THE ML BAG THE BAG WAS FILLED WITH WATER THE RETURNED CL DEVICE WAS THEN ATTACHED TO THE BAG BY ATTACHING THE BAG PORT ALL THE WAY FLUSH OVER THE SHOULDER OF THE CL SPIKE THE PACKAGED CL WITH THE ROLLER CLAMP IN THE OFF POSITION WAS THEN ATTACHED TO THE CL THE ROLLER CLAMP WAS THEN OPENED AND WATER FLOWED THROUGH THE CLCL UNTIL EMPTY NO LEAKS OR BAG SEPARATION OCCURRED FINAL ANALYSIS SUMMARY THE REPORTED PRODUCT PROBLEM FOR CL SEPARATINGLEAKING FROM THE ML BAG PORT COULD NOT BE REPLICATEDCONFIRMED THERE WAS NO SEPARATIONLEAKAGE FOUND BETWEEN THE SPIKEBAG PORT CONNECTION AS LONG AS THE PORT WAS SECURELY ATTACHED OVER THE SHOULDER FLUSH WITH THE BASE OF THE SPIKE OF THE CL ICU MEDICAL WILL MONITOR AND TREND | Leakage/Seal | Medical Evaluation result code
6932423 | 89955485 | N | ZIMMER BIOMET COMPLAINT B COMPLAINT SAMPLE WAS EVALUATED AND THE REPORTED EVENT WAS CONFIRMED WE CAN CONCLUDE FROM THE INVESTIGATION THAT THE IMPLANT WAS CORRECT TO PACKAGING PROTOCOL AND CONFORMING THE INNER PACKAGE BEING TEAR IS MOSTLY LIKELY DUE TO INAPPROPRIATE HANDLING AND SHIPPING OPERATIONSREVIEW OF DEVICE HISTORY RECORDS FOUND THESE UNITS WERE RELEASED TO DISTRIBUTION WITH NO DEVIATIONS OR ANOMALIES IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | Tear, Rip or Hole in Device Packaging | Device Problem Code
6932423 | 89955485 | N | ZIMMER BIOMET COMPLAINT B COMPLAINT SAMPLE WAS EVALUATED AND THE REPORTED EVENT WAS CONFIRMED WE CAN CONCLUDE FROM THE INVESTIGATION THAT THE IMPLANT WAS CORRECT TO PACKAGING PROTOCOL AND CONFORMING THE INNER PACKAGE BEING TEAR IS MOSTLY LIKELY DUE TO INAPPROPRIATE HANDLING AND SHIPPING OPERATIONSREVIEW OF DEVICE HISTORY RECORDS FOUND THESE UNITS WERE RELEASED TO DISTRIBUTION WITH NO DEVIATIONS OR ANOMALIES IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION A SUPPLEMENTAL WILL BE FILED ACCORDINGLY ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS | Packaging Compromised | Medical Evaluation result code
6929680 | 89955683 | N | B THE DEVICE WAS RETURNED FOR EVALUATION AND VISUALLY AND FUNCTIONALLY TESTED THE COMPLAINT WAS CONFIRMED DISSECTION OF THE HANDLE REVEALED THAT THE CABLES HAD BECOME CAUGHT BETWEEN THE HANDLE HALVES DURING ASSEMBLY | Mechanical Jam | Device problem code
6929680 | 89955683 | N | B THE DEVICE WAS RETURNED FOR EVALUATION AND VISUALLY AND FUNCTIONALLY TESTED THE COMPLAINT WAS CONFIRMED DISSECTION OF THE HANDLE REVEALED THAT THE CABLES HAD BECOME CAUGHT BETWEEN THE HANDLE HALVES DURING ASSEMBLY | Mechanical Problem Identified | Medical Evaluation result code
6968856 | 89958924 | D | IT WAS REPORTED THE PATIENT EXPERIENCED ONE INCIDENT OF A BURNING SENSATION AT IPG SITE CLINICAL SPECIALIST COULD NOT REPLICATE THE EVENT WHEN TROUBLESHOOTING PATIENT UNDERWENT SCS SYSTEM EXPLANT EXACT EXPLANT DATE IS UNKNOWN ALSO SEE MFR REPORT | Insufficient Information | Device problem code
6968856 | 89958924 | D | IT WAS REPORTED THE PATIENT EXPERIENCED ONE INCIDENT OF A BURNING SENSATION AT IPG SITE CLINICAL SPECIALIST COULD NOT REPLICATE THE EVENT WHEN TROUBLESHOOTING PATIENT UNDERWENT SCS SYSTEM EXPLANT EXACT EXPLANT DATE IS UNKNOWN ALSO SEE MFR REPORT | Burning Sensation | Patient problem code
6969003 | 89969427 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Material Discolored | Device problem code
6969003 | 89969427 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Stress Problem Identified | Medical Evaluation result code
6969522 | 89980676 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX MECHANISM WAS RETRACTED ELECTRICAL TESTING AND XRAY ANALYSIS DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS DETAILED ANALYSIS CONFIRMED THAT LEAD MEETS SPECIFICATIONS | Retraction Problem | Device problem code
6969522 | 89980676 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX MECHANISM WAS RETRACTED ELECTRICAL TESTING AND XRAY ANALYSIS DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS DETAILED ANALYSIS CONFIRMED THAT LEAD MEETS SPECIFICATIONS | No Device Problem Found | Medical Evaluation result code
6969342 | 89984627 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY TESTING WAS UNABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS OF PERFORATION AND PACING THRESHOLDS THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | Material Perforation | Device problem code
6969342 | 89984627 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY TESTING WAS UNABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS OF PERFORATION AND PACING THRESHOLDS THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | Retraction Problem | Device problem code
6969342 | 89984627 | N | THIS INGEVITY LEAD WAS RETURNED TO BOSTON SCIENTIFICS POST MARKET QUALITY ASSURANCE LABORATORY WITH THE HELIX MECHANISM IN A RETRACTED POSITION SUBSEQUENT ELECTRICAL AND MECHANICAL TESTING DID NOT IDENTIFY ANY PRODUCT ABNORMALITIES THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS LABORATORY TESTING WAS UNABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS OF PERFORATION AND PACING THRESHOLDS THE LEAD TIP AND HELIX APPEARS INTACT AND UNDAMAGED | No Device Problem Found | Medical Evaluation result code
6964669 | 89989690 | N | THE REPORTED INCIDENT WAS REPRODUCED DURING THE HARDWARE EVALUATION THE CAUSE OF THE REPORTED COMMUNICATION ISSUE WAS ISOLATED TO AN INTERMITTENT ABNORMAL FUNCTION OF THE SINGLE BOARD COMPUTER MODULE THE ISSUE WAS RESOLVED AFTER REPLACING THE SINGLE BOARD COMPUTER MODULE WITH A KNOWN GOOD UNIT AND THE AMPLIFIER PASSED THE EXTENDED COMMUNICATION TEST WITHOUT FURTHER ISSUE | Communication or Transmission Problem | Device problem code
6964669 | 89989690 | N | THE REPORTED INCIDENT WAS REPRODUCED DURING THE HARDWARE EVALUATION THE CAUSE OF THE REPORTED COMMUNICATION ISSUE WAS ISOLATED TO AN INTERMITTENT ABNORMAL FUNCTION OF THE SINGLE BOARD COMPUTER MODULE THE ISSUE WAS RESOLVED AFTER REPLACING THE SINGLE BOARD COMPUTER MODULE WITH A KNOWN GOOD UNIT AND THE AMPLIFIER PASSED THE EXTENDED COMMUNICATION TEST WITHOUT FURTHER ISSUE | Communications Problem Identified | Medical Evaluation result code
6969643 | 89992000 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF INDURATED ABSCESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Migration or Expulsion of Device | Device problem code
6969643 | 89992000 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF INDURATED ABSCESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Leak / Splash | Device problem code
6969643 | 89992000 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF INDURATED ABSCESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Material Discolored | Device problem code
6969643 | 89992000 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF INDURATED ABSCESS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING ADVERSE EVENTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA ISRS THAT LASTED BEYOND THE DAY DIARIES WERE CONSIDERED ADVERSE EVENTS ADVERSE EVENTS WERE ALSO REPORTED BY THE TREATING INVESTIGATOR AT FOLLOWUP VISITS AFTER INITIAL TREATMENT OR TOUCHUP TREATMENT IF PERFORMED A TOTAL OF TREATMENTRELATED ADVERSE EVENTS WERE REPORTED IN OF SUBJECTS IN GENERAL AES WERE MILD OR MODERATE RESOLVED WITHOUT SEQUELAE AND REQUIRED NO ACTION AES TYPICALLY RESOLVED WITHIN MONTHS TREATMENTRELATED ADVERSE EVENTS THAT OCCURRED IN OF SUBJECTS WERE INJECTION SITE MASS INDURATION DISCOLORATION PAIN BRUISING SWELLING ERYTHEMA AND REACTION SIMILAR AES WERE REPORTED AFTER REPEAT TREATMENT | Stress Problem Identified | Medical Evaluation result code
6918495 | 90036130 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE WITH ROUGH BUTTONS AND ADHERENCE IN THE WELLS | Connection Problem | Device problem code
6918495 | 90036130 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE WITH ROUGH BUTTONS AND ADHERENCE IN THE WELLS | Interoperability Problem Identified | Medical Evaluation result code
6934839 | 90041542 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS EXTENDED AND DRIED BLOODTISSUE WAS PRESENT AROUND THE HELIX ADDITIONALLY STRETCHED CONDUCTOR COILS WERE OBSERVED IN THE NECK AREA RESISTANCE TESTING FOUND THE LEAD WAS NOT ELECTRICALLY CONTINUOUS DETAILED ANALYSIS CONFIRMED THAT THE INNER CONDUCTOR COIL HAD A BREAK AT THE TIP END NEAR WHERE THE STRECHED COILS WERE OBSERVED | Break | Device problem code
6934839 | 90041542 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS EXTENDED AND DRIED BLOODTISSUE WAS PRESENT AROUND THE HELIX ADDITIONALLY STRETCHED CONDUCTOR COILS WERE OBSERVED IN THE NECK AREA RESISTANCE TESTING FOUND THE LEAD WAS NOT ELECTRICALLY CONTINUOUS DETAILED ANALYSIS CONFIRMED THAT THE INNER CONDUCTOR COIL HAD A BREAK AT THE TIP END NEAR WHERE THE STRECHED COILS WERE OBSERVED | Stretched | Device problem code
6934839 | 90041542 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION NOTED THE HELIX WAS EXTENDED AND DRIED BLOODTISSUE WAS PRESENT AROUND THE HELIX ADDITIONALLY STRETCHED CONDUCTOR COILS WERE OBSERVED IN THE NECK AREA RESISTANCE TESTING FOUND THE LEAD WAS NOT ELECTRICALLY CONTINUOUS DETAILED ANALYSIS CONFIRMED THAT THE INNER CONDUCTOR COIL HAD A BREAK AT THE TIP END NEAR WHERE THE STRECHED COILS WERE OBSERVED | Stress Problem Identified | Medical Evaluation result code
6936104 | 90046079 | N | B RESULTS OF INVESTIGATION THE VYAIRE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE TO EVALUATE THE SUSPECT DEVICE THE FSR DISCOVERED THAT THE CALIBRATIONS OF THE INSPIRATORY PRESSURE TRANSDUCER WERE CONSIDERABLY OUT OF SPECIFICATIONS THE FSR RECALIBRATED THE DEVICE AND RESOLVED THE CUSTOMERS REPORTED ISSUE RETURNING THE DEVICE TO THE CUSTOMER OPERATING TO MANUFACTURER SPECIFICATIONS | Calibration Problem | Device problem code
6936104 | 90046079 | N | B RESULTS OF INVESTIGATION THE VYAIRE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE TO EVALUATE THE SUSPECT DEVICE THE FSR DISCOVERED THAT THE CALIBRATIONS OF THE INSPIRATORY PRESSURE TRANSDUCER WERE CONSIDERABLY OUT OF SPECIFICATIONS THE FSR RECALIBRATED THE DEVICE AND RESOLVED THE CUSTOMERS REPORTED ISSUE RETURNING THE DEVICE TO THE CUSTOMER OPERATING TO MANUFACTURER SPECIFICATIONS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6936273 | 90047158 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THIS LEAD WAS PERFORMED THIS LEAD FAILED TO PASS THE WIRE INSERTION TEST DUE TO THE PRESENCE OF FOREIGN MATERIAL INSIDE THE LUMEN THE LEAD WAS CUT AND VISUAL INSPECTION SHOWED THAT THE BLOCKAGE WAS CAUSED BY AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION | Partial Blockage | Device problem code
6936273 | 90047158 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THIS LEAD WAS PERFORMED THIS LEAD FAILED TO PASS THE WIRE INSERTION TEST DUE TO THE PRESENCE OF FOREIGN MATERIAL INSIDE THE LUMEN THE LEAD WAS CUT AND VISUAL INSPECTION SHOWED THAT THE BLOCKAGE WAS CAUSED BY AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION | Device Contamination with Chemical or Other Material | Device problem code
6936273 | 90047158 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THIS LEAD WAS PERFORMED THIS LEAD FAILED TO PASS THE WIRE INSERTION TEST DUE TO THE PRESENCE OF FOREIGN MATERIAL INSIDE THE LUMEN THE LEAD WAS CUT AND VISUAL INSPECTION SHOWED THAT THE BLOCKAGE WAS CAUSED BY AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION | No Device Problem Found | Medical Evaluation result code
6936745 | 90050561 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6936745 | 90050561 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6937410 | 90056732 | N | RESULTS THE SECOND SMART COIL EVALUATED HAD MINOR OFFSET COIL WINDS CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | Failure to Advance | Device problem code
6937410 | 90056732 | N | RESULTS THE SECOND SMART COIL EVALUATED HAD MINOR OFFSET COIL WINDS CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | Failure to Align | Device problem code
6937410 | 90056732 | N | RESULTS THE SECOND SMART COIL EVALUATED HAD MINOR OFFSET COIL WINDS CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | No Device Problem Found | Medical Evaluation result code
6937425 | 90058343 | N | RESULTS THE FIRST PENUMBRA SMART COIL SMART COIL EVALUATED HAD OFFSET COIL WINDS ON THE PROXIMAL END OF ITS EMBOLIZATION COIL CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | Failure to Advance | Device problem code
6937425 | 90058343 | N | RESULTS THE FIRST PENUMBRA SMART COIL SMART COIL EVALUATED HAD OFFSET COIL WINDS ON THE PROXIMAL END OF ITS EMBOLIZATION COIL CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | Failure to Align | Device problem code
6937425 | 90058343 | N | RESULTS THE FIRST PENUMBRA SMART COIL SMART COIL EVALUATED HAD OFFSET COIL WINDS ON THE PROXIMAL END OF ITS EMBOLIZATION COIL CONCLUSIONS EVALUATION OF THE RETURNED SMART COILS REVEALED OFFSET COIL WINDS ON THEIR EMBOLIZATION COILS IF THE INTRODUCER SHEATH IS NOT ALIGNED PROPERLY WITHIN THE HUB PRIOR TO INSERTION RESISTANCE MAY BE EXPERIENCED AND OFFSET COIL WINDS MAY RESULT THESE OFFSET COIL WINDS CONTRIBUTED TO INABILITY TO ADVANCE INTO THE DEMONSTRATION CATHETER DURING FUNCTIONAL TESTING AND LIKELY CONTRIBUTED TO THE RESISTANCE EXPERIENCED BY THE PHYSICIAN PENUMBRA COILS ARE INSPECTED DURING INPROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES DESIGN OR QUALITY CONCERNS | No Device Problem Found | Medical Evaluation result code
6966980 | 90061151 | N | THE CUSTOMER CALLED STATING THE ECG WAVE DIDNT SHOW ALL OF THE DATA CORRECTLY BPM ACTUAL BUT ONLY BPM WAS SHOWN ON THE STRIP AFTER A DRUG WAS ADMINISTERED NO HARM TO PATIENT WAS REPORTED B CONTACTED THE CUSTOMER ON B CUSTOMER STATES THE ECG TRACING AND HR DO NOT CORRELATE UPON REVIEWING THE DATA B FELT THAT ONE TRACING WAS DOUBLE COUNTING DUE TO THE QRS AND T WAVE BEING THE SAME HEIGHT THE SECOND WAVE THAT SHOWED AN ECG TRACING OF SR IN THE S BUT THE HR NUMERIC OF WAS UNEXPLAINABLE B HAD THE CUSTOMER CHANGE SOME CONFIGURATIONS AND SHE WAS SUPPOSED TO FOLLOW UP WITH HER THE FOLLOWING DAY CUSTOMER NEVER CALLED AND NKTS LEFT A MESSAGE FOR THE CUSTOMER DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Device Displays Incorrect Message | Device problem code
6966980 | 90061151 | N | THE CUSTOMER CALLED STATING THE ECG WAVE DIDNT SHOW ALL OF THE DATA CORRECTLY BPM ACTUAL BUT ONLY BPM WAS SHOWN ON THE STRIP AFTER A DRUG WAS ADMINISTERED NO HARM TO PATIENT WAS REPORTED B CONTACTED THE CUSTOMER ON B CUSTOMER STATES THE ECG TRACING AND HR DO NOT CORRELATE UPON REVIEWING THE DATA B FELT THAT ONE TRACING WAS DOUBLE COUNTING DUE TO THE QRS AND T WAVE BEING THE SAME HEIGHT THE SECOND WAVE THAT SHOWED AN ECG TRACING OF SR IN THE S BUT THE HR NUMERIC OF WAS UNEXPLAINABLE B HAD THE CUSTOMER CHANGE SOME CONFIGURATIONS AND SHE WAS SUPPOSED TO FOLLOW UP WITH HER THE FOLLOWING DAY CUSTOMER NEVER CALLED AND NKTS LEFT A MESSAGE FOR THE CUSTOMER DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6926506 | 90066171 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND THE INSTRUMENT TO BE OPERATING AS EXPECTED THE FSE PROACTIVELY INSTALLED A NEW SAMPLE PROBE | Connection Problem | Device problem code
6926506 | 90066171 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND FOUND THE INSTRUMENT TO BE OPERATING AS EXPECTED THE FSE PROACTIVELY INSTALLED A NEW SAMPLE PROBE | Interoperability Problem Identified | Medical Evaluation result code
6972702 | 90069745 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE WAS THE TRANSECTION TAKEN IN ONE OR MULTIPLE FIRINGS COULD YOU PLEASE CONFIRM THE PRODUCT CODE WHAT IS THE CURRENT PATIENT STATUS | Fire | Device problem code
6972702 | 90069745 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE WAS THE TRANSECTION TAKEN IN ONE OR MULTIPLE FIRINGS COULD YOU PLEASE CONFIRM THE PRODUCT CODE WHAT IS THE CURRENT PATIENT STATUS | Misfire | Device problem code
6972702 | 90069745 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE WAS THE TRANSECTION TAKEN IN ONE OR MULTIPLE FIRINGS COULD YOU PLEASE CONFIRM THE PRODUCT CODE WHAT IS THE CURRENT PATIENT STATUS | Difficult to Open or Close | Device problem code
6972702 | 90069745 | N | B THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE WAS THE TRANSECTION TAKEN IN ONE OR MULTIPLE FIRINGS COULD YOU PLEASE CONFIRM THE PRODUCT CODE WHAT IS THE CURRENT PATIENT STATUS | Operational Problem Identified | Medical Evaluation result code
6972809 | 90071084 | N | INVESTIGATION DURING CUSTOMER FOLLOWUP THE CUSTOMER STATED THAT THEY HAD A DISCUSSION WITHTHE RN WHO PERFORMED THE PROCEDURE THE RN STATED THAT THE SALINE ROLLER CLAMPS WERE CLOSEDAND THEY RECEIVED NO ALARMS PRIOR TO THE START OF THE PROCEDURE THE RN ALSO STATED THAT THEY DIDNOT BOLUS THROUGH THE MACHINE AND THE MACHINE WAS TURNED OFF PRIOR TO OPENING THE CLAMPSTHE RUN DATA FILE RDF WAS ANALYZED FOR THIS EVENT SIGNALS IN THE RDF SHOWED THAT THE OPTIASYSTEM OPERATED AS INTENDEDINVESTIGATION IS IN PROCESS A FOLLOWUP REPORT WILL BE PROVIDED | Device Alarm System | Device problem code
6938250 | 90084484 | N | ONE CONTROLLER AND TWO BATTERIES WERE RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE RETURNED CONTROLLER REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED THREE CRITICAL BATTERY ALARMS INVOLVING BATTERIES B THESE ALARMS WERE MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES BATTERIES B WERE NOT TESTED BASED ON THE RESULTS OF THE INTERNAL INVESTIGATION NO ADDITIONAL INVESTIGATION OF THIS EVENT IS REQUIRED AT THIS TIME THE MOST LIKELY CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA BATTERY B MANUFACTURING DATE EXPIRATION DATE FDA CODES B BATTERY B MANUFACTURING DATE EXPIRATION DATE FDA CODES B | Communication or Transmission Problem | Device problem code
6938250 | 90084484 | N | ONE CONTROLLER AND TWO BATTERIES WERE RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT ANALYSIS OF THE RETURNED CONTROLLER REVEALED THAT THE DEVICE MET SPECIFICATIONS THE DEVICE PASSED VISUAL EXAMINATION AND FUNCTIONAL TESTING THE REPORTED EVENT WAS CONFIRMED VIA REVIEW OF THE CONTROLLER LOG FILES WHICH REVEALED THREE CRITICAL BATTERY ALARMS INVOLVING BATTERIES B THESE ALARMS WERE MOST LIKELY CAUSED BY A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND BATTERIES HEARTWARE HAS OPENED AN INTERNAL INVESTIGATION TO EVALUATE THESE TYPES OF ISSUES BATTERIES B WERE NOT TESTED BASED ON THE RESULTS OF THE INTERNAL INVESTIGATION NO ADDITIONAL INVESTIGATION OF THIS EVENT IS REQUIRED AT THIS TIME THE MOST LIKELY CAUSE OF THE REPORTED EVENT MAY BE ATTRIBUTED TO A COMMUNICATION ERROR BETWEEN THE CONTROLLER AND THE BATTERIES THE INSTRUCTIONS FOR USE IFU AND PATIENT MANUAL INCLUDE A REFERENCE GUIDE FOR BOTH VISUAL AND TONE ALARMS INCLUDING POTENTIAL CAUSES AND ACTIONS TO TAKE ADDITIONALLY THERE IS A WARNING TO KEEP SPARE FULLY CHARGED BATTERIES AND BACK UP CONTROLLER AVAILABLE AT ALL TIME THE STEPS FOR EXCHANGE OF BATTERIES AND CONTROLLERS ARE OUTLINED THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT THIS EVENT WAS ASSESSED AND IS BEING REPORTED AS PART OF A RETROSPECTIVE REVIEW OF EVENTS WHICH WAS IN RESPONSE TO AN UPDATE TO THE MDR DECISION CRITERIA BATTERY B MANUFACTURING DATE EXPIRATION DATE FDA CODES B BATTERY B MANUFACTURING DATE EXPIRATION DATE FDA CODES B | Communications Problem Identified | Medical Evaluation result code
6973274 | 90084722 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE DRINKS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS COLLAPSING AND LOSS OF CONSCIOUSNESS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE CUSTOMER THINKS THE RECALLED SETS MAY HAVE CAUSED THEIR LOWS THE CUSTOMER ALSO HAD OTHER LOW INCIDENTS IN THE WEEK PRIOR TO THE CURRENT EVENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6973274 | 90084722 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE DRINKS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS COLLAPSING AND LOSS OF CONSCIOUSNESS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE CUSTOMER THINKS THE RECALLED SETS MAY HAVE CAUSED THEIR LOWS THE CUSTOMER ALSO HAD OTHER LOW INCIDENTS IN THE WEEK PRIOR TO THE CURRENT EVENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Loss Of Consciousness | Patient problem code
6973274 | 90084722 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE DRINKS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS COLLAPSING AND LOSS OF CONSCIOUSNESS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE CUSTOMER THINKS THE RECALLED SETS MAY HAVE CAUSED THEIR LOWS THE CUSTOMER ALSO HAD OTHER LOW INCIDENTS IN THE WEEK PRIOR TO THE CURRENT EVENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Collapse | Patient problem code
6973274 | 90084722 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN GLUCOSE DRINKS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS COLLAPSING AND LOSS OF CONSCIOUSNESS THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE CUSTOMER THINKS THE RECALLED SETS MAY HAVE CAUSED THEIR LOWS THE CUSTOMER ALSO HAD OTHER LOW INCIDENTS IN THE WEEK PRIOR TO THE CURRENT EVENT THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient problem code
6973452 | 90088096 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Calcified | Device problem code
6973452 | 90088096 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Inadequacy of Device Shape and/or Size | Device Problem Code
6973452 | 90088096 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Deflation Problem | Device problem code
6973507 | 90090116 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Deflation Problem | Device problem code
6973507 | 90090116 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Inadequacy of Device Shape and/or Size | Device problem code
6973507 | 90090116 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Calcified | Device problem code
6973507 | 90090116 | N | THE EVENT OF VISIBILITYPALPABILITY IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY UNSATISFACTORY RESULTS PATIENTS SHOULD BE INFORMED THAT DISSATISFACTION WITH COSMETIC RESULTS RELATED TO SUCH THINGS AS SCAR DEFORMITY HYPERTROPHIC SCARRING CAPSULAR CONTRACTURE ASYMMETRY WRINKLING IMPLANT DISPLACEMENTMIGRATION INCORRECT SIZE AND IMPLANT PALPABILITYVISIBILITY MAY OCCUR CAREFUL SURGICAL PLANNING AND TECHNIQUE CAN MINIMIZE BUT NOT PRECLUDE THE RISK OF SUCH RESULTS PREEXISTING ASYMMETRY MAY NOT BE ENTIRELY CORRECTABLE REVISION SURGERY MAY BE INDICATED TO MAINTAIN PATIENT SATISFACTION BUT CARRIES ADDITIONAL CONSIDERATIONS AND RISKS | Inadequacy of Device Shape and/or Size | Device Problem Code
6925626 | 90090361 | N | INITIAL REPORTER PHONE B INVESTIGATION DHR REVIEW FOR BATCH PN MANUFACTURING DATES ALL VISUAL INSPECTIONS WERE PERFORMED AS PER REQUIREMENT WITH NO QUALITY NOTIFICATIONS RELATED TO THE COMPLAINT DEFECT BATCH WAS INSPECTED AND ACCEPTED BASED ON MEETING OUR INSPECTION CONTROL PLAN AND SUBSEQUENTLY APPROVED FOR SHIPMENT SAMPLE EVALUATION VARIOUS SAMPLES IN THEIR ORIGINAL SEALED PACKAGING RECEIVED AND EVALUATED TWO ML PLASTIPAK SYRINGES FROM LOT P PN NOT CANAAN PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE BD ECLIPSE NEEDLES WITH SMARTSLIP FROM LOT PN NOT CANAAN PRODUCT VISUAL INSPECTION DID NOT REVEAL ANY DEFECTS IN ANY OF THE SAMPLES RECEIVED A BD ECLIPSE NEEDLE WAS THEN ATTACHED TO DIFFERENT SYRINGE SAMPLES INCLUDING ML ML AND ML IT WAS OBSERVED THAT WHEN BEING ATTACHED TO THE ML SAMPLES THE HUB TIGHTENED SOONER THAN WHEN ATTACHING TO ML OR ML SAMPLES THE CONNECTION WAS STILL SECURE THERE WAS NO DIFFICULTY CONNECTING THE NEEDLE UP TO THE POINT IT FELT TIGHT FOR ANY OF THE SAMPLES THE INVESTIGATION DISCOVERED THAT THE LUER COLLAR WAS THE SAME IN ALL OF THE SAMPLES AND HAS NO TAPER THE DIFFERENCE BETWEEN THE SAMPLES WAS THE SYRINGE TIP WHICH HAS A TAPER AN ISO RING GAGE TOOL WAS USED TO VERIFY THE TAPER ON ALL OF THE SAMPLES THE ML SAMPLES WERE AT THE LOWER END OF THE RANGE THAN THE ML AND ML SAMPLES WHICH MEANS THE ML SAMPLES HAVE A SLIGHTLY HIGHER DEGREE OF TAPER THE TIP WIDENS FASTER THIS WOULD EXPLAIN THE OBSERVED DIFFERENCE WHEN CONNECTING TO NEEDLE HUBS AND OTHER LUER LOCK DEVICES THE RESULT WOULD BE THE TIPS ON THE ML SYRINGES COMPLETING THE CONNECTION SOONER THAN ML OR ML USED IN THIS INVESTIGATION HOWEVER ALL SAMPLES WERE WITHIN THE SPECIFIED ISO RANGE AND ALL OF THE SAMPLES ARE ACCEPTABLE FOR USE IT IS NOT RECOMMENDED TO FORCE ANY DEVICES PAST THE POINT OF A SNUG SECURE CONNECTION BASED ON THE SAMPLE EVALUATION UNCONFIRMED BD CANAAN WAS NOT ABLE TO DUPLICATE OR CONFIRM THE CUSTOMERS INDICATED FAILURE BASED ON THE FACT THAT NO DEFECT WAS FOUND CAPA IS NOT REQUIRED AT THIS TIME UDI B | Mechanical Jam | Device problem code
6925626 | 90090361 | N | INITIAL REPORTER PHONE B INVESTIGATION DHR REVIEW FOR BATCH PN MANUFACTURING DATES ALL VISUAL INSPECTIONS WERE PERFORMED AS PER REQUIREMENT WITH NO QUALITY NOTIFICATIONS RELATED TO THE COMPLAINT DEFECT BATCH WAS INSPECTED AND ACCEPTED BASED ON MEETING OUR INSPECTION CONTROL PLAN AND SUBSEQUENTLY APPROVED FOR SHIPMENT SAMPLE EVALUATION VARIOUS SAMPLES IN THEIR ORIGINAL SEALED PACKAGING RECEIVED AND EVALUATED TWO ML PLASTIPAK SYRINGES FROM LOT P PN NOT CANAAN PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE BD ECLIPSE NEEDLES WITH SMARTSLIP FROM LOT PN NOT CANAAN PRODUCT VISUAL INSPECTION DID NOT REVEAL ANY DEFECTS IN ANY OF THE SAMPLES RECEIVED A BD ECLIPSE NEEDLE WAS THEN ATTACHED TO DIFFERENT SYRINGE SAMPLES INCLUDING ML ML AND ML IT WAS OBSERVED THAT WHEN BEING ATTACHED TO THE ML SAMPLES THE HUB TIGHTENED SOONER THAN WHEN ATTACHING TO ML OR ML SAMPLES THE CONNECTION WAS STILL SECURE THERE WAS NO DIFFICULTY CONNECTING THE NEEDLE UP TO THE POINT IT FELT TIGHT FOR ANY OF THE SAMPLES THE INVESTIGATION DISCOVERED THAT THE LUER COLLAR WAS THE SAME IN ALL OF THE SAMPLES AND HAS NO TAPER THE DIFFERENCE BETWEEN THE SAMPLES WAS THE SYRINGE TIP WHICH HAS A TAPER AN ISO RING GAGE TOOL WAS USED TO VERIFY THE TAPER ON ALL OF THE SAMPLES THE ML SAMPLES WERE AT THE LOWER END OF THE RANGE THAN THE ML AND ML SAMPLES WHICH MEANS THE ML SAMPLES HAVE A SLIGHTLY HIGHER DEGREE OF TAPER THE TIP WIDENS FASTER THIS WOULD EXPLAIN THE OBSERVED DIFFERENCE WHEN CONNECTING TO NEEDLE HUBS AND OTHER LUER LOCK DEVICES THE RESULT WOULD BE THE TIPS ON THE ML SYRINGES COMPLETING THE CONNECTION SOONER THAN ML OR ML USED IN THIS INVESTIGATION HOWEVER ALL SAMPLES WERE WITHIN THE SPECIFIED ISO RANGE AND ALL OF THE SAMPLES ARE ACCEPTABLE FOR USE IT IS NOT RECOMMENDED TO FORCE ANY DEVICES PAST THE POINT OF A SNUG SECURE CONNECTION BASED ON THE SAMPLE EVALUATION UNCONFIRMED BD CANAAN WAS NOT ABLE TO DUPLICATE OR CONFIRM THE CUSTOMERS INDICATED FAILURE BASED ON THE FACT THAT NO DEFECT WAS FOUND CAPA IS NOT REQUIRED AT THIS TIME UDI B | Defective Device | Device problem code
6925626 | 90090361 | N | INITIAL REPORTER PHONE B INVESTIGATION DHR REVIEW FOR BATCH PN MANUFACTURING DATES ALL VISUAL INSPECTIONS WERE PERFORMED AS PER REQUIREMENT WITH NO QUALITY NOTIFICATIONS RELATED TO THE COMPLAINT DEFECT BATCH WAS INSPECTED AND ACCEPTED BASED ON MEETING OUR INSPECTION CONTROL PLAN AND SUBSEQUENTLY APPROVED FOR SHIPMENT SAMPLE EVALUATION VARIOUS SAMPLES IN THEIR ORIGINAL SEALED PACKAGING RECEIVED AND EVALUATED TWO ML PLASTIPAK SYRINGES FROM LOT P PN NOT CANAAN PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE BD ECLIPSE NEEDLES WITH SMARTSLIP FROM LOT PN NOT CANAAN PRODUCT VISUAL INSPECTION DID NOT REVEAL ANY DEFECTS IN ANY OF THE SAMPLES RECEIVED A BD ECLIPSE NEEDLE WAS THEN ATTACHED TO DIFFERENT SYRINGE SAMPLES INCLUDING ML ML AND ML IT WAS OBSERVED THAT WHEN BEING ATTACHED TO THE ML SAMPLES THE HUB TIGHTENED SOONER THAN WHEN ATTACHING TO ML OR ML SAMPLES THE CONNECTION WAS STILL SECURE THERE WAS NO DIFFICULTY CONNECTING THE NEEDLE UP TO THE POINT IT FELT TIGHT FOR ANY OF THE SAMPLES THE INVESTIGATION DISCOVERED THAT THE LUER COLLAR WAS THE SAME IN ALL OF THE SAMPLES AND HAS NO TAPER THE DIFFERENCE BETWEEN THE SAMPLES WAS THE SYRINGE TIP WHICH HAS A TAPER AN ISO RING GAGE TOOL WAS USED TO VERIFY THE TAPER ON ALL OF THE SAMPLES THE ML SAMPLES WERE AT THE LOWER END OF THE RANGE THAN THE ML AND ML SAMPLES WHICH MEANS THE ML SAMPLES HAVE A SLIGHTLY HIGHER DEGREE OF TAPER THE TIP WIDENS FASTER THIS WOULD EXPLAIN THE OBSERVED DIFFERENCE WHEN CONNECTING TO NEEDLE HUBS AND OTHER LUER LOCK DEVICES THE RESULT WOULD BE THE TIPS ON THE ML SYRINGES COMPLETING THE CONNECTION SOONER THAN ML OR ML USED IN THIS INVESTIGATION HOWEVER ALL SAMPLES WERE WITHIN THE SPECIFIED ISO RANGE AND ALL OF THE SAMPLES ARE ACCEPTABLE FOR USE IT IS NOT RECOMMENDED TO FORCE ANY DEVICES PAST THE POINT OF A SNUG SECURE CONNECTION BASED ON THE SAMPLE EVALUATION UNCONFIRMED BD CANAAN WAS NOT ABLE TO DUPLICATE OR CONFIRM THE CUSTOMERS INDICATED FAILURE BASED ON THE FACT THAT NO DEFECT WAS FOUND CAPA IS NOT REQUIRED AT THIS TIME UDI B | Connection Problem | Device problem code
6925626 | 90090361 | N | INITIAL REPORTER PHONE B INVESTIGATION DHR REVIEW FOR BATCH PN MANUFACTURING DATES ALL VISUAL INSPECTIONS WERE PERFORMED AS PER REQUIREMENT WITH NO QUALITY NOTIFICATIONS RELATED TO THE COMPLAINT DEFECT BATCH WAS INSPECTED AND ACCEPTED BASED ON MEETING OUR INSPECTION CONTROL PLAN AND SUBSEQUENTLY APPROVED FOR SHIPMENT SAMPLE EVALUATION VARIOUS SAMPLES IN THEIR ORIGINAL SEALED PACKAGING RECEIVED AND EVALUATED TWO ML PLASTIPAK SYRINGES FROM LOT P PN NOT CANAAN PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE ML PLASTIPAK SYRINGES FROM LOT PN CANAANS PRODUCT THREE BD ECLIPSE NEEDLES WITH SMARTSLIP FROM LOT PN NOT CANAAN PRODUCT VISUAL INSPECTION DID NOT REVEAL ANY DEFECTS IN ANY OF THE SAMPLES RECEIVED A BD ECLIPSE NEEDLE WAS THEN ATTACHED TO DIFFERENT SYRINGE SAMPLES INCLUDING ML ML AND ML IT WAS OBSERVED THAT WHEN BEING ATTACHED TO THE ML SAMPLES THE HUB TIGHTENED SOONER THAN WHEN ATTACHING TO ML OR ML SAMPLES THE CONNECTION WAS STILL SECURE THERE WAS NO DIFFICULTY CONNECTING THE NEEDLE UP TO THE POINT IT FELT TIGHT FOR ANY OF THE SAMPLES THE INVESTIGATION DISCOVERED THAT THE LUER COLLAR WAS THE SAME IN ALL OF THE SAMPLES AND HAS NO TAPER THE DIFFERENCE BETWEEN THE SAMPLES WAS THE SYRINGE TIP WHICH HAS A TAPER AN ISO RING GAGE TOOL WAS USED TO VERIFY THE TAPER ON ALL OF THE SAMPLES THE ML SAMPLES WERE AT THE LOWER END OF THE RANGE THAN THE ML AND ML SAMPLES WHICH MEANS THE ML SAMPLES HAVE A SLIGHTLY HIGHER DEGREE OF TAPER THE TIP WIDENS FASTER THIS WOULD EXPLAIN THE OBSERVED DIFFERENCE WHEN CONNECTING TO NEEDLE HUBS AND OTHER LUER LOCK DEVICES THE RESULT WOULD BE THE TIPS ON THE ML SYRINGES COMPLETING THE CONNECTION SOONER THAN ML OR ML USED IN THIS INVESTIGATION HOWEVER ALL SAMPLES WERE WITHIN THE SPECIFIED ISO RANGE AND ALL OF THE SAMPLES ARE ACCEPTABLE FOR USE IT IS NOT RECOMMENDED TO FORCE ANY DEVICES PAST THE POINT OF A SNUG SECURE CONNECTION BASED ON THE SAMPLE EVALUATION UNCONFIRMED BD CANAAN WAS NOT ABLE TO DUPLICATE OR CONFIRM THE CUSTOMERS INDICATED FAILURE BASED ON THE FACT THAT NO DEFECT WAS FOUND CAPA IS NOT REQUIRED AT THIS TIME UDI B | Mechanical Problem Identified | Medical Evaluation result code
6966970 | 90102041 | N | THE NURSE REPORTED TO THE BIOMEDICAL ENGINEER BME THAT THE CENTRAL NURSES STATION CNS WAS NOT ALARMING PROPERLY NIHON KOHDEN TECHNICAL SUPPORT CALLED THE NURSE WHO REPORTED THE ISSUE SHE HAD AN OVERVIEW ALARM GOING ON IN THE BACKGROUND ONCE THE OVERVIEW ALARM WAS REMOVED THE NURSE FELT CONFIDENT THE ALARMS ARE FUNCTIONING PROPERLY | Device Alarm System | Device problem code
6966970 | 90102041 | N | THE NURSE REPORTED TO THE BIOMEDICAL ENGINEER BME THAT THE CENTRAL NURSES STATION CNS WAS NOT ALARMING PROPERLY NIHON KOHDEN TECHNICAL SUPPORT CALLED THE NURSE WHO REPORTED THE ISSUE SHE HAD AN OVERVIEW ALARM GOING ON IN THE BACKGROUND ONCE THE OVERVIEW ALARM WAS REMOVED THE NURSE FELT CONFIDENT THE ALARMS ARE FUNCTIONING PROPERLY | Mechanical Problem Identified | Medical Evaluation result code
6967207 | 90102168 | N | THE CUSTOMER REPORTED THAT THE CENTRAL NURSES STATION CNS SCREEN WENT BLACK AFTER NIHON KOHDEN TECHNICAL SUPPORT NKTS ASKED IF THE POWER WAS INTERRUPTED THE CUSTOMER CONFIRMED THE CNS DID NOT HAVE A UPS CONNECTED SO THAT SHUT DOWN THE CNS AFTER POWER CAME BACK THE CNS STARTED AND COULD MONITOR PATIENTS ISSUE RESOLVED | No Display / Image | Device Problem Code
6967207 | 90102168 | N | THE CUSTOMER REPORTED THAT THE CENTRAL NURSES STATION CNS SCREEN WENT BLACK AFTER NIHON KOHDEN TECHNICAL SUPPORT NKTS ASKED IF THE POWER WAS INTERRUPTED THE CUSTOMER CONFIRMED THE CNS DID NOT HAVE A UPS CONNECTED SO THAT SHUT DOWN THE CNS AFTER POWER CAME BACK THE CNS STARTED AND COULD MONITOR PATIENTS ISSUE RESOLVED | Image Artifact | Medical Evaluation result code
6966969 | 90102683 | N | THE CUSTOMER B STATED THAT AROUND AM ON B THE NURSES REPORTED THAT THEIR PURA CENTRAL NURSES STATION FROZE WHEN B CHECKED THE UNIT HE NOTICED THAT BOTH THE MONITORS WERE COMPLETELY BLANK AND HE NOTICED THAT ONE OF THE TWO DC ADAPTERS THAT GOES TO THEIR VANCRYST SWITCH WHERE VIDEO INTERFACE COMES IN WAS UNPLUGGED NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD B POWER CYCLE THE CENTRAL NURSES STATION CNS BY HOLDING THE POWER BUTTON ON THE CPU TO DISCHARGE THE ENERGY AND TURN IT ON HE WAS ABLE TO GET AN IMAGE AGAIN ON ONE MONITOR AND THE CNS PROGRAM LOADED FINE BUT THE OTHER MONITOR THAT HAD THE DC ADAPTER UNPLUGGED TO THE VANCRYST WOULD STILL NOT COME BACK ON BUT HE WOULD STILL SEE MESSAGES ON IT IF HE UNPLUGS THE DVI CABLE THAT THE CABLE IS UNPLUGGED PLUGGING THE DC ADAPTER BACK IN DIDNT WORK TO POWER THE MONITOR WE NARROWED IT DOWN TO THIS VANCRYST SWITCH AND THE DC ADAPTER BEING THE ISSUE AND THAT THE MONITOR WASNT GETTING THE POWER IT NEEDS NKTS RECOMMENDED TO B TO GET THEIR IT DEPARTMENT INVOLVED TO FURTHER TROUBLESHOOT THE VANCRYST ISSUE WHICH HE AGREED TO DO HE ALSO STATED THAT A NETWORK PERSON FROM NIHON KOHDEN IS COMING OUT NEXT WEEK AND WILL BE WORKING WITH HIM ON THIS ISSUE DURING INVESTIGATION B SAID HE NARROWED THE PROBLEM DOWN TO THE VANCRYST SWITCH NKTS ADVISED THAT THIS IS NOT A PART WE SELL OR HAVE IN STOCK | Material Separation | Device problem code
6966969 | 90102683 | N | THE CUSTOMER B STATED THAT AROUND AM ON B THE NURSES REPORTED THAT THEIR PURA CENTRAL NURSES STATION FROZE WHEN B CHECKED THE UNIT HE NOTICED THAT BOTH THE MONITORS WERE COMPLETELY BLANK AND HE NOTICED THAT ONE OF THE TWO DC ADAPTERS THAT GOES TO THEIR VANCRYST SWITCH WHERE VIDEO INTERFACE COMES IN WAS UNPLUGGED NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD B POWER CYCLE THE CENTRAL NURSES STATION CNS BY HOLDING THE POWER BUTTON ON THE CPU TO DISCHARGE THE ENERGY AND TURN IT ON HE WAS ABLE TO GET AN IMAGE AGAIN ON ONE MONITOR AND THE CNS PROGRAM LOADED FINE BUT THE OTHER MONITOR THAT HAD THE DC ADAPTER UNPLUGGED TO THE VANCRYST WOULD STILL NOT COME BACK ON BUT HE WOULD STILL SEE MESSAGES ON IT IF HE UNPLUGS THE DVI CABLE THAT THE CABLE IS UNPLUGGED PLUGGING THE DC ADAPTER BACK IN DIDNT WORK TO POWER THE MONITOR WE NARROWED IT DOWN TO THIS VANCRYST SWITCH AND THE DC ADAPTER BEING THE ISSUE AND THAT THE MONITOR WASNT GETTING THE POWER IT NEEDS NKTS RECOMMENDED TO B TO GET THEIR IT DEPARTMENT INVOLVED TO FURTHER TROUBLESHOOT THE VANCRYST ISSUE WHICH HE AGREED TO DO HE ALSO STATED THAT A NETWORK PERSON FROM NIHON KOHDEN IS COMING OUT NEXT WEEK AND WILL BE WORKING WITH HIM ON THIS ISSUE DURING INVESTIGATION B SAID HE NARROWED THE PROBLEM DOWN TO THE VANCRYST SWITCH NKTS ADVISED THAT THIS IS NOT A PART WE SELL OR HAVE IN STOCK | Leak / Splash | Device problem code
6966969 | 90102683 | N | THE CUSTOMER B STATED THAT AROUND AM ON B THE NURSES REPORTED THAT THEIR PURA CENTRAL NURSES STATION FROZE WHEN B CHECKED THE UNIT HE NOTICED THAT BOTH THE MONITORS WERE COMPLETELY BLANK AND HE NOTICED THAT ONE OF THE TWO DC ADAPTERS THAT GOES TO THEIR VANCRYST SWITCH WHERE VIDEO INTERFACE COMES IN WAS UNPLUGGED NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD B POWER CYCLE THE CENTRAL NURSES STATION CNS BY HOLDING THE POWER BUTTON ON THE CPU TO DISCHARGE THE ENERGY AND TURN IT ON HE WAS ABLE TO GET AN IMAGE AGAIN ON ONE MONITOR AND THE CNS PROGRAM LOADED FINE BUT THE OTHER MONITOR THAT HAD THE DC ADAPTER UNPLUGGED TO THE VANCRYST WOULD STILL NOT COME BACK ON BUT HE WOULD STILL SEE MESSAGES ON IT IF HE UNPLUGS THE DVI CABLE THAT THE CABLE IS UNPLUGGED PLUGGING THE DC ADAPTER BACK IN DIDNT WORK TO POWER THE MONITOR WE NARROWED IT DOWN TO THIS VANCRYST SWITCH AND THE DC ADAPTER BEING THE ISSUE AND THAT THE MONITOR WASNT GETTING THE POWER IT NEEDS NKTS RECOMMENDED TO B TO GET THEIR IT DEPARTMENT INVOLVED TO FURTHER TROUBLESHOOT THE VANCRYST ISSUE WHICH HE AGREED TO DO HE ALSO STATED THAT A NETWORK PERSON FROM NIHON KOHDEN IS COMING OUT NEXT WEEK AND WILL BE WORKING WITH HIM ON THIS ISSUE DURING INVESTIGATION B SAID HE NARROWED THE PROBLEM DOWN TO THE VANCRYST SWITCH NKTS ADVISED THAT THIS IS NOT A PART WE SELL OR HAVE IN STOCK | Leakage/Seal | Medical Evaluation result code
6966969 | 90102683 | N | THE CUSTOMER B STATED THAT AROUND AM ON B THE NURSES REPORTED THAT THEIR PURA CENTRAL NURSES STATION FROZE WHEN B CHECKED THE UNIT HE NOTICED THAT BOTH THE MONITORS WERE COMPLETELY BLANK AND HE NOTICED THAT ONE OF THE TWO DC ADAPTERS THAT GOES TO THEIR VANCRYST SWITCH WHERE VIDEO INTERFACE COMES IN WAS UNPLUGGED NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD B POWER CYCLE THE CENTRAL NURSES STATION CNS BY HOLDING THE POWER BUTTON ON THE CPU TO DISCHARGE THE ENERGY AND TURN IT ON HE WAS ABLE TO GET AN IMAGE AGAIN ON ONE MONITOR AND THE CNS PROGRAM LOADED FINE BUT THE OTHER MONITOR THAT HAD THE DC ADAPTER UNPLUGGED TO THE VANCRYST WOULD STILL NOT COME BACK ON BUT HE WOULD STILL SEE MESSAGES ON IT IF HE UNPLUGS THE DVI CABLE THAT THE CABLE IS UNPLUGGED PLUGGING THE DC ADAPTER BACK IN DIDNT WORK TO POWER THE MONITOR WE NARROWED IT DOWN TO THIS VANCRYST SWITCH AND THE DC ADAPTER BEING THE ISSUE AND THAT THE MONITOR WASNT GETTING THE POWER IT NEEDS NKTS RECOMMENDED TO B TO GET THEIR IT DEPARTMENT INVOLVED TO FURTHER TROUBLESHOOT THE VANCRYST ISSUE WHICH HE AGREED TO DO HE ALSO STATED THAT A NETWORK PERSON FROM NIHON KOHDEN IS COMING OUT NEXT WEEK AND WILL BE WORKING WITH HIM ON THIS ISSUE DURING INVESTIGATION B SAID HE NARROWED THE PROBLEM DOWN TO THE VANCRYST SWITCH NKTS ADVISED THAT THIS IS NOT A PART WE SELL OR HAVE IN STOCK | Fracture Problem | Medical Evaluation result code
6973659 | 90102742 | N | B THE GORE EXCLUDER AAA ENDOPROSTHESIS FEATURING C DELIVERY SYSTEM INSTRUCTIONS FOR USE IFU STATES THAT ADVERSE EVENTS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK AND ANEURYSM ENLARGEMENT FURTHER USERS ARE MADE AWARE OF THE RISKS ASSOCIATED WITH TYPE II ENDOLEAKS IN THE IFU AND ARE INSTRUCTED TO CONSIDER THE RISKS AND BENEFITS DISCUSSED IN THE IFU FOR EACH PATIENT BEFORE USING THE DEVICES | No Device Problem Found | Medical Evaluation result code
6972937 | 90102865 | N | THE BIOMEDICAL ENGINEER BME STATES THAT A PATIENT HAD THREE PAUSES BUT ONLY TWO ALARMED THE FIRST PAUSE WAS SECONDS LONG AND ALARMED AS BRADYCARDIA THE SECOND WAS SECONDS LONG AND ALARMED ASYSTOLE THE THIRD PAUSE WAS SECONDS LONG AND IT DID NOT ALARM ON THE TRANSMITTER DEVICE WHEN REVIEWING IN FULL DISCLOSURE THE CUSTOMER STATES THAT IN ALL THREE EPISODES THERE WAS A LEAD OFF BUT THE FIRST TWO STILL CAUSED AN ALARM SETTINGS ARE ALL APPROPRIATE REVIEWED GOOD ELECTRODE SKIN PREP PLACEMENT AND CHANGING EVER HOURS | Device Alarm System | Device problem code
6972937 | 90102865 | N | THE BIOMEDICAL ENGINEER BME STATES THAT A PATIENT HAD THREE PAUSES BUT ONLY TWO ALARMED THE FIRST PAUSE WAS SECONDS LONG AND ALARMED AS BRADYCARDIA THE SECOND WAS SECONDS LONG AND ALARMED ASYSTOLE THE THIRD PAUSE WAS SECONDS LONG AND IT DID NOT ALARM ON THE TRANSMITTER DEVICE WHEN REVIEWING IN FULL DISCLOSURE THE CUSTOMER STATES THAT IN ALL THREE EPISODES THERE WAS A LEAD OFF BUT THE FIRST TWO STILL CAUSED AN ALARM SETTINGS ARE ALL APPROPRIATE REVIEWED GOOD ELECTRODE SKIN PREP PLACEMENT AND CHANGING EVER HOURS | Mechanical Problem Identified | Medical Evaluation result code
6967203 | 90103317 | N | THE BIOMEDICAL ENGINEER STATED THAT THE SPO ADVISORY ALARMS ARE BLUE INSTEAD OF YELLOW THE BEDSIDE MONITORS BSM ARE WORKING NORMALLY BUT THE ISSUE IS WITH THE TELEMETRY UNITS THAT ARE NOT ALARMING THE CORRECT ADVISORY COLOR NIHON KOHDEN TECHNICAL SUPPORT ADVISED THE CUSTOMER TO GO TO SYSTEM SET UP PUT IN PASSWORD AND GO TO THE ORG TABALARM LEVELS AND CHANGE COLOR ALSO CHECK UNDER SYSTEM CONFIGURATION THAT THE ALARM ADVISORY COLOR WAS SET TO YELLOW CUSTOMER WAS ADVISED TO MAKE THE CHANGES TO THE OTHER CENTRAL NURSES STATIONS CNS ISSUE RESOLVED | Telemetry Discrepancy | Device problem code
6967203 | 90103317 | N | THE BIOMEDICAL ENGINEER STATED THAT THE SPO ADVISORY ALARMS ARE BLUE INSTEAD OF YELLOW THE BEDSIDE MONITORS BSM ARE WORKING NORMALLY BUT THE ISSUE IS WITH THE TELEMETRY UNITS THAT ARE NOT ALARMING THE CORRECT ADVISORY COLOR NIHON KOHDEN TECHNICAL SUPPORT ADVISED THE CUSTOMER TO GO TO SYSTEM SET UP PUT IN PASSWORD AND GO TO THE ORG TABALARM LEVELS AND CHANGE COLOR ALSO CHECK UNDER SYSTEM CONFIGURATION THAT THE ALARM ADVISORY COLOR WAS SET TO YELLOW CUSTOMER WAS ADVISED TO MAKE THE CHANGES TO THE OTHER CENTRAL NURSES STATIONS CNS ISSUE RESOLVED | Device Alarm System | Device problem code
6967203 | 90103317 | N | THE BIOMEDICAL ENGINEER STATED THAT THE SPO ADVISORY ALARMS ARE BLUE INSTEAD OF YELLOW THE BEDSIDE MONITORS BSM ARE WORKING NORMALLY BUT THE ISSUE IS WITH THE TELEMETRY UNITS THAT ARE NOT ALARMING THE CORRECT ADVISORY COLOR NIHON KOHDEN TECHNICAL SUPPORT ADVISED THE CUSTOMER TO GO TO SYSTEM SET UP PUT IN PASSWORD AND GO TO THE ORG TABALARM LEVELS AND CHANGE COLOR ALSO CHECK UNDER SYSTEM CONFIGURATION THAT THE ALARM ADVISORY COLOR WAS SET TO YELLOW CUSTOMER WAS ADVISED TO MAKE THE CHANGES TO THE OTHER CENTRAL NURSES STATIONS CNS ISSUE RESOLVED | Mechanical Problem Identified | Medical Evaluation result code
6960820 | 90107808 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Insufficient Information | Device problem code
6960820 | 90107808 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6960838 | 90107951 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6960838 | 90107951 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6943623 | 90108002 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED MICROSCOPIC VISUAL INSPECTION FOUND TRACES OF DRIED BLOOD IN THE LUMEN A NEW STYLET WAS RETRIEVED AND AN ATTEMPT WAS MADE TO PASS IT THROUGH THE LEAD HOWEVER THE STYLET STOPPED AT THE SPIRAL FIXATION AREA CLOSE TO THE ELECTRODE TIP IT WAS DETERMINED THAT THERE WAS WHAT APPEARED TO BE FOREIGN MATERIAL CAUSING A BLOCKAGE IN THE LUMEN THE MATERIAL WAS REMOVED AND DETAIL ANALYSIS DETERMINED IT TO BE AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION THAT OCCURRED DURING THE IMPLANT PROCEDURE LABORATORY ANALYSIS WAS ABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS | Device Contamination with Body Fluid | Device problem code
6943623 | 90108002 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED MICROSCOPIC VISUAL INSPECTION FOUND TRACES OF DRIED BLOOD IN THE LUMEN A NEW STYLET WAS RETRIEVED AND AN ATTEMPT WAS MADE TO PASS IT THROUGH THE LEAD HOWEVER THE STYLET STOPPED AT THE SPIRAL FIXATION AREA CLOSE TO THE ELECTRODE TIP IT WAS DETERMINED THAT THERE WAS WHAT APPEARED TO BE FOREIGN MATERIAL CAUSING A BLOCKAGE IN THE LUMEN THE MATERIAL WAS REMOVED AND DETAIL ANALYSIS DETERMINED IT TO BE AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION THAT OCCURRED DURING THE IMPLANT PROCEDURE LABORATORY ANALYSIS WAS ABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS | Device Contamination with Chemical or Other Material | Device problem code
6943623 | 90108002 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED MICROSCOPIC VISUAL INSPECTION FOUND TRACES OF DRIED BLOOD IN THE LUMEN A NEW STYLET WAS RETRIEVED AND AN ATTEMPT WAS MADE TO PASS IT THROUGH THE LEAD HOWEVER THE STYLET STOPPED AT THE SPIRAL FIXATION AREA CLOSE TO THE ELECTRODE TIP IT WAS DETERMINED THAT THERE WAS WHAT APPEARED TO BE FOREIGN MATERIAL CAUSING A BLOCKAGE IN THE LUMEN THE MATERIAL WAS REMOVED AND DETAIL ANALYSIS DETERMINED IT TO BE AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION THAT OCCURRED DURING THE IMPLANT PROCEDURE LABORATORY ANALYSIS WAS ABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS | Partial Blockage | Device problem code
6943623 | 90108002 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED MICROSCOPIC VISUAL INSPECTION FOUND TRACES OF DRIED BLOOD IN THE LUMEN A NEW STYLET WAS RETRIEVED AND AN ATTEMPT WAS MADE TO PASS IT THROUGH THE LEAD HOWEVER THE STYLET STOPPED AT THE SPIRAL FIXATION AREA CLOSE TO THE ELECTRODE TIP IT WAS DETERMINED THAT THERE WAS WHAT APPEARED TO BE FOREIGN MATERIAL CAUSING A BLOCKAGE IN THE LUMEN THE MATERIAL WAS REMOVED AND DETAIL ANALYSIS DETERMINED IT TO BE AN AGGLOMERATE OF BLOODBODY FLUID AND POLYMER FROM THE LEADGUIDEWIRE INTERACTION THAT OCCURRED DURING THE IMPLANT PROCEDURE LABORATORY ANALYSIS WAS ABLE TO CONFIRM THE REPORTED CLINICAL OBSERVATIONS | Operational Problem Identified | Medical Evaluation result code
6943664 | 90114363 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION OF THE TERMINAL PIN ASSEMBLY LEAD BODY AND ELECTRODE TIP FOUND DRIED BODY FLUIDBLOOD IN THE TIP AREA OF THE OPEN LUMEN THE LEAD WAS RETURNED WITH THE GUIDEWIRE FULLY INSERTED PROLAPSED AT THE DISTAL END OF THE LEAD TIP ANALYSIS CONFIRMED THE REPORTED CLINICAL OBSERVATION NOTING THE GUIDEWIRE LIKELY PROLAPSED DISTAL TO THE LEAD TIP AND WAS THEN PULLED BACK INTO THE LEAD IN A PROLAPSED STATE CAUSING IT TO BECOME STUCK | Mechanical Jam | Device problem code
6943664 | 90114363 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION OF THE TERMINAL PIN ASSEMBLY LEAD BODY AND ELECTRODE TIP FOUND DRIED BODY FLUIDBLOOD IN THE TIP AREA OF THE OPEN LUMEN THE LEAD WAS RETURNED WITH THE GUIDEWIRE FULLY INSERTED PROLAPSED AT THE DISTAL END OF THE LEAD TIP ANALYSIS CONFIRMED THE REPORTED CLINICAL OBSERVATION NOTING THE GUIDEWIRE LIKELY PROLAPSED DISTAL TO THE LEAD TIP AND WAS THEN PULLED BACK INTO THE LEAD IN A PROLAPSED STATE CAUSING IT TO BECOME STUCK | Mechanical Problem Identified | Medical Evaluation result code
6965796 | 90119053 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A COUPLE HOURS AFTER THE EVENT | Device Dislodged or Dislocated | Device problem code
6965796 | 90119053 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE A NEW POD WAS APPLIED A COUPLE HOURS AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
6960852 | 90120148 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6960852 | 90120148 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6974318 | 90122041 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION ISSUE WAS ALLEGED AGAINST TWO SENSORS REPORTED UNDER AND | Material Discolored | Device problem code
6974318 | 90122041 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION ISSUE WAS ALLEGED AGAINST TWO SENSORS REPORTED UNDER AND | Stress Problem Identified | Medical Evaluation result code
6974314 | 90124831 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION ISSUE WAS ALLEGED AGAINST TWO SENSORS REPORTED UNDER THIS REPORT AND | Material Discolored | Device problem code
6974314 | 90124831 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION ISSUE WAS ALLEGED AGAINST TWO SENSORS REPORTED UNDER THIS REPORT AND | Stress Problem Identified | Medical Evaluation result code
6975516 | 90151561 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6975516 | 90151561 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
6974959 | 90151918 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
6974959 | 90151918 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Material Discolored | Device problem code
6974959 | 90151918 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
6974959 | 90151918 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Device Migration | Medical Evaluation result code
6974959 | 90151918 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
6974966 | 90151987 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS INJECTION PROCEDURE REACTION TO JUVÉDERM ULTRA INJECTABLE GEL HAS BEEN OBSERVED AS CONSISTING MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND WITH LESS THAN DAYS DURATION REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION B OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Leak / Splash | Device problem code
6974966 | 90151987 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS INJECTION PROCEDURE REACTION TO JUVÉDERM ULTRA INJECTABLE GEL HAS BEEN OBSERVED AS CONSISTING MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND WITH LESS THAN DAYS DURATION REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION B OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Migration or Expulsion of Device | Device problem code
6974966 | 90151987 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS INJECTION PROCEDURE REACTION TO JUVÉDERM ULTRA INJECTABLE GEL HAS BEEN OBSERVED AS CONSISTING MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND WITH LESS THAN DAYS DURATION REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION B OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Material Discolored | Device problem code
6974966 | 90151987 | N | B THE EVENTS OF EDEMA ERYTHEMA INDURATION ITCHING INFLAMMATORY RESPONSE AND HYPERSENSITIVITY ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING WARNINGS INJECTION PROCEDURE REACTION TO JUVÉDERM ULTRA INJECTABLE GEL HAS BEEN OBSERVED AS CONSISTING MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND WITH LESS THAN DAYS DURATION REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF TREATED SUBJECTS POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM ULTRA INCLUDE REDNESS PAINTENDERNESS FIRMNESS SWELLING LUMPSBUMPS BRUISING ITCHING AND DISCOLORATION B OTHER SAFETY DATA POSTMARKET SURVEILLANCE THE FOLLOWING ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM ULTRA AND ULTRA PLUS WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL ADVERSE EVENTS OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED LACK OR LOSS OF CORRECTION INFLAMMATORY REACTION ALLERGIC REACTION INFECTION MIGRATION PARESTHESIA VASCULAR OCCLUSION NECROSIS ABSCESS FLULIKE SYMPTOMS HEADACHE MALAISE VISION ABNORMALITIES SCARRING NAUSEA DRAINAGE DYSPNEA BEADING SYNCOPE DIZZINESS ANXIETY DEEPER WRINKLE AND GRANULOMA IN MANY CASES THE SYMPTOMS RESOLVED WITHOUT ANY TREATMENT REPORTED TREATMENTS HAVE INCLUDED ANTIBIOTICS STEROIDS STEROIDAL CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION AND DRAINAGE ULTRASOUND THERAPY ANALGESICS ANTIVIRAL EXCISION EYE DROPS HYPERBARIC OXYGEN LASER RESURFACING TISSUE DEBRIDEMENT SURGICAL SCAR REVISION ICE MASSAGE AND WARM COMPRESS | Stress Problem Identified | Medical Evaluation result code
6929590 | 90160583 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED | Connection Problem | Device problem code
6929590 | 90160583 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGE IN QUESTION THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
6933738 | 90166324 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR | Component Incompatible | Device problem code
6933738 | 90166324 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY WE ARE WAITING FOR ADDITIONAL INFORMATION FROM DISTRIBUTOR | Interoperability Problem Identified | Medical Evaluation result code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Mechanical Jam | Device problem code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Material Twisted / Bent | Device problem code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Partial Blockage | Device problem code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Material Separation | Device problem code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Mechanical Problem Identified | Medical Evaluation result code
6942928 | 90169543 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT AND PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING ADDRESSES THE REPORTED EVENTS AS FOLLOWS METHOD OF USEPOSOLOGY REMOVE TIP CAP BY PULLING IT STRAIGHT OFF THE SYRINGE THEN FIRMLY PUSH THE NEEDLE PROVIDED IN THE BOX INTO THE SYRINGE SCREWING IT GENTLY CLOCKWISE TWIST ONCE MORE UNTIL IT IS FULLY LOCKED NEXT REMOVE THE PROTECTIVE CAP BY HOLDING THE BODY OF THE SYRINGE IN ONE HAND THE PROTECTIVE CAP IN THE OTHER AND PULLING THE TWO HANDS IN OPPOSITE DIRECTIONS PRIOR TO INJECTING DEPRESS THE PLUNGER ROD UNTIL THE PRODUCT FLOWS OUT OF THE NEEDLE INJECT SLOWLY AND APPLY THE LEAST AMOUNT OF PRESSURE NECESSARY IF THE NEEDLE IS BLOCKED DO NOT INCREASE THE PRESSURE ON THE PLUNGER ROD INSTEAD STOP THE INJECTION AND REPLACE THE NEEDLE WARNINGS IN THE EVENT THAT THE CONTENT OF A SYRINGE SHOWS SIGNS OF SEPARATION ANDOR APPEARS CLOUDY DO NOT USE THE SYRINGE STORAGE CONDITIONS STORE BETWEEN C AND C | Stress Problem Identified | Medical Evaluation result code
6963098 | 90173791 | D | IT WAS REPORTED THAT CLIPS ARE FALLING OUT OF THE JAW IT DID NOT FALL INTO THE PATIENTS BODY THE PATIENTS CONDITION IS FINE | Device Fell | Device problem code
6963098 | 90173791 | D | IT WAS REPORTED THAT CLIPS ARE FALLING OUT OF THE JAW IT DID NOT FALL INTO THE PATIENTS BODY THE PATIENTS CONDITION IS FINE | No Consequences Or Impact To Patient | Patient problem code
6936450 | 90179107 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP FOR CUSTOMER COMPLAINT OF LOW VACUUM ALARMS THE FSE VERIFIED THE ALARMS IN THE ERROR LOGS AND COMPLETELY TESTED THE IABP AND WAS UNABLE TO RECREATE ANY FAILURE WITH THE OPERATION OF THE IABP THE PERFORMANCE AND CALIBRATION PROCEDURES PERFORMED SATISFACTORILY AND THE IABP WAS APPROVED FOR CLINICAL USE AND RETURNED TO THE CUSTOMER | Calibration Problem | Device problem code
6936450 | 90179107 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP FOR CUSTOMER COMPLAINT OF LOW VACUUM ALARMS THE FSE VERIFIED THE ALARMS IN THE ERROR LOGS AND COMPLETELY TESTED THE IABP AND WAS UNABLE TO RECREATE ANY FAILURE WITH THE OPERATION OF THE IABP THE PERFORMANCE AND CALIBRATION PROCEDURES PERFORMED SATISFACTORILY AND THE IABP WAS APPROVED FOR CLINICAL USE AND RETURNED TO THE CUSTOMER | Decrease in Suction | Device problem code
6936450 | 90179107 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE EVALUATED THE IABP FOR CUSTOMER COMPLAINT OF LOW VACUUM ALARMS THE FSE VERIFIED THE ALARMS IN THE ERROR LOGS AND COMPLETELY TESTED THE IABP AND WAS UNABLE TO RECREATE ANY FAILURE WITH THE OPERATION OF THE IABP THE PERFORMANCE AND CALIBRATION PROCEDURES PERFORMED SATISFACTORILY AND THE IABP WAS APPROVED FOR CLINICAL USE AND RETURNED TO THE CUSTOMER | No Findings Available | Medical Evaluation result code
6971465 | 90183514 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY SHOWED COLORED VERTICAL LINES AND WAS NOT LEGIBLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6971465 | 90183514 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY SHOWED COLORED VERTICAL LINES AND WAS NOT LEGIBLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Erratic or Intermittent Display | Device Problem Code
6939704 | 90186437 | N | B BATCH PKD THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DID ANY PIECE OR PIECES OF THE FIRING KNOB FALL INTO THE PATIENT NO INFORMATION WILL ALL PIECES BE RETURNED WITH THE DEVICE NO INFORMATION | Insufficient Information | Device problem code
6939704 | 90186437 | N | B BATCH PKD THE LOT HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE LOT ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DID ANY PIECE OR PIECES OF THE FIRING KNOB FALL INTO THE PATIENT NO INFORMATION WILL ALL PIECES BE RETURNED WITH THE DEVICE NO INFORMATION | No Device Problem Found | Medical Evaluation result code
6944835 | 90187104 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944835 | 90187104 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944796 | 90187919 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6944796 | 90187919 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Findings Available | Medical Evaluation result code
6944807 | 90188001 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944807 | 90188001 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944802 | 90188055 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944802 | 90188055 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944810 | 90188152 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944810 | 90188152 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944811 | 90188419 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944811 | 90188419 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6945881 | 90206694 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6945881 | 90206694 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6955027 | 90214974 | N | NO PATIENT INFORMATION PROVIDED AS NO PATIENT WAS INVOLVED IN THIS CONCERN THE SUSPECT VERTEK ARM HAS BEEN RECEIVED BY MANUFACTURER FOR EVALUATION THE RETURNED VERTEK ARM FAILED FORCE TEST THE ARM SHOULD HOLD A DOWNWARD FORCE UOF UPTO LBS BUT THE RETURNED VERTEK ARM WAS HOLDING LBS THE ARM SHOULD ALSO HOLD SIDE WARD FORCE UP TO LBS BUT FAILED AT LBS | No Device Problem Found | Medical Evaluation result code
6947333 | 90216257 | D | IT WAS REPORTED THAT THE JIG WOULD NOT HOLD THE SET POSITION WHEN THE SURGEON SELECTED HIS ROTATION NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE MALFUNCTION | Activation, Positioning or Separation Problem | Device Problem Code
6947333 | 90216257 | D | IT WAS REPORTED THAT THE JIG WOULD NOT HOLD THE SET POSITION WHEN THE SURGEON SELECTED HIS ROTATION NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE MALFUNCTION | No Information | Patient Problem Code
6976660 | 90218737 | N | B PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6976660 | 90218737 | N | B PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
6952445 | 90219147 | N | PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASEVIAL VISUAL INSPECTION FOUND THE HAPTIC TORNBROKEN AND A PIECE MISSING B | Component Missing | Device problem code
6952445 | 90219147 | N | PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASEVIAL VISUAL INSPECTION FOUND THE HAPTIC TORNBROKEN AND A PIECE MISSING B | Break | Device problem code
6952445 | 90219147 | N | PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASEVIAL VISUAL INSPECTION FOUND THE HAPTIC TORNBROKEN AND A PIECE MISSING B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6975511 | 90221162 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6975511 | 90221162 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
6944823 | 90222483 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6944823 | 90222483 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6972306 | 90223922 | N | THE REPORTED NEEDLE INSERTION DIFFICULTY WAS CONFIRMED RESISTANCE WAS NOTED NEAR THE DISTAL END OF THE DILATOR WHEN A BRK NEEDLESTYLET ASSEMBLY FROM CURRENT INVENTORY WAS ADVANCED THROUGH THE RETURNED DILATOR AND SHEATH THE DILATOR TUBING WAS CROSSSECTIONED AND A BUILDUP OF SKIVED MATERIAL WAS NOTED AT THE DISTAL END OF THE DILATOR THE CAUSE OF THE REPORTED COMPLAINT IS CONSISTENT WITH THE FAILURE TO USE A STYLETSIMILAR COMPONENT DURING NEEDLE INSERTION THROUGH THE DILATOR THE SWARTZ TRANSSEPTAL GUIDING INTRODUCER INSTRUCTIONS FOR USE IFU STATES DURING INSERTION ALWAYS USE THE STYLET TO FACILITATE NEEDLE PASSAGE THROUGH THE DILATORSHEATH ASSEMBLY FAILURE TO USE THE STYLET MAY INHIBIT ADVANCEMENT OF THE NEEDLE AND MAY RESULT IN INADVERTENT NEEDLE PUNCTURE OF THE DILATORSHEATH ASSEMBLY OR SKIVING OF MATERIAL FROM THE INNER SURFACE OF THE DILATOR THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED | Material Puncture / Hole | Device Problem Code
6972306 | 90223922 | N | THE REPORTED NEEDLE INSERTION DIFFICULTY WAS CONFIRMED RESISTANCE WAS NOTED NEAR THE DISTAL END OF THE DILATOR WHEN A BRK NEEDLESTYLET ASSEMBLY FROM CURRENT INVENTORY WAS ADVANCED THROUGH THE RETURNED DILATOR AND SHEATH THE DILATOR TUBING WAS CROSSSECTIONED AND A BUILDUP OF SKIVED MATERIAL WAS NOTED AT THE DISTAL END OF THE DILATOR THE CAUSE OF THE REPORTED COMPLAINT IS CONSISTENT WITH THE FAILURE TO USE A STYLETSIMILAR COMPONENT DURING NEEDLE INSERTION THROUGH THE DILATOR THE SWARTZ TRANSSEPTAL GUIDING INTRODUCER INSTRUCTIONS FOR USE IFU STATES DURING INSERTION ALWAYS USE THE STYLET TO FACILITATE NEEDLE PASSAGE THROUGH THE DILATORSHEATH ASSEMBLY FAILURE TO USE THE STYLET MAY INHIBIT ADVANCEMENT OF THE NEEDLE AND MAY RESULT IN INADVERTENT NEEDLE PUNCTURE OF THE DILATORSHEATH ASSEMBLY OR SKIVING OF MATERIAL FROM THE INNER SURFACE OF THE DILATOR THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED | Stress Problem Identified | Medical Evaluation result code
6944813 | 90224781 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6944813 | 90224781 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6956581 | 90229509 | N | DATE OF EVENT B | Insufficient Information | Device Problem Code
6958268 | 90235373 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Component or Accessory Incompatibility | Device Problem Code
6958268 | 90235373 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE B | Interoperability Problem Identified | Medical Evaluation result code
6977200 | 90237393 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6965463 | 90237404 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE WEEK LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965463 | 90237404 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B ONE WEEK LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6977184 | 90239151 | N | UNIQUE COMPLAINT ID B SENT NON SLIP PADS FOR CUSTOMER ON B AND WERE DELIVERED ON B CUSTOMER CONFIRMS RECEIPT OF NON SLIP PADS CUSTOMER WILL HAVE HER SON IN LAW INSTALL NON SLIP PADS ON THE BASE NO FURTHER ASSISTANCE IS NEEDED THE CUSTOMER HAS NOT INDICATED A MALFUNCTION A DEFECT OF THE DEVICE OR AN ADVERSE EVENT CUSTOMER CONFIRMED PART RECEIPT | No Device Problem Found | Medical Evaluation result code
6969223 | 90239269 | N | VYAIRE FILE IDENTIFICATION B ANY ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER WILL BE INCLUDED IN A FOLLOWUP REPORT THE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE THE EVALUATE THE SUSPECT DEVICE THE FSR DETERMINED THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS THE EXHALATION FLOW TRANSDUCER BEING OUT OF CALIBRATION AFTER PERFORMING CALIBRATION THE ISSUE WAS RESOLVED THE FSR PERFORMED THE OPERATIONAL VERIFICATION PROCEDURE AND REPORTED THE UNIT MEETS SERVICE SPECIFICATIONS | Calibration Problem | Device problem code
6969223 | 90239269 | N | VYAIRE FILE IDENTIFICATION B ANY ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER WILL BE INCLUDED IN A FOLLOWUP REPORT THE FIELD SERVICE REPRESENTATIVE FSR WENT ONSITE THE EVALUATE THE SUSPECT DEVICE THE FSR DETERMINED THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS THE EXHALATION FLOW TRANSDUCER BEING OUT OF CALIBRATION AFTER PERFORMING CALIBRATION THE ISSUE WAS RESOLVED THE FSR PERFORMED THE OPERATIONAL VERIFICATION PROCEDURE AND REPORTED THE UNIT MEETS SERVICE SPECIFICATIONS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6948545 | 90247805 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6948545 | 90247805 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6928926 | 90251515 | N | B THE FIELD SERVICE REPRESENTATIVE FSR PROVIDED THE CUSTOMER WITH LITERATURE THAT INCLUDED TECHNIQUES FROM A CLINICAL STANDPOINT TO POTENTIALLY MINIMIZE ARTIFACT ON MONITORS | Signal Artifact | Device problem code
6928926 | 90251515 | N | B THE FIELD SERVICE REPRESENTATIVE FSR PROVIDED THE CUSTOMER WITH LITERATURE THAT INCLUDED TECHNIQUES FROM A CLINICAL STANDPOINT TO POTENTIALLY MINIMIZE ARTIFACT ON MONITORS | Image Artifact | Medical Evaluation result code
6931387 | 90251726 | N | B THE FIELD SERVICE REPRESENTATIVE FSR COULD NOT VERIFY THE DISPLAY BLINKING A LOANER PUMP WAS INSTALLED THE UNIT OPERATED TO THE MANUFACTURERS SPECIFICATIONS THE SUSPECT DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION | Display or Visual Feedback Problem | Device problem code
6931387 | 90251726 | N | B THE FIELD SERVICE REPRESENTATIVE FSR COULD NOT VERIFY THE DISPLAY BLINKING A LOANER PUMP WAS INSTALLED THE UNIT OPERATED TO THE MANUFACTURERS SPECIFICATIONS THE SUSPECT DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION | Results Pending Completion of Investigation | Medical Evaluation result code
6954420 | 90252291 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6954420 | 90252291 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6954649 | 90254275 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENTS FATHER OBSERVED THE CANNULA THROUGH THE VIEWING WINDOW AND WAS UNSURE IF IT WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6954649 | 90254275 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN THE PATIENTS FATHER OBSERVED THE CANNULA THROUGH THE VIEWING WINDOW AND WAS UNSURE IF IT WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
6950254 | 90254506 | N | B ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL THE FACILITY DECLINED TO PROVIDE PATIENT INFORMATION NO TESTSLABORATORY DATA WAS AVAILABLE THERE IS NO SERIAL NUMBER FOR THIS DEVICE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Apparent Adverse Event | Device problem code
6950254 | 90254506 | N | B ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL THE FACILITY DECLINED TO PROVIDE PATIENT INFORMATION NO TESTSLABORATORY DATA WAS AVAILABLE THERE IS NO SERIAL NUMBER FOR THIS DEVICE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Findings Available | Medical Evaluation result code
6977877 | 90257539 | D | THE CUSTOMER REPORTED VIA PHONE CALL THEY WERE EXPERIENCING HIGH BLOOD GLUCOSE LEVELS OF MGDL THE CUSTOMER STATED THAT THEY MIGHT HAVE INSERTED THEIR MIO SET INCORRECTLY THE INFUSION SET WAS NOT EXPIRED THE CUSTOMER TREATED HIGH BLOOD GLUCOSE LEVELS WITH PUMP THE CUSTOMER DECLINED TROUBLESHOOTING AND INDICATED THAT THEY WOULD TREAT ELEVATED BLOOD GLUCOSE WITH THE PUMP AS NECESSARY THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Activation, Positioning or Separation Problem | Device problem code
6977877 | 90257539 | D | THE CUSTOMER REPORTED VIA PHONE CALL THEY WERE EXPERIENCING HIGH BLOOD GLUCOSE LEVELS OF MGDL THE CUSTOMER STATED THAT THEY MIGHT HAVE INSERTED THEIR MIO SET INCORRECTLY THE INFUSION SET WAS NOT EXPIRED THE CUSTOMER TREATED HIGH BLOOD GLUCOSE LEVELS WITH PUMP THE CUSTOMER DECLINED TROUBLESHOOTING AND INDICATED THAT THEY WOULD TREAT ELEVATED BLOOD GLUCOSE WITH THE PUMP AS NECESSARY THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Death | Patient problem code
6977877 | 90257539 | D | THE CUSTOMER REPORTED VIA PHONE CALL THEY WERE EXPERIENCING HIGH BLOOD GLUCOSE LEVELS OF MGDL THE CUSTOMER STATED THAT THEY MIGHT HAVE INSERTED THEIR MIO SET INCORRECTLY THE INFUSION SET WAS NOT EXPIRED THE CUSTOMER TREATED HIGH BLOOD GLUCOSE LEVELS WITH PUMP THE CUSTOMER DECLINED TROUBLESHOOTING AND INDICATED THAT THEY WOULD TREAT ELEVATED BLOOD GLUCOSE WITH THE PUMP AS NECESSARY THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
6978474 | 90284745 | N | THERE IS NO DEFECT OR MALFUNCTION OF THE DEVICE ASSOCIATED WITH THIS EVENT ALL PERTINENT INFORMATION AVAILABLE TO SECOND SIGHT MEDICAL PRODUCTS HAS BEEN SUBMITTED THE COMPANY IS SUBMITTING THIS MDR TO ENSURE FULL COMPLIANCE WITH CFR PART | Defective Device | Device problem code
6978474 | 90284745 | N | THERE IS NO DEFECT OR MALFUNCTION OF THE DEVICE ASSOCIATED WITH THIS EVENT ALL PERTINENT INFORMATION AVAILABLE TO SECOND SIGHT MEDICAL PRODUCTS HAS BEEN SUBMITTED THE COMPANY IS SUBMITTING THIS MDR TO ENSURE FULL COMPLIANCE WITH CFR PART | No Findings Available | Medical Evaluation result code
6948284 | 90291462 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948284 | 90291462 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6931969 | 90292769 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM AC POWER SUPPLY HAD A CONNECTION ISSUE IT DID NOT PREVENT THE FREEDOM DRIVER FROM PERFORMING ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER HAS A REDUNDANT POWER SOURCE OF ONBOARD BATTERIES THE FREEDOM AC POWER SUPPLY HAS BEEN RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B INITIAL | Connection Problem | Device problem code
6931969 | 90292769 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM AC POWER SUPPLY HAD A CONNECTION ISSUE IT DID NOT PREVENT THE FREEDOM DRIVER FROM PERFORMING ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER HAS A REDUNDANT POWER SOURCE OF ONBOARD BATTERIES THE FREEDOM AC POWER SUPPLY HAS BEEN RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B INITIAL | No Device Problem Found | Medical Evaluation result code
6931965 | 90293014 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED FAULT ALARMS IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B INITIAL | Defective Alarm | Device problem code
6931965 | 90293014 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE FREEDOM DRIVER EXHIBITED FAULT ALARMS IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE FREEDOM DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOWUP MDR B INITIAL | Mechanical Problem Identified | Medical Evaluation result code
6979120 | 90293429 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
6979120 | 90293429 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
6940799 | 90294421 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT PROCEDURE WAS BEING PERFORMED WHEN THE DEVICE CUT BEFORE CLAMPING AND DID NOT FIRE STAPLES WHAT WAS DONE TO ADDRESS THIS ISSUE HOW WAS THE PROCEDURE COMPLETED WERE THERE ANY PATIENT CONSEQUENCES AS A RESULT OF THE EVENT | Misfire | Device problem code
6940799 | 90294421 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT PROCEDURE WAS BEING PERFORMED WHEN THE DEVICE CUT BEFORE CLAMPING AND DID NOT FIRE STAPLES WHAT WAS DONE TO ADDRESS THIS ISSUE HOW WAS THE PROCEDURE COMPLETED WERE THERE ANY PATIENT CONSEQUENCES AS A RESULT OF THE EVENT | Fire | Device problem code
6940799 | 90294421 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION TO DATE NO RESPONSE HAS BEEN PROVIDED IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT PROCEDURE WAS BEING PERFORMED WHEN THE DEVICE CUT BEFORE CLAMPING AND DID NOT FIRE STAPLES WHAT WAS DONE TO ADDRESS THIS ISSUE HOW WAS THE PROCEDURE COMPLETED WERE THERE ANY PATIENT CONSEQUENCES AS A RESULT OF THE EVENT | Operational Problem Identified | Medical Evaluation result code
6979221 | 90297360 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
6979221 | 90297360 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
6948198 | 90298643 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6948198 | 90298643 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6948200 | 90298696 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948200 | 90298696 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6949720 | 90301599 | N | B MANUFACTURING SITE EVALUATION SAMPLES RECEIVED OPEN CASSETTE ANALYSIS AND RESULTS THERE ARE NO PREVIOUS COMPLAINTS OF THE SAME CODEBATCH MANUFACTURED AND DISTRIBUTED IN THE MARKET B UNITS OF THIS CODE BATCH THERE ARE NO UNITS IN STOCK WE HAVE RECEIVED ONE OPEN AND USED CASSETTE TESTED THE KNOT PULL TENSILE STRENGTH OF THE SAMPLE RECEIVED AND THE RESULTS FULFILL THE OEM REQUIREMENTS FINAL CONCLUSION ALTHOUGH THE RESULTS OF THE SAMPLE RECEIVED FULFILL THE OEM SPECIFICATIONS NOTE OF THIS INCIDENCE IN ORDER TO ASSESS IF NEW OR ADDITIONAL ACTIONS ARE NEEDED CORRECTIVEPREVENTIVE ACTIONS ACCORDING TO OUR INTERNAL PROCEDURES THERE IS NO NEED TO ESTABLISH CORRECTIVE OR PREVENTIVE ACTIONS NEVERTHELESS THIS COMPLAINT IS RECORDED FOR TRENDING ANALYSIS TO ASSESS IF ACTIONS ARE NEEDED IN THE FUTURE | Insufficient Information | Device Problem Code
6949720 | 90301599 | N | B MANUFACTURING SITE EVALUATION SAMPLES RECEIVED OPEN CASSETTE ANALYSIS AND RESULTS THERE ARE NO PREVIOUS COMPLAINTS OF THE SAME CODEBATCH MANUFACTURED AND DISTRIBUTED IN THE MARKET B UNITS OF THIS CODE BATCH THERE ARE NO UNITS IN STOCK WE HAVE RECEIVED ONE OPEN AND USED CASSETTE TESTED THE KNOT PULL TENSILE STRENGTH OF THE SAMPLE RECEIVED AND THE RESULTS FULFILL THE OEM REQUIREMENTS FINAL CONCLUSION ALTHOUGH THE RESULTS OF THE SAMPLE RECEIVED FULFILL THE OEM SPECIFICATIONS NOTE OF THIS INCIDENCE IN ORDER TO ASSESS IF NEW OR ADDITIONAL ACTIONS ARE NEEDED CORRECTIVEPREVENTIVE ACTIONS ACCORDING TO OUR INTERNAL PROCEDURES THERE IS NO NEED TO ESTABLISH CORRECTIVE OR PREVENTIVE ACTIONS NEVERTHELESS THIS COMPLAINT IS RECORDED FOR TRENDING ANALYSIS TO ASSESS IF ACTIONS ARE NEEDED IN THE FUTURE | Results Pending Completion of Investigation | Medical Evaluation result code
6948277 | 90303294 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948277 | 90303294 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6950300 | 90303484 | N | DEVICE EVALUATION SUMMARY THE EVERFLEX ENTRUST STENT WAS INSPECTED THE DISTAL END OF THE CATHETER SHOWED A CURVE BUT NO CREASES AT THE DISTAL END KINKS TO THE CATHETER SHAFT WITH A CREASE WERE NOTED AT CM AND CM FROM THE DISTAL TIP IT WAS DISCOVERED THE BLUE ISOLATION SHEATH DETACHED FROM THE HANDLE ASSEMBLY THE STRAIN RELIEF WAS CUT FROM THE CATHETER AND NOTED THE CLEAR CONNECTOR WITH TRACES OF DRIED ADHESIVE AROUND THE ENTIRE DIAMETER OF THE CONNECTOR IT SHOULD BE NOTED THE STENT HAS NOT BEEN DEPLOYED AND IS CURRENTLY LOADED THE CATHETER TRACES OF DRIED BLOOD WAS NOTED WITHIN THE CATHETER AT THE PROXIMAL AREA OF THE STENT THE ROTATING WHEEL ON THE HANDLE ASSEMBLY WAS TURNED THE STENT DID NOT DEPLOY RESISTANCE WAS ENCOUNTERED | Assembly Problem Identified | Medical Evaluation result code
6950300 | 90303484 | N | DEVICE EVALUATION SUMMARY THE EVERFLEX ENTRUST STENT WAS INSPECTED THE DISTAL END OF THE CATHETER SHOWED A CURVE BUT NO CREASES AT THE DISTAL END KINKS TO THE CATHETER SHAFT WITH A CREASE WERE NOTED AT CM AND CM FROM THE DISTAL TIP IT WAS DISCOVERED THE BLUE ISOLATION SHEATH DETACHED FROM THE HANDLE ASSEMBLY THE STRAIN RELIEF WAS CUT FROM THE CATHETER AND NOTED THE CLEAR CONNECTOR WITH TRACES OF DRIED ADHESIVE AROUND THE ENTIRE DIAMETER OF THE CONNECTOR IT SHOULD BE NOTED THE STENT HAS NOT BEEN DEPLOYED AND IS CURRENTLY LOADED THE CATHETER TRACES OF DRIED BLOOD WAS NOTED WITHIN THE CATHETER AT THE PROXIMAL AREA OF THE STENT THE ROTATING WHEEL ON THE HANDLE ASSEMBLY WAS TURNED THE STENT DID NOT DEPLOY RESISTANCE WAS ENCOUNTERED | Activation, Positioning or Separation Problem | Device problem code
6950300 | 90303484 | N | DEVICE EVALUATION SUMMARY THE EVERFLEX ENTRUST STENT WAS INSPECTED THE DISTAL END OF THE CATHETER SHOWED A CURVE BUT NO CREASES AT THE DISTAL END KINKS TO THE CATHETER SHAFT WITH A CREASE WERE NOTED AT CM AND CM FROM THE DISTAL TIP IT WAS DISCOVERED THE BLUE ISOLATION SHEATH DETACHED FROM THE HANDLE ASSEMBLY THE STRAIN RELIEF WAS CUT FROM THE CATHETER AND NOTED THE CLEAR CONNECTOR WITH TRACES OF DRIED ADHESIVE AROUND THE ENTIRE DIAMETER OF THE CONNECTOR IT SHOULD BE NOTED THE STENT HAS NOT BEEN DEPLOYED AND IS CURRENTLY LOADED THE CATHETER TRACES OF DRIED BLOOD WAS NOTED WITHIN THE CATHETER AT THE PROXIMAL AREA OF THE STENT THE ROTATING WHEEL ON THE HANDLE ASSEMBLY WAS TURNED THE STENT DID NOT DEPLOY RESISTANCE WAS ENCOUNTERED | Detachment of Device or device Component | Device problem code
6940256 | 90304961 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Defective Device | Device problem code
6940256 | 90304961 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Moisture Damage | Device problem code
6940256 | 90304961 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Storage Problem Identified | Medical Evaluation result code
6940252 | 90305067 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Defective Device | Device problem code
6940252 | 90305067 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Moisture Damage | Device problem code
6940252 | 90305067 | N | THIS INVESTIGATION WAS INITIATED IN RESPONSE TO A CUSTOMER COMPLAINTS REPORTING HOLES IN THE WHITE CARD POUCHES OF SOME CARDS AT THE STITCH SEAM THE ROOT CAUSE OF THE DEFECT WAS LINKED TO THE DESIGN OF THE WHEEL THAT APPLIES THE STITCH SEAL TO THE POUCH ON B STITCH WHEELS WITH A NEW DESIGN WERE INSTALLED THAT RESOLVE THE ISSUE A CUSTOMER COMMUNICATION HAS BEEN CREATED APR AND WAS DISTRIBUTED TO CUSTOMERS WHO RECEIVED IMPACTED CARD LOTS THE CUSTOMER LETTER WILL INCLUDE INSTRUCTIONS FOR INSPECTING POUCHES PRIOR TO USE IN ORDER TO DETECT BREACHES AND REDUCE THE LIKELIHOOD OF USING CARDS EXPOSED TO MOISTURE THE CUSTOMER LETTER WILL ALSO EXPLAIN THE POTENTIAL EFFECTS OF MOISTURE EXPOSURE ON CARD PERFORMANCE | Storage Problem Identified | Medical Evaluation result code
6948279 | 90306541 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948279 | 90306541 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948272 | 90306561 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948272 | 90306561 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948281 | 90306565 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948281 | 90306565 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948234 | 90306567 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948234 | 90306567 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948269 | 90306644 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948269 | 90306644 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948318 | 90306646 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948318 | 90306646 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948223 | 90306673 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948223 | 90306673 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948315 | 90306696 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948315 | 90306696 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948228 | 90306723 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948228 | 90306723 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948236 | 90306774 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948236 | 90306774 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948354 | 90306777 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948354 | 90306777 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948331 | 90306784 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6948331 | 90306784 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | Insufficient Information | Device problem code
6948323 | 90306817 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948323 | 90306817 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948328 | 90306850 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948328 | 90306850 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948338 | 90306892 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948338 | 90306892 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948203 | 90306912 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948203 | 90306912 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948346 | 90306951 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948346 | 90306951 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948283 | 90307036 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948283 | 90307036 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948363 | 90307045 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948363 | 90307045 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6948266 | 90307114 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6948266 | 90307114 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6945763 | 90309833 | N | WE HAVE EXTENSIVELY TESTED VITRECTOME OUR R D DEPARTMENT HAS MADE A CHANGE AS WAS ALREADY INDICATED IN THE SNF THE NEW VITRECTOME WILL SOON BE AVAILABLE OUR CSC DEPARTMENT CAN TELL WHEN THE NEW SERIES IS AVAILABLE SINCE THE BUBBLES ARE FORMED BY THE INSTRUMENT DURING THE SURGERY WHICH CAN LEAD TO MICROBIAL INFECTION THIS COMPLAINT IS CONCLUDED AS REPORTABLE PERFORMANCE VARIES IN AND OUTOFSPECIFICATION THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER DETAILS OF THIS ACTIVITY WERE DISCUSSED WITH CDRH OFFICE OF COMPLIANCE B DURING A TELECONFERENCE ON SINCE THIS IS A RETROSPECTIVE REVIEW THERE IS LIMITED INFORMATION THERE IS NO INFORMATION ON THE STATUS OF THE PATIENT | Insufficient Information | Device problem code
6945763 | 90309833 | N | WE HAVE EXTENSIVELY TESTED VITRECTOME OUR R D DEPARTMENT HAS MADE A CHANGE AS WAS ALREADY INDICATED IN THE SNF THE NEW VITRECTOME WILL SOON BE AVAILABLE OUR CSC DEPARTMENT CAN TELL WHEN THE NEW SERIES IS AVAILABLE SINCE THE BUBBLES ARE FORMED BY THE INSTRUMENT DURING THE SURGERY WHICH CAN LEAD TO MICROBIAL INFECTION THIS COMPLAINT IS CONCLUDED AS REPORTABLE PERFORMANCE VARIES IN AND OUTOFSPECIFICATION THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER DETAILS OF THIS ACTIVITY WERE DISCUSSED WITH CDRH OFFICE OF COMPLIANCE B DURING A TELECONFERENCE ON SINCE THIS IS A RETROSPECTIVE REVIEW THERE IS LIMITED INFORMATION THERE IS NO INFORMATION ON THE STATUS OF THE PATIENT | Assembly Problem Identified | Medical Evaluation result code
6936482 | 90311169 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS TEST WELL IMAGES IN QUESTION WHICH SHOWED THEY WERE VISUALLY NEGATIVE THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON OCT WHICH PERFORMED AS EXPECTED THE IMMUCOR LABORATORY ALSO TESTED RETURNED BLOOD SAMPLES FROM THE CUSTOMER SITE ON OCT WHICH YIELDED UNEXPECTED NEGATIVE OUTCOMES AND THEREFORE THE CUSTOMERS BLOOD SAMPLE TESTING OBSERVATIONS WERE REPRODUCED AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND DETERMINED THAT THE INSTRUMENT WAS OPERATING AS EXPECTED | Connection Problem | Device problem code
6936482 | 90311169 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS TEST WELL IMAGES IN QUESTION WHICH SHOWED THEY WERE VISUALLY NEGATIVE THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON OCT WHICH PERFORMED AS EXPECTED THE IMMUCOR LABORATORY ALSO TESTED RETURNED BLOOD SAMPLES FROM THE CUSTOMER SITE ON OCT WHICH YIELDED UNEXPECTED NEGATIVE OUTCOMES AND THEREFORE THE CUSTOMERS BLOOD SAMPLE TESTING OBSERVATIONS WERE REPRODUCED AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON SEP TO ASSESS THE TESTING INSTRUMENT IN QUESTION AND DETERMINED THAT THE INSTRUMENT WAS OPERATING AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
6949649 | 90312173 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS THE DISTRIBUTOR PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE MYLAR FLAP OF THE EXPIRATORY FLOW SENSOR WAS STUCK TO THE TOP OF THE FLOW SENSOR HOUSING A QUOTE WAS ISSUED BY THE DISTRIBUTOR AND NOT ACCEPTED AT THIS TIME | Mechanical Jam | Device problem code
6949649 | 90312173 | N | NO REPORT OF PATIENT INVOLVEMENT THE INITIAL REPORTER IS LOCATED OUTSIDE THE US AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS THE DISTRIBUTOR PERFORMED A CHECKOUT OF THE EQUIPMENT AND CONFIRMED THE MYLAR FLAP OF THE EXPIRATORY FLOW SENSOR WAS STUCK TO THE TOP OF THE FLOW SENSOR HOUSING A QUOTE WAS ISSUED BY THE DISTRIBUTOR AND NOT ACCEPTED AT THIS TIME | Mechanical Problem Identified | Medical Evaluation result code
6944248 | 90312245 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON B WHICH APPEARED VISUALLY NEGATIVE FOR THE B BLOOD SAMPLE WHEN RUN ON BOTH B THE INSTRUMENT TEST WELLS FOR THE OTHER BLOOD SAMPLE FROM B WERE VISUALLY POSITIVE | Connection Problem | Device problem code
6944248 | 90312245 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON B WHICH APPEARED VISUALLY NEGATIVE FOR THE B BLOOD SAMPLE WHEN RUN ON BOTH B THE INSTRUMENT TEST WELLS FOR THE OTHER BLOOD SAMPLE FROM B WERE VISUALLY POSITIVE | Interoperability Problem Identified | Medical Evaluation result code
6979516 | 90320756 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6979613 | 90322246 | D | IT WAS REPORTED THAT THE PATIENT HAD ARRHYTHMIA AN UNSPECIFIED ANGIOJET WAS SELECTED FOR A THROMBECTOMY PROCEDURE IN THE INFERIOR VENA CAVA THE ANGIOJET WAS USED IN A NONBSC FILTER DURING THE PROCEDURE THE PATIENT BRADY DOWN AND HAD SOME ARRHYTHMIA AFTER STEPPING OFF OF THE ANGIOJET PEDAL THE PATIENT WOULD GO BACK TO NORMAL THERE WERE NO DEVICE ISSUES THE PROCEDURE WAS COMPLETED WITH THIS DEVICE AND THE PATIENT WAS FINE AND DISCHARGED | Adverse Event Without Identified Device or Use Problem | Device problem code
6979613 | 90322246 | D | IT WAS REPORTED THAT THE PATIENT HAD ARRHYTHMIA AN UNSPECIFIED ANGIOJET WAS SELECTED FOR A THROMBECTOMY PROCEDURE IN THE INFERIOR VENA CAVA THE ANGIOJET WAS USED IN A NONBSC FILTER DURING THE PROCEDURE THE PATIENT BRADY DOWN AND HAD SOME ARRHYTHMIA AFTER STEPPING OFF OF THE ANGIOJET PEDAL THE PATIENT WOULD GO BACK TO NORMAL THERE WERE NO DEVICE ISSUES THE PROCEDURE WAS COMPLETED WITH THIS DEVICE AND THE PATIENT WAS FINE AND DISCHARGED | No Consequences Or Impact To Patient | Patient problem code
6979646 | 90323382 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6980293 | 90340266 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT THE PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST NO DETAILS SURROUNDING THE PATIENTS DEATH ARE KNOWN A REVIEW OF DEVICE DOWNLOAD INFORMATION REVEALS THAT THE LIFEVEST DETECTED FOUR ARRHYTHMIAS AT AND RESPECTIVELY THE PATIENTS CONTINUOUS ECG RECORDING SHOWS THAT BEGINNING AT THE PATIENT WAS IN SINUS RHYTHM AT BPM THE PATIENTS RHYTHM THEN TRANSITIONED TO VF THE PATIENT REMAINED IN VF UNTIL THE DEVICE WAS SHUT DOWN AT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE DETECTIONS WHICH DELAYED THE INITIATION OF THE TREATMENT SEQUENCE IT IS UNKNOWN WHO PRESSED THE RESPONSE BUTTONS THE LIFEVEST DETECTED THE VF THE DEVICE EXITED ARRHYTHMIA DETECTIONS DUE TO VARYING AMPLITUDE ECG SIGNAL THE FOURTH DETECTION WAS INTERRUPTED BY THE DEVICE SHUTDOWN IT IS UNKNOWN WHO SHUT DOWN THE DEVICE | Death | Patient problem code
6980293 | 90340266 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT THE PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST NO DETAILS SURROUNDING THE PATIENTS DEATH ARE KNOWN A REVIEW OF DEVICE DOWNLOAD INFORMATION REVEALS THAT THE LIFEVEST DETECTED FOUR ARRHYTHMIAS AT AND RESPECTIVELY THE PATIENTS CONTINUOUS ECG RECORDING SHOWS THAT BEGINNING AT THE PATIENT WAS IN SINUS RHYTHM AT BPM THE PATIENTS RHYTHM THEN TRANSITIONED TO VF THE PATIENT REMAINED IN VF UNTIL THE DEVICE WAS SHUT DOWN AT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE DETECTIONS WHICH DELAYED THE INITIATION OF THE TREATMENT SEQUENCE IT IS UNKNOWN WHO PRESSED THE RESPONSE BUTTONS THE LIFEVEST DETECTED THE VF THE DEVICE EXITED ARRHYTHMIA DETECTIONS DUE TO VARYING AMPLITUDE ECG SIGNAL THE FOURTH DETECTION WAS INTERRUPTED BY THE DEVICE SHUTDOWN IT IS UNKNOWN WHO SHUT DOWN THE DEVICE | Arrhythmia | Patient problem code
6980293 | 90340266 | D | A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT THE PATIENT PASSED AWAY ON B WHILE WEARING THE LIFEVEST NO DETAILS SURROUNDING THE PATIENTS DEATH ARE KNOWN A REVIEW OF DEVICE DOWNLOAD INFORMATION REVEALS THAT THE LIFEVEST DETECTED FOUR ARRHYTHMIAS AT AND RESPECTIVELY THE PATIENTS CONTINUOUS ECG RECORDING SHOWS THAT BEGINNING AT THE PATIENT WAS IN SINUS RHYTHM AT BPM THE PATIENTS RHYTHM THEN TRANSITIONED TO VF THE PATIENT REMAINED IN VF UNTIL THE DEVICE WAS SHUT DOWN AT THE RESPONSE BUTTONS WERE PRESSED INTERMITTENTLY THROUGHOUT THE DETECTIONS WHICH DELAYED THE INITIATION OF THE TREATMENT SEQUENCE IT IS UNKNOWN WHO PRESSED THE RESPONSE BUTTONS THE LIFEVEST DETECTED THE VF THE DEVICE EXITED ARRHYTHMIA DETECTIONS DUE TO VARYING AMPLITUDE ECG SIGNAL THE FOURTH DETECTION WAS INTERRUPTED BY THE DEVICE SHUTDOWN IT IS UNKNOWN WHO SHUT DOWN THE DEVICE | Component or Accessory Incompatibility | Device problem code
6952349 | 90351392 | N | THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6952349 | 90351392 | N | THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6982194 | 90386033 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER WAS IN EMERGENCY ROOM AS A RESULT OF HIGH BLOOD GLUCOSE THE CUSTOMERS MOTHER REPORTED THAT THEY TREATED THE HIGH BLOOD GLUCOSE WITH INSULIN PUMP CUSTOMER WAS WEARING THE INSULIN PUMP AT TIME OF HOSPITALIZATION THE DRIVE SUPPORT CAP WAS NORMAL THE CUSTOMERS MOTHER STATED THAT THEY WERE NO AIR BUBBLES THE CUSTOMERS MOTHER DECLINED TO COMPLETE TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Adverse Event Without Identified Device or Use Problem | Device problem code
6982194 | 90386033 | D | THE CUSTOMERS MOTHER REPORTED VIA PHONE CALL THAT WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMERS BLOOD GLUCOSE WAS MGDL AT THE TIME OF CALL THE CUSTOMER WAS IN EMERGENCY ROOM AS A RESULT OF HIGH BLOOD GLUCOSE THE CUSTOMERS MOTHER REPORTED THAT THEY TREATED THE HIGH BLOOD GLUCOSE WITH INSULIN PUMP CUSTOMER WAS WEARING THE INSULIN PUMP AT TIME OF HOSPITALIZATION THE DRIVE SUPPORT CAP WAS NORMAL THE CUSTOMERS MOTHER STATED THAT THEY WERE NO AIR BUBBLES THE CUSTOMERS MOTHER DECLINED TO COMPLETE TROUBLESHOOTING THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
6981521 | 90386105 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER REPORTED SYMPTOMS SUCH AS NAUSEA AND VOMITING OF THEIR BLOOD GLUCOSE LEVELS THE CUSTOMER TREATED WITH THE INJECTIONS CUSTOMERS BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE INCIDENT CUSTOMERS CURRENT BLOOD GLUCOSE VALUE WAS MGDL THE CUSTOMER REPORTED THEY WERE IN HOSPITAL TWICE THIS YEAR FOR BLOOD GLUCOSE LEVEL OVER MGDL SHE WAS ADMITTED WITH MGDL CUSTOMER GOT HOME YESTERDAY AND WENT IN LAST FRIDAY OR SATURDAY NIGHT CALLER DIDNT REPORT FIRST EVENT BUT IT WAS SIMILAR TO THIS ONE THE FIRST TIME CUSTOMER REPORTED THAT THEY WERE NOT FEELING WELL SO HUSBAND CHECKED BLOOD GLUCOSE AND IT WAS OVER MGDL SHE WAS THROWING UP BLACK SO HE CALLED AND WAS TAKEN TO HOSPITAL SHE WAS SEPTIC AND CUSTOMER WAS VERY ILL SHE WAS HOSPITALIZED FOR ALMOST DAYS THEN THIS TIME SHE STARTED TALKING FUNNY HUSBAND CHECKED BLOOD GLUCOSE IT READ HIGH SO HE TOOK HER TO EMERGENCY ROOM AND SHE WAS ADMITTED AND TAKEN OFF PUMP THEY TRIED TO PUT HER BACK ON PUMP BUT SHE WAS TO CONFUSED SO SHE WAS TAKEN BACK OFF AND PUT BACK ON IT DAY SHE WAS DISCHARGED TROUBLESHOOTING WAS PERFORMED THE CUSTOMER HAD VISITED TO EMERGENCY ROOM AND WAS HOSPITALIZED ON B THE BLOOD GLUCOSE VALUE WHEN ADMITTED TO HOSPITAL WAS MGDL THIS ONE IT WAS MGDL THE CUSTOMER COMPLAINED ABOUT HYPERGLYCEMIA THE CUSTOMER CAUSE OF HOSPITALIZATION WAS T WAS POSSIBLY DUE TO HEART ATTACK FIRST EVENT UNKNOWN SHE HAD TO GET TRANSFUSION WITH THIS EVENT CUSTOMER WEARING THE PUMP AT TIME OF HOSPITALIZATION THE PRODUCT WAS NOT RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
6981521 | 90386105 | D | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE LEVEL OF MGDL THE CUSTOMER REPORTED SYMPTOMS SUCH AS NAUSEA AND VOMITING OF THEIR BLOOD GLUCOSE LEVELS THE CUSTOMER TREATED WITH THE INJECTIONS CUSTOMERS BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE INCIDENT CUSTOMERS CURRENT BLOOD GLUCOSE VALUE WAS MGDL THE CUSTOMER REPORTED THEY WERE IN HOSPITAL TWICE THIS YEAR FOR BLOOD GLUCOSE LEVEL OVER MGDL SHE WAS ADMITTED WITH MGDL CUSTOMER GOT HOME YESTERDAY AND WENT IN LAST FRIDAY OR SATURDAY NIGHT CALLER DIDNT REPORT FIRST EVENT BUT IT WAS SIMILAR TO THIS ONE THE FIRST TIME CUSTOMER REPORTED THAT THEY WERE NOT FEELING WELL SO HUSBAND CHECKED BLOOD GLUCOSE AND IT WAS OVER MGDL SHE WAS THROWING UP BLACK SO HE CALLED AND WAS TAKEN TO HOSPITAL SHE WAS SEPTIC AND CUSTOMER WAS VERY ILL SHE WAS HOSPITALIZED FOR ALMOST DAYS THEN THIS TIME SHE STARTED TALKING FUNNY HUSBAND CHECKED BLOOD GLUCOSE IT READ HIGH SO HE TOOK HER TO EMERGENCY ROOM AND SHE WAS ADMITTED AND TAKEN OFF PUMP THEY TRIED TO PUT HER BACK ON PUMP BUT SHE WAS TO CONFUSED SO SHE WAS TAKEN BACK OFF AND PUT BACK ON IT DAY SHE WAS DISCHARGED TROUBLESHOOTING WAS PERFORMED THE CUSTOMER HAD VISITED TO EMERGENCY ROOM AND WAS HOSPITALIZED ON B THE BLOOD GLUCOSE VALUE WHEN ADMITTED TO HOSPITAL WAS MGDL THIS ONE IT WAS MGDL THE CUSTOMER COMPLAINED ABOUT HYPERGLYCEMIA THE CUSTOMER CAUSE OF HOSPITALIZATION WAS T WAS POSSIBLY DUE TO HEART ATTACK FIRST EVENT UNKNOWN SHE HAD TO GET TRANSFUSION WITH THIS EVENT CUSTOMER WEARING THE PUMP AT TIME OF HOSPITALIZATION THE PRODUCT WAS NOT RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
6982762 | 90390300 | N | NO INFORMATION ABOUT RETURN OF DEVICE | Insufficient Information | Device problem code
6981450 | 90392754 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF ANGIOEDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS THE FOLLOWING REPORTED ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLBELLA XC FOR LIP AUGMENTATION OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THESE ADVERSE EVENTS WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION INFECTION PARESTHESIA HERPES MIGRATION ANGIOEDEMA AND NECROSIS | Migration or Expulsion of Device | Device problem code
6981450 | 90392754 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF ANGIOEDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING PRECAUTIONS PATIENTS MAY EXPERIENCE LATE ONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLBELLA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS THE FOLLOWING REPORTED ADVERSE EVENTS WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLBELLA XC FOR LIP AUGMENTATION OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THESE ADVERSE EVENTS WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION INFECTION PARESTHESIA HERPES MIGRATION ANGIOEDEMA AND NECROSIS | Device Migration | Medical Evaluation result code
6947388 | 90397814 | N | B THE FIELD SERVICE REPRESENTATIVE FSR VERIFIED THAT THE WATER IN THE HEATER COOLER WAS NOT CLOUDY OR DISCOLORED THE FSR CHECKED FOR DRIPPING WATER AND NONE WAS NOTED THE INTERIOR OF THE UNIT HAD NO LEAKS ALL MODES OF OPERATION WERE WORKING PROPERLY THE FSR SUSPECTED THE HIGH HUMIDITY AND CONDENSATION WAS LIKELY THE CAUSE OF THE WATER DRIP THE UNIT OPERATED TO THE MANUFACTURERS SPECIFICATIONS IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY | Material Discolored | Device problem code
6947388 | 90397814 | N | B THE FIELD SERVICE REPRESENTATIVE FSR VERIFIED THAT THE WATER IN THE HEATER COOLER WAS NOT CLOUDY OR DISCOLORED THE FSR CHECKED FOR DRIPPING WATER AND NONE WAS NOTED THE INTERIOR OF THE UNIT HAD NO LEAKS ALL MODES OF OPERATION WERE WORKING PROPERLY THE FSR SUSPECTED THE HIGH HUMIDITY AND CONDENSATION WAS LIKELY THE CAUSE OF THE WATER DRIP THE UNIT OPERATED TO THE MANUFACTURERS SPECIFICATIONS IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY | Leak / Splash | Device problem code
6947388 | 90397814 | N | B THE FIELD SERVICE REPRESENTATIVE FSR VERIFIED THAT THE WATER IN THE HEATER COOLER WAS NOT CLOUDY OR DISCOLORED THE FSR CHECKED FOR DRIPPING WATER AND NONE WAS NOTED THE INTERIOR OF THE UNIT HAD NO LEAKS ALL MODES OF OPERATION WERE WORKING PROPERLY THE FSR SUSPECTED THE HIGH HUMIDITY AND CONDENSATION WAS LIKELY THE CAUSE OF THE WATER DRIP THE UNIT OPERATED TO THE MANUFACTURERS SPECIFICATIONS IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY | Leakage/Seal | Medical Evaluation result code
6949258 | 90402240 | N | B PER SUBSIDIARY THE USER ENTERED A CONCENTRATION VALUE OF K TO RECALIBRATE HOWEVER THE MEASURED VALUE COULD NOT BE ADJUSTED AND AN IMPROPER VALUE WAS INDICATED ON THE SCREEN THE VALUE OF K ON THE SCREEN CONTINUED TO BE GRAY IN COLOR | Calibration Problem | Device problem code
6949258 | 90402240 | N | B PER SUBSIDIARY THE USER ENTERED A CONCENTRATION VALUE OF K TO RECALIBRATE HOWEVER THE MEASURED VALUE COULD NOT BE ADJUSTED AND AN IMPROPER VALUE WAS INDICATED ON THE SCREEN THE VALUE OF K ON THE SCREEN CONTINUED TO BE GRAY IN COLOR | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6983124 | 90409853 | N | B CUSTOMER WAS SENT SETS OF NON SLIP PADS AND THEY WERE DELIVERED ON CUSTOMER STATED THAT THE ISSUE WAS NOT RESOLVED SO ANOTHER SETS OF NON SLIP PADS WERE SENT TO CUSTOMER AND DELIVERED ON CUSTOMER STATED TODAY THAT THIS HAS NOT RESOLVED THE ISSUE CUSTOMER PREFERS THAT SHE NO LONGER BE CONTACTED IN REGARDS TO THIS ISSUE THE CUSTOMER REQUIRES NO FURTHER ASSISTANCE | No Device Problem Found | Medical Evaluation result code
6954022 | 90410919 | N | EVALUATION SUMMARY SHEATH AND STYLETTE WAS RETURNED LOADED IN DEVICE DUE TO THE DEFORMATION OF THE STENT SHEATH COULDNT BE FULLY LOADED OVER STENT THE STENT WAS POSITIONED ON THE BALLOON BETWEEN THE MARKER BANDS AS PER SPECIFICATION DEFORMATION WAS EVIDENT TO THE ST AND ND DISTAL STENT WRAPS WITH STRUTS RAISED AND BUNCHED NO DEFORMATION WAS EVIDENT TO THE DISTAL TIP | Stress Problem Identified | Medical Evaluation result code
6946222 | 90411915 | N | WE HAVE EXAMINED THE INSTRUMENT THE PACKAGE WAS OPENED BY THE CUSTOMER WE HAVE JUST OBSERVED DUST THIS CAN NOT BE DONE DURING PRODUCTION SINCE EVERYTHING IS CLEANED IN THE CLEANROOM WE CAN NOT TRACE THE CAUSE OF THE COMPLAINT THE WHITE PARTICLE MAY LEAD TO FOREIGN BODY REACTION AND IF USED CAN LEAD TO PATIENT HARM HENCE CAN BE CONCLUDED AS REPORTABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER B SINCE THIS IS A RETROSPECTIVE REVIEW THERE IS LIMITED INFORMATION | Environmental Particulates | Device problem code
6946222 | 90411915 | N | WE HAVE EXAMINED THE INSTRUMENT THE PACKAGE WAS OPENED BY THE CUSTOMER WE HAVE JUST OBSERVED DUST THIS CAN NOT BE DONE DURING PRODUCTION SINCE EVERYTHING IS CLEANED IN THE CLEANROOM WE CAN NOT TRACE THE CAUSE OF THE COMPLAINT THE WHITE PARTICLE MAY LEAD TO FOREIGN BODY REACTION AND IF USED CAN LEAD TO PATIENT HARM HENCE CAN BE CONCLUDED AS REPORTABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER B SINCE THIS IS A RETROSPECTIVE REVIEW THERE IS LIMITED INFORMATION | Environment Problem Identified | Medical Evaluation result code
6947145 | 90417030 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6947145 | 90417030 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | No Device Problem Found | Medical Evaluation result code
6957437 | 90419077 | N | ON B THE ACCOUNT REPORTED THAT THEY HAD A CATHETER THAT THE STYLET WAS NOT ABLE TO BE PULLED OUT OF THE DEVICE THE ACCOUNT REPORTED THEY HAD EXPERIENCED THIS ISSUE IN THE PAST BUT PROVIDED NO ADDITIONAL INFORMATION ON B THE ACCOUNT RETURNED THREE CATHETERS ONE OF WHICH WAS MISSING APPROXIMATELY CM OF THE OUTER FEP COATING FROM THE CATHETERS DISTAL END THE INTERNAL SPRING WAS LEFT FULLY INTACT NEEDLE AND PATIENT INFO WERE NOT PROVIDED TO EPIMED BY THE ACCOUNT NO INFO WAS PROVIDED TO EPIMED REGARDING THE LOCATION OF THE MISSING CM OF THE CATHETER COATING THE BATCH HISTORY RECORDS WERE REVIEWED FOR THE REPORTED DEVICE LOT REVEALING NO ABNORMALITIES WITHIN THE LOT EPIMED HAS CONTACTED THE ACCOUNT REQUESTING ADDITIONAL INFORMATION REGARDING THE REPORTED COMPLAINT ONCE THE ACCOUNT REPLIES TO EPIMED AN UPDATED REPORT WILL BE PROVIDED | Detachment of Device or device Component | Device problem code
6957437 | 90419077 | N | ON B THE ACCOUNT REPORTED THAT THEY HAD A CATHETER THAT THE STYLET WAS NOT ABLE TO BE PULLED OUT OF THE DEVICE THE ACCOUNT REPORTED THEY HAD EXPERIENCED THIS ISSUE IN THE PAST BUT PROVIDED NO ADDITIONAL INFORMATION ON B THE ACCOUNT RETURNED THREE CATHETERS ONE OF WHICH WAS MISSING APPROXIMATELY CM OF THE OUTER FEP COATING FROM THE CATHETERS DISTAL END THE INTERNAL SPRING WAS LEFT FULLY INTACT NEEDLE AND PATIENT INFO WERE NOT PROVIDED TO EPIMED BY THE ACCOUNT NO INFO WAS PROVIDED TO EPIMED REGARDING THE LOCATION OF THE MISSING CM OF THE CATHETER COATING THE BATCH HISTORY RECORDS WERE REVIEWED FOR THE REPORTED DEVICE LOT REVEALING NO ABNORMALITIES WITHIN THE LOT EPIMED HAS CONTACTED THE ACCOUNT REQUESTING ADDITIONAL INFORMATION REGARDING THE REPORTED COMPLAINT ONCE THE ACCOUNT REPLIES TO EPIMED AN UPDATED REPORT WILL BE PROVIDED | Component Missing | Device problem code
6957437 | 90419077 | N | ON B THE ACCOUNT REPORTED THAT THEY HAD A CATHETER THAT THE STYLET WAS NOT ABLE TO BE PULLED OUT OF THE DEVICE THE ACCOUNT REPORTED THEY HAD EXPERIENCED THIS ISSUE IN THE PAST BUT PROVIDED NO ADDITIONAL INFORMATION ON B THE ACCOUNT RETURNED THREE CATHETERS ONE OF WHICH WAS MISSING APPROXIMATELY CM OF THE OUTER FEP COATING FROM THE CATHETERS DISTAL END THE INTERNAL SPRING WAS LEFT FULLY INTACT NEEDLE AND PATIENT INFO WERE NOT PROVIDED TO EPIMED BY THE ACCOUNT NO INFO WAS PROVIDED TO EPIMED REGARDING THE LOCATION OF THE MISSING CM OF THE CATHETER COATING THE BATCH HISTORY RECORDS WERE REVIEWED FOR THE REPORTED DEVICE LOT REVEALING NO ABNORMALITIES WITHIN THE LOT EPIMED HAS CONTACTED THE ACCOUNT REQUESTING ADDITIONAL INFORMATION REGARDING THE REPORTED COMPLAINT ONCE THE ACCOUNT REPLIES TO EPIMED AN UPDATED REPORT WILL BE PROVIDED | Missing or Inadequate Safety Measures | Medical Evaluation result code
6983454 | 90422674 | N | B ALMA LASERS HAS REQUESTED THE CUSTOMER TO SEND THE SUSPECTED DEVICE FOR EVALUATION AND HAS ALSO PROVIDED THEM THE UPS RETURN LABEL DEVICE EVALUATION WILL BE PERFORMED ONCE THE DEVICE IS RETURNED TO ALMA LASERS HOWEVER AN ASSESSMENT OF WORLDWIDE ADVERSE INCIDENTS FOR THE SUBSTANTIALLY SIMILAR DEVICES WAS PERFORMED TO DETERMINE WHETHER THE ALLEGED INJURY IS ATTRIBUTABLE TO THE DEVICE NO OTHER SIMILAR INCIDENTS WERE NOTED AND THE COMPANY BELIEVES THIS IS A SINGULAR INCIDENT IN ALL GOOD FAITH EFFORTS ALMA LASERS IS REPORTING THIS INCIDENT TO THE FDA SHOULD ANY ADDITIONAL INFORMATION BECOMES AVAILABLE ALMA LASERS LTD WILL FILE A FOLLOWUP REPORT WITHIN TIME LINES PUBLISHED BY THE FDA BASED ON THE INJURY REPORT TREATMENT PARAMETERS WERE EVALUATED AND DETERMINED TO BE WITHIN RECOMMENDED RANGEUPPER LIMIT HOWEVER IT WAS NOTED THAT THE PATIENT UNDERWENT TREATMENTS WHICH IS EXCESSIVE COMPARED TO ALMA LASERS RECOMMENDATION OF TO TREATMENT SESSIONS THIS EXCESSIVE NUMBER OF TREATMENTS MAY EXPLAIN THE ROOT CAUSE OF THIS ADVERSE REACTION THOUGH THERE IS NO INFORMATION AVAILABLE ON HOW THE PROCEDURE WAS DONE BUT IT IS ALSO POSSIBLE THAT A SMALL SECTION WAS OVER TREATED LEADING TO EXCESSIVE ENERGY ACCUMULATION THAT COULD HAVE REACHED THE DEEPER SKIN LAYER A LETTER TO THE FACILITY WILL BE SENT TO EMPHASIZE ON FOLLOWING THE MANUFACTURERS RECOMMENDATIONS ON B THE CLINICAL EXPERT CONTACTED BY ALMA LASERS DIAGNOSED THIS DENT AS POSSIBLE LIPOATROPHY BASED ON AVAILABLE PHOTOGRAPHS THIS IS PERMANENT AND MAY REQUIRE DERMAL FILLER ALMA LASERS ATTEMPTED MULTIPLE FOLLOWUPS ON THE PATIENTS CONDITION AS PER THE FOLLOWUP PERFORMED ON B THE CUSTOMER REPORTED THAT THE PATIENT IS GETTING BETTER HOWEVER NO ADDITIONAL PHOTOGRAPHS WERE MADE AVAILABLE FOR AN EXPERTS ASSESSMENT IN LACK OF ADDITIONAL PHOTOS EVEN WHEN THE PATIENT IS GETTING BETTER ALMA LASERS CANNOT DETERMINE IF THE INJURY IS PERMANENT OR NOT THE PATIENT IS SCHEDULED TO RETURN TO THE CUSTOMERS FACILITY NEXT WEEK ALMA LASERS HAS REQUESTED ADDITIONAL PHOTOGRAPHS FOR THE PURPOSE OF FURTHER MEDICAL ASSESSMENT DURING THIS VISIT IN ALL GOOD FAITH EFFORTS ALMA LASERS IS REPORTING THIS INCIDENT TO THE FDA SHOULD ANY ADDITIONAL INFORMATION BECOMES AVAILABLE ALMA LASERS LTD WILL FILE A FOLLOWUP REPORT WITHIN TIME LINES PUBLISHED BY THE FDA | Dent in Material | Device problem code
6983454 | 90422674 | N | B ALMA LASERS HAS REQUESTED THE CUSTOMER TO SEND THE SUSPECTED DEVICE FOR EVALUATION AND HAS ALSO PROVIDED THEM THE UPS RETURN LABEL DEVICE EVALUATION WILL BE PERFORMED ONCE THE DEVICE IS RETURNED TO ALMA LASERS HOWEVER AN ASSESSMENT OF WORLDWIDE ADVERSE INCIDENTS FOR THE SUBSTANTIALLY SIMILAR DEVICES WAS PERFORMED TO DETERMINE WHETHER THE ALLEGED INJURY IS ATTRIBUTABLE TO THE DEVICE NO OTHER SIMILAR INCIDENTS WERE NOTED AND THE COMPANY BELIEVES THIS IS A SINGULAR INCIDENT IN ALL GOOD FAITH EFFORTS ALMA LASERS IS REPORTING THIS INCIDENT TO THE FDA SHOULD ANY ADDITIONAL INFORMATION BECOMES AVAILABLE ALMA LASERS LTD WILL FILE A FOLLOWUP REPORT WITHIN TIME LINES PUBLISHED BY THE FDA BASED ON THE INJURY REPORT TREATMENT PARAMETERS WERE EVALUATED AND DETERMINED TO BE WITHIN RECOMMENDED RANGEUPPER LIMIT HOWEVER IT WAS NOTED THAT THE PATIENT UNDERWENT TREATMENTS WHICH IS EXCESSIVE COMPARED TO ALMA LASERS RECOMMENDATION OF TO TREATMENT SESSIONS THIS EXCESSIVE NUMBER OF TREATMENTS MAY EXPLAIN THE ROOT CAUSE OF THIS ADVERSE REACTION THOUGH THERE IS NO INFORMATION AVAILABLE ON HOW THE PROCEDURE WAS DONE BUT IT IS ALSO POSSIBLE THAT A SMALL SECTION WAS OVER TREATED LEADING TO EXCESSIVE ENERGY ACCUMULATION THAT COULD HAVE REACHED THE DEEPER SKIN LAYER A LETTER TO THE FACILITY WILL BE SENT TO EMPHASIZE ON FOLLOWING THE MANUFACTURERS RECOMMENDATIONS ON B THE CLINICAL EXPERT CONTACTED BY ALMA LASERS DIAGNOSED THIS DENT AS POSSIBLE LIPOATROPHY BASED ON AVAILABLE PHOTOGRAPHS THIS IS PERMANENT AND MAY REQUIRE DERMAL FILLER ALMA LASERS ATTEMPTED MULTIPLE FOLLOWUPS ON THE PATIENTS CONDITION AS PER THE FOLLOWUP PERFORMED ON B THE CUSTOMER REPORTED THAT THE PATIENT IS GETTING BETTER HOWEVER NO ADDITIONAL PHOTOGRAPHS WERE MADE AVAILABLE FOR AN EXPERTS ASSESSMENT IN LACK OF ADDITIONAL PHOTOS EVEN WHEN THE PATIENT IS GETTING BETTER ALMA LASERS CANNOT DETERMINE IF THE INJURY IS PERMANENT OR NOT THE PATIENT IS SCHEDULED TO RETURN TO THE CUSTOMERS FACILITY NEXT WEEK ALMA LASERS HAS REQUESTED ADDITIONAL PHOTOGRAPHS FOR THE PURPOSE OF FURTHER MEDICAL ASSESSMENT DURING THIS VISIT IN ALL GOOD FAITH EFFORTS ALMA LASERS IS REPORTING THIS INCIDENT TO THE FDA SHOULD ANY ADDITIONAL INFORMATION BECOMES AVAILABLE ALMA LASERS LTD WILL FILE A FOLLOWUP REPORT WITHIN TIME LINES PUBLISHED BY THE FDA | No Device Problem Found | Medical Evaluation result code
6957261 | 90423949 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6957261 | 90423949 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6944370 | 90424225 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION DURING INVESTIGATION A LIP WAS FELT PRIOR TO ENTERING STRAIN RELIEF WITH A GUIDEWIRE THERE WAS AN OBSTRUCTION UNDER THE STRAIN RELIEF AND THE GUIDEWIRE WOULD NOT PASS THE GUIDEWIRE LUMEN WAS STILL PATENT AIR WOULD PASS THROUGH THE BALLOON INFLATED HELD AIR AND DEFLATED SUCCESSFULLY OBSTRUCTION UNDER THE HUB SEEN DUE TO THE GLUE | Deflation Problem | Device problem code
6944370 | 90424225 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION DURING INVESTIGATION A LIP WAS FELT PRIOR TO ENTERING STRAIN RELIEF WITH A GUIDEWIRE THERE WAS AN OBSTRUCTION UNDER THE STRAIN RELIEF AND THE GUIDEWIRE WOULD NOT PASS THE GUIDEWIRE LUMEN WAS STILL PATENT AIR WOULD PASS THROUGH THE BALLOON INFLATED HELD AIR AND DEFLATED SUCCESSFULLY OBSTRUCTION UNDER THE HUB SEEN DUE TO THE GLUE | Obstruction of Flow | Device problem code
6944370 | 90424225 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION DURING INVESTIGATION A LIP WAS FELT PRIOR TO ENTERING STRAIN RELIEF WITH A GUIDEWIRE THERE WAS AN OBSTRUCTION UNDER THE STRAIN RELIEF AND THE GUIDEWIRE WOULD NOT PASS THE GUIDEWIRE LUMEN WAS STILL PATENT AIR WOULD PASS THROUGH THE BALLOON INFLATED HELD AIR AND DEFLATED SUCCESSFULLY OBSTRUCTION UNDER THE HUB SEEN DUE TO THE GLUE | Operational Problem Identified | Medical Evaluation result code
6945806 | 90424633 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON SEP WHICH APPEARED VISUALLY NEGATIVE FOR THE TESTING ON BOTH THE GALILEO ECHO AND THE GALILEO NEO INSTRUMENTS THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED THE IMMUCOR LABORATORY ALSO TESTED RETURNED PATIENT BLOOD SAMPLE AND RETURNED PRODUCT FROM THE CUSTOMER SITE ON OCT WITH THE FOLLOWING OUTCOMES THE RETURNED BLOOD SAMPLE WAS TESTED AGAINST R RETENTION STRIPS ON BOTH AN IMMUCOR LABORATORY GALILEO ECHO AND A GALILEO NEO BOTH OF WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE THE RETURNED R CUSTOMER TEST STRIPS WERE TESTED AGAINST A KNOWN SOURCE OF ANTIE WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE THE RETURNED BLOOD SAMPLE WAS TESTED AGAINST THE RETURNED R CUSTOMER STRIPS ON BOTH AN IMMUCOR LABORATORY GALILEO ECHO AND A GALILEO NEO BOTH OF WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE | Connection Problem | Device problem code
6945806 | 90424633 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON SEP WHICH APPEARED VISUALLY NEGATIVE FOR THE TESTING ON BOTH THE GALILEO ECHO AND THE GALILEO NEO INSTRUMENTS THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON SEP WHICH PERFORMED AS EXPECTED THE IMMUCOR LABORATORY ALSO TESTED RETURNED PATIENT BLOOD SAMPLE AND RETURNED PRODUCT FROM THE CUSTOMER SITE ON OCT WITH THE FOLLOWING OUTCOMES THE RETURNED BLOOD SAMPLE WAS TESTED AGAINST R RETENTION STRIPS ON BOTH AN IMMUCOR LABORATORY GALILEO ECHO AND A GALILEO NEO BOTH OF WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE THE RETURNED R CUSTOMER TEST STRIPS WERE TESTED AGAINST A KNOWN SOURCE OF ANTIE WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE THE RETURNED BLOOD SAMPLE WAS TESTED AGAINST THE RETURNED R CUSTOMER STRIPS ON BOTH AN IMMUCOR LABORATORY GALILEO ECHO AND A GALILEO NEO BOTH OF WHICH GENERATED THE EXPECTED POSITIVE OUTCOMES CONSISTENT WITH ANTIE | Interoperability Problem Identified | Medical Evaluation result code
6952222 | 90426172 | N | THE MGPS WAS RETURNED FOR FAILURE ANALYSIS AND THE REPORTED FAILURE WAS REPRODUCED THE UNIT WAS INSTALLED AND FAILED TO POWER ON DURING FUNCTIONAL TESTING THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6952222 | 90426172 | N | THE MGPS WAS RETURNED FOR FAILURE ANALYSIS AND THE REPORTED FAILURE WAS REPRODUCED THE UNIT WAS INSTALLED AND FAILED TO POWER ON DURING FUNCTIONAL TESTING THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Energy Storage System Problem | Medical Evaluation result code
6958253 | 90426579 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6958253 | 90426579 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6958261 | 90426662 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6958261 | 90426662 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6958272 | 90426906 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6958272 | 90426906 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6958263 | 90427373 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6958263 | 90427373 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6954580 | 90428047 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6954580 | 90428047 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | Insufficient Information | Device problem code
6983669 | 90437543 | D | IT WAS REPORTED THAT ON YEAR FOLLOW UP CT SCAN A TYPE III ENDOLEAK AT THE OVERLAP OF THE BRANCH GRAFT AND THE STENT WAS NOTED A TYPE II LUMBAR ENDOLEAK WAS ALSO PRESENT WITHIN THE INFRARENAL ANEURYSM SAC | Hemorrhage | Patient problem code
6983669 | 90437543 | D | IT WAS REPORTED THAT ON YEAR FOLLOW UP CT SCAN A TYPE III ENDOLEAK AT THE OVERLAP OF THE BRANCH GRAFT AND THE STENT WAS NOTED A TYPE II LUMBAR ENDOLEAK WAS ALSO PRESENT WITHIN THE INFRARENAL ANEURYSM SAC | Insufficient Information | Device problem code
6983669 | 90437543 | D | IT WAS REPORTED THAT ON YEAR FOLLOW UP CT SCAN A TYPE III ENDOLEAK AT THE OVERLAP OF THE BRANCH GRAFT AND THE STENT WAS NOTED A TYPE II LUMBAR ENDOLEAK WAS ALSO PRESENT WITHIN THE INFRARENAL ANEURYSM SAC | Aneurysm | Patient problem code
6957622 | 90442021 | N | THE DISTRIBUTOR REPORTED BLEEDING DISPLAY AND ALARM THE SERVICE CHECKED THE PUMP WHICH WAS FOUND WORKING ACCORDING TO ITS SPECIFICATIONS AND INTENDED USE NO EVIDENCE OF MALFUNCTION WAS FOUND BY SERVICE WHICH PROCEEDED WITH THE REGULAR MAINTENANCE SERVICE NO PATIENT INVOLVEMENT | No Device Problem Found | Medical Evaluation result code
6955653 | 90444784 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY B | Leak / Splash | Device problem code
6955653 | 90444784 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY B | Leakage/Seal | Medical Evaluation result code
6955653 | 90444784 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY B | Device Displays Incorrect Message | Device Problem Code
6955653 | 90444784 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY B | Results Pending Completion of Investigation | Medical Evaluation result code
6984416 | 90458300 | N | THE CUSTOMERS REPORT OF A HEPARIN INFUSION STOPPING WITHOUT ALARMING OR TRANSITIONING TO KVO WAS CONFIRMED THE PCU EVENT LOG SHOWS THAT AT PM ON B THE PUMP MODULE WAS PROGRAMMED TO INFUSE DRUGID AT A RATE OF MLHR WITH A VTBI OF ML AT AM ON B THE USER PROGRAMMED A DELAY FOR MINUTES WITH NO CALLBACK AFTER PROGRAMMED AT AM THE DELAY COMPLETED AND THE INFUSION RESTARTED THE RESUMED INFUSION COMPLETED AT AM ON B THE ROOT CAUSE OF THE HEPARIN INFUSION STOPPING WITHOUT ALARMING OR TRANSITIONING TO KVO WAS A USER PROGRAMMING ISSUE WHEN A DELAY IS PROGRAMMED WITH NO CALLBACK AFTER THE INFUSION DOES NOT ALARM WHEN THE RESUMED INFUSION COMPLETES BY DESIGN A DELAYED INFUSION DOES NOT TRANSITION TO KVO WHEN THE RESUMED INFUSION COMPLETES AS IT IS INTENDED FOR USE WHEN THERE IS ANOTHER INFUSION RUNNING TO KEEP THE VEIN OPEN | Component or Accessory Incompatibility | Device Problem Code
6984416 | 90458300 | N | THE CUSTOMERS REPORT OF A HEPARIN INFUSION STOPPING WITHOUT ALARMING OR TRANSITIONING TO KVO WAS CONFIRMED THE PCU EVENT LOG SHOWS THAT AT PM ON B THE PUMP MODULE WAS PROGRAMMED TO INFUSE DRUGID AT A RATE OF MLHR WITH A VTBI OF ML AT AM ON B THE USER PROGRAMMED A DELAY FOR MINUTES WITH NO CALLBACK AFTER PROGRAMMED AT AM THE DELAY COMPLETED AND THE INFUSION RESTARTED THE RESUMED INFUSION COMPLETED AT AM ON B THE ROOT CAUSE OF THE HEPARIN INFUSION STOPPING WITHOUT ALARMING OR TRANSITIONING TO KVO WAS A USER PROGRAMMING ISSUE WHEN A DELAY IS PROGRAMMED WITH NO CALLBACK AFTER THE INFUSION DOES NOT ALARM WHEN THE RESUMED INFUSION COMPLETES BY DESIGN A DELAYED INFUSION DOES NOT TRANSITION TO KVO WHEN THE RESUMED INFUSION COMPLETES AS IT IS INTENDED FOR USE WHEN THERE IS ANOTHER INFUSION RUNNING TO KEEP THE VEIN OPEN | Interoperability Problem Identified | Medical Evaluation result code
6970253 | 90461666 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6970253 | 90461666 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6964589 | 90463617 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE MECHANISM DEPLOYED LATER THAN INTENDED DURING ACTIVATION | Insufficient Information | Device problem code
6964589 | 90463617 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE NEEDLE MECHANISM DEPLOYED LATER THAN INTENDED DURING ACTIVATION | No Known Impact Or Consequence To Patient | Patient Problem Code
6946274 | 90466301 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THEY WERE VISUALLY NEGATIVE THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON OCT WHICH PERFORMED AS EXPECTED | Connection Problem | Device problem code
6946274 | 90466301 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS TEST WELL IMAGES IN QUESTION WHICH SHOWED THAT THEY WERE VISUALLY NEGATIVE THE IMMUCOR LABORATORY TESTED RETENTION PRODUCT ON OCT WHICH PERFORMED AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
6960861 | 90467980 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD IT WAS ALSO REPORTED THAT THE NEEDLE DEPLOYED EARLY INDICATING A NEEDLE MECHANISM FAILURE | Mechanical Problem | Device problem code
6960861 | 90467980 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD IT WAS ALSO REPORTED THAT THE NEEDLE DEPLOYED EARLY INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6960824 | 90468386 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6960824 | 90468386 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6947144 | 90470786 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
6947144 | 90470786 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
6958613 | 90479691 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THAT THE BATTERY WOULD FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE THAT THE BATTERY WAS RECEIVED INOPERABLE SUPPLEMENTAL TESTING WAS PERFORMED AND REVEALED THAT A THERMAL CUTOFF FUSE WAS OPEN A THERMAL CUTOFF FUSE INTERRUPTS AN ELECTRICAL CONNECTION WHEN A CERTAIN TEMPERATURE CONDITION IS MET BASED ON THE AVAILABLE INFORMATION A POSSIBLE ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN THERMAL CUTOFF FUSE PREVENTING THE BATTERY FROM CHARGING OR PROVIDING POWER TO A CONTROLLER THIS MOST LIKELY CONTRIBUTED THE BATTERY TO FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE RETURNED BATTERY HOWEVER DID NOT EXHIBIT A RED FLASHING WHEN ANALYZED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Flare or Flash | Device problem code
6958613 | 90479691 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THAT THE BATTERY WOULD FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE THAT THE BATTERY WAS RECEIVED INOPERABLE SUPPLEMENTAL TESTING WAS PERFORMED AND REVEALED THAT A THERMAL CUTOFF FUSE WAS OPEN A THERMAL CUTOFF FUSE INTERRUPTS AN ELECTRICAL CONNECTION WHEN A CERTAIN TEMPERATURE CONDITION IS MET BASED ON THE AVAILABLE INFORMATION A POSSIBLE ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN THERMAL CUTOFF FUSE PREVENTING THE BATTERY FROM CHARGING OR PROVIDING POWER TO A CONTROLLER THIS MOST LIKELY CONTRIBUTED THE BATTERY TO FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE RETURNED BATTERY HOWEVER DID NOT EXHIBIT A RED FLASHING WHEN ANALYZED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Charging Problem | Device problem code
6958613 | 90479691 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THAT THE BATTERY WOULD FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE THAT THE BATTERY WAS RECEIVED INOPERABLE SUPPLEMENTAL TESTING WAS PERFORMED AND REVEALED THAT A THERMAL CUTOFF FUSE WAS OPEN A THERMAL CUTOFF FUSE INTERRUPTS AN ELECTRICAL CONNECTION WHEN A CERTAIN TEMPERATURE CONDITION IS MET BASED ON THE AVAILABLE INFORMATION A POSSIBLE ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN THERMAL CUTOFF FUSE PREVENTING THE BATTERY FROM CHARGING OR PROVIDING POWER TO A CONTROLLER THIS MOST LIKELY CONTRIBUTED THE BATTERY TO FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE RETURNED BATTERY HOWEVER DID NOT EXHIBIT A RED FLASHING WHEN ANALYZED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Connection Problem | Device problem code
6958613 | 90479691 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS IT WAS REPORTED THAT THE BATTERY WOULD FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE BATTERY WAS RETURNED FOR EVALUATION VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE DEVICE PASSED VISUAL EXAMINATION BUT FAILED FUNCTIONAL TESTING DUE THAT THE BATTERY WAS RECEIVED INOPERABLE SUPPLEMENTAL TESTING WAS PERFORMED AND REVEALED THAT A THERMAL CUTOFF FUSE WAS OPEN A THERMAL CUTOFF FUSE INTERRUPTS AN ELECTRICAL CONNECTION WHEN A CERTAIN TEMPERATURE CONDITION IS MET BASED ON THE AVAILABLE INFORMATION A POSSIBLE ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN OPEN THERMAL CUTOFF FUSE PREVENTING THE BATTERY FROM CHARGING OR PROVIDING POWER TO A CONTROLLER THIS MOST LIKELY CONTRIBUTED THE BATTERY TO FLASH RED WHEN CONNECTED TO A BATTERY CHARGER THE RETURNED BATTERY HOWEVER DID NOT EXHIBIT A RED FLASHING WHEN ANALYZED HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Energy Storage System Problem | Medical Evaluation result code
6985906 | 90509543 | N | FOLLOWUP INFORMATION INDICATES THAT THE STENT THAT WAS PLACED OVER THE PERFORATION HAS BEEN REMOVED AND THE PATIENT LEFT THE HOSPITAL | Material Perforation | Device problem code
6985906 | 90509543 | N | FOLLOWUP INFORMATION INDICATES THAT THE STENT THAT WAS PLACED OVER THE PERFORATION HAS BEEN REMOVED AND THE PATIENT LEFT THE HOSPITAL | Stress Problem Identified | Medical Evaluation result code
6986385 | 90509629 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FEELS THAT THEY ARENT EMPTYING THEIR BLADDER ALL THE WAY NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6986385 | 90509629 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT FEELS THAT THEY ARENT EMPTYING THEIR BLADDER ALL THE WAY NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6986349 | 90510467 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN ALLEGING THAT THEIR ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH THE FOLLOWING COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN THE PATIENT REPORTED OBTAINING A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT STATED THAT ON B THEIR DOCTOR STATED THAT THIS READING WAS TOO HIGH COMPARED TO THEIR HBAC READING LIFESCAN DOES NOT CONSIDER THIS TO BE A VALID COMPARISON THE PATIENT MANAGES THEIR DIABETES WITH A COMBINATION OF MEDICATION INCLUDING ACTOS AND METFORMIN THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN BUT STATED THAT THEY HAD NOT MADE ANY CHANGES TO THEIR USUAL DIABETES MANAGEMENT REGIMEN THE PATIENT STATED THAT ON B THEY OBTAINED A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT REPORTED TAKING UNITS OF LANTUS AND UNITS OF NOVALOG BASED ON THE METER READING THE PATIENT STATED THAT LATER THAT DAY THEY DEVELOPED SYMPTOM OF CONFUSION AND SELFTREATED THIS SYMPTOM WITH ORANGE JUICE THE PATIENT STATED THAT THEY DID NOT FEEL BETTER AFTER THIS AND THEY WERE TAKEN TO THE EMERGENCY ROOM BY THEIR NEIGHBOR THE PATIENT REPORTED RECEIVING HCP TREATMENT OF IV FLUIDS THE PATIENT REPORTED TESTING THEIR BLOOD GLUCOSE AT THE HOSPITAL THEY DID NOT REPORT ANY EXACT READING BUT STATED THAT THEY WERE NORMAL THE PATIENT STATED THAT THEY WERE DISCHARGED FROM HOSPITAL ON THE SAME DAY AND WERE ADVISED TO ONLY TAKE UNITS OF LANTUS AND UNITS OF NOVALOG AT THE TIME OF TROUBLESHOOTING THE CCA VERIFIED THAT THE TEST STRIPS WERE STORED CORRECTLY PER OWNERS BOOKLET RECOMMENDATION REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED ISSUE BEGAN | Injury | Patient problem code
6986349 | 90510467 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN ALLEGING THAT THEIR ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH THE FOLLOWING COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN THE PATIENT REPORTED OBTAINING A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT STATED THAT ON B THEIR DOCTOR STATED THAT THIS READING WAS TOO HIGH COMPARED TO THEIR HBAC READING LIFESCAN DOES NOT CONSIDER THIS TO BE A VALID COMPARISON THE PATIENT MANAGES THEIR DIABETES WITH A COMBINATION OF MEDICATION INCLUDING ACTOS AND METFORMIN THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN BUT STATED THAT THEY HAD NOT MADE ANY CHANGES TO THEIR USUAL DIABETES MANAGEMENT REGIMEN THE PATIENT STATED THAT ON B THEY OBTAINED A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT REPORTED TAKING UNITS OF LANTUS AND UNITS OF NOVALOG BASED ON THE METER READING THE PATIENT STATED THAT LATER THAT DAY THEY DEVELOPED SYMPTOM OF CONFUSION AND SELFTREATED THIS SYMPTOM WITH ORANGE JUICE THE PATIENT STATED THAT THEY DID NOT FEEL BETTER AFTER THIS AND THEY WERE TAKEN TO THE EMERGENCY ROOM BY THEIR NEIGHBOR THE PATIENT REPORTED RECEIVING HCP TREATMENT OF IV FLUIDS THE PATIENT REPORTED TESTING THEIR BLOOD GLUCOSE AT THE HOSPITAL THEY DID NOT REPORT ANY EXACT READING BUT STATED THAT THEY WERE NORMAL THE PATIENT STATED THAT THEY WERE DISCHARGED FROM HOSPITAL ON THE SAME DAY AND WERE ADVISED TO ONLY TAKE UNITS OF LANTUS AND UNITS OF NOVALOG AT THE TIME OF TROUBLESHOOTING THE CCA VERIFIED THAT THE TEST STRIPS WERE STORED CORRECTLY PER OWNERS BOOKLET RECOMMENDATION REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED ISSUE BEGAN | Hyperglycemia | Patient Problem Code
6986349 | 90510467 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN ALLEGING THAT THEIR ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH THE FOLLOWING COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN THE PATIENT REPORTED OBTAINING A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT STATED THAT ON B THEIR DOCTOR STATED THAT THIS READING WAS TOO HIGH COMPARED TO THEIR HBAC READING LIFESCAN DOES NOT CONSIDER THIS TO BE A VALID COMPARISON THE PATIENT MANAGES THEIR DIABETES WITH A COMBINATION OF MEDICATION INCLUDING ACTOS AND METFORMIN THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN BUT STATED THAT THEY HAD NOT MADE ANY CHANGES TO THEIR USUAL DIABETES MANAGEMENT REGIMEN THE PATIENT STATED THAT ON B THEY OBTAINED A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT REPORTED TAKING UNITS OF LANTUS AND UNITS OF NOVALOG BASED ON THE METER READING THE PATIENT STATED THAT LATER THAT DAY THEY DEVELOPED SYMPTOM OF CONFUSION AND SELFTREATED THIS SYMPTOM WITH ORANGE JUICE THE PATIENT STATED THAT THEY DID NOT FEEL BETTER AFTER THIS AND THEY WERE TAKEN TO THE EMERGENCY ROOM BY THEIR NEIGHBOR THE PATIENT REPORTED RECEIVING HCP TREATMENT OF IV FLUIDS THE PATIENT REPORTED TESTING THEIR BLOOD GLUCOSE AT THE HOSPITAL THEY DID NOT REPORT ANY EXACT READING BUT STATED THAT THEY WERE NORMAL THE PATIENT STATED THAT THEY WERE DISCHARGED FROM HOSPITAL ON THE SAME DAY AND WERE ADVISED TO ONLY TAKE UNITS OF LANTUS AND UNITS OF NOVALOG AT THE TIME OF TROUBLESHOOTING THE CCA VERIFIED THAT THE TEST STRIPS WERE STORED CORRECTLY PER OWNERS BOOKLET RECOMMENDATION REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED ISSUE BEGAN | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6986349 | 90510467 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN ALLEGING THAT THEIR ONETOUCH ULTRAMINI METER WAS READING INACCURATELY HIGH THE FOLLOWING COMPLAINT WAS CLASSIFIED BASED ON INFORMATION OBTAINED FROM THE CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN THE PATIENT REPORTED OBTAINING A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT STATED THAT ON B THEIR DOCTOR STATED THAT THIS READING WAS TOO HIGH COMPARED TO THEIR HBAC READING LIFESCAN DOES NOT CONSIDER THIS TO BE A VALID COMPARISON THE PATIENT MANAGES THEIR DIABETES WITH A COMBINATION OF MEDICATION INCLUDING ACTOS AND METFORMIN THE PATIENT WAS UNSURE WHEN THE ALLEGED ISSUE BEGAN BUT STATED THAT THEY HAD NOT MADE ANY CHANGES TO THEIR USUAL DIABETES MANAGEMENT REGIMEN THE PATIENT STATED THAT ON B THEY OBTAINED A BLOOD GLUCOSE READING OF MGDL ON THE SUBJECT METER THE PATIENT REPORTED TAKING UNITS OF LANTUS AND UNITS OF NOVALOG BASED ON THE METER READING THE PATIENT STATED THAT LATER THAT DAY THEY DEVELOPED SYMPTOM OF CONFUSION AND SELFTREATED THIS SYMPTOM WITH ORANGE JUICE THE PATIENT STATED THAT THEY DID NOT FEEL BETTER AFTER THIS AND THEY WERE TAKEN TO THE EMERGENCY ROOM BY THEIR NEIGHBOR THE PATIENT REPORTED RECEIVING HCP TREATMENT OF IV FLUIDS THE PATIENT REPORTED TESTING THEIR BLOOD GLUCOSE AT THE HOSPITAL THEY DID NOT REPORT ANY EXACT READING BUT STATED THAT THEY WERE NORMAL THE PATIENT STATED THAT THEY WERE DISCHARGED FROM HOSPITAL ON THE SAME DAY AND WERE ADVISED TO ONLY TAKE UNITS OF LANTUS AND UNITS OF NOVALOG AT THE TIME OF TROUBLESHOOTING THE CCA VERIFIED THAT THE TEST STRIPS WERE STORED CORRECTLY PER OWNERS BOOKLET RECOMMENDATION REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS SUGGESTIVE OF A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED ISSUE BEGAN | Confusion/Disorientation | Patient problem code
6946383 | 90514224 | N | INVESTIGATION SHOWED THAT THE OUTER KNIFE WAS LOOSE WHICH CAUSED THE AIR BUBBLES AND INSUFFICIENT ASPIRATION THIS INCIDENT HAS OCCURRED DURING SURGERY FORMATION OF BUBBLES DURING SURGERY MIGHT LEAD TO PATIENT HARM HENCE THIS COMPLAINT IS CONCLUDED AS REPORTABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER DETAILS OF THIS ACTIVITY WERE DISCUSSED WITH CDRH OFFICE OF COMPLIANCE B DURING A TELECONFERENCE ON THERE IS LIMITED INFORMATION ON THIS INCIDENT AND NO INFORMATION ON THE PATIENT OUTCOME OR IF THEY WERE HARMED | Detachment of Device or device Component | Device problem code
6946383 | 90514224 | N | INVESTIGATION SHOWED THAT THE OUTER KNIFE WAS LOOSE WHICH CAUSED THE AIR BUBBLES AND INSUFFICIENT ASPIRATION THIS INCIDENT HAS OCCURRED DURING SURGERY FORMATION OF BUBBLES DURING SURGERY MIGHT LEAD TO PATIENT HARM HENCE THIS COMPLAINT IS CONCLUDED AS REPORTABLE THIS REPORTABLE EVENT WAS IDENTIFIED DURING A VOLUNTARY RETROSPECTIVE REVIEW OF ALL COMPLAINTS SINCE BY THE MANUFACTURER DETAILS OF THIS ACTIVITY WERE DISCUSSED WITH CDRH OFFICE OF COMPLIANCE B DURING A TELECONFERENCE ON THERE IS LIMITED INFORMATION ON THIS INCIDENT AND NO INFORMATION ON THE PATIENT OUTCOME OR IF THEY WERE HARMED | Fracture Problem | Medical Evaluation result code
6968460 | 90521096 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SELECT ENERGY LEVEL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6968460 | 90521096 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO SELECT ENERGY LEVEL COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Energy Output Problem | Device Problem Code
6961572 | 90529528 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component Incompatible | Device Problem Code
6961572 | 90529528 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6951379 | 90531046 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION A KINK FROM THE PACKAGING COULD BE SEEN AT THE DISTAL END OF THE STRAIN RELIEF ADDITIONAL VISUAL INVESTIGATION AND FUNCTIONAL TESTING FOUND THAT WHEN THE DEVICE WAS NOT IN THE STRAIGHT POSITION A SLIGHT OBSTRUCTION COULD BE FELT ON THE DISTAL SIDE OF THE STRAIN RELIEF BUT A GUIDEWIRE WAS STILL ABLE TO BE PASSED THE BALLOON SUCCESSFULLY INFLATED MAINTAINED VOLUME AND DEFLATED A FAINT WITNESS MARK COULD BE SEEN WHERE THE STRAIN RELIEF INTERACTS WITH THE SHAFT | Deflation Problem | Device problem code
6951379 | 90531046 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION A KINK FROM THE PACKAGING COULD BE SEEN AT THE DISTAL END OF THE STRAIN RELIEF ADDITIONAL VISUAL INVESTIGATION AND FUNCTIONAL TESTING FOUND THAT WHEN THE DEVICE WAS NOT IN THE STRAIGHT POSITION A SLIGHT OBSTRUCTION COULD BE FELT ON THE DISTAL SIDE OF THE STRAIN RELIEF BUT A GUIDEWIRE WAS STILL ABLE TO BE PASSED THE BALLOON SUCCESSFULLY INFLATED MAINTAINED VOLUME AND DEFLATED A FAINT WITNESS MARK COULD BE SEEN WHERE THE STRAIN RELIEF INTERACTS WITH THE SHAFT | Fracture Problem | Medical Evaluation result code
6956398 | 90531579 | N | THE REPORTED MONITOR ARTIFACT ISSUE WAS VERBALLY CONFIRMED ON OCTOBER A DIFFERENT CAMERA SETUP WAS USED AND THERE IS NO ISSUE WITH THE LASER | Signal Artifact | Device problem code
6956398 | 90531579 | N | THE REPORTED MONITOR ARTIFACT ISSUE WAS VERBALLY CONFIRMED ON OCTOBER A DIFFERENT CAMERA SETUP WAS USED AND THERE IS NO ISSUE WITH THE LASER | Image Artifact | Medical Evaluation result code
6974386 | 90531644 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | Positioning Problem | Device problem code
6974386 | 90531644 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6974388 | 90531750 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6974388 | 90531750 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | Positioning Problem | Device problem code
6961266 | 90534460 | N | EVALUATION SUMMARY THE BALLOON FAILED NEGATIVE PREP ON PRESSURISATION OF THE DEVICE A BURST SITE WAS OBSERVED ON THE BALLOONS WORKING LENGTH THE BALLOON FAILED TO MAINTAIN PRESSURE UPON VISUAL INSPECTION OF THE DEVICE THERE WAS A LARGE RADIAL TEAR ON THE BALLOON MATERIAL IMMEDIATELY PROXIMAL TO THE BALLOON DISTAL MARKER BAND THE BALLOON MATERIAL WAS JAGGED AND UNEVEN AT THE TEAR SITE NO DEFORMATION WAS VISIBLE TO THE DISTAL TIP INNER LUMEN PATENCY WAS VERIFIED B | Material Deformation | Device problem code
6961266 | 90534460 | N | EVALUATION SUMMARY THE BALLOON FAILED NEGATIVE PREP ON PRESSURISATION OF THE DEVICE A BURST SITE WAS OBSERVED ON THE BALLOONS WORKING LENGTH THE BALLOON FAILED TO MAINTAIN PRESSURE UPON VISUAL INSPECTION OF THE DEVICE THERE WAS A LARGE RADIAL TEAR ON THE BALLOON MATERIAL IMMEDIATELY PROXIMAL TO THE BALLOON DISTAL MARKER BAND THE BALLOON MATERIAL WAS JAGGED AND UNEVEN AT THE TEAR SITE NO DEFORMATION WAS VISIBLE TO THE DISTAL TIP INNER LUMEN PATENCY WAS VERIFIED B | Break | Device problem code
6961266 | 90534460 | N | EVALUATION SUMMARY THE BALLOON FAILED NEGATIVE PREP ON PRESSURISATION OF THE DEVICE A BURST SITE WAS OBSERVED ON THE BALLOONS WORKING LENGTH THE BALLOON FAILED TO MAINTAIN PRESSURE UPON VISUAL INSPECTION OF THE DEVICE THERE WAS A LARGE RADIAL TEAR ON THE BALLOON MATERIAL IMMEDIATELY PROXIMAL TO THE BALLOON DISTAL MARKER BAND THE BALLOON MATERIAL WAS JAGGED AND UNEVEN AT THE TEAR SITE NO DEFORMATION WAS VISIBLE TO THE DISTAL TIP INNER LUMEN PATENCY WAS VERIFIED B | Burst Container or Vessel | Device problem code
6961266 | 90534460 | N | EVALUATION SUMMARY THE BALLOON FAILED NEGATIVE PREP ON PRESSURISATION OF THE DEVICE A BURST SITE WAS OBSERVED ON THE BALLOONS WORKING LENGTH THE BALLOON FAILED TO MAINTAIN PRESSURE UPON VISUAL INSPECTION OF THE DEVICE THERE WAS A LARGE RADIAL TEAR ON THE BALLOON MATERIAL IMMEDIATELY PROXIMAL TO THE BALLOON DISTAL MARKER BAND THE BALLOON MATERIAL WAS JAGGED AND UNEVEN AT THE TEAR SITE NO DEFORMATION WAS VISIBLE TO THE DISTAL TIP INNER LUMEN PATENCY WAS VERIFIED B | Stress Problem Identified | Medical Evaluation result code
6952414 | 90534692 | N | B AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMEDTHE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE DID ANY PIECE OR PIECES OF THE FIRING KNOB FALL INTO THE PATIENT NO INFORMATION WILL ALL PIECES BE RETURNED WITH THE DEVICE NO INFORMATION | Insufficient Information | Device problem code
6957054 | 90534873 | N | ISI HAS RECEIVED THE DEVICE INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE REPORTED FAILURE MODE VISUAL INSPECTION WAS PERFORMED AND THE UNIT WAS STILL IN GOOD CONDITION THE UNIT WAS INSTALLED INTO PCA SYSTEM AND AT STARTUP THERE WAS NO POWER FROM THE UNIT AND AN ERROR APPEARED THIS REPORT IS BEING REPORTED DUE A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Electrical Power Problem | Device problem code
6957054 | 90534873 | N | ISI HAS RECEIVED THE DEVICE INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS INVESTIGATION WAS ABLE TO REPRODUCE THE REPORTED FAILURE MODE VISUAL INSPECTION WAS PERFORMED AND THE UNIT WAS STILL IN GOOD CONDITION THE UNIT WAS INSTALLED INTO PCA SYSTEM AND AT STARTUP THERE WAS NO POWER FROM THE UNIT AND AN ERROR APPEARED THIS REPORT IS BEING REPORTED DUE A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Energy Storage System Problem | Medical Evaluation result code
6962609 | 90540813 | N | DEVICE ANALYSIS DID NOT DETECT ANY EVIDENCE OF A MANUFACTURING MATERIAL OR DESIGN ISSUE THE RATE OF OCCURRENCE IS WITHIN PREDICTED LEVELS IN THE RISK MANAGEMENT DOCUMENT THERE WERE NO ALLEGATIONS OR EVIDENCE OF PARTICULATES OR POTENTIAL FOR PARTICULATES THIS REPORT IS BEING SUBMITTED SINCE THE DEVICE INVOLVED SHARES THE SAME PRODUCT FAMILY BUT WITH DESIGN DIFFERENCES AS OTHER DEVICES THAT WERE MDR REPORTED | Particulates | Device problem code
6962609 | 90540813 | N | DEVICE ANALYSIS DID NOT DETECT ANY EVIDENCE OF A MANUFACTURING MATERIAL OR DESIGN ISSUE THE RATE OF OCCURRENCE IS WITHIN PREDICTED LEVELS IN THE RISK MANAGEMENT DOCUMENT THERE WERE NO ALLEGATIONS OR EVIDENCE OF PARTICULATES OR POTENTIAL FOR PARTICULATES THIS REPORT IS BEING SUBMITTED SINCE THE DEVICE INVOLVED SHARES THE SAME PRODUCT FAMILY BUT WITH DESIGN DIFFERENCES AS OTHER DEVICES THAT WERE MDR REPORTED | No Findings Available | Medical Evaluation result code
6958302 | 90549397 | N | THE EVENT DESCRIPTION STATES THAT THE HUB SEPARATED FROM THE SHAFT THERE WAS NO LOT NUMBER PROVIDED FOR THIS EVENT OR RETURNED PRODUCT TO EVALUATE THE ROOT CAUSE FOR THIS FAILURE IS UNDETERMINABLE | Fracture Problem | Medical Evaluation result code
6958231 | 90549593 | N | THE EVENT DESCRIPTION STATES THAT THE HUB SEPARATED FROM THE SHAFT THERE WAS NO LOT NUMBER PROVIDED FOR THIS EVENT OR RETURNED PRODUCT TO EVALUATE THE ROOT CAUSE FOR THIS FAILURE IS UNDETERMINABLE | Material Separation | Device problem code
6958231 | 90549593 | N | THE EVENT DESCRIPTION STATES THAT THE HUB SEPARATED FROM THE SHAFT THERE WAS NO LOT NUMBER PROVIDED FOR THIS EVENT OR RETURNED PRODUCT TO EVALUATE THE ROOT CAUSE FOR THIS FAILURE IS UNDETERMINABLE | Fracture Problem | Medical Evaluation result code
6944666 | 90549847 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN INSTRUMENT PERIPUMP AND PROBE WAS SENT TO THE CUSTOMER SITE FROM IMMUCOR FOR AN IMMUCOR FIELD SERVICE ENGINEER FSE TO VISIT THE CUSTOMER SITE AND WORK ON THE INSTRUMENT | Connection Problem | Device problem code
6944666 | 90549847 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON B TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED VISUALLY POSITIVE AN INSTRUMENT PERIPUMP AND PROBE WAS SENT TO THE CUSTOMER SITE FROM IMMUCOR FOR AN IMMUCOR FIELD SERVICE ENGINEER FSE TO VISIT THE CUSTOMER SITE AND WORK ON THE INSTRUMENT | Interoperability Problem Identified | Medical Evaluation result code
6987144 | 90549924 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT MENTIONED THEY HAVE THE URGE TO VOID BUT THEY HAVE NOT BEEN ABLE TO SINCE STARTING THE TRIAL TODAY THEY ALSO MENTIONED THE DEVICE IS UNCOMFORTABLE AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Discomfort | Patient problem code
6987144 | 90549924 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT MENTIONED THEY HAVE THE URGE TO VOID BUT THEY HAVE NOT BEEN ABLE TO SINCE STARTING THE TRIAL TODAY THEY ALSO MENTIONED THE DEVICE IS UNCOMFORTABLE AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6951333 | 90552053 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION A KINK COULD BE SEEN ON THE DISTAL SIDE OF THE STRAIN RELIEF FROM THE PACKAGING RESISTANCE COULD BE FELT DISTAL TO THE STRAIN RELIEF OTHER THAN SLIGHT RESISTANCE DEVICE WAS FULLY PATENT WITH THE GUIDEWIRE THE BALLOON INFLATED MAINTAINED AIR AND DEFLATED SUCCESSFULLY NO WITNESS MARK NOTICED WHEN THE STRAIN RELIEF WAS REMOVED | Deflation Problem | Device problem code
6951333 | 90552053 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION A KINK COULD BE SEEN ON THE DISTAL SIDE OF THE STRAIN RELIEF FROM THE PACKAGING RESISTANCE COULD BE FELT DISTAL TO THE STRAIN RELIEF OTHER THAN SLIGHT RESISTANCE DEVICE WAS FULLY PATENT WITH THE GUIDEWIRE THE BALLOON INFLATED MAINTAINED AIR AND DEFLATED SUCCESSFULLY NO WITNESS MARK NOTICED WHEN THE STRAIN RELIEF WAS REMOVED | No Device Problem Found | Medical Evaluation result code
6951259 | 90552096 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE COMPANION DRIVER EXHIBITED A LEFT PRESSURE INCORRECT ALARM IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE COMPANION DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE EVALUATION WILL BE PROVIDED IN A FOLLOWUP MDR B | False Alarm | Device problem code
6951259 | 90552096 | N | THIS ALLEGED FAILURE MODE POSES A LOW RISK TO THE PATIENT BECAUSE ALTHOUGH THE COMPANION DRIVER EXHIBITED A LEFT PRESSURE INCORRECT ALARM IT CONTINUED TO PERFORM ITS LIFESUSTAINING FUNCTIONS THE COMPANION DRIVER WILL BE RETURNED TO SYNCARDIA FOR EVALUATION THE RESULTS OF THE EVALUATION WILL BE PROVIDED IN A FOLLOWUP MDR B | Mechanical Problem Identified | Medical Evaluation result code
6961609 | 90561677 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961609 | 90561677 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961558 | 90561752 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961558 | 90561752 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987417 | 90561781 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SURGICAL INCISION THAT WAS IRRITATED BY WEARING THE LIFEVEST THE PATIENTS PHYSICIAN ALLOWED THE PATIENT TO REMOVE THE LIFEVEST TO LET HER SKIN HEAL FOLLOWUP INDICATED THAT THE IRRITATION WAS IMPROVING | Insufficient Information | Device problem code
6987417 | 90561781 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD A SURGICAL INCISION THAT WAS IRRITATED BY WEARING THE LIFEVEST THE PATIENTS PHYSICIAN ALLOWED THE PATIENT TO REMOVE THE LIFEVEST TO LET HER SKIN HEAL FOLLOWUP INDICATED THAT THE IRRITATION WAS IMPROVING | Irritation | Patient problem code
6987411 | 90562396 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A RED ITCHY RASH WITH WELTS ALL OVER HER BODY WHILE WEARING THE LIFEVEST THE PATIENT WAS GIVEN A PRESCRIPTION FOR AN ORAL MEDICATION AND LOTION FROM HER PHYSICIAN THE OUTCOME OF THE IRRITATION IS NOT KNOWN | Rash | Patient problem code
6987411 | 90562396 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A RED ITCHY RASH WITH WELTS ALL OVER HER BODY WHILE WEARING THE LIFEVEST THE PATIENT WAS GIVEN A PRESCRIPTION FOR AN ORAL MEDICATION AND LOTION FROM HER PHYSICIAN THE OUTCOME OF THE IRRITATION IS NOT KNOWN | Insufficient Information | Device problem code
6987411 | 90562396 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A RED ITCHY RASH WITH WELTS ALL OVER HER BODY WHILE WEARING THE LIFEVEST THE PATIENT WAS GIVEN A PRESCRIPTION FOR AN ORAL MEDICATION AND LOTION FROM HER PHYSICIAN THE OUTCOME OF THE IRRITATION IS NOT KNOWN | Irritation | Patient problem code
6987411 | 90562396 | D | A US DISTRIBUTOR REPORTED THAT A PATIENT HAD DEVELOPED A RED ITCHY RASH WITH WELTS ALL OVER HER BODY WHILE WEARING THE LIFEVEST THE PATIENT WAS GIVEN A PRESCRIPTION FOR AN ORAL MEDICATION AND LOTION FROM HER PHYSICIAN THE OUTCOME OF THE IRRITATION IS NOT KNOWN | Itching | Patient problem code
6961573 | 90562446 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961573 | 90562446 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961577 | 90562646 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961577 | 90562646 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961576 | 90562844 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961576 | 90562844 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961565 | 90563005 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961565 | 90563005 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961600 | 90563065 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961600 | 90563065 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961575 | 90564634 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961575 | 90564634 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961621 | 90566231 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961621 | 90566231 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961616 | 90566489 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961616 | 90566489 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987613 | 90572652 | N | B DEXCOM LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6987969 | 90581829 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS FROM THE PROCEDURE THE NEXT DAY PATIENT MENTIONED THEY ARENT COMPLETELY EMPTYING THEIR BLADDER BECAUSE THEY STILL PRESSES DOWN ON THEIR STOMACH THEY ALSO MENTIONED ITS A HARD TIME GETTING TO SLEEP BECAUSE OF THE WIRES ON B PATIENT MENTIONED WHEN THEY GET OFF THE ANTIBIOTICS THEY WILL GET A YEAST INFECTION AND THEY WANTED TO KNOW WHO SHOULD THEY SPEAK WITH AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED BACK TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6987969 | 90581829 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS FROM THE PROCEDURE THE NEXT DAY PATIENT MENTIONED THEY ARENT COMPLETELY EMPTYING THEIR BLADDER BECAUSE THEY STILL PRESSES DOWN ON THEIR STOMACH THEY ALSO MENTIONED ITS A HARD TIME GETTING TO SLEEP BECAUSE OF THE WIRES ON B PATIENT MENTIONED WHEN THEY GET OFF THE ANTIBIOTICS THEY WILL GET A YEAST INFECTION AND THEY WANTED TO KNOW WHO SHOULD THEY SPEAK WITH AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED BACK TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient Problem Code
6987969 | 90581829 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS FROM THE PROCEDURE THE NEXT DAY PATIENT MENTIONED THEY ARENT COMPLETELY EMPTYING THEIR BLADDER BECAUSE THEY STILL PRESSES DOWN ON THEIR STOMACH THEY ALSO MENTIONED ITS A HARD TIME GETTING TO SLEEP BECAUSE OF THE WIRES ON B PATIENT MENTIONED WHEN THEY GET OFF THE ANTIBIOTICS THEY WILL GET A YEAST INFECTION AND THEY WANTED TO KNOW WHO SHOULD THEY SPEAK WITH AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED BACK TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Fungal | Patient problem code
6987969 | 90581829 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT EXPERIENCED SORENESS FROM THE PROCEDURE THE NEXT DAY PATIENT MENTIONED THEY ARENT COMPLETELY EMPTYING THEIR BLADDER BECAUSE THEY STILL PRESSES DOWN ON THEIR STOMACH THEY ALSO MENTIONED ITS A HARD TIME GETTING TO SLEEP BECAUSE OF THE WIRES ON B PATIENT MENTIONED WHEN THEY GET OFF THE ANTIBIOTICS THEY WILL GET A YEAST INFECTION AND THEY WANTED TO KNOW WHO SHOULD THEY SPEAK WITH AS A RESULT OF WHAT WAS REPORTED THE PATIENT WAS REFERRED BACK TO THEIR HEALTHCARE PROVIDER HCP NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6960259 | 90595475 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED ULTRASOUND GEL TO PROTECT THE PATIENTS SKIN | Leak / Splash | Device problem code
6960259 | 90595475 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED ULTRASOUND GEL TO PROTECT THE PATIENTS SKIN | Component Incompatible | Device problem code
6960259 | 90595475 | N | THE NEEDLE WAS RETURNED FOR INVESTIGATION THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELATED DEVIATION OR CONCERNS WERE FOUND THE REPORTED FROST ON THE NEEDLE SHAFT WAS CONFIRMED AS THE NEEDLE FAILED INVESTIGATIVE TESTING THE NEEDLE WAS DISASSEMBLED AND INSPECTED NO VISIBLE SOLDERING GAPS OR VOIDS WERE FOUND MICROSCOPIC INSPECTION OF THE VACUUM SLEEVE WAS CONDUCTED AND NO CORROSIVE PITS OR SOLDERING FLUX RESIDUALS WERE IDENTIFIED ON THE EXTERNAL SURFACE OF THE VACUUM SLEEVE A BUBBLE TEST OF THE VACUUM SLEEVE WAS CONDUCTED NO LEAKS WERE IDENTIFIED BASED ON THE INVESTIGATION RESULTS THE CUSTOMER COMPLAINT WAS VERIFIED AND THE ROOT CAUSE WAS DETERMINED TO BE A COMPONENT FAILURE NO RELATIONSHIP WAS IDENTIFIED BETWEEN THE NEEDLE ASSEMBLY PROCESSES AND THE VACUUM LOSS PHENOMENA THE TREATMENT WAS COMPLETED WITH NO INJURY TO THE PATIENT AS THE PHYSICIAN USED ULTRASOUND GEL TO PROTECT THE PATIENTS SKIN | No Device Problem Found | Medical Evaluation result code
6988929 | 90624702 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6988929 | 90624702 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6988648 | 90626331 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
6988648 | 90626331 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6989600 | 90627294 | D | ON B A REPORTER FOR THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT THE PATIENTS ONETOUCH ULTRAMINI METER READ INACCURATELY LOW IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE REPORTER STATED THAT ON B THE PATIENT OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY LOW METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT DOES NOT TAKE ANY MEDICATION TO MANAGE HER DIABETES AND THE REPORTER STATED THAT THE PATIENT CONTINUED TO FOLLOW HER USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE REPORTER DENIED THAT THE PATIENT DEVELOPED ANY SYMPTOMS AS A RESULT OF THE METER ISSUE HOWEVER STATED THAT ON B AT AM THE PATIENT VISITED AN EMERGENCY ROOM ER WHERE SHE WAS TREATED WITH IV FLUIDS AND AT AM THE PATIENT HAD HER BLOOD GLUCOSE TESTED ON AN ER METER WHICH GAVE A RESULT OF MGDL DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Insufficient Information | Device problem code
6989600 | 90627294 | D | ON B A REPORTER FOR THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT THE PATIENTS ONETOUCH ULTRAMINI METER READ INACCURATELY LOW IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE REPORTER STATED THAT ON B THE PATIENT OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY LOW METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT DOES NOT TAKE ANY MEDICATION TO MANAGE HER DIABETES AND THE REPORTER STATED THAT THE PATIENT CONTINUED TO FOLLOW HER USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE REPORTER DENIED THAT THE PATIENT DEVELOPED ANY SYMPTOMS AS A RESULT OF THE METER ISSUE HOWEVER STATED THAT ON B AT AM THE PATIENT VISITED AN EMERGENCY ROOM ER WHERE SHE WAS TREATED WITH IV FLUIDS AND AT AM THE PATIENT HAD HER BLOOD GLUCOSE TESTED ON AN ER METER WHICH GAVE A RESULT OF MGDL DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Hypoglycemia | Patient Problem Code
6989600 | 90627294 | D | ON B A REPORTER FOR THE LAY USERPATIENT CONTACTED LIFESCAN LFS ALLEGING THAT THE PATIENTS ONETOUCH ULTRAMINI METER READ INACCURATELY LOW IN COMPARISON TO HER FEELINGS OR NORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE REPORTER STATED THAT ON B THE PATIENT OBTAINED A RESULT OF MGDL ON THE SUBJECT METER WHICH SHE FELT WAS INACCURATELY LOW METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT DOES NOT TAKE ANY MEDICATION TO MANAGE HER DIABETES AND THE REPORTER STATED THAT THE PATIENT CONTINUED TO FOLLOW HER USUAL DIABETES MANAGEMENT ROUTINE IN RESPONSE TO THE ALLEGED ISSUE THE REPORTER DENIED THAT THE PATIENT DEVELOPED ANY SYMPTOMS AS A RESULT OF THE METER ISSUE HOWEVER STATED THAT ON B AT AM THE PATIENT VISITED AN EMERGENCY ROOM ER WHERE SHE WAS TREATED WITH IV FLUIDS AND AT AM THE PATIENT HAD HER BLOOD GLUCOSE TESTED ON AN ER METER WHICH GAVE A RESULT OF MGDL DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS HAD NOT EXPIRED OR BEEN STORED INCORRECTLY PRODUCT WAS REPLACED AND REQUESTED BACK THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
6989057 | 90627875 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AT THE INJECTION SITE AND THREE HARD NODULES IN THE LIPS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS INJECTION SITE RESPONSES CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND LASTING DAYS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATEONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLLURE XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Leak / Splash | Device problem code
6989057 | 90627875 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AT THE INJECTION SITE AND THREE HARD NODULES IN THE LIPS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS INJECTION SITE RESPONSES CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND LASTING DAYS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATEONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLLURE XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Material Discolored | Device problem code
6989057 | 90627875 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF SWELLING AT THE INJECTION SITE AND THREE HARD NODULES IN THE LIPS ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS INJECTION SITE RESPONSES CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS STARTING EARLY AFTER TREATMENT AND LASTING DAYS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS PATIENTS MAY EXPERIENCE LATEONSET ADVERSE EVENTS WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLLURE XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS PER TABLE INJECTION SITE RESPONSES BY MAXIMUM SEVERITY IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE INJECTION SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLLURE XC INCLUDE FIRMNESS SWELLING TENDERNESS TO TOUCH LUMPSBUMPS REDNESS PAIN AFTER INJECTION BRUISING ITCHING AND DISCOLORATION POSTMARKET SURVEILLANCE JUVÉDERM VOLLURE XC HAS BEEN MARKETED OUTSIDE THE US SINCE AS JUVÉDERM VOLIFT WITH LIDOCAINE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE ON THE USE OF JUVÉDERM VOLLURE XC OUTSIDE THE UNITED STATES AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS THESE AES WITH A FREQUENCY OF EVENTS OR MORE ARE LISTED IN ORDER OF PREVALENCE INFLAMMATORY REACTION SKIN NODULE LOSSLACK OF CORRECTION HEMATOMA ALLERGIC REACTION NECROSIS INFECTION FLULIKE SYMPTOMS PARESTHESIA MIGRATION OF DEVICE ABSCESS DRAINAGE HERPES MALAISE HEADACHE AND ANXIETY IN ADDITION REPORT OF BLURRY VISION AFTER INJECTION IN THE PERIORBITAL AREA REPORT OF BLURRY VISION AFTER INJECTION IN AN UNSPECIFIED AREA AND REPORT OF STROKE AFTER INJECTIONS IN AN UNSPECIFIED AREA WITH MULTIPLE DERMAL FILLERS WERE REPORTED | Migration or Expulsion of Device | Device problem code
6989202 | 90628375 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6989202 | 90628375 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6989199 | 90628424 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6989199 | 90628424 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6988800 | 90628704 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6988800 | 90628704 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6965413 | 90630306 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965413 | 90630306 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965408 | 90630369 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6965408 | 90630369 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6965425 | 90630657 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965425 | 90630657 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965403 | 90630846 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965403 | 90630846 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965416 | 90631059 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6965416 | 90631059 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6965459 | 90631471 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965459 | 90631471 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965404 | 90631834 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DOESNT TAKE AND FAILED TO INTEGRATE THESE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THESE TWO IMPLANTS FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965404 | 90631834 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DOESNT TAKE AND FAILED TO INTEGRATE THESE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THESE TWO IMPLANTS FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965453 | 90631954 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON AND REMOVED ON B | Loss of Osseointegration | Device problem code
6965453 | 90631954 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6965426 | 90632075 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965426 | 90632075 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965414 | 90632201 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965414 | 90632201 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965389 | 90632314 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6965389 | 90632314 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6957022 | 90632458 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO INVESTIGATE THE FSE EVALUATED THE IABP UNIT AND VERIFIED THAT THE IABP UNIT SHUT DOWN WHILE PLUGGED INTO AC WHILE TRYING TO RETRIEVE ERROR CODES UPON BOOT UP THE UNIT DISPLAYED MAINTENANCE REQUIRED CODE THE POWER SUPPLY WAS OCCLUDED WITH LINT THE FSE CLEANEDVACUUMED THE POWER SUPPLY AND PERFORMED FULL FUNCTIONAL TEST IN ADDITION TO SOLENOID AND GASKET FIELD ACTIONS THE UNIT CYCLED OVERNIGHT IN BIOMED SHOP ON TEST BALLOON AND SIMULATOR WITHOUT ANY ERRORS TO REPORT THE IABP UNIT WAS RETURNED FOR CLINICAL USE | Device Displays Incorrect Message | Device Problem Code
6957022 | 90632458 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER GETINGE STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED TO INVESTIGATE THE FSE EVALUATED THE IABP UNIT AND VERIFIED THAT THE IABP UNIT SHUT DOWN WHILE PLUGGED INTO AC WHILE TRYING TO RETRIEVE ERROR CODES UPON BOOT UP THE UNIT DISPLAYED MAINTENANCE REQUIRED CODE THE POWER SUPPLY WAS OCCLUDED WITH LINT THE FSE CLEANEDVACUUMED THE POWER SUPPLY AND PERFORMED FULL FUNCTIONAL TEST IN ADDITION TO SOLENOID AND GASKET FIELD ACTIONS THE UNIT CYCLED OVERNIGHT IN BIOMED SHOP ON TEST BALLOON AND SIMULATOR WITHOUT ANY ERRORS TO REPORT THE IABP UNIT WAS RETURNED FOR CLINICAL USE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6965467 | 90632682 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965467 | 90632682 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965466 | 90632772 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965466 | 90632772 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6965464 | 90632833 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6965464 | 90632833 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961610 | 90632879 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961610 | 90632879 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961617 | 90633105 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961617 | 90633105 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961602 | 90633160 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961602 | 90633160 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6961603 | 90633291 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6961603 | 90633291 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6955646 | 90634622 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY THIS COMPLAINT IS RELATED TO B MEDWATCH | Leak / Splash | Device problem code
6955646 | 90634622 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY THIS COMPLAINT IS RELATED TO B MEDWATCH | Device Displays Incorrect Message | Device problem code
6955646 | 90634622 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY THIS COMPLAINT IS RELATED TO B MEDWATCH | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6955646 | 90634622 | N | B AS PER FIELD SERVICE REPRESENTATIVE A PUMP NOT PRIMED ERROR MESSAGE WAS DISPLAYED AND THE PUMP NOT PRIME LIGHT WAS ILLUMINATED THERE WERE NO ISSUES WITH HEATING OR COOLING AND NO LEAKS WERE OBSERVED THE CUSTOMER CHOSE NOT TO HAVE THE UNIT SERVICED IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS ANDOR CONCLUSION A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY THIS COMPLAINT IS RELATED TO B MEDWATCH | Leakage/Seal | Medical Evaluation result code
6960787 | 90635397 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS UNSURE IF THE CANNULA HAD DEPLOYED INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6960787 | 90635397 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE BACK THE PATIENT WAS UNSURE IF THE CANNULA HAD DEPLOYED INDICATING A NEEDLE MECHANISM FAILURE | Activation, Positioning or Separation Problem | Device problem code
6960274 | 90635585 | N | THE CUSTOMER WAS ABLE TO RETURN UNITS FROM THE SAME LOT TO NATUS THE FAILURE WAS CONFIRMED IN A FRACTION OF THE UNITS RETURNED NATUS INSTRUCTIONS INCLUDE SELECTING THE APPROPRIATE SIZE EAR TIP FOR THE PATIENT AS WELL AS INSTRUCTIONS REGARDING USING LARGER SIZE FOAM TIPS FOR INSERTION INTO THE EAR CANAL NATUS IS INVESTIGATING THIS ISSUE THROUGH CORRECTIVE ACTIONS AT THIS TIME | Inadequacy of Device Shape and/or Size | Device problem code
6960274 | 90635585 | N | THE CUSTOMER WAS ABLE TO RETURN UNITS FROM THE SAME LOT TO NATUS THE FAILURE WAS CONFIRMED IN A FRACTION OF THE UNITS RETURNED NATUS INSTRUCTIONS INCLUDE SELECTING THE APPROPRIATE SIZE EAR TIP FOR THE PATIENT AS WELL AS INSTRUCTIONS REGARDING USING LARGER SIZE FOAM TIPS FOR INSERTION INTO THE EAR CANAL NATUS IS INVESTIGATING THIS ISSUE THROUGH CORRECTIVE ACTIONS AT THIS TIME | Stress Problem Identified | Medical Evaluation result code
6989807 | 90637631 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THEY HAVE PHANTOM SYMPTOMS THEY FEEL LIKE THEY NEED TO GO BUT WHEN THEY TRY NOTHING COMES OUT NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6989807 | 90637631 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A BASIC EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT SAID THEY HAVE PHANTOM SYMPTOMS THEY FEEL LIKE THEY NEED TO GO BUT WHEN THEY TRY NOTHING COMES OUT NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6963801 | 90641773 | N | THE BIOMEDICAL ENGINEER BME REPORTED THAT HIS CENTRAL NURSES STATION CNS JUST WENT BLACK AND STOPPED WORKING THE BME WILL NOT BE SENDING IN THE UNIT HE PLANS TO BUY A NEW ONE DUE TO THE SCREEN BEING OUT OF WARRANTY NIHON KOHDEN TECHNICAL SUPPORT PROVIDED THE CUSTOMER WITH THE PART NUMBER FOR THE CNS SCREEN | No Display / Image | Device Problem Code
6963801 | 90641773 | N | THE BIOMEDICAL ENGINEER BME REPORTED THAT HIS CENTRAL NURSES STATION CNS JUST WENT BLACK AND STOPPED WORKING THE BME WILL NOT BE SENDING IN THE UNIT HE PLANS TO BUY A NEW ONE DUE TO THE SCREEN BEING OUT OF WARRANTY NIHON KOHDEN TECHNICAL SUPPORT PROVIDED THE CUSTOMER WITH THE PART NUMBER FOR THE CNS SCREEN | Image Artifact | Medical Evaluation result code
6960438 | 90641787 | N | B ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL ATTEMPTS TO OBTAIN THE PATIENT INFORMATION WERE MADE IN PERSON NO TESTSLABORATORY DATA WAS AVAILABLE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Apparent Adverse Event | Device problem code
6960438 | 90641787 | N | B ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL ATTEMPTS TO OBTAIN THE PATIENT INFORMATION WERE MADE IN PERSON NO TESTSLABORATORY DATA WAS AVAILABLE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Device Problem Found | Medical Evaluation result code
6960391 | 90642389 | N | B THIS CASE WAS REVIEWED AND INVESTIGATED ACCORDING TO THE MANUFACTURERS POLICY ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL ATTEMPTS TO OBTAIN THE PATIENT INFORMATION WERE MADE VIA EMAIL AND PHONE NO TESTSLABORATORY DATA WAS AVAILABLE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Apparent Adverse Event | Device problem code
6960391 | 90642389 | N | B THIS CASE WAS REVIEWED AND INVESTIGATED ACCORDING TO THE MANUFACTURERS POLICY ATTEMPTS TO OBTAIN PATIENT INFORMATION WAS UNSUCCESSFUL ATTEMPTS TO OBTAIN THE PATIENT INFORMATION WERE MADE VIA EMAIL AND PHONE NO TESTSLABORATORY DATA WAS AVAILABLE THE IMPLANT OR EXPLANT DATES ARE NOT APPLICABLE TO THIS DEVICE | No Device Problem Found | Medical Evaluation result code
6958445 | 90643377 | N | THE VENTILATOR WAS INVESTIGATED ONSITE BY OUR FIELD SERVICE ENGINEER NO DAMAGES WERE FOUND ON THE VENTILATOR ELECTRICAL SAFETY TEST WAS PERFORMED WITH SUCCESSFUL RESULT AND THE VENTILATOR WAS THEREFORE RETURNED TO SERVICE ACCORDING TO RECEIVED INFORMATION THE WHEELS OF THE VENTILATOR WERE NOT LOCKED AND THE VENTILATOR WAS PLACED INSIDE THE SAFETY LINE AT THE TIME OF THE EVENT THE VENTILATOR OPERATES IN A FIELD OF UP TO MT GAUSS FOR TUNNEL SCANNERS AND UP TO MT GAUSS FOR OPEN SCANNERS BEFORE USING THE VENTILATOR IN THE MR ENVIRONMENT CONDITIONS MUST BE MET INCLUDING LOCKING THE TWO FRONT WHEELS AND PLACING THE VENTILATOR OUTSIDE THE GAUSS MT SAFETY LINE WHICH MUST BE MARKED AROUND THE MR SCANNER THESE CONDITIONS ARE STATED IN THE DECLARATION FOR THE USE OF THE VENTILATOR IN THE MR ENVIRONMENT THE GAUSS SAFETY LINE WAS CORRECTLY MARKED ON THE FLOOR OUR CONCLUSION IS THAT THE INCIDENT WAS CAUSED BY THE USER NOT FOLLOWING THE REQUIREMENTS FOR USE OF THE VENTILATOR IN MR ENVIRONMENT | Mechanical Jam | Device problem code
6958445 | 90643377 | N | THE VENTILATOR WAS INVESTIGATED ONSITE BY OUR FIELD SERVICE ENGINEER NO DAMAGES WERE FOUND ON THE VENTILATOR ELECTRICAL SAFETY TEST WAS PERFORMED WITH SUCCESSFUL RESULT AND THE VENTILATOR WAS THEREFORE RETURNED TO SERVICE ACCORDING TO RECEIVED INFORMATION THE WHEELS OF THE VENTILATOR WERE NOT LOCKED AND THE VENTILATOR WAS PLACED INSIDE THE SAFETY LINE AT THE TIME OF THE EVENT THE VENTILATOR OPERATES IN A FIELD OF UP TO MT GAUSS FOR TUNNEL SCANNERS AND UP TO MT GAUSS FOR OPEN SCANNERS BEFORE USING THE VENTILATOR IN THE MR ENVIRONMENT CONDITIONS MUST BE MET INCLUDING LOCKING THE TWO FRONT WHEELS AND PLACING THE VENTILATOR OUTSIDE THE GAUSS MT SAFETY LINE WHICH MUST BE MARKED AROUND THE MR SCANNER THESE CONDITIONS ARE STATED IN THE DECLARATION FOR THE USE OF THE VENTILATOR IN THE MR ENVIRONMENT THE GAUSS SAFETY LINE WAS CORRECTLY MARKED ON THE FLOOR OUR CONCLUSION IS THAT THE INCIDENT WAS CAUSED BY THE USER NOT FOLLOWING THE REQUIREMENTS FOR USE OF THE VENTILATOR IN MR ENVIRONMENT | Mechanical Problem Identified | Medical Evaluation result code
6958445 | 90643378 | D | IT WAS REPORTED THAT WHEN CONNECTED TO A PATIENT THE VENTILATOR WAS PULLED INTO AN MR SCANNER SINCE THE TWO FRONT WHEELS WERE UNLOCKED AND THE VENTILATOR WAS PLACED TOO CLOSE THERE WAS NO PATIENT HARM B | No Consequences Or Impact To Patient | Patient problem code
6958445 | 90643378 | D | IT WAS REPORTED THAT WHEN CONNECTED TO A PATIENT THE VENTILATOR WAS PULLED INTO AN MR SCANNER SINCE THE TWO FRONT WHEELS WERE UNLOCKED AND THE VENTILATOR WAS PLACED TOO CLOSE THERE WAS NO PATIENT HARM B | Difficult to Open or Close | Device problem code
6963597 | 90644620 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6963597 | 90644620 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6962761 | 90652824 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
6962761 | 90652824 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | Device Fell | Device problem code
6962746 | 90652960 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | Device Fell | Device problem code
6962746 | 90652960 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
6963057 | 90653937 | D | IT WAS REPORTED THAT CLIPS ARE FALLING OUT OF THE JAW IT DID NOT FALL INTO THE PATIENTS BODY THE PATIENTS CONDITION IS FINE | No Consequences Or Impact To Patient | Patient problem code
6963057 | 90653937 | D | IT WAS REPORTED THAT CLIPS ARE FALLING OUT OF THE JAW IT DID NOT FALL INTO THE PATIENTS BODY THE PATIENTS CONDITION IS FINE | Insufficient Information | Device problem code
6962798 | 90654882 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
6962798 | 90654882 | D | IT WAS REPORTED THAT THE CLIPS ARE FALLING OUT OF THE JAW BUT THEY DID NOT FALL INTO THE PATIENTS BODY THERE WAS NO PATIENT INJURY | Device Fell | Device problem code
6984809 | 90656610 | N | THE NEXT DAY WE RECEIVED WORD FROM THE BIOMED THAT HE HAD MINOR BURNS ON HIS ARM FROM TRYING TO CLOSE THE TANK HE HEARD A LOUD POP THEN THE REGULATOR JUST BLEW APART AND WAS ON FIRE HE SENT PICTURES AND THE OXYGEN REGULATOR PICTURE SHOWS THAT THE ENTIRE TOP OF THE REGULATOR IS BLOWN OFF OTHER PICTURES SHOW EVIDENCE OF PARTS INVOLVED IN THE FLAMES THE DEVICES WERE SHIPPED FROM PRECISION FLUIDICS ON THE BIOMED FURTHER REPORTS THAT THE BURNS ARE HEALING WELL HE IS FOLLOWING TREATMENT INSTRUCTIONS FROM A DOCTOR AT THE HOSPITAL AND HE IS BACK TO WORK PRECISION FLUIDICS HAS RECEIVED THE EBLOCK ASSEMBLY OF THE CART FOR OUR ANALYSIS PICTURES OF THE SUBJECT REGULATOR HAVE BEEN TAKEN AND THE ASSEMBLY IS SECURED AT THE B MANUFACTURING FACILITY PENDING OUR DETAILED INVESTIGATIONS WE HAVE NO FURTHER ANALYSIS OF THE INCIDENT BUT WILL CONDUCT A COMPLAINT ANALYSIS AND PRODUCE NECESSARY MEDICAL AND FDA INTERNAL DOCUMENTATION WE HAVE NO RECORDS OF ANY SIMILAR INCIDENTS OF FIRE OR REGULATORS BLOWING UP INVOLVING OUR PRODUCTS PARTICULARLY THE REGULATOR MODELS WE HAVE USED FROM OUR SUPPLIER FOR OVER YEARS SUCCESSFULLY IN THE APPLICATION | Fire | Device problem code
6984809 | 90656610 | N | THE NEXT DAY WE RECEIVED WORD FROM THE BIOMED THAT HE HAD MINOR BURNS ON HIS ARM FROM TRYING TO CLOSE THE TANK HE HEARD A LOUD POP THEN THE REGULATOR JUST BLEW APART AND WAS ON FIRE HE SENT PICTURES AND THE OXYGEN REGULATOR PICTURE SHOWS THAT THE ENTIRE TOP OF THE REGULATOR IS BLOWN OFF OTHER PICTURES SHOW EVIDENCE OF PARTS INVOLVED IN THE FLAMES THE DEVICES WERE SHIPPED FROM PRECISION FLUIDICS ON THE BIOMED FURTHER REPORTS THAT THE BURNS ARE HEALING WELL HE IS FOLLOWING TREATMENT INSTRUCTIONS FROM A DOCTOR AT THE HOSPITAL AND HE IS BACK TO WORK PRECISION FLUIDICS HAS RECEIVED THE EBLOCK ASSEMBLY OF THE CART FOR OUR ANALYSIS PICTURES OF THE SUBJECT REGULATOR HAVE BEEN TAKEN AND THE ASSEMBLY IS SECURED AT THE B MANUFACTURING FACILITY PENDING OUR DETAILED INVESTIGATIONS WE HAVE NO FURTHER ANALYSIS OF THE INCIDENT BUT WILL CONDUCT A COMPLAINT ANALYSIS AND PRODUCE NECESSARY MEDICAL AND FDA INTERNAL DOCUMENTATION WE HAVE NO RECORDS OF ANY SIMILAR INCIDENTS OF FIRE OR REGULATORS BLOWING UP INVOLVING OUR PRODUCTS PARTICULARLY THE REGULATOR MODELS WE HAVE USED FROM OUR SUPPLIER FOR OVER YEARS SUCCESSFULLY IN THE APPLICATION | Operational Problem Identified | Medical Evaluation result code
6972591 | 90659503 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX CLOGGED WITH DRIED BLOOD VISUAL AND XRAY EXAMINATION REVEALED THE OUTER INSULATION WAS TWISTED DUE TO THE INNER COIL OVERTORQUED AT THE CONNECTOR REGION AFTER CLEANING AND CUTTING THE LEAD THE HELIX EXTENDED AND RETRACTED WHEN APPLYING TORQUE DIRECTLY AT THE INNER COIL ELECTRICAL TESTING REVEALED NO CONDUCTOR FRACTURES OR INTERNAL SHORTS THE EVENT OF TWISTED LEAD WAS CONFIRMED | Complete Blockage | Device problem code
6972591 | 90659503 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX CLOGGED WITH DRIED BLOOD VISUAL AND XRAY EXAMINATION REVEALED THE OUTER INSULATION WAS TWISTED DUE TO THE INNER COIL OVERTORQUED AT THE CONNECTOR REGION AFTER CLEANING AND CUTTING THE LEAD THE HELIX EXTENDED AND RETRACTED WHEN APPLYING TORQUE DIRECTLY AT THE INNER COIL ELECTRICAL TESTING REVEALED NO CONDUCTOR FRACTURES OR INTERNAL SHORTS THE EVENT OF TWISTED LEAD WAS CONFIRMED | Material Twisted / Bent | Device problem code
6972591 | 90659503 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX CLOGGED WITH DRIED BLOOD VISUAL AND XRAY EXAMINATION REVEALED THE OUTER INSULATION WAS TWISTED DUE TO THE INNER COIL OVERTORQUED AT THE CONNECTOR REGION AFTER CLEANING AND CUTTING THE LEAD THE HELIX EXTENDED AND RETRACTED WHEN APPLYING TORQUE DIRECTLY AT THE INNER COIL ELECTRICAL TESTING REVEALED NO CONDUCTOR FRACTURES OR INTERNAL SHORTS THE EVENT OF TWISTED LEAD WAS CONFIRMED | Retraction Problem | Device problem code
6972591 | 90659503 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX CLOGGED WITH DRIED BLOOD VISUAL AND XRAY EXAMINATION REVEALED THE OUTER INSULATION WAS TWISTED DUE TO THE INNER COIL OVERTORQUED AT THE CONNECTOR REGION AFTER CLEANING AND CUTTING THE LEAD THE HELIX EXTENDED AND RETRACTED WHEN APPLYING TORQUE DIRECTLY AT THE INNER COIL ELECTRICAL TESTING REVEALED NO CONDUCTOR FRACTURES OR INTERNAL SHORTS THE EVENT OF TWISTED LEAD WAS CONFIRMED | Stress Problem Identified | Medical Evaluation result code
6962672 | 90662974 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6962672 | 90662974 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6990200 | 90663945 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS UNABLE TO MAKE IT TO THE RESTROOM ONCE THEY GET THERE THEY CANT EMPTY NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
6990200 | 90663945 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA A B TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT IS UNABLE TO MAKE IT TO THE RESTROOM ONCE THEY GET THERE THEY CANT EMPTY NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6959561 | 90665226 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE GETINGE FIELD SERVICE ENGINEER REPORTED THAT THE IABP BATTERIES WERE CHARGING AND VERIFIED THE LINE VOLTAGE TO THE POWER SUPPLY WERE WORKING TO FACTORY SPECIFICATION THE FSE DISASSEMBLED THE IABP AND INSTALLED A NEW POWER MANAGEMENT BOARD THE IABP WAS THEN REASSEMBLED AND THEN RAN THROUGH COMPLETE CALIBRATION AND FUNCTIONAL TESTS ALL TESTING PASSED AS PER FACTORY SPECIFICATIONS THE BATTERIES ARE NOW CHARGING AS THEY SHOULD AND THE IABP WAS CLEARED FOR CLINICAL USE B | Calibration Problem | Device problem code
6959561 | 90665226 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE GETINGE FIELD SERVICE ENGINEER REPORTED THAT THE IABP BATTERIES WERE CHARGING AND VERIFIED THE LINE VOLTAGE TO THE POWER SUPPLY WERE WORKING TO FACTORY SPECIFICATION THE FSE DISASSEMBLED THE IABP AND INSTALLED A NEW POWER MANAGEMENT BOARD THE IABP WAS THEN REASSEMBLED AND THEN RAN THROUGH COMPLETE CALIBRATION AND FUNCTIONAL TESTS ALL TESTING PASSED AS PER FACTORY SPECIFICATIONS THE BATTERIES ARE NOW CHARGING AS THEY SHOULD AND THE IABP WAS CLEARED FOR CLINICAL USE B | Charging Problem | Device problem code
6959561 | 90665226 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE GETINGE FIELD SERVICE ENGINEER REPORTED THAT THE IABP BATTERIES WERE CHARGING AND VERIFIED THE LINE VOLTAGE TO THE POWER SUPPLY WERE WORKING TO FACTORY SPECIFICATION THE FSE DISASSEMBLED THE IABP AND INSTALLED A NEW POWER MANAGEMENT BOARD THE IABP WAS THEN REASSEMBLED AND THEN RAN THROUGH COMPLETE CALIBRATION AND FUNCTIONAL TESTS ALL TESTING PASSED AS PER FACTORY SPECIFICATIONS THE BATTERIES ARE NOW CHARGING AS THEY SHOULD AND THE IABP WAS CLEARED FOR CLINICAL USE B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6959561 | 90665226 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE THE GETINGE FIELD SERVICE ENGINEER REPORTED THAT THE IABP BATTERIES WERE CHARGING AND VERIFIED THE LINE VOLTAGE TO THE POWER SUPPLY WERE WORKING TO FACTORY SPECIFICATION THE FSE DISASSEMBLED THE IABP AND INSTALLED A NEW POWER MANAGEMENT BOARD THE IABP WAS THEN REASSEMBLED AND THEN RAN THROUGH COMPLETE CALIBRATION AND FUNCTIONAL TESTS ALL TESTING PASSED AS PER FACTORY SPECIFICATIONS THE BATTERIES ARE NOW CHARGING AS THEY SHOULD AND THE IABP WAS CLEARED FOR CLINICAL USE B | Energy Storage System Problem | Medical Evaluation result code
6990281 | 90668403 | D | THIS REPORT IS RELATED TO THE CLINICAL PATIENT IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO CHEST PAIN LEFT HEART CATHETERIZATION WAS PERFORMED THE PATIENT WAS DISCHARGED ON DIURETICS AND RANEXA THE ISSUE MAY POSSIBLY RELATED TO THE DEVICE | Insufficient Information | Device problem code
6990281 | 90668403 | D | THIS REPORT IS RELATED TO THE CLINICAL PATIENT IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO CHEST PAIN LEFT HEART CATHETERIZATION WAS PERFORMED THE PATIENT WAS DISCHARGED ON DIURETICS AND RANEXA THE ISSUE MAY POSSIBLY RELATED TO THE DEVICE | Chest Pain | Patient problem code
6990281 | 90668403 | D | THIS REPORT IS RELATED TO THE CLINICAL PATIENT IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO CHEST PAIN LEFT HEART CATHETERIZATION WAS PERFORMED THE PATIENT WAS DISCHARGED ON DIURETICS AND RANEXA THE ISSUE MAY POSSIBLY RELATED TO THE DEVICE | Pain | Patient problem code
6990376 | 90670352 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | No Device Problem Found | Medical Evaluation result code
6990382 | 90673834 | N | ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE SURGEON JUST PERFORMED A POLY SWAP NO ADDITIONAL CUTS MADE ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS INSERT REVISION IN TKA YEAR AFTER PRIMARY FOR JOINT INSTABILITY THE INSERT WAS EXCHANGED TO A THICKER ONE THE DEVELOPMENTAL INSTABILITY IS A RATHER COMMONLY DESCRIBED POSSIBLE COMPLICATION FOLLOWING TKA BECAUSE SOFT TISSUE MAY STRETCH AND PROVIDE INSUFFICIENT STABILITY IT IS THEREFORE COMMON PRACTICE TO RESORT TO A THICKER INSERT AND RECREATE SOFT TISSUE TENSION THIS IS NOT A PROBLEM DUE TO ANY IMPLANT DEFECT OR MALFUNCTION BATCH REVIEW PERFORMED ON OCTOBER B | Defective Device | Device problem code
6990382 | 90673834 | N | ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE SURGEON JUST PERFORMED A POLY SWAP NO ADDITIONAL CUTS MADE ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS INSERT REVISION IN TKA YEAR AFTER PRIMARY FOR JOINT INSTABILITY THE INSERT WAS EXCHANGED TO A THICKER ONE THE DEVELOPMENTAL INSTABILITY IS A RATHER COMMONLY DESCRIBED POSSIBLE COMPLICATION FOLLOWING TKA BECAUSE SOFT TISSUE MAY STRETCH AND PROVIDE INSUFFICIENT STABILITY IT IS THEREFORE COMMON PRACTICE TO RESORT TO A THICKER INSERT AND RECREATE SOFT TISSUE TENSION THIS IS NOT A PROBLEM DUE TO ANY IMPLANT DEFECT OR MALFUNCTION BATCH REVIEW PERFORMED ON OCTOBER B | Unstable | Device problem code
6990382 | 90673834 | N | ADDITIONAL INFORMATION RECEIVED ON OCTOBER AND INCLUDES THE SURGEON JUST PERFORMED A POLY SWAP NO ADDITIONAL CUTS MADE ON OCTOBER THE MEDICAL AFFAIRS DIRECTOR PERFORMED A CLINICAL EVALUATION AND COMMENTED AS FOLLOWS INSERT REVISION IN TKA YEAR AFTER PRIMARY FOR JOINT INSTABILITY THE INSERT WAS EXCHANGED TO A THICKER ONE THE DEVELOPMENTAL INSTABILITY IS A RATHER COMMONLY DESCRIBED POSSIBLE COMPLICATION FOLLOWING TKA BECAUSE SOFT TISSUE MAY STRETCH AND PROVIDE INSUFFICIENT STABILITY IT IS THEREFORE COMMON PRACTICE TO RESORT TO A THICKER INSERT AND RECREATE SOFT TISSUE TENSION THIS IS NOT A PROBLEM DUE TO ANY IMPLANT DEFECT OR MALFUNCTION BATCH REVIEW PERFORMED ON OCTOBER B | Incorrect Dimension | Medical Evaluation result code
6990398 | 90674039 | N | ON OCTOBER THE MANUFACTURER OF THE CERAMIC COMPONENT INVOLVED IN THE COMPLAINT NOT MARKETED IN US PROVIDED A DOCUMENT REVIEW REPORTING AS FOLLOWS THE COMPONENT PROPERTIES AND THE MICROSTRUCTURE AS OBTAINED FROM THE QUALITY DOCUMENTS ACCOMPLISH THE REQUIREMENTS AS SPECIFIED AT THE TIME OF PRODUCTION THERE ARE NO INDICATIONS OF ANY PREEXISTING MATERIAL DEFECT OR NONCONFORMITIES REGARDING STERILIZATION DUE TO LACK OF CERAMIC PARTS FURTHER INVESTIGATIONS CANNOT BE DONE BATCH REVIEWS PERFORMED ON OCTOBER B TO DATE ALL ITEMS OF THE SAME LOT HAVE BEEN ALREADY SOLD WITHOUT ANY SIMILAR REPORTED EVENT | Defective Device | Device problem code
6990398 | 90674039 | N | ON OCTOBER THE MANUFACTURER OF THE CERAMIC COMPONENT INVOLVED IN THE COMPLAINT NOT MARKETED IN US PROVIDED A DOCUMENT REVIEW REPORTING AS FOLLOWS THE COMPONENT PROPERTIES AND THE MICROSTRUCTURE AS OBTAINED FROM THE QUALITY DOCUMENTS ACCOMPLISH THE REQUIREMENTS AS SPECIFIED AT THE TIME OF PRODUCTION THERE ARE NO INDICATIONS OF ANY PREEXISTING MATERIAL DEFECT OR NONCONFORMITIES REGARDING STERILIZATION DUE TO LACK OF CERAMIC PARTS FURTHER INVESTIGATIONS CANNOT BE DONE BATCH REVIEWS PERFORMED ON OCTOBER B TO DATE ALL ITEMS OF THE SAME LOT HAVE BEEN ALREADY SOLD WITHOUT ANY SIMILAR REPORTED EVENT | Incorrect Dimension | Medical Evaluation result code
6964424 | 90675327 | N | SIEMENS HEALTHCARE DIAGNOSTICS INC IS INVESTIGATING THE CAUSE OF THE DISCORDANT LOW INNOVANCE HEPARIN PATIENT RESULT THAT WAS GENERATED ON THE CS COAGULATION ANALYZER | Coagulation in Device or device Ingredient | Device problem code
6964424 | 90675327 | N | SIEMENS HEALTHCARE DIAGNOSTICS INC IS INVESTIGATING THE CAUSE OF THE DISCORDANT LOW INNOVANCE HEPARIN PATIENT RESULT THAT WAS GENERATED ON THE CS COAGULATION ANALYZER | No Device Problem Found | Medical Evaluation result code
6964358 | 90675485 | N | NO REPORT OF PATIENT INVOLVEMENT B THE GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE SYSTEM AND RETIGHTENED CONNECTIONS TO ELIMINATE POSSIBLE LEAKS | Leak / Splash | Device problem code
6964358 | 90675485 | N | NO REPORT OF PATIENT INVOLVEMENT B THE GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE SYSTEM AND RETIGHTENED CONNECTIONS TO ELIMINATE POSSIBLE LEAKS | Leakage/Seal | Medical Evaluation result code
6972029 | 90676663 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST DETAIL TRACE ANALYSIS CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT CONNECTOR BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON CONNECTOR BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW B | Electrical Power Problem | Device problem code
6972029 | 90676663 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST DETAIL TRACE ANALYSIS CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT CONNECTOR BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON CONNECTOR BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW B | Disconnection | Device problem code
6972029 | 90676663 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST DETAIL TRACE ANALYSIS CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT CONNECTOR BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON CONNECTOR BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW B | Scratched Material | Device problem code
6972029 | 90676663 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST DETAIL TRACE ANALYSIS CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT CONNECTOR BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON CONNECTOR BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE INSULIN PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW B | Energy Storage System Problem | Medical Evaluation result code
6962790 | 90683027 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE UNIT WAS TESTED IN VENT MODE FOR SEVEN DAYS WITHOUT ISSUE THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6962790 | 90683027 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE UNIT WAS TESTED IN VENT MODE FOR SEVEN DAYS WITHOUT ISSUE THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6963435 | 90683155 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE CENTRAL NURSES STATION CNS HAD NO AUDIO SOUND AFTER INSTALLATION AND SETUP WHILE TROUBLESHOOTING IT WAS DISCOVERED THAT THE AUDIO CABLE WAS NOT PLUGGED INTO THE CORRECT JACK ON THE CNS PLUGGING IT INTO THE CORRECT JACK RESOLVED THE ISSUE NO PATIENT ALARMS WERE MISSED AND NO PATIENT HARM WAS REPORTED | Noise, Audible | Device problem code
6963435 | 90683155 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE CENTRAL NURSES STATION CNS HAD NO AUDIO SOUND AFTER INSTALLATION AND SETUP WHILE TROUBLESHOOTING IT WAS DISCOVERED THAT THE AUDIO CABLE WAS NOT PLUGGED INTO THE CORRECT JACK ON THE CNS PLUGGING IT INTO THE CORRECT JACK RESOLVED THE ISSUE NO PATIENT ALARMS WERE MISSED AND NO PATIENT HARM WAS REPORTED | No Audible Alarm | Device problem code
6963435 | 90683155 | N | THE BIOMEDICAL ENGINEER REPORTED THAT THE CENTRAL NURSES STATION CNS HAD NO AUDIO SOUND AFTER INSTALLATION AND SETUP WHILE TROUBLESHOOTING IT WAS DISCOVERED THAT THE AUDIO CABLE WAS NOT PLUGGED INTO THE CORRECT JACK ON THE CNS PLUGGING IT INTO THE CORRECT JACK RESOLVED THE ISSUE NO PATIENT ALARMS WERE MISSED AND NO PATIENT HARM WAS REPORTED | No Device Problem Found | Medical Evaluation result code
6964836 | 90684171 | N | THE DATA LOG FILES FROM THE AQUIOS CL FLOW CYTOMETER WERE INVESTIGATED BY THE SOFTWARE ENGINEERING TEAM AND THEY FOUND THREE DIFFERENT PATTERNS O ONE SAMPLE IDTWO DIFFERENT BLOOD TUBES O TWO DIFFERENT SAMPLE IDSSAME BLOOD TUBE O TWO SAME SAMPLE IDSSAME BLOOD TUBE DATA ANALYSIS FOR THE INSTRUMENT SERIAL B WAS BEEN COMPLETED FOR A TOTAL OF SAMPLES RANGING FROM AUGUST THROUGH AUGUST THE ASSESSMENT COMPLETED BY BECKMAN COULTER SOFTWARE ENGINEERING REPORTED THAT A TOTAL OF SIX INSTANCES WERE IDENTIFIED THAT LED TO ERRONEOUS RESULTS TO DATE THE ANALYSIS HAS CONFIRMED THAT ONLY THE SAMPLE IDS LISTED BELOW WERE AFFECTED BASED ON THE INFORMATION BY THE SOFTWARE TEAM ASSESSMENT THE FAILURE MODE IS CAUSED BY AN AQUIOS CL SOFTWARE MALFUNCTION B RELATED EVENTS B B | Incorrect, Inadequate or Imprecise Result or Readings | Device problem code
6972002 | 90686187 | N | B THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES DETAILED TRACE ANALYSIS SHOWS THAT PUMP ALARMED LOW BATTERY AND THEN POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT J PCB THE LOADED VOLTAGE LOADED V LITH DISPLAY AT THE POWER GRAPH MANAGEMENT TOOL SHOWS ABNORMAL BEHAVIOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCB THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6972002 | 90686187 | N | B THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES DETAILED TRACE ANALYSIS SHOWS THAT PUMP ALARMED LOW BATTERY AND THEN POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT J PCB THE LOADED VOLTAGE LOADED V LITH DISPLAY AT THE POWER GRAPH MANAGEMENT TOOL SHOWS ABNORMAL BEHAVIOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCB THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Electrical Power Problem | Device problem code
6972002 | 90686187 | N | B THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES DETAILED TRACE ANALYSIS SHOWS THAT PUMP ALARMED LOW BATTERY AND THEN POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT J PCB THE LOADED VOLTAGE LOADED V LITH DISPLAY AT THE POWER GRAPH MANAGEMENT TOOL SHOWS ABNORMAL BEHAVIOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCB THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Disconnection | Device problem code
6972002 | 90686187 | N | B THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DAT TEST AT INCHES DETAILED TRACE ANALYSIS SHOWS THAT PUMP ALARMED LOW BATTERY AND THEN POWER ERROR ALARM WAS TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AT J PCB THE LOADED VOLTAGE LOADED V LITH DISPLAY AT THE POWER GRAPH MANAGEMENT TOOL SHOWS ABNORMAL BEHAVIOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCB THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED PUMP WAS RECEIVED WITH SCRATCHED CASE PILLOWING KEYPAD OVERLAY AND MINOR SCRATCHED LCD WINDOW | Energy Storage System Problem | Medical Evaluation result code
6963792 | 90686996 | N | THE CUSTOMER B STATED THAT HE JUST INSTALLED THIS CENTRAL NURSES STATION CNS AND WAS MOVING THINGS AROUND AND HAD THE CNS PLUGGED INTO THE UPS ONLY THE UPS RAN OUT OF POWER AND THE CNS SHUT DOWN WHEN HE PLUGGED IT BACK THE CNS WAS STUCK IN THE BLACK CNS BOOTUP SCREEN NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD HIM PULL OUT THE DRIVE AND BOOT UP THE CNS THE CNS CAME BACK UP CORRECTLY NKTS THEN HAD B REINSTALL THE DRIVE TO LET IT REBUILD THE VOLUME THIS SOLVED THE ISSUE | Mechanical Jam | Device problem code
6963792 | 90686996 | N | THE CUSTOMER B STATED THAT HE JUST INSTALLED THIS CENTRAL NURSES STATION CNS AND WAS MOVING THINGS AROUND AND HAD THE CNS PLUGGED INTO THE UPS ONLY THE UPS RAN OUT OF POWER AND THE CNS SHUT DOWN WHEN HE PLUGGED IT BACK THE CNS WAS STUCK IN THE BLACK CNS BOOTUP SCREEN NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD HIM PULL OUT THE DRIVE AND BOOT UP THE CNS THE CNS CAME BACK UP CORRECTLY NKTS THEN HAD B REINSTALL THE DRIVE TO LET IT REBUILD THE VOLUME THIS SOLVED THE ISSUE | No Display / Image | Device Problem Code
6963792 | 90686996 | N | THE CUSTOMER B STATED THAT HE JUST INSTALLED THIS CENTRAL NURSES STATION CNS AND WAS MOVING THINGS AROUND AND HAD THE CNS PLUGGED INTO THE UPS ONLY THE UPS RAN OUT OF POWER AND THE CNS SHUT DOWN WHEN HE PLUGGED IT BACK THE CNS WAS STUCK IN THE BLACK CNS BOOTUP SCREEN NIHON KOHDEN TECHNICAL SUPPORT NKTS HAD HIM PULL OUT THE DRIVE AND BOOT UP THE CNS THE CNS CAME BACK UP CORRECTLY NKTS THEN HAD B REINSTALL THE DRIVE TO LET IT REBUILD THE VOLUME THIS SOLVED THE ISSUE | Image Artifact | Medical Evaluation result code
6963601 | 90687699 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6963601 | 90687699 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6966974 | 90688548 | N | THE BIOMEDICAL ENGINEER BME REPORTED THAT WHEN A PATIENT WAS EXPERIENCING VFIB THE CENTRAL NURSES STATION MONITOR DID NOT SOUND AN ALARM THE BME WOULD WANTED NIHON KOHDEN TO VERIFY THE LOGS FROM TO TO SEE IF THE MONITOR DID ALERT TO THE PATIENT EVENT NIHON KOHDEN TECHNICAL SUPPORT NKTS CONTACTED NIHON KOHDEN QUALITY ASSURANCE NKQA REGARDING THE ISSUE FOR EVALUATION NKQA EVALUATED THE LOGS AND DETERMINED THAT THE DEVICE WAS FUNCTIONAL AND ALL VITALS WERE SENT TO THE CENTRAL NURSES STATION THIS ISSUE WAS CAUSED BY THE CNS DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | No Audible Prompt / Feedback | Device problem code
6966974 | 90688548 | N | THE BIOMEDICAL ENGINEER BME REPORTED THAT WHEN A PATIENT WAS EXPERIENCING VFIB THE CENTRAL NURSES STATION MONITOR DID NOT SOUND AN ALARM THE BME WOULD WANTED NIHON KOHDEN TO VERIFY THE LOGS FROM TO TO SEE IF THE MONITOR DID ALERT TO THE PATIENT EVENT NIHON KOHDEN TECHNICAL SUPPORT NKTS CONTACTED NIHON KOHDEN QUALITY ASSURANCE NKQA REGARDING THE ISSUE FOR EVALUATION NKQA EVALUATED THE LOGS AND DETERMINED THAT THE DEVICE WAS FUNCTIONAL AND ALL VITALS WERE SENT TO THE CENTRAL NURSES STATION THIS ISSUE WAS CAUSED BY THE CNS DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Noise, Audible | Device problem code
6966974 | 90688548 | N | THE BIOMEDICAL ENGINEER BME REPORTED THAT WHEN A PATIENT WAS EXPERIENCING VFIB THE CENTRAL NURSES STATION MONITOR DID NOT SOUND AN ALARM THE BME WOULD WANTED NIHON KOHDEN TO VERIFY THE LOGS FROM TO TO SEE IF THE MONITOR DID ALERT TO THE PATIENT EVENT NIHON KOHDEN TECHNICAL SUPPORT NKTS CONTACTED NIHON KOHDEN QUALITY ASSURANCE NKQA REGARDING THE ISSUE FOR EVALUATION NKQA EVALUATED THE LOGS AND DETERMINED THAT THE DEVICE WAS FUNCTIONAL AND ALL VITALS WERE SENT TO THE CENTRAL NURSES STATION THIS ISSUE WAS CAUSED BY THE CNS DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Results Pending Completion of Investigation | Medical Evaluation result code
6966985 | 90688569 | N | THE CUSTOMER SAID THEY DID A GENERATOR TEST AND THE CNS SHUT OFF AND NOW GETTING ERROR PROTECTED KEYS NOT CONNECTED WHEN STARTING THE CNS NIHON KOHDEN TECHNICAL SUPPORT HAD HIM POWER OFF THE CENTRAL NURSES STATION CNS RESEAT THE DONGLE AND POWER IT BACK ON THE CNS SOFTWARE STARTED AND EVERYTHING WORKED AS IT SHOULD THIS RESOLVED THEIR ISSUE | Connection Problem | Device problem code
6966985 | 90688569 | N | THE CUSTOMER SAID THEY DID A GENERATOR TEST AND THE CNS SHUT OFF AND NOW GETTING ERROR PROTECTED KEYS NOT CONNECTED WHEN STARTING THE CNS NIHON KOHDEN TECHNICAL SUPPORT HAD HIM POWER OFF THE CENTRAL NURSES STATION CNS RESEAT THE DONGLE AND POWER IT BACK ON THE CNS SOFTWARE STARTED AND EVERYTHING WORKED AS IT SHOULD THIS RESOLVED THEIR ISSUE | Interoperability Problem Identified | Medical Evaluation result code
6968573 | 90689097 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BG LEVELS RETURNED TO NORMAL | Hyperglycemia | Patient problem code
6968573 | 90689097 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BG LEVELS RETURNED TO NORMAL | Insufficient Information | Device problem code
6975178 | 90689110 | D | THE CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP HAD MULTIPLE SAFETY ALARMS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE INSULIN PUMP HAD A RECURRING SAFETY ALARMS THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP WAS NORMAL AND THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS AND WAS ABLE TO COMPLETE THE REWIND PROCESS CUSTOMER ALSO REPORTED HIGH BLOOD GLUCOSE OF ALMOST MGDL AD DECLINED TROUBLESHOOTING THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP IS BEING REPLACED AND IS EXPECTED TO RETURN FOR ANALYSIS | Defective Device | Device problem code
6975178 | 90689110 | D | THE CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP HAD MULTIPLE SAFETY ALARMS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS UNKNOWN AT THE TIME OF THE INCIDENT THE CUSTOMER STATED THAT THE INSULIN PUMP HAD A RECURRING SAFETY ALARMS THE CUSTOMER STATED THAT THE DRIVE SUPPORT CAP WAS NORMAL AND THAT THE INSULIN PUMP WAS NOT EXPOSED TO HIGH MAGNETIC FIELDS AND WAS ABLE TO COMPLETE THE REWIND PROCESS CUSTOMER ALSO REPORTED HIGH BLOOD GLUCOSE OF ALMOST MGDL AD DECLINED TROUBLESHOOTING THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND TO REVERT TO THE BACKUP PLAN THE INSULIN PUMP IS BEING REPLACED AND IS EXPECTED TO RETURN FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient Problem Code
6962776 | 90690038 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE UNIT WAS TESTED FOR EIGHTEEN HOURS WITHOUT ISSUE THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE DURING THE DISPLAY OF THE TRIANGLE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6962776 | 90690038 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE UNIT WAS TESTED FOR EIGHTEEN HOURS WITHOUT ISSUE THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE DURING THE DISPLAY OF THE TRIANGLE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6973253 | 90690415 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973253 | 90690415 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973254 | 90690687 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6973254 | 90690687 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | No Device Problem Found | Medical Evaluation result code
6990953 | 90692549 | N | THE EMT WAS TREATED FOR THE SCRATCHED WRIST AND WAS PUT ON LIGHT DUTY FOR THE ELBOW INJURY LATER RECEIVING AN MRI AND PHYSICAL THERAPY AS A RESULT OF THIS ALLEGED EVENT A VISUAL AND FUNCTIONAL INSPECTION WAS PERFORMED AND IT WAS IDENTIFIED THAT THERE WERE NO MALFUNCTIONS OR DEFECTS ON THE DEVICE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED EVENT | Scratched Material | Device problem code
6990953 | 90692549 | N | THE EMT WAS TREATED FOR THE SCRATCHED WRIST AND WAS PUT ON LIGHT DUTY FOR THE ELBOW INJURY LATER RECEIVING AN MRI AND PHYSICAL THERAPY AS A RESULT OF THIS ALLEGED EVENT A VISUAL AND FUNCTIONAL INSPECTION WAS PERFORMED AND IT WAS IDENTIFIED THAT THERE WERE NO MALFUNCTIONS OR DEFECTS ON THE DEVICE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED EVENT | Stress Problem Identified | Medical Evaluation result code
6962781 | 90698386 | N | THE UNIT WAS REBOOTED WHICH RESOLVED THE ISSUE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYED USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR B | Computer Software Problem | Device problem code
6962781 | 90698386 | N | THE UNIT WAS REBOOTED WHICH RESOLVED THE ISSUE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYED USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR B | Software Problem Identified | Medical Evaluation result code
6961113 | 90698537 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Apparent Adverse Event | Device problem code
6961113 | 90698537 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Device Problem Found | Medical Evaluation result code
6967938 | 90702690 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT THEIR ONETOUCH VERIO IQ METER HAD READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON ADDITIONAL INFORMATION OBTAINED BY THE CUSTOMER CARE ADVOCATE CCA DURING A FOLLOW UP CALL WITH THE PATIENT THE PATIENT REPORTED THAT THE ALLEGED INACCURACY FIRST OCCURRED AT AM ON B AT THIS TIME THE PATIENT OBTAINED BLOOD GLUCOSE RESULTS OF AND MGDL WITH THE SUBJECT METER WITHIN MINUTES OF ONE ANOTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE OF THESE GLUCOSE RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES THEIR DIABETES WITH MG GLUCOPHAGE EVERY MORNING AND DID NOT MAKE ANY CHANGES TO THEIR NORMAL DIABETES MANAGEMENT REGIMEN AS A RESULT OF THE ALLEGED PRODUCT ISSUE THE PATIENT REPORTED THAT ONE HOUR LATER THEY DEVELOPED THE SYMPTOM OF VERY WEAK AND WAS SUBSEQUENTLY TAKEN TO THE DOCTORS AT PM THE SAME DAY THE PATIENT ONLY RECEIVED A BLOOD GLUCOSE TEST AT THE DOCTORS AND A RESULT OF MGDL WAS OBTAINED ON THE DOCTORS METER AT THIS TIME NO FURTHER TREATMENT WAS REQUIRED AS A RESULT OF THIS AT THE TIME OF TROUBLESHOOTING THE CCA NOTED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER AND AN APPROVED SAMPLE SITE WAS BEING USED TO OBTAIN THE BLOOD GLUCOSE RESULTS THE PATIENT DID NOT HAVE CONTROL SOLUTION AVAILABLE TO WALK THROUGH A CONTROL SOLUTION TEST WITH THE CCA THE PATIENTS PRODUCTS WERE REPLACED BASED ON THE INFORMATION PROVIDED THERE IS NO INDICATION THAT THE SUBJECT METER CAUSED OR CONTRIBUTED TO A SERIOUS INJURY ADVERSE EVENT THE PATIENT DID NOT DEVELOP SIGNSSYMPTOMS SUGGESTIVE OF A SERIOUS INJURY NOR DID THEY RECEIVE MEDICAL INTERVENTION FOR AN ACUTE BLOOD GLUCOSE EXCURSION HOWEVER THIS COMPLAINT IS BEING REPORTED AS A MALFUNCTION BECAUSE THE CALCULATED DIFFERENCE BETWEEN THE REPORTED RESULTS FALLS OUTSIDE LFS PRECISION CRITERIA AND THE ALLEGED INACCURACY WAS NOT RESOLVED DURING TROUBLESHOOTING | Injury | Patient problem code
6967938 | 90702690 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT THEIR ONETOUCH VERIO IQ METER HAD READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON ADDITIONAL INFORMATION OBTAINED BY THE CUSTOMER CARE ADVOCATE CCA DURING A FOLLOW UP CALL WITH THE PATIENT THE PATIENT REPORTED THAT THE ALLEGED INACCURACY FIRST OCCURRED AT AM ON B AT THIS TIME THE PATIENT OBTAINED BLOOD GLUCOSE RESULTS OF AND MGDL WITH THE SUBJECT METER WITHIN MINUTES OF ONE ANOTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE OF THESE GLUCOSE RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES THEIR DIABETES WITH MG GLUCOPHAGE EVERY MORNING AND DID NOT MAKE ANY CHANGES TO THEIR NORMAL DIABETES MANAGEMENT REGIMEN AS A RESULT OF THE ALLEGED PRODUCT ISSUE THE PATIENT REPORTED THAT ONE HOUR LATER THEY DEVELOPED THE SYMPTOM OF VERY WEAK AND WAS SUBSEQUENTLY TAKEN TO THE DOCTORS AT PM THE SAME DAY THE PATIENT ONLY RECEIVED A BLOOD GLUCOSE TEST AT THE DOCTORS AND A RESULT OF MGDL WAS OBTAINED ON THE DOCTORS METER AT THIS TIME NO FURTHER TREATMENT WAS REQUIRED AS A RESULT OF THIS AT THE TIME OF TROUBLESHOOTING THE CCA NOTED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER AND AN APPROVED SAMPLE SITE WAS BEING USED TO OBTAIN THE BLOOD GLUCOSE RESULTS THE PATIENT DID NOT HAVE CONTROL SOLUTION AVAILABLE TO WALK THROUGH A CONTROL SOLUTION TEST WITH THE CCA THE PATIENTS PRODUCTS WERE REPLACED BASED ON THE INFORMATION PROVIDED THERE IS NO INDICATION THAT THE SUBJECT METER CAUSED OR CONTRIBUTED TO A SERIOUS INJURY ADVERSE EVENT THE PATIENT DID NOT DEVELOP SIGNSSYMPTOMS SUGGESTIVE OF A SERIOUS INJURY NOR DID THEY RECEIVE MEDICAL INTERVENTION FOR AN ACUTE BLOOD GLUCOSE EXCURSION HOWEVER THIS COMPLAINT IS BEING REPORTED AS A MALFUNCTION BECAUSE THE CALCULATED DIFFERENCE BETWEEN THE REPORTED RESULTS FALLS OUTSIDE LFS PRECISION CRITERIA AND THE ALLEGED INACCURACY WAS NOT RESOLVED DURING TROUBLESHOOTING | Weakness | Patient problem code
6967938 | 90702690 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS USA ALLEGING THAT THEIR ONETOUCH VERIO IQ METER HAD READ INACCURATELY ERRATIC THE COMPLAINT WAS CLASSIFIED BASED ON ADDITIONAL INFORMATION OBTAINED BY THE CUSTOMER CARE ADVOCATE CCA DURING A FOLLOW UP CALL WITH THE PATIENT THE PATIENT REPORTED THAT THE ALLEGED INACCURACY FIRST OCCURRED AT AM ON B AT THIS TIME THE PATIENT OBTAINED BLOOD GLUCOSE RESULTS OF AND MGDL WITH THE SUBJECT METER WITHIN MINUTES OF ONE ANOTHER BASED ON STATISTICAL METHODOLOGY THE CALCULATED DIFFERENCE OF THESE GLUCOSE RESULTS EXCEEDS LFS CRITERIA FOR PRECISION THE PATIENT MANAGES THEIR DIABETES WITH MG GLUCOPHAGE EVERY MORNING AND DID NOT MAKE ANY CHANGES TO THEIR NORMAL DIABETES MANAGEMENT REGIMEN AS A RESULT OF THE ALLEGED PRODUCT ISSUE THE PATIENT REPORTED THAT ONE HOUR LATER THEY DEVELOPED THE SYMPTOM OF VERY WEAK AND WAS SUBSEQUENTLY TAKEN TO THE DOCTORS AT PM THE SAME DAY THE PATIENT ONLY RECEIVED A BLOOD GLUCOSE TEST AT THE DOCTORS AND A RESULT OF MGDL WAS OBTAINED ON THE DOCTORS METER AT THIS TIME NO FURTHER TREATMENT WAS REQUIRED AS A RESULT OF THIS AT THE TIME OF TROUBLESHOOTING THE CCA NOTED THAT THE UNIT OF MEASURE WAS SET CORRECTLY ON THE SUBJECT METER AND AN APPROVED SAMPLE SITE WAS BEING USED TO OBTAIN THE BLOOD GLUCOSE RESULTS THE PATIENT DID NOT HAVE CONTROL SOLUTION AVAILABLE TO WALK THROUGH A CONTROL SOLUTION TEST WITH THE CCA THE PATIENTS PRODUCTS WERE REPLACED BASED ON THE INFORMATION PROVIDED THERE IS NO INDICATION THAT THE SUBJECT METER CAUSED OR CONTRIBUTED TO A SERIOUS INJURY ADVERSE EVENT THE PATIENT DID NOT DEVELOP SIGNSSYMPTOMS SUGGESTIVE OF A SERIOUS INJURY NOR DID THEY RECEIVE MEDICAL INTERVENTION FOR AN ACUTE BLOOD GLUCOSE EXCURSION HOWEVER THIS COMPLAINT IS BEING REPORTED AS A MALFUNCTION BECAUSE THE CALCULATED DIFFERENCE BETWEEN THE REPORTED RESULTS FALLS OUTSIDE LFS PRECISION CRITERIA AND THE ALLEGED INACCURACY WAS NOT RESOLVED DURING TROUBLESHOOTING | Incorrect, Inadequate or Imprecise Result or Readings | Device Problem Code
6968458 | 90702797 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE ANESTHETIC AGENT CASSETTE NOTED IT TO BE LEAKING THE HOSPITAL BIOMED WILL REPLACE THE CASSETTE | Leak / Splash | Device problem code
6968458 | 90702797 | N | A GEHC SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE EQUIPMENT INSPECTION OF THE ANESTHETIC AGENT CASSETTE NOTED IT TO BE LEAKING THE HOSPITAL BIOMED WILL REPLACE THE CASSETTE | Leakage/Seal | Medical Evaluation result code
6991369 | 90704781 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED PAIN FROM STIMULATION AND HAS CONSTANT URINARY TRACT INFECTIONS UTI AS A RESULT OF WHAT WAS REPORTED THEY WERE ADVISED TO SPEAK WITH THEIR HEALTHCARE PROVIDER HCP THE NEXT DAY PATIENT MENTIONED PAIN AGAIN ON B PATIENT WAS FEELING STIMULATION SIX DAYS LATER PATIENT WENT TO THEIR HCP ON B TO FIND THEY WERE RETAINING CCS AT THAT TIME THEY HAVE TO BEAR DOWN TO USE THE RESTROOM THE PATIENT IS FRUSTRATED AS THEY DIDNT REALIZE THEY HAD RETENTION THEY WERENT SURE ON THEIR IMPROVEMENT LEVEL SO STIMULATION WAS INCREASED TO MA AT THE TIME OF THE CALL ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WILL BE DONE WITH THEIR ANTIBIOTICS ON B THEY ALSO HAD A DISPUTE ABOUT THE ANESTHESIA ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WERE TAKING MEDICATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Adverse Event Without Identified Device or Use Problem | Device problem code
6991369 | 90704781 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED PAIN FROM STIMULATION AND HAS CONSTANT URINARY TRACT INFECTIONS UTI AS A RESULT OF WHAT WAS REPORTED THEY WERE ADVISED TO SPEAK WITH THEIR HEALTHCARE PROVIDER HCP THE NEXT DAY PATIENT MENTIONED PAIN AGAIN ON B PATIENT WAS FEELING STIMULATION SIX DAYS LATER PATIENT WENT TO THEIR HCP ON B TO FIND THEY WERE RETAINING CCS AT THAT TIME THEY HAVE TO BEAR DOWN TO USE THE RESTROOM THE PATIENT IS FRUSTRATED AS THEY DIDNT REALIZE THEY HAD RETENTION THEY WERENT SURE ON THEIR IMPROVEMENT LEVEL SO STIMULATION WAS INCREASED TO MA AT THE TIME OF THE CALL ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WILL BE DONE WITH THEIR ANTIBIOTICS ON B THEY ALSO HAD A DISPUTE ABOUT THE ANESTHESIA ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WERE TAKING MEDICATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Anesthesia, Insufficient, Light Or Patchy | Patient problem code
6991369 | 90704781 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED PAIN FROM STIMULATION AND HAS CONSTANT URINARY TRACT INFECTIONS UTI AS A RESULT OF WHAT WAS REPORTED THEY WERE ADVISED TO SPEAK WITH THEIR HEALTHCARE PROVIDER HCP THE NEXT DAY PATIENT MENTIONED PAIN AGAIN ON B PATIENT WAS FEELING STIMULATION SIX DAYS LATER PATIENT WENT TO THEIR HCP ON B TO FIND THEY WERE RETAINING CCS AT THAT TIME THEY HAVE TO BEAR DOWN TO USE THE RESTROOM THE PATIENT IS FRUSTRATED AS THEY DIDNT REALIZE THEY HAD RETENTION THEY WERENT SURE ON THEIR IMPROVEMENT LEVEL SO STIMULATION WAS INCREASED TO MA AT THE TIME OF THE CALL ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WILL BE DONE WITH THEIR ANTIBIOTICS ON B THEY ALSO HAD A DISPUTE ABOUT THE ANESTHESIA ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WERE TAKING MEDICATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection, Urinary Tract | Patient problem code
6991369 | 90704781 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED PAIN FROM STIMULATION AND HAS CONSTANT URINARY TRACT INFECTIONS UTI AS A RESULT OF WHAT WAS REPORTED THEY WERE ADVISED TO SPEAK WITH THEIR HEALTHCARE PROVIDER HCP THE NEXT DAY PATIENT MENTIONED PAIN AGAIN ON B PATIENT WAS FEELING STIMULATION SIX DAYS LATER PATIENT WENT TO THEIR HCP ON B TO FIND THEY WERE RETAINING CCS AT THAT TIME THEY HAVE TO BEAR DOWN TO USE THE RESTROOM THE PATIENT IS FRUSTRATED AS THEY DIDNT REALIZE THEY HAD RETENTION THEY WERENT SURE ON THEIR IMPROVEMENT LEVEL SO STIMULATION WAS INCREASED TO MA AT THE TIME OF THE CALL ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WILL BE DONE WITH THEIR ANTIBIOTICS ON B THEY ALSO HAD A DISPUTE ABOUT THE ANESTHESIA ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WERE TAKING MEDICATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Pain | Patient problem code
6991369 | 90704781 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP VIA AN ADVANCED EVALUATION TRIAL PATIENT REGARDING AN EXTERNAL NEUROSTIMULATOR ENS PATIENT REPORTED PAIN FROM STIMULATION AND HAS CONSTANT URINARY TRACT INFECTIONS UTI AS A RESULT OF WHAT WAS REPORTED THEY WERE ADVISED TO SPEAK WITH THEIR HEALTHCARE PROVIDER HCP THE NEXT DAY PATIENT MENTIONED PAIN AGAIN ON B PATIENT WAS FEELING STIMULATION SIX DAYS LATER PATIENT WENT TO THEIR HCP ON B TO FIND THEY WERE RETAINING CCS AT THAT TIME THEY HAVE TO BEAR DOWN TO USE THE RESTROOM THE PATIENT IS FRUSTRATED AS THEY DIDNT REALIZE THEY HAD RETENTION THEY WERENT SURE ON THEIR IMPROVEMENT LEVEL SO STIMULATION WAS INCREASED TO MA AT THE TIME OF THE CALL ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WILL BE DONE WITH THEIR ANTIBIOTICS ON B THEY ALSO HAD A DISPUTE ABOUT THE ANESTHESIA ON B PATIENT MENTIONED THEY SUFFER FROM UTIS AND WERE TAKING MEDICATION NO DEVICE ISSUE AND NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Infection | Patient problem code
6973429 | 90708616 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
6973429 | 90708616 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
6969214 | 90715258 | N | THE HANDLE AND BAG ARE PULLED INTO THE PURPLE TUBE AS PART OF THE PREPARATION FOR THE PROCEDURE IT IS BELIEVED THAT THE TOGGLE RELEASE BUTTON WAS INADVERTENTLY DEPRESSED DURING THE PREPARATION PROCESS THIS WOULD CAUSE THE BAG TO DISENGAGE FROM THE HANDLE PREMATURELY THE RETURNED DEVICE WAS REVIEWED AS PART OF THE COMPLAINT INVESTIGATION THE CONDITION OF THE RETURNED DEVICE WAS CONSISTENT WITH PREDEPLOYMENT OF THE TOGGLE BUTTON | Detachment of Device or device Component | Device problem code
6969214 | 90715258 | N | THE HANDLE AND BAG ARE PULLED INTO THE PURPLE TUBE AS PART OF THE PREPARATION FOR THE PROCEDURE IT IS BELIEVED THAT THE TOGGLE RELEASE BUTTON WAS INADVERTENTLY DEPRESSED DURING THE PREPARATION PROCESS THIS WOULD CAUSE THE BAG TO DISENGAGE FROM THE HANDLE PREMATURELY THE RETURNED DEVICE WAS REVIEWED AS PART OF THE COMPLAINT INVESTIGATION THE CONDITION OF THE RETURNED DEVICE WAS CONSISTENT WITH PREDEPLOYMENT OF THE TOGGLE BUTTON | Incorrect Dimension | Medical Evaluation result code
6991727 | 90716352 | D | IT WAS REPORTED THAT THE SAFETY SHIELD FELL OFF DURING USE ON A UNKNOWN ECLIPSE NEEDLE RESULTING IN NEEDLE STICK INJURY MEDICAL INTERVENTION UNKNOWN | Injury | Patient problem code
6991727 | 90716352 | D | IT WAS REPORTED THAT THE SAFETY SHIELD FELL OFF DURING USE ON A UNKNOWN ECLIPSE NEEDLE RESULTING IN NEEDLE STICK INJURY MEDICAL INTERVENTION UNKNOWN | Physical Resistance / Sticking | Device problem code
6991727 | 90716352 | D | IT WAS REPORTED THAT THE SAFETY SHIELD FELL OFF DURING USE ON A UNKNOWN ECLIPSE NEEDLE RESULTING IN NEEDLE STICK INJURY MEDICAL INTERVENTION UNKNOWN | Device Fell | Device problem code
6991807 | 90739222 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF INTERMITTENT EDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS | Scratched Material | Device problem code
6991807 | 90739222 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF INTERMITTENT EDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS | Deformation Due to Compressive Stress | Device problem code
6991807 | 90739222 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF INTERMITTENT EDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS | Material Discolored | Device problem code
6991807 | 90739222 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF INTERMITTENT EDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS | Stress Problem Identified | Medical Evaluation result code
6991807 | 90739222 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING THE EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENT OF INTERMITTENT EDEMA IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT DEVICE LABELING ADVERSE EVENTS PER TABLE TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT POSSIBLE TREATMENT SITE RESPONSES POST INJECTION WITH JUVÉDERM VOLUMA XC INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING TREATMENT SITE RESPONSES REPORTED BY OF SUBJECTS INCLUDED ACHE ACNE BULGE BUMPS CHEEK LARGER UPON WAKING UP DRY PATCH FINE WRINKLES INJECTIONNEEDLE MARKS NUMBNESS PIGMENTATION FROM TREATMENT PUFFINESS RASH SCRATCH NEAR INJECTION POINT SORENESS TIGHTNESS AND YELLOWNESS | Deformation Problem | Medical Evaluation result code
6991800 | 90739985 | N | CURRENTLY IT IS UNKNOWN TO WHAT EXTENT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT TO DATE NO MEDICAL RECORDS HAVE BEEN PROVIDED BASED ON THE INFORMATION PROVIDED IT IS UNCLEAR AT THIS TIME IF BOTH OR ONLY ONE DMAX MESH DEVICE WAS REMOVED IN ADDITION IT IS UNCLEAR IF THE HERNIA REPAIR DONE IN B WAS A REPAIR OF THE ORIGINAL LEFT ORIGINAL RIGHT DEFECT OR BOTH WITH THE CURRENT INFORMATION AVAILABLE NO DEFINITIVE CONCLUSION CAN BE MADE IF ADDITIONAL EVENT ANDOR EVALUATION INFORMATION IS OBTAINED A FOLLOW UP MDR WILL BE SUBMITTED A SECOND EMDR HAS BEEN SUBMITTED FOR THE OTHER DMAX DEVICE IMPLANTED | Defective Device | Device problem code
6991800 | 90739985 | N | CURRENTLY IT IS UNKNOWN TO WHAT EXTENT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT TO DATE NO MEDICAL RECORDS HAVE BEEN PROVIDED BASED ON THE INFORMATION PROVIDED IT IS UNCLEAR AT THIS TIME IF BOTH OR ONLY ONE DMAX MESH DEVICE WAS REMOVED IN ADDITION IT IS UNCLEAR IF THE HERNIA REPAIR DONE IN B WAS A REPAIR OF THE ORIGINAL LEFT ORIGINAL RIGHT DEFECT OR BOTH WITH THE CURRENT INFORMATION AVAILABLE NO DEFINITIVE CONCLUSION CAN BE MADE IF ADDITIONAL EVENT ANDOR EVALUATION INFORMATION IS OBTAINED A FOLLOW UP MDR WILL BE SUBMITTED A SECOND EMDR HAS BEEN SUBMITTED FOR THE OTHER DMAX DEVICE IMPLANTED | Incorrect Dimension | Medical Evaluation result code
6991910 | 90741201 | N | B ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE WHICH VESSEL WAS BLEEDING WAS THE SURGEON HAVING HEMOSTASIS ISSUES WITH MAIN VESSEL OR DURING MOBILIZATION ON WHAT VESSEL WAS THE DEVICE USED DESCRIBE THE SURGEONS STEPS FOR USE TO INCLUDE THE ADVANCEMENT OF KNIFE BLADE IN CORRELATION TO TONE TWO WHAT IS MEANT BY NOT GETTING CORRECT TONES WHAT TONES WERE HEARD ANY GENERATOR ALERT SCREENS WHAT IS APPROXIMATE BLOOD LOSS INTRAOPERATIVELY POSTOPERATIVELY WHAT WAS FOUND AT REOP WERE ANY SUBSTITUTE BLOOD PRODUCTS USED WHAT WAS THE PREOP POSTOP HEMATINIC AND HEMOGLOBIN LEVELS PLEASE CONFIRM IF THE RETURN DEVICE WAS THE FIRST OR THE SECOND DEVICE WHAT FACTORS DOES THE SURGEON BELIEVE WERE INSTRUMENTAL IN LEADING TO THE PATIENTS DEATH DOES THE PATIENT HAVE ANY KNOWN COAGULOPOLY DISORDER WAS THIS PATIENT GIVEN ANY ANTICOAGULATE TO PREVENT POSTOPERATIVE COMPLICATIONS | No Device Problem Found | Medical Evaluation result code
6991910 | 90741201 | N | B ADDITIONAL INFORMATION REQUESTED BUT UNAVAILABLE WHICH VESSEL WAS BLEEDING WAS THE SURGEON HAVING HEMOSTASIS ISSUES WITH MAIN VESSEL OR DURING MOBILIZATION ON WHAT VESSEL WAS THE DEVICE USED DESCRIBE THE SURGEONS STEPS FOR USE TO INCLUDE THE ADVANCEMENT OF KNIFE BLADE IN CORRELATION TO TONE TWO WHAT IS MEANT BY NOT GETTING CORRECT TONES WHAT TONES WERE HEARD ANY GENERATOR ALERT SCREENS WHAT IS APPROXIMATE BLOOD LOSS INTRAOPERATIVELY POSTOPERATIVELY WHAT WAS FOUND AT REOP WERE ANY SUBSTITUTE BLOOD PRODUCTS USED WHAT WAS THE PREOP POSTOP HEMATINIC AND HEMOGLOBIN LEVELS PLEASE CONFIRM IF THE RETURN DEVICE WAS THE FIRST OR THE SECOND DEVICE WHAT FACTORS DOES THE SURGEON BELIEVE WERE INSTRUMENTAL IN LEADING TO THE PATIENTS DEATH DOES THE PATIENT HAVE ANY KNOWN COAGULOPOLY DISORDER WAS THIS PATIENT GIVEN ANY ANTICOAGULATE TO PREVENT POSTOPERATIVE COMPLICATIONS | Incorrect Dimension | Medical Evaluation result code
6992290 | 90741628 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Component Missing | Device problem code
6992290 | 90741628 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT | Missing or Inadequate Safety Measures | Medical Evaluation result code
6973276 | 90746170 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973276 | 90746170 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO WEEKS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973281 | 90746246 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973281 | 90746246 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973278 | 90746331 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973278 | 90746331 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973226 | 90746344 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973226 | 90746344 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973269 | 90746384 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973269 | 90746384 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973272 | 90746399 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973272 | 90746399 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6970313 | 90747953 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6970313 | 90747953 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6980839 | 90748347 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE DIDNT DELIVER ENERGY AND DISCHARGED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Insufficient Information | Device problem code
6980839 | 90748347 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN THE DEVICE DIDNT DELIVER ENERGY AND DISCHARGED ON ITS OWN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6970218 | 90748650 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6970218 | 90748650 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6973259 | 90748816 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973259 | 90748816 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973264 | 90749132 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973264 | 90749132 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6973270 | 90749308 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6973270 | 90749308 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6970294 | 90749312 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B | Component Missing | Device problem code
6970294 | 90749312 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION HEARTWARE WILL SUBMIT A SUPPLEMENTAL REPORT WHEN NEW FACTS ARISES WHICH MATERIALLY ALTERS INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B BATTERY B DE MANUFACTURING DATE EXPIRATION DATE B | Assembly Problem Identified | Medical Evaluation result code
6970301 | 90750855 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | Device Dislodged or Dislocated | Device problem code
6970301 | 90750855 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6967157 | 90754142 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION THROUGH VISUAL INVESTIGATION IT WAS FOUND THAT AN OBSTRUCTION WAS FELT UNDER THE STRAIN RELIEF A GUIDEWIRE WILL PASS WITH FORCE THE LUMEN WAS STILL PATENT THE BALLOON PROTECTOR WAS REMOVED THE BALLOON INFLATED HELD ITS DIAMETER AND DEFLATED EXCESSIVE GLUE WAS FOUND DISTAL TO THE LUER CONSTRICTING THE INNER LUMEN | Obstruction of Flow | Device problem code
6967157 | 90754142 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION THROUGH VISUAL INVESTIGATION IT WAS FOUND THAT AN OBSTRUCTION WAS FELT UNDER THE STRAIN RELIEF A GUIDEWIRE WILL PASS WITH FORCE THE LUMEN WAS STILL PATENT THE BALLOON PROTECTOR WAS REMOVED THE BALLOON INFLATED HELD ITS DIAMETER AND DEFLATED EXCESSIVE GLUE WAS FOUND DISTAL TO THE LUER CONSTRICTING THE INNER LUMEN | Deflation Problem | Device problem code
6967157 | 90754142 | N | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION THROUGH VISUAL INVESTIGATION IT WAS FOUND THAT AN OBSTRUCTION WAS FELT UNDER THE STRAIN RELIEF A GUIDEWIRE WILL PASS WITH FORCE THE LUMEN WAS STILL PATENT THE BALLOON PROTECTOR WAS REMOVED THE BALLOON INFLATED HELD ITS DIAMETER AND DEFLATED EXCESSIVE GLUE WAS FOUND DISTAL TO THE LUER CONSTRICTING THE INNER LUMEN | Operational Problem Identified | Medical Evaluation result code
6973143 | 90754356 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT WAS HAVING ADHESIVE ISSUES WHICH LED TO THE CANNULA DISLODGING FROM THE INFUSION SITE AND THE POD FALLING OFF TO TREAT THE PATIENTS ELEVATED BG LEVELS A NEW POD WAS APPLIED | Device Dislodged or Dislocated | Device problem code
6973143 | 90754356 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT WAS HAVING ADHESIVE ISSUES WHICH LED TO THE CANNULA DISLODGING FROM THE INFUSION SITE AND THE POD FALLING OFF TO TREAT THE PATIENTS ELEVATED BG LEVELS A NEW POD WAS APPLIED | No Known Impact Or Consequence To Patient | Patient Problem Code
6973426 | 90755380 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY INSERTED INTO THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6973426 | 90755380 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY INSERTED INTO THE INFUSION SITE ARM | Positioning Problem | Device problem code
6967536 | 90757573 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE B A GETINGE FIELD SERVICE ENGINEER FSE REPORTED THAT THE BATTERIES WERE COMPLETELY DEPLETED THE FSE HAD TO INSERT THEM INTO THE PORTABLE BATTERY CHARGER IN ORDER TO GET PAST THE CHARGE AT THAT POINT THE FSE INSTALLED THEM BACK INTO THE IABP AND THEY WERE ABLE TO RECHARGE COMPLETELY THE IABP WAS RETURNED TO THE CUSTOMER AND CLEARED FOR CLINICAL USE | Battery Problem | Device Problem Code
6967536 | 90757573 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP WAS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE B A GETINGE FIELD SERVICE ENGINEER FSE REPORTED THAT THE BATTERIES WERE COMPLETELY DEPLETED THE FSE HAD TO INSERT THEM INTO THE PORTABLE BATTERY CHARGER IN ORDER TO GET PAST THE CHARGE AT THAT POINT THE FSE INSTALLED THEM BACK INTO THE IABP AND THEY WERE ABLE TO RECHARGE COMPLETELY THE IABP WAS RETURNED TO THE CUSTOMER AND CLEARED FOR CLINICAL USE | Energy Storage System Problem | Medical Evaluation result code
6974674 | 90758267 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Component Missing | Device problem code
6974674 | 90758267 | N | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY ENDSTAGE LEFT VENTRICULAR HEART FAILURE THE SYSTEM IS DESIGNED FOR INHOSPITAL AND OUTOFHOSPITAL SETTINGS MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS A BLANK UNKNOWN AND AS NO INFORMATION NI THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION | Missing or Inadequate Safety Measures | Medical Evaluation result code
6968188 | 90758365 | N | B BATCH UNK AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND UNAVAILABLE DID ANY PIECE OR PIECES OF THE FIRING KNOB FALL INTO THE PATIENT NO INFORMATION | Insufficient Information | Device problem code
6968188 | 90758365 | N | B BATCH UNK AS A LOTBATCH WAS NOT PROVIDED A DEVICE HISTORY COULD NOT BE PERFORMED ADDITIONAL INFORMATION WAS REQUESTED AND UNAVAILABLE DID ANY PIECE OR PIECES OF THE FIRING KNOB FALL INTO THE PATIENT NO INFORMATION | No Device Problem Found | Medical Evaluation result code
6988485 | 90764788 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE ILLUMINATOR WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH ERROR AND THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6988485 | 90764788 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE ILLUMINATOR WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH ERROR AND THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Loss of Power | Medical Evaluation result code
6988528 | 90764870 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE RAC WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH ERROR AND THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Failure to Power Up | Device problem code
6988528 | 90764870 | N | INTUITIVE SURGICAL INC ISI RECEIVED THE UNIT INVOLVED WITH THIS COMPLAINT AND COMPLETED THE DEVICE EVALUATION FAILURE ANALYSIS CONFIRMED THE CUSTOMER REPORTED FAILURE THE RAC WAS INSTALLED INTO A TEST SYSTEM AND IT FAILED TO POWER ON WITH ERROR AND THIS COMPLAINT IS BEING REPORTED DUE TO A DA VINCI SYSTEM MALFUNCTION RENDERING THE DA VINCI SYSTEM UNAVAILABLE FOR USE AFTER THE START OF A SURGICAL PROCEDURE ALTHOUGH NO PATIENT HARM OCCURRED IF THIS MALFUNCTION WERE TO RECUR IT COULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT | Loss of Power | Medical Evaluation result code
6970649 | 90792884 | N | THE EVENT DOCUMENTED IN THIS REPORT IS A REPEAT INSTANCE OF A MISSING HIGH VOLTAGE PUMP COVER A CAPA WAS OPENED TO ADDRESS A PRIOR REPORTABLE EVENT WHICH LED TO PREVENTIVE ACTION IN THE FORM OF A PROCEDURE UPDATE FOR PERIODIC MAINTENANCE BECAUSE OF THIS PROCEDURE UPDATE FIELD SERVICE WAS ABLE TO IDENTIFY AND RECTIFY THIS ISSUE AS PRESCRIBED BY THE NEWLY INCLUDED INSTRUCTION TO ENSURE MODULATOR SAFETY COVERS WERE INSTALLED | Component Missing | Device problem code
6970649 | 90792884 | N | THE EVENT DOCUMENTED IN THIS REPORT IS A REPEAT INSTANCE OF A MISSING HIGH VOLTAGE PUMP COVER A CAPA WAS OPENED TO ADDRESS A PRIOR REPORTABLE EVENT WHICH LED TO PREVENTIVE ACTION IN THE FORM OF A PROCEDURE UPDATE FOR PERIODIC MAINTENANCE BECAUSE OF THIS PROCEDURE UPDATE FIELD SERVICE WAS ABLE TO IDENTIFY AND RECTIFY THIS ISSUE AS PRESCRIBED BY THE NEWLY INCLUDED INSTRUCTION TO ENSURE MODULATOR SAFETY COVERS WERE INSTALLED | Component or Accessory Incompatibility | Device Problem Code
6970649 | 90792884 | N | THE EVENT DOCUMENTED IN THIS REPORT IS A REPEAT INSTANCE OF A MISSING HIGH VOLTAGE PUMP COVER A CAPA WAS OPENED TO ADDRESS A PRIOR REPORTABLE EVENT WHICH LED TO PREVENTIVE ACTION IN THE FORM OF A PROCEDURE UPDATE FOR PERIODIC MAINTENANCE BECAUSE OF THIS PROCEDURE UPDATE FIELD SERVICE WAS ABLE TO IDENTIFY AND RECTIFY THIS ISSUE AS PRESCRIBED BY THE NEWLY INCLUDED INSTRUCTION TO ENSURE MODULATOR SAFETY COVERS WERE INSTALLED | Missing or Inadequate Safety Measures | Medical Evaluation result code
6978098 | 90796864 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6978098 | 90796864 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6978146 | 90804192 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6978146 | 90804192 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6979313 | 90813971 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979313 | 90813971 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE AND A HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6979317 | 90814284 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979317 | 90814284 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6972933 | 90840725 | N | THE CUSTOMER STATED THE BSM HAD AN INTERMITTENT POWER CYCLE ISSUE HE STATES MAYBE TWICE A MONTH THE DEVICE WILL POWER OFF ON ITS OWN IT WILL POWER BACK ON BY ITSELF IT HAS ALSO BECOME FROZEN BEFORE THE CUSTOMER POWER CYCLED THE UNIT AND IT CAME BACK UP THE POWER CYCLING ISSUE HAS OCCURRED APPROXIMATELY TIMES IN THE PAST YEAR THE BIOMEDICAL ENGINEER STATES THIS DEVICE IS AT B HOSPITAL IN B DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Power Conditioning Problem | Device problem code
6972933 | 90840725 | N | THE CUSTOMER STATED THE BSM HAD AN INTERMITTENT POWER CYCLE ISSUE HE STATES MAYBE TWICE A MONTH THE DEVICE WILL POWER OFF ON ITS OWN IT WILL POWER BACK ON BY ITSELF IT HAS ALSO BECOME FROZEN BEFORE THE CUSTOMER POWER CYCLED THE UNIT AND IT CAME BACK UP THE POWER CYCLING ISSUE HAS OCCURRED APPROXIMATELY TIMES IN THE PAST YEAR THE BIOMEDICAL ENGINEER STATES THIS DEVICE IS AT B HOSPITAL IN B DUE TO THE AGE OF THIS COMPLAINT ADDITIONAL INFORMATION NECESSARY TO CONDUCT AN INVESTIGATION IS NOT READILY AVAILABLE | Loss of Power | Medical Evaluation result code
6973654 | 90843305 | N | INVESTIGATION OF THE DEVICE INCLUDING THE LOGS CONFIRMED THE REPORTED COMPLAINT DURING THE ISSUE OF A YELLOW TRIANGLE MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6973654 | 90843305 | N | INVESTIGATION OF THE DEVICE INCLUDING THE LOGS CONFIRMED THE REPORTED COMPLAINT DURING THE ISSUE OF A YELLOW TRIANGLE MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6992150 | 90844107 | N | FINAL ANALYSIS FOUND BURN MARKS ON THE CIRCUIT BOARDPRINTED CIRCUIT BOARD PCB WHICH RESULTED IN POWER UP ANOMALY | Power Problem | Device problem code
6992150 | 90844107 | N | FINAL ANALYSIS FOUND BURN MARKS ON THE CIRCUIT BOARDPRINTED CIRCUIT BOARD PCB WHICH RESULTED IN POWER UP ANOMALY | Power Source Problem Identified | Medical Evaluation result code
6975119 | 90849107 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6975119 | 90849107 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Thermal Problem | Medical Evaluation result code
6975119 | 90849107 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT THE DEVICE PASSED ALL OTHER TESTS AND WAS RETURNED TO PATIENT SERVICE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6980653 | 90851432 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION REVEALED CUTS IN THE INSULATION THE PROXIMAL END OF THE DISTAL SPRING ELECTRODE WAS SEPARATED FROM THE LEAD BODY INSULATION AND MEDICAL ADHESIVE MA WAS NOTED THE HELIX WAS EXTENDED WITH TISSUE ENTWINED IN THE HELIX AND AN ARC MARK MELTED METAL WAS NOTED ON THE DISTAL SPRING ELECTRODE MILLIMETERS MM FROM THE TIP RESISTANCE TESTS WERE COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS AN ARC MARK ON THIS PORTION OF THE LEAD INDICATED THAT THE LEADS DISTAL COIL WAS IN CLOSE PROXIMITY TO THE DEVICE CAN WHEN THE DEVICE DELIVERED THERAPY | Melted | Device problem code
6980653 | 90851432 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION REVEALED CUTS IN THE INSULATION THE PROXIMAL END OF THE DISTAL SPRING ELECTRODE WAS SEPARATED FROM THE LEAD BODY INSULATION AND MEDICAL ADHESIVE MA WAS NOTED THE HELIX WAS EXTENDED WITH TISSUE ENTWINED IN THE HELIX AND AN ARC MARK MELTED METAL WAS NOTED ON THE DISTAL SPRING ELECTRODE MILLIMETERS MM FROM THE TIP RESISTANCE TESTS WERE COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS AN ARC MARK ON THIS PORTION OF THE LEAD INDICATED THAT THE LEADS DISTAL COIL WAS IN CLOSE PROXIMITY TO THE DEVICE CAN WHEN THE DEVICE DELIVERED THERAPY | Material Separation | Device problem code
6980653 | 90851432 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY VISUAL INSPECTION REVEALED CUTS IN THE INSULATION THE PROXIMAL END OF THE DISTAL SPRING ELECTRODE WAS SEPARATED FROM THE LEAD BODY INSULATION AND MEDICAL ADHESIVE MA WAS NOTED THE HELIX WAS EXTENDED WITH TISSUE ENTWINED IN THE HELIX AND AN ARC MARK MELTED METAL WAS NOTED ON THE DISTAL SPRING ELECTRODE MILLIMETERS MM FROM THE TIP RESISTANCE TESTS WERE COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS AN ARC MARK ON THIS PORTION OF THE LEAD INDICATED THAT THE LEADS DISTAL COIL WAS IN CLOSE PROXIMITY TO THE DEVICE CAN WHEN THE DEVICE DELIVERED THERAPY | Insufficient Information | Device Problem Code
6970295 | 90852067 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION WHEN WAS THE PROCEDURE WHEN THE TISSUE PAD WAS FOUND INSIDE THE PATIENT ON B DID THE TISSUE PAD COMPLETELY DETACH COMPLETELY MISSING FROM THE CLAMP ARM OF THE DEVICE THAT WAS BEING USED DURING THIS SURGICAL PROCEDURE ONLY THE TISSUE PAD WAS FOUND AND THE DEVICE WAS ASSUMED TO BE USED IN THE HARTMANNS OPERATION IN B OR THE HEPATECTOMY IN B | Detachment of Device or device Component | Device problem code
6970295 | 90852067 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION WHEN WAS THE PROCEDURE WHEN THE TISSUE PAD WAS FOUND INSIDE THE PATIENT ON B DID THE TISSUE PAD COMPLETELY DETACH COMPLETELY MISSING FROM THE CLAMP ARM OF THE DEVICE THAT WAS BEING USED DURING THIS SURGICAL PROCEDURE ONLY THE TISSUE PAD WAS FOUND AND THE DEVICE WAS ASSUMED TO BE USED IN THE HARTMANNS OPERATION IN B OR THE HEPATECTOMY IN B | Component Missing | Device problem code
6970295 | 90852067 | N | B WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION WHEN WAS THE PROCEDURE WHEN THE TISSUE PAD WAS FOUND INSIDE THE PATIENT ON B DID THE TISSUE PAD COMPLETELY DETACH COMPLETELY MISSING FROM THE CLAMP ARM OF THE DEVICE THAT WAS BEING USED DURING THIS SURGICAL PROCEDURE ONLY THE TISSUE PAD WAS FOUND AND THE DEVICE WAS ASSUMED TO BE USED IN THE HARTMANNS OPERATION IN B OR THE HEPATECTOMY IN B | Missing or Inadequate Safety Measures | Medical Evaluation result code
6973095 | 90852074 | N | THIS KNOWN FAILURE MODE WAS ANALYZED AT THE PARENT COMPANY IN B AND ALL PARTS MET DESIGN SPECIFICATIONS HOWEVER IT WAS DETERMINED THAT THESE PARTS HAD FAILED BECAUSE THEY HAD SEEN UNUSUAL STRESS WHICH ALLOWED THEM TO FATIGUE AND BREAK IT WAS DETERMINED THAT A CHANGE IN THE MATERIAL FOR THIS COMPONENT WOULD BE ABLE TO WITHSTAND MORE SEVERE USE AND NOT BE PRONE TO FATIGUE A NEW DESIGN OF THIS PART HAS BEEN IMPLEMENTED INTO PRODUCTION AND THE DEALER HAS BEEN SUPPLIED THE REDESIGNED PART TO REPAIR THE DEVICE THE DHR FOR THIS DEVICE WAS REVIEWED AND THE WHEELCHAIR MET SPECIFICATIONS PRIOR TO DISTRIBUTION | Break | Device problem code
6973095 | 90852074 | N | THIS KNOWN FAILURE MODE WAS ANALYZED AT THE PARENT COMPANY IN B AND ALL PARTS MET DESIGN SPECIFICATIONS HOWEVER IT WAS DETERMINED THAT THESE PARTS HAD FAILED BECAUSE THEY HAD SEEN UNUSUAL STRESS WHICH ALLOWED THEM TO FATIGUE AND BREAK IT WAS DETERMINED THAT A CHANGE IN THE MATERIAL FOR THIS COMPONENT WOULD BE ABLE TO WITHSTAND MORE SEVERE USE AND NOT BE PRONE TO FATIGUE A NEW DESIGN OF THIS PART HAS BEEN IMPLEMENTED INTO PRODUCTION AND THE DEALER HAS BEEN SUPPLIED THE REDESIGNED PART TO REPAIR THE DEVICE THE DHR FOR THIS DEVICE WAS REVIEWED AND THE WHEELCHAIR MET SPECIFICATIONS PRIOR TO DISTRIBUTION | Fracture Problem | Medical Evaluation result code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Component or Accessory Incompatibility | Device problem code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Device Alarm System | Device problem code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Disconnection | Device problem code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Scratched Material | Device problem code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Stress Problem Identified | Medical Evaluation result code
6976738 | 90855324 | N | B THE INSULIN PUMP WAS RECEIVED WITH OPERATING CURRENTS WITHIN SPEC THE PUMP PASSED SELF TEST REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR AND POWER LOSS ALARMS WERE TRIGGERED WHEN BACKUP BATTERY LOADED VOLTAGE LOADED V LITH WAS LESS THAT V FOR CONSECUTIVE HOURS DUE TO RESISTANCE ISSUE AT J CONNECTOR AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J PCBA PUMP WAS MONITORED AND FUNCTIONED PROPERLY THE PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Interoperability Problem Identified | Medical Evaluation result code
6974746 | 90856062 | N | THE CUSTOMER REPORTED THAT THE DISPLAY ON THE CENTRAL NURSES STATION CNS WENT BLANK POWERCYCLING THE CNS RESOLVED THE ISSUE | No Display / Image | Device Problem Code
6974746 | 90856062 | N | THE CUSTOMER REPORTED THAT THE DISPLAY ON THE CENTRAL NURSES STATION CNS WENT BLANK POWERCYCLING THE CNS RESOLVED THE ISSUE | Image Artifact | Medical Evaluation result code
6974745 | 90856251 | N | THE CUSTOMER REPORTED THAT AFTER A POWER OUTAGE THE UPS WOULD NOT HOLD A CHARGE WHICH CAUSED THE CENTRAL NURSES STATION CNS TO SHUT DOWN THE CENTRAL NURSES STATION CNS THEN POWERED UP TO A BLUE SCREEN THE CUSTOMER WILL CALL BACK IF THEY DECIDE TO SEND THE UNIT IN FOR REPAIR | Charging Problem | Device problem code
6974745 | 90856251 | N | THE CUSTOMER REPORTED THAT AFTER A POWER OUTAGE THE UPS WOULD NOT HOLD A CHARGE WHICH CAUSED THE CENTRAL NURSES STATION CNS TO SHUT DOWN THE CENTRAL NURSES STATION CNS THEN POWERED UP TO A BLUE SCREEN THE CUSTOMER WILL CALL BACK IF THEY DECIDE TO SEND THE UNIT IN FOR REPAIR | Energy Storage System Problem | Medical Evaluation result code
6974738 | 90856875 | N | THE CUSTOMER REPORTED THAT AFTER A POWER OUTAGE THE UPS WOULD NOT HOLD A CHARGE WHICH CAUSED THE CENTRAL NURSES STATION CNS TO SHUT DOWN THE CENTRAL NURSES STATION CNS THEN POWERED UP TO A BLUE SCREEN THE CUSTOMER LATER REPORTED THAT THE HARD DISK DRIVE HDD WAS SALVAGEABLE AND THEY WERE ABLE TO RUN A SCAN DISC RECOVERY THE UNIT IS NOW WORKING WITH NO FURTHER ISSUES | Charging Problem | Device problem code
6974738 | 90856875 | N | THE CUSTOMER REPORTED THAT AFTER A POWER OUTAGE THE UPS WOULD NOT HOLD A CHARGE WHICH CAUSED THE CENTRAL NURSES STATION CNS TO SHUT DOWN THE CENTRAL NURSES STATION CNS THEN POWERED UP TO A BLUE SCREEN THE CUSTOMER LATER REPORTED THAT THE HARD DISK DRIVE HDD WAS SALVAGEABLE AND THEY WERE ABLE TO RUN A SCAN DISC RECOVERY THE UNIT IS NOW WORKING WITH NO FURTHER ISSUES | Energy Storage System Problem | Medical Evaluation result code
6982349 | 90857242 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON AC POWER ADOPTER AND THE MAIN CIRCUIT BOARDPRINTED CIRCUIT BOARD PCB WHICH RESULTED IN A NO POWER ISSUE | Power Source Problem Identified | Medical Evaluation result code
6982349 | 90857242 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON AC POWER ADOPTER AND THE MAIN CIRCUIT BOARDPRINTED CIRCUIT BOARD PCB WHICH RESULTED IN A NO POWER ISSUE | Electrical Power Problem | Device problem code
6982349 | 90857242 | N | NO COMPLAINT WAS RECEIVED WITH RETURN OF THE DEVICE FAILURE EVENT OBSERVED DURING ANALYSIS FINAL ANALYSIS FOUND BURN MARKS ON AC POWER ADOPTER AND THE MAIN CIRCUIT BOARDPRINTED CIRCUIT BOARD PCB WHICH RESULTED IN A NO POWER ISSUE | Loss of Power | Medical Evaluation result code
6983152 | 90866533 | N | B PRELIMINARY INVESTIGATION OF THE RETURNED DEVICE SAMPLE INDICATES A SHEATH VALVE ASSEMBLY SEPARATION | Fracture Problem | Medical Evaluation result code
6979506 | 90867027 | N | B RESULTS OF INVESTIGATION VYAIRE MEDICAL WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE DURING TESTING AND EVALUATION DURING AN INITIAL BENCH TEST OF THE LTV SPRINTPACK POWER MANAGER THE SPRINTPACK POWER MANAGER WOULD NOT POWER UP THE GOLDEN TEST VENTILATOR FROM THE EXTERNAL POWER CONNECTOR THE SPRINTPACK POWER MANAGER FAILED THE BATTERY CHARGING TEST BUT WOULD STILL POWER UP NORMALLY FROM THE BATTERIES VISUAL INSPECTION DID NOT REVEAL ANY ANOMALIES WITH THE SEAL HOWEVER FURTHER INTERNAL INVESTIGATION FOUND THE MICRO POWER BOARD OF THE INPUT PCBA HAD OVER HEATED AND BURNED | Charging Problem | Device problem code
6979506 | 90867027 | N | B RESULTS OF INVESTIGATION VYAIRE MEDICAL WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE DURING TESTING AND EVALUATION DURING AN INITIAL BENCH TEST OF THE LTV SPRINTPACK POWER MANAGER THE SPRINTPACK POWER MANAGER WOULD NOT POWER UP THE GOLDEN TEST VENTILATOR FROM THE EXTERNAL POWER CONNECTOR THE SPRINTPACK POWER MANAGER FAILED THE BATTERY CHARGING TEST BUT WOULD STILL POWER UP NORMALLY FROM THE BATTERIES VISUAL INSPECTION DID NOT REVEAL ANY ANOMALIES WITH THE SEAL HOWEVER FURTHER INTERNAL INVESTIGATION FOUND THE MICRO POWER BOARD OF THE INPUT PCBA HAD OVER HEATED AND BURNED | Power Problem | Device problem code
6979506 | 90867027 | N | B RESULTS OF INVESTIGATION VYAIRE MEDICAL WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE DURING TESTING AND EVALUATION DURING AN INITIAL BENCH TEST OF THE LTV SPRINTPACK POWER MANAGER THE SPRINTPACK POWER MANAGER WOULD NOT POWER UP THE GOLDEN TEST VENTILATOR FROM THE EXTERNAL POWER CONNECTOR THE SPRINTPACK POWER MANAGER FAILED THE BATTERY CHARGING TEST BUT WOULD STILL POWER UP NORMALLY FROM THE BATTERIES VISUAL INSPECTION DID NOT REVEAL ANY ANOMALIES WITH THE SEAL HOWEVER FURTHER INTERNAL INVESTIGATION FOUND THE MICRO POWER BOARD OF THE INPUT PCBA HAD OVER HEATED AND BURNED | Failure to Power Up | Device problem code
6979506 | 90867027 | N | B RESULTS OF INVESTIGATION VYAIRE MEDICAL WAS ABLE TO VERIFY THE CUSTOMERS REPORTED ISSUE DURING TESTING AND EVALUATION DURING AN INITIAL BENCH TEST OF THE LTV SPRINTPACK POWER MANAGER THE SPRINTPACK POWER MANAGER WOULD NOT POWER UP THE GOLDEN TEST VENTILATOR FROM THE EXTERNAL POWER CONNECTOR THE SPRINTPACK POWER MANAGER FAILED THE BATTERY CHARGING TEST BUT WOULD STILL POWER UP NORMALLY FROM THE BATTERIES VISUAL INSPECTION DID NOT REVEAL ANY ANOMALIES WITH THE SEAL HOWEVER FURTHER INTERNAL INVESTIGATION FOUND THE MICRO POWER BOARD OF THE INPUT PCBA HAD OVER HEATED AND BURNED | Thermal Problem | Medical Evaluation result code
6978075 | 90869608 | D | PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6978075 | 90869608 | D | PATIENTS BLOOD GLUCOSE BG REACHED GREATER THAN MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6974806 | 90869633 | N | THE PATIENT HAD TWO X SOLSTICE SCREW DISENGAGE ON THE RIGHT C LEVEL AND THE LEFT C | Detachment of Device or device Component | Device problem code
6974806 | 90869633 | N | THE PATIENT HAD TWO X SOLSTICE SCREW DISENGAGE ON THE RIGHT C LEVEL AND THE LEFT C | Interoperability Problem Identified | Medical Evaluation result code
6983174 | 90871995 | N | B PRELIMINARY INVESTIGATION OF THE RETURNED DEVICE SAMPLE INDICATES A SHEATH VALVE ASSEMBLY SEPARATION | Material Separation | Device problem code
6983174 | 90871995 | N | B PRELIMINARY INVESTIGATION OF THE RETURNED DEVICE SAMPLE INDICATES A SHEATH VALVE ASSEMBLY SEPARATION | Fracture Problem | Medical Evaluation result code
6979301 | 90872221 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6979301 | 90872221 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Non-Functional Defect | Medical Evaluation result code
6979400 | 90872931 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979400 | 90872931 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6979326 | 90873715 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979326 | 90873715 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6975096 | 90879087 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT DURING THE DISPLAY OF THE TRIANGLE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6975096 | 90879087 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT DURING THE DISPLAY OF THE TRIANGLE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6979309 | 90879179 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6979309 | 90879179 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Interoperability Problem Identified | Medical Evaluation result code
6979444 | 90879214 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANTS FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979444 | 90879214 | N | ACCORDING TO THE PRACTITIONER TWO IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE TWO IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANTS FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6979319 | 90879289 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979319 | 90879289 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO AND HALF MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6982558 | 90881388 | N | FINAL ANALYSIS FOUND THAT EPOXY DEBRIS WITHIN THE CONNECTOR BLOCK OF THE HEADER WAS THE CAUSE OF THE ATRIAL SETSCREW BEING FIRMLY STUCK AND LOCKED IN PLACE THE EPOXY DEBRIS FOUND WITHIN THE THREADS OF THE SET SCREW IS A RESULT OF A MANUFACTURING ANOMALY | Mechanical Jam | Device problem code
6982558 | 90881388 | N | FINAL ANALYSIS FOUND THAT EPOXY DEBRIS WITHIN THE CONNECTOR BLOCK OF THE HEADER WAS THE CAUSE OF THE ATRIAL SETSCREW BEING FIRMLY STUCK AND LOCKED IN PLACE THE EPOXY DEBRIS FOUND WITHIN THE THREADS OF THE SET SCREW IS A RESULT OF A MANUFACTURING ANOMALY | Mechanical Problem Identified | Medical Evaluation result code
6983116 | 90882306 | N | B PRELIMINARY INVESTIGATION OF THE RETURNED SAMPLE INDICATES THAT THE SHEATH VALVE LEAKED IN USE | Leak / Splash | Device problem code
6983116 | 90882306 | N | B PRELIMINARY INVESTIGATION OF THE RETURNED SAMPLE INDICATES THAT THE SHEATH VALVE LEAKED IN USE | Leakage/Seal | Medical Evaluation result code
6977907 | 90882476 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | Device Dislodged or Dislocated | Device problem code
6977907 | 90882476 | D | PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS THE POD FELL OFF WHICH RESULTED IN THE CANNULA DISLODGING FROM THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6978144 | 90883395 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6978144 | 90883395 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6973534 | 90888072 | N | THE DEVICE LOGS WERE REVIEWED AND THE REPORTED ISSUE WAS CONFIRMED AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6973534 | 90888072 | N | THE DEVICE LOGS WERE REVIEWED AND THE REPORTED ISSUE WAS CONFIRMED AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6983316 | 90891041 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN DURING FLIGHT THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT INDICATED THAT THE CLINICIAN RESEATED THE BATTERY TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION COMPLAINANT INDICATED THAT SUBSEQUENT TESTING WAS ABLE TO INTERMITTENTLY DUPLICATE THE REPORTED MALFUNCTION | Defective Device | Device Problem Code
6983316 | 90891041 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE GENDER UNKNOWN DURING FLIGHT THE DEVICE WOULD RESTARTREBOOT ITSELF COMPLAINANT INDICATED THAT THE CLINICIAN RESEATED THE BATTERY TO CONTINUE TREATING THE PATIENT COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION COMPLAINANT INDICATED THAT SUBSEQUENT TESTING WAS ABLE TO INTERMITTENTLY DUPLICATE THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
6985093 | 90893329 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY WAS FUZZY AND NO CLINICAL DATA COULD BE SEEN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Display / Image | Device Problem Code
6985093 | 90893329 | D | COMPLAINANT ALLEGED THAT THE DEVICES DISPLAY WAS FUZZY AND NO CLINICAL DATA COULD BE SEEN COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6979246 | 90894281 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Loss of Osseointegration | Device problem code
6979246 | 90894281 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT IS LOST LOSS OF OSSEOINTEGRATION AFTER IMPLANT HAS BEEN RESTORED THE IMPLANT HAS BEEN PLACED IN POSITION ON B AND REMOVED ON B | Interoperability Problem Identified | Medical Evaluation result code
6973277 | 90905798 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND FOUND FAULT CODES AND IN THE FAULT LOGS THE FSE RELOADED THE SOFTWARE TO RESOLVE THIS ISSUE THE IABP WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED | Component or Accessory Incompatibility | Device Problem Code
6973277 | 90905798 | N | THE PRODUCTION DEVICE HISTORY RECORD DHR FOR THIS INTRAAORTIC BALLOON PUMP IABP IS NOT REQUIRED TO BE REVIEWED PER COMPANY STANDARD OPERATING PROCEDURE SINCE THE DEVICE MANUFACTURE DATE IS GREATER THAN ONE YEAR FROM THE EVENT DATE A GETINGE FIELD SERVICE ENGINEER FSE WAS DISPATCHED AND FOUND FAULT CODES AND IN THE FAULT LOGS THE FSE RELOADED THE SOFTWARE TO RESOLVE THIS ISSUE THE IABP WAS RETURNED TO THE CUSTOMER CLEARED FOR CLINICAL USE IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED | Interoperability Problem Identified | Medical Evaluation result code
6978149 | 90909726 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6978149 | 90909726 | D | PATIENTS BLOOD GLUCOSE REACHED BETWEEN AND MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6979386 | 90943090 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979386 | 90943090 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6979396 | 90943174 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6979396 | 90943174 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6977786 | 90949566 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM AS PAIN WAS SENSED | No Known Impact Or Consequence To Patient | Patient Problem Code
6977786 | 90949566 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM AS PAIN WAS SENSED | Positioning Problem | Device problem code
6985221 | 90954160 | N | NO PATIENT INFORMATION PROVIDED AS NO PATIENT WAS INVOLVED IN THIS CONCERN A MEDTRONIC REPRESENTATIVE INSPECTED THE IMAGING SYSTEM ONSITE AND THE SYSTEMS GENERATOR DID NOT ACHIEVE A READY STATE UPON STARTUP UPON STARTUP OF THE SYSTEM A CONSTANT AUDIBLE BEEPING WAS PRESENT FROM THE IMAGE ACQUISITION SYSTEM IAS GENERATOR CABINET THE XRAY INDICATOR ON THE PENDANT DID NOT ACHIEVE A READY STATE REMOTE DESKTOP SHOWED GENERATOR STATUS AS NOT READY THE GENERATOR WAS ACCESSED AND TROUBLE SHOOTING REVEALED THAT THE J CONNECTOR WAS PARTIALLY UNSEATED THE J CONNECTOR WAS RESEATED ON THE ATP PCB AND SUBSEQUENT STARTUP OF SYSTEM WAS SUCCESSFUL BOOTING INTO D MODE WITH NO BEEPING HEARD FROM GENERATOR MULTIPLE SYSTEM STARTUPS YIELDED SAME SUCCESSFUL RESULTS SYSTEM CHECKOUT PERFORMED WITH SATISFACTORY RESULTS THE SYSTEM RETURNED TO SERVICE NO PARTS WERE RETURNED TO THE MANUFACTURER FOR EVALUATION | Noise, Audible | Device problem code
6985221 | 90954160 | N | NO PATIENT INFORMATION PROVIDED AS NO PATIENT WAS INVOLVED IN THIS CONCERN A MEDTRONIC REPRESENTATIVE INSPECTED THE IMAGING SYSTEM ONSITE AND THE SYSTEMS GENERATOR DID NOT ACHIEVE A READY STATE UPON STARTUP UPON STARTUP OF THE SYSTEM A CONSTANT AUDIBLE BEEPING WAS PRESENT FROM THE IMAGE ACQUISITION SYSTEM IAS GENERATOR CABINET THE XRAY INDICATOR ON THE PENDANT DID NOT ACHIEVE A READY STATE REMOTE DESKTOP SHOWED GENERATOR STATUS AS NOT READY THE GENERATOR WAS ACCESSED AND TROUBLE SHOOTING REVEALED THAT THE J CONNECTOR WAS PARTIALLY UNSEATED THE J CONNECTOR WAS RESEATED ON THE ATP PCB AND SUBSEQUENT STARTUP OF SYSTEM WAS SUCCESSFUL BOOTING INTO D MODE WITH NO BEEPING HEARD FROM GENERATOR MULTIPLE SYSTEM STARTUPS YIELDED SAME SUCCESSFUL RESULTS SYSTEM CHECKOUT PERFORMED WITH SATISFACTORY RESULTS THE SYSTEM RETURNED TO SERVICE NO PARTS WERE RETURNED TO THE MANUFACTURER FOR EVALUATION | No Device Problem Found | Medical Evaluation result code
6978038 | 90955549 | D | PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK PATIENT SELFTREATED ELEVATED BG LEVELS BY CHANGING TO A NEW POD | No Known Impact Or Consequence To Patient | Patient Problem Code
6978038 | 90955549 | D | PATIENTS BLOOD GLUCOSE BG REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE BACK PATIENT SELFTREATED ELEVATED BG LEVELS BY CHANGING TO A NEW POD | Positioning Problem | Device problem code
6991428 | 90960643 | N | B INITIAL EVALUATION OF RETURNED DEVICE SAMPLE INDICATES THAT THE SPRING WIRE GUIDE UNRAVELED | Unraveled Material | Device problem code
6991428 | 90960643 | N | B INITIAL EVALUATION OF RETURNED DEVICE SAMPLE INDICATES THAT THE SPRING WIRE GUIDE UNRAVELED | No Device Problem Found | Medical Evaluation result code
6982449 | 90965260 | N | THE DEVICE WAS RETURNED TO RESMED AND AN EVALUATION WAS PERFORMED THE EVALUATION DETERMINED THAT THE ERROR MESSAGES SF SF SF AND SF WAS DUE TO A CONTAMINATED PNEUMATIC BLOCK RESMEDS RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE RESMED REFERENCE B | Device Displays Incorrect Message | Device problem code
6982449 | 90965260 | N | THE DEVICE WAS RETURNED TO RESMED AND AN EVALUATION WAS PERFORMED THE EVALUATION DETERMINED THAT THE ERROR MESSAGES SF SF SF AND SF WAS DUE TO A CONTAMINATED PNEUMATIC BLOCK RESMEDS RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE RESMED REFERENCE B | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6981875 | 90981460 | N | THIS KNOWN FAILURE MODE WAS ANALYZED AT THE PARENT COMPANY IN SWEDEN AND ALL PARTS MET DESIGN SPECIFICATIONS IT WAS HOWEVER NOTED THAT THESE PARTS HAD FAILED BECAUSE THEY HAD SEEN UNUSUAL STRESS WHICH ALLOWED THEM TO FATIGUE AND BREAK OVER TIME IT WAS DECIDED THAT THE PART COULD BE CHANGED TO A DIFFERENT MATERIAL THAT WOULD BE ABLE TO WITHSTAND MORE SEVERE USE AND NOT FATIGUE A NEW DESIGN OF THIS PART HAS BEEN IMPLEMENTED INTO PRODUCTION THE REDESIGNED PART HAS BEEN ISSUED TO THE CLIENT TO ADDRESS THIS REPORTED FAILURE THE DHR FOR THIS DEVICE WAS REVIEWED AND CHAIR MET SPECIFICATIONS PRIOR TO DISTRIBUTION | Break | Device problem code
6981875 | 90981460 | N | THIS KNOWN FAILURE MODE WAS ANALYZED AT THE PARENT COMPANY IN SWEDEN AND ALL PARTS MET DESIGN SPECIFICATIONS IT WAS HOWEVER NOTED THAT THESE PARTS HAD FAILED BECAUSE THEY HAD SEEN UNUSUAL STRESS WHICH ALLOWED THEM TO FATIGUE AND BREAK OVER TIME IT WAS DECIDED THAT THE PART COULD BE CHANGED TO A DIFFERENT MATERIAL THAT WOULD BE ABLE TO WITHSTAND MORE SEVERE USE AND NOT FATIGUE A NEW DESIGN OF THIS PART HAS BEEN IMPLEMENTED INTO PRODUCTION THE REDESIGNED PART HAS BEEN ISSUED TO THE CLIENT TO ADDRESS THIS REPORTED FAILURE THE DHR FOR THIS DEVICE WAS REVIEWED AND CHAIR MET SPECIFICATIONS PRIOR TO DISTRIBUTION | Fracture Problem | Medical Evaluation result code
6987684 | 90981692 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987684 | 90981692 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987746 | 90981830 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B SIX MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987746 | 90981830 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B SIX MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987749 | 90982805 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B SIX MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987749 | 90982805 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B SIX MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987674 | 90983390 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987674 | 90983390 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B FIVE DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987771 | 90983409 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987771 | 90983409 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B THREE MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987727 | 90983620 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987727 | 90983620 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987679 | 90983971 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON SEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987679 | 90983971 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON SEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6989530 | 90987861 | D | IT WAS REPORTED THAT A SPECTRUM PUMP FAILED DOWNSTREAM OCCLUSION TESTING THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | Component or Accessory Incompatibility | Device problem code
6989530 | 90987861 | D | IT WAS REPORTED THAT A SPECTRUM PUMP FAILED DOWNSTREAM OCCLUSION TESTING THERE WAS NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Patient Involvement | Patient problem code
6979576 | 91007150 | N | AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS CONFIRMED THE REPORTED COMPLAINT DURING THE DISPLAY OF THE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6979576 | 91007150 | N | AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS CONFIRMED THE REPORTED COMPLAINT DURING THE DISPLAY OF THE ERROR MESSAGE THE ARKON AUTOMATICALLY ACTIVATES THE EMERGENCY OXYGEN AND THE LICENSED CLINICIAN WHO IS IN CONSTANT ATTENDANCE CAN CONTINUE TO MANUALLY VENTILATE THE PATIENT WHILE POWER CYCLING THE DEVICE WHICH RESOLVES THE ISSUE THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM TO THE PATIENT AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6985492 | 91011134 | N | B THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHED LCD WINDOW | Component Missing | Device problem code
6985492 | 91011134 | N | B THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6985492 | 91011134 | N | B THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHED LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
6985492 | 91011134 | N | B THE INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE RETAINER MISSING RESERVOIR TUBE ORING AND MINOR SCRATCHED LCD WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
6979113 | 91022010 | N | THE AFFECTED DEVICE WAS PRODUCED IN MAY AND STILL HAD THE INITIAL SOFTWARE VERSION INSTALLED THE REPORTED ERROR CODE COULD BE CONFIRMED BY THE LOG ENTRIES AND INDICATES AN ISOLATED FAILURE OF A FLASH EPROM ELECTRONIC MEMORY MODULE DUE TO THIS ERROR THE PCB CPU WAS NO LONGER WITHIN TOLERANCE THIS WAS DETECTED AS SPECIFIED BY THE INTERNAL SAFETY SOFTWARE OF THE DEVICE AND INITIATED WARM START IN AN ATTEMPT TO CORRECT THE ISSUE DURING A WARM START THE DEVICE OPENS THE AIRWAY SYSTEM TO AMBIENT TO ALLOW SPONTANEOUS BREATHING THIS IS ACCOMPANIED BY AN ACOUSTIC ALARM | Flare or Flash | Device problem code
6979113 | 91022010 | N | THE AFFECTED DEVICE WAS PRODUCED IN MAY AND STILL HAD THE INITIAL SOFTWARE VERSION INSTALLED THE REPORTED ERROR CODE COULD BE CONFIRMED BY THE LOG ENTRIES AND INDICATES AN ISOLATED FAILURE OF A FLASH EPROM ELECTRONIC MEMORY MODULE DUE TO THIS ERROR THE PCB CPU WAS NO LONGER WITHIN TOLERANCE THIS WAS DETECTED AS SPECIFIED BY THE INTERNAL SAFETY SOFTWARE OF THE DEVICE AND INITIATED WARM START IN AN ATTEMPT TO CORRECT THE ISSUE DURING A WARM START THE DEVICE OPENS THE AIRWAY SYSTEM TO AMBIENT TO ALLOW SPONTANEOUS BREATHING THIS IS ACCOMPANIED BY AN ACOUSTIC ALARM | Device Displays Incorrect Message | Device problem code
6979113 | 91022010 | N | THE AFFECTED DEVICE WAS PRODUCED IN MAY AND STILL HAD THE INITIAL SOFTWARE VERSION INSTALLED THE REPORTED ERROR CODE COULD BE CONFIRMED BY THE LOG ENTRIES AND INDICATES AN ISOLATED FAILURE OF A FLASH EPROM ELECTRONIC MEMORY MODULE DUE TO THIS ERROR THE PCB CPU WAS NO LONGER WITHIN TOLERANCE THIS WAS DETECTED AS SPECIFIED BY THE INTERNAL SAFETY SOFTWARE OF THE DEVICE AND INITIATED WARM START IN AN ATTEMPT TO CORRECT THE ISSUE DURING A WARM START THE DEVICE OPENS THE AIRWAY SYSTEM TO AMBIENT TO ALLOW SPONTANEOUS BREATHING THIS IS ACCOMPANIED BY AN ACOUSTIC ALARM | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6983234 | 91075204 | D | THE PATIENT REPORTED THAT SHE DID NOT BELIEVE THE CANNULA WAS PROPERLY INSERTED DUE TO HER BLOOD GLUCOSE BEING HIGH ALL DAY HER BLOOD GLUCOSE REACHED ABOVE MGDL WHILE WEARING THE POD ON HER LEG | Positioning Problem | Device problem code
6983234 | 91075204 | D | THE PATIENT REPORTED THAT SHE DID NOT BELIEVE THE CANNULA WAS PROPERLY INSERTED DUE TO HER BLOOD GLUCOSE BEING HIGH ALL DAY HER BLOOD GLUCOSE REACHED ABOVE MGDL WHILE WEARING THE POD ON HER LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
6982784 | 91075366 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS DUE TO THE SMELL OF INSULIN AND ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | Positioning Problem | Device problem code
6982784 | 91075366 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS DUE TO THE SMELL OF INSULIN AND ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
6989595 | 91077148 | D | A PATIENTS BLOOD GLUCOSE BG REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT SELFTREATED ELEVATED BG LEVELS BY CHANGING TO A NEW POD AND INJECTING UNITS OF INSULIN CORRECTION THE CANNULA DID NOT COMPLETELY DEPLOY FROM THE POD INDICATING A NEEDLE MECHANISM FAILURE | No Known Impact Or Consequence To Patient | Patient Problem Code
6989595 | 91077148 | D | A PATIENTS BLOOD GLUCOSE BG REACHED THE S MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT SELFTREATED ELEVATED BG LEVELS BY CHANGING TO A NEW POD AND INJECTING UNITS OF INSULIN CORRECTION THE CANNULA DID NOT COMPLETELY DEPLOY FROM THE POD INDICATING A NEEDLE MECHANISM FAILURE | Activation, Positioning or Separation Problem | Device problem code
6986516 | 91081519 | N | OF THE ALLEGATIONS OF A MALFUNCTION PUMPS WERE RETURNED TO ANIMAS AND INVESTIGATED OF THE INVESTIGATED PUMPS WERE FOUND TO BE MALFUNCTIONING DUE TO A BOLUS BUTTON DEFECT WITH MOISTURE IN THE PUMP DURING INVESTIGATION FOR UNRELATED COMPLAINTS THERE WERE ADDITIONAL FAILURES FOUND B | Defective Device | Device problem code
6986516 | 91081519 | N | OF THE ALLEGATIONS OF A MALFUNCTION PUMPS WERE RETURNED TO ANIMAS AND INVESTIGATED OF THE INVESTIGATED PUMPS WERE FOUND TO BE MALFUNCTIONING DUE TO A BOLUS BUTTON DEFECT WITH MOISTURE IN THE PUMP DURING INVESTIGATION FOR UNRELATED COMPLAINTS THERE WERE ADDITIONAL FAILURES FOUND B | Moisture or Humidity Problem | Device problem code
6986516 | 91081519 | N | OF THE ALLEGATIONS OF A MALFUNCTION PUMPS WERE RETURNED TO ANIMAS AND INVESTIGATED OF THE INVESTIGATED PUMPS WERE FOUND TO BE MALFUNCTIONING DUE TO A BOLUS BUTTON DEFECT WITH MOISTURE IN THE PUMP DURING INVESTIGATION FOR UNRELATED COMPLAINTS THERE WERE ADDITIONAL FAILURES FOUND B | Storage Problem Identified | Medical Evaluation result code
6989592 | 91082092 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | Positioning Problem | Device problem code
6989592 | 91082092 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS STILL PROPERLY INSERTED INTO THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
6989624 | 91085409 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
6989624 | 91085409 | D | A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
6983031 | 91097746 | D | IT WAS REPORTED THAT THE DEVICE LACKED LUBRICITY THE PATIENT EXPERIENCED A BURNING SENSATION MEDICAL INTERVENTION WAS NOT NEEDED | Insufficient Information | Device problem code
6983031 | 91097746 | D | IT WAS REPORTED THAT THE DEVICE LACKED LUBRICITY THE PATIENT EXPERIENCED A BURNING SENSATION MEDICAL INTERVENTION WAS NOT NEEDED | Burning Sensation | Patient problem code
6976448 | 91099097 | N | B BATCH PJY DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PLEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Material Twisted / Bent | Device problem code
6976448 | 91099097 | N | B BATCH PJY DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PLEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Mechanical Jam | Device problem code
6976448 | 91099097 | N | B BATCH PJY DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PLEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Device Dislodged or Dislocated | Device problem code
6976448 | 91099097 | N | B BATCH PJY DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PLEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Mechanical Problem Identified | Medical Evaluation result code
6991074 | 91103187 | N | B UNIT RECEIVED WITH DOG CHEW MARKS AT RESERVOIR TUBE AREA AND ONAROUND LCD DISPLAY WINDOW AREA UNIT RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING DEEP MINOR SCRATCHES ON LCD WINDOW MISSING DISPLAY WINDOW COVER PILLOWING KEYPAD OVERLAY AND SLIGHTLY PEELING END CAP ADDRESS LABEL | Scratched Material | Device problem code
6991074 | 91103187 | N | B UNIT RECEIVED WITH DOG CHEW MARKS AT RESERVOIR TUBE AREA AND ONAROUND LCD DISPLAY WINDOW AREA UNIT RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING DEEP MINOR SCRATCHES ON LCD WINDOW MISSING DISPLAY WINDOW COVER PILLOWING KEYPAD OVERLAY AND SLIGHTLY PEELING END CAP ADDRESS LABEL | Peeled / Delaminated | Device problem code
6991074 | 91103187 | N | B UNIT RECEIVED WITH DOG CHEW MARKS AT RESERVOIR TUBE AREA AND ONAROUND LCD DISPLAY WINDOW AREA UNIT RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING DEEP MINOR SCRATCHES ON LCD WINDOW MISSING DISPLAY WINDOW COVER PILLOWING KEYPAD OVERLAY AND SLIGHTLY PEELING END CAP ADDRESS LABEL | Component Missing | Device problem code
6991074 | 91103187 | N | B UNIT RECEIVED WITH DOG CHEW MARKS AT RESERVOIR TUBE AREA AND ONAROUND LCD DISPLAY WINDOW AREA UNIT RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER UNIT RECEIVED WITH MISSING RESERVOIR TUBE ORING SCRATCHED CASING DEEP MINOR SCRATCHES ON LCD WINDOW MISSING DISPLAY WINDOW COVER PILLOWING KEYPAD OVERLAY AND SLIGHTLY PEELING END CAP ADDRESS LABEL | Missing or Inadequate Safety Measures | Medical Evaluation result code
6981874 | 91112190 | N | SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND CONFIRMED THE REPORTED COMPLAINT POWER CYCLING THE DEVICE RESOLVED THE ISSUE AND THE DEVICE WAS RETURNED TO USE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6981874 | 91112190 | N | SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND CONFIRMED THE REPORTED COMPLAINT POWER CYCLING THE DEVICE RESOLVED THE ISSUE AND THE DEVICE WAS RETURNED TO USE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6981864 | 91113147 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND CONFIRMED THE REPORTED COMPLAINT POWER CYCLING THE DEVICE RESOLVED THE ISSUE AND THE DEVICE WAS RETURNED TO USE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6981864 | 91113147 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND CONFIRMED THE REPORTED COMPLAINT POWER CYCLING THE DEVICE RESOLVED THE ISSUE AND THE DEVICE WAS RETURNED TO USE AS THE PROBLEM WAS DISCOVERED PRIOR TO USE AND A WARNING MESSAGE DISPLAYS USE OF THE DEVICE IS PREVENTED UNTIL THE ISSUE IS RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6979537 | 91113319 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT RESETTING THE DEVICE AND PERFORMING THE CHECKOUT PROCEDURE AGAIN RESOLVED THE ISSUE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Device Displays Incorrect Message | Device problem code
6979537 | 91113319 | N | A SPACELABS FIELD SERVICE ENGINEER PERFORMED AN INVESTIGATION OF THE DEVICE INCLUDING THE LOGS AND BASED UPON THE LOG DATA CONFIRMED THE REPORTED COMPLAINT RESETTING THE DEVICE AND PERFORMING THE CHECKOUT PROCEDURE AGAIN RESOLVED THE ISSUE AS THE DISPLAY OF THE TRIANGLE ERROR MESSAGE ALERTED THE USER TO THE DEVICES CONDITION PRIOR TO USE IT PREVENTS USE OF THE DEVICE UNTIL RESOLVED THE INVESTIGATION FINDINGS SHOWED NO RISK OF SERIOUS HARM AND NO SERIOUS RISK SHOULD THE EVENT RECUR HOWEVER SPACELABS IS FILING THIS REPORT AS REQUESTED BY THE FDA LETTER DATED OCTOBER | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6987732 | 91120042 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B B LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987732 | 91120042 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B B LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987724 | 91120114 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987724 | 91120114 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987658 | 91120132 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987658 | 91120132 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWO MONTHS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6980572 | 91120773 | N | NO PATIENT INVOLVEMENT DEVICE DISPLAYS ERROR MESSAGE GAS DELIVERY SYSTEM THE ISSUE WAS DISCOVERED DURING TESTING OF THE DEVICE | Device Displays Incorrect Message | Device problem code
6980572 | 91120773 | N | NO PATIENT INVOLVEMENT DEVICE DISPLAYS ERROR MESSAGE GAS DELIVERY SYSTEM THE ISSUE WAS DISCOVERED DURING TESTING OF THE DEVICE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
6989919 | 91123836 | D | ACCORDING TO THE REPORTER DURING LAPAROSCOPIC NISSEN PROCEDURE THE DEVICE WAS DIFFICULT TO TOGGLE AND IT WOULD NOT HOLD NEEDLE NEW DEVICE WAS USED TO COMPLETE THE CASE THERE WAS NO INJURY CAUSED TO THE PATIENT AND NO MEDICAL INTERVENTION WAS REQUIRED | Activation, Positioning or Separation Problem | Device problem code
6989919 | 91123836 | D | ACCORDING TO THE REPORTER DURING LAPAROSCOPIC NISSEN PROCEDURE THE DEVICE WAS DIFFICULT TO TOGGLE AND IT WOULD NOT HOLD NEEDLE NEW DEVICE WAS USED TO COMPLETE THE CASE THERE WAS NO INJURY CAUSED TO THE PATIENT AND NO MEDICAL INTERVENTION WAS REQUIRED | No Consequences Or Impact To Patient | Patient problem code
6794714 | 91125640 | N | THE STAFF MEMBER WAS CLEANING THE STRETCHER AND WALKING AROUND IT THE IV POLE WAS STILL EXTENDED BUT LAYING HORIZONTALLY TO BE STOWED THE STAFF MEMBERS PANT LEG GOT CAUGHT ON THE IV POLE HOOK AND THE STAFF MEMBER TRIPPED | Mechanical Jam | Device problem code
6794714 | 91125640 | N | THE STAFF MEMBER WAS CLEANING THE STRETCHER AND WALKING AROUND IT THE IV POLE WAS STILL EXTENDED BUT LAYING HORIZONTALLY TO BE STOWED THE STAFF MEMBERS PANT LEG GOT CAUGHT ON THE IV POLE HOOK AND THE STAFF MEMBER TRIPPED | Mechanical Problem Identified | Medical Evaluation result code
6975905 | 91129893 | N | B BATCH PJP DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PCEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Device Dislodged or Dislocated | Device problem code
6975905 | 91129893 | N | B BATCH PJP DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PCEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Material Twisted / Bent | Device problem code
6975905 | 91129893 | N | B BATCH PJP DEVICE EVALUATION THE ANALYSIS RESULTS THAT ONE PCEEA DEVICE WAS RETURNED WITH NO VISUAL NONCONFORMANCES AND WITHOUT A CARTRIDGE RELOAD PRESENT IN ADDITION A CARTRIDGE PAN WAS FOUND LODGED INSIDE THE CHANNEL THE PAN WAS REMOVED FROM THE CHANNEL AND THE DEVICE WAS TESTED FOR FUNCTIONALITY WITH A TEST RELOAD THE DEVICE FIRED CUT AND FORMED ALL THE STAPLES AS INTENDED THE STAPLE LINE WAS COMPLETE AND THE STAPLES WERE NOTED TO HAVE THE PROPER BFORMED SHAPE NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE CARTRIDGE PAN TO DISLODGE ONE POSSIBLE CAUSE COULD BE IMPROPER UNLOADING TECHNIQUE PLEASE MAKE SURE THE DEVICE IS UNLOADED BY PUSHING THE CARTRIDGE UPWARD TOWARD THE ANVIL TO UNSNAP THE RELOAD FROM THE CARTRIDGE JAW ANY TWISTING OR OFF CENTER PUSHING CAN LEAD TO THIS ISSUE IT SHOULD BE NOTED THAT A INSPECTIONS TAKES PLACE DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS THE BATCH HISTORY RECORD WAS REVIEWED AND NO DEFECTS NCRS OR PROTOCOLS RELATED TO THE COMPLAINT WERE FOUND DURING THE MANUFACTURING PROCESS | Mechanical Jam | Device problem code
6541756 | 91132093 | N | IT WAS ALSO IDENTIFIED DURING EVALUATION THAT THE SIDERAIL WOULD NOT LATCH | Mechanics Altered | Device problem code
6541756 | 91132093 | N | IT WAS ALSO IDENTIFIED DURING EVALUATION THAT THE SIDERAIL WOULD NOT LATCH | Mechanical Problem Identified | Medical Evaluation result code
6987672 | 91158581 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY SEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987672 | 91158581 | N | ACCORDING TO THE PRACTITIONER THE IMPLANT DIDNT TAKE AND FAILED TO INTEGRATE THE IMPLANT HAS BEEN PLACED IN POSITION ON B TWENTY SEVEN DAYS LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANT FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6987716 | 91189398 | N | ACCORDING TO THE PRACTITIONER FOUR IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE THE FOUR IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANTS FAILED TO INTEGRATE | Component or Accessory Incompatibility | Device Problem Code
6987716 | 91189398 | N | ACCORDING TO THE PRACTITIONER FOUR IMPLANTS DIDNT TAKE AND FAILED TO INTEGRATE THE FOUR IMPLANTS HAVE BEEN PLACED IN AND POSITION ON B ONE MONTH LATER AFTER THE IMPLANTATION THE PRACTITIONER HAS FOUND THAT THE IMPLANTS FAILED TO INTEGRATE | Interoperability Problem Identified | Medical Evaluation result code
6991475 | 91320357 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON OCT WHICH APPEARED AS VISUALLY POSITIVE | Connection Problem | Device problem code
6991475 | 91320357 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON OCT WHICH APPEARED AS VISUALLY POSITIVE | Interoperability Problem Identified | Medical Evaluation result code
6984781 | 91328898 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGE IN QUESTION ON OCT WHICH APPEARED AS VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON OCT TO ASSESS THE INSTRUMENT IN QUESTION WHICH APPEARED AS EXPECTED | Connection Problem | Device problem code
6984781 | 91328898 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGE IN QUESTION ON OCT WHICH APPEARED AS VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON OCT TO ASSESS THE INSTRUMENT IN QUESTION WHICH APPEARED AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
6984790 | 91372576 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGE IN QUESTION ON B WHICH APPEARED AS VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON B TO ASSESS THE INSTRUMENT IN QUESTION WHICH APPEARED AS EXPECTED | Connection Problem | Device problem code
6984790 | 91372576 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGE IN QUESTION ON B WHICH APPEARED AS VISUALLY POSITIVE AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON B TO ASSESS THE INSTRUMENT IN QUESTION WHICH APPEARED AS EXPECTED | Interoperability Problem Identified | Medical Evaluation result code
7026109 | 91896263 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Injury | Patient problem code
7026109 | 91896263 | D | IT IS ALLEGED THAT THE PLAINTIFF RECEIVED A GUNTHER FILTER ON B IT IS ALLEGED THAT THE PLAINTIFF WAS INJURED WITHOUT FURTHER EXPLANATION HOSPITAL AND MEDICAL RECORDS HAVE BEEN REQUESTED BUT NOT YET PROVIDED | Insufficient Information | Device Problem Code
6857606 | 92927317 | N | THERE WAS NO FINDING OF A DISPLAY ISSUE DURING THE PRODUCT INVESTIGATION | Display or Visual Feedback Problem | Device problem code
6857606 | 92927317 | N | THERE WAS NO FINDING OF A DISPLAY ISSUE DURING THE PRODUCT INVESTIGATION | No Device Problem Found | Medical Evaluation result code
6763275 | 92945219 | N | BASED ON THE PROVIDED INFORMATION AND TESTS PERFORMED THERE IS NO INDICATION OF A MALFUNCTION OF THE MR SYSTEM THE INJURY IS CONSISTENT WITH PATIENT HEATING INCIDENTS CAUSED BY A TOO CLOSE PROXIMITY TO THE COIL CABLE IT WAS CONFIRMED THAT THE PATIENTS ARM WAS MOST LIKELY POSITIONED ON TOP OF THE COIL CABLE INSPECTION OF THE COIL REVEALED THAT ONE OF THE BALUNS IN THE BASEPLATE WAS OUT OF SPECIFICATION THIS IS A SINGLE FAULT CONDITION AND THE CABLE BALUN IS STILL ABLE TO KEEP THE CABLE CURRENTS TO A LOW ENOUGH VALUE IN THIS CASE INSUFFICIENT DISTANCE WAS KEPT BETWEEN THE CABLE OF THE COIL AND THE ARM OF THE PATIENT FURTHERMORE THE USED SCAN PROTOCOLS WITH HIGH SAR MAY HAVE CONTRIBUTED TO THE EVENT THE INSTRUCTIONS FOR USE CONTAIN WARNINGS ABOUT COIL POSITIONING AND HIGH SAR SCANNING | Insufficient Information | Device Problem Code
6763275 | 92945219 | N | BASED ON THE PROVIDED INFORMATION AND TESTS PERFORMED THERE IS NO INDICATION OF A MALFUNCTION OF THE MR SYSTEM THE INJURY IS CONSISTENT WITH PATIENT HEATING INCIDENTS CAUSED BY A TOO CLOSE PROXIMITY TO THE COIL CABLE IT WAS CONFIRMED THAT THE PATIENTS ARM WAS MOST LIKELY POSITIONED ON TOP OF THE COIL CABLE INSPECTION OF THE COIL REVEALED THAT ONE OF THE BALUNS IN THE BASEPLATE WAS OUT OF SPECIFICATION THIS IS A SINGLE FAULT CONDITION AND THE CABLE BALUN IS STILL ABLE TO KEEP THE CABLE CURRENTS TO A LOW ENOUGH VALUE IN THIS CASE INSUFFICIENT DISTANCE WAS KEPT BETWEEN THE CABLE OF THE COIL AND THE ARM OF THE PATIENT FURTHERMORE THE USED SCAN PROTOCOLS WITH HIGH SAR MAY HAVE CONTRIBUTED TO THE EVENT THE INSTRUCTIONS FOR USE CONTAIN WARNINGS ABOUT COIL POSITIONING AND HIGH SAR SCANNING | No Device Problem Found | Medical Evaluation result code
7051376 | 93206997 | D | IT WAS REPORTED TO THE SALES REP THAT THE FACILITY HAD A CM HEMOSTAR DIALYSIS CATHETER THAT WAS BLEEDING FROM THE EXIT SITE IT BEGAN BLEEDING ONLY ON DIALYSIS THE RATE WAS DECREASED TO ML | Hemorrhage | Patient problem code
7051376 | 93206997 | D | IT WAS REPORTED TO THE SALES REP THAT THE FACILITY HAD A CM HEMOSTAR DIALYSIS CATHETER THAT WAS BLEEDING FROM THE EXIT SITE IT BEGAN BLEEDING ONLY ON DIALYSIS THE RATE WAS DECREASED TO ML | Insufficient Information | Device Problem Code
6917843 | 93384849 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6917843 | 93384849 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6961993 | 93756945 | N | THE INVESTIGATION HAS BEEN COMPLETED BASED ON THE ANALYSIS THE ALLEGED BATTERY MALFUNCTION WAS VERIFIED THE ALLEGED FILL ESTIMATE MALFUNCTION COULD NOT BE VERIFIED | Battery Problem | Device problem code
6961993 | 93756945 | N | THE INVESTIGATION HAS BEEN COMPLETED BASED ON THE ANALYSIS THE ALLEGED BATTERY MALFUNCTION WAS VERIFIED THE ALLEGED FILL ESTIMATE MALFUNCTION COULD NOT BE VERIFIED | Energy Storage System Problem | Medical Evaluation result code
6978741 | 93828349 | N | ACCORDING TO THE INITIAL REPORTER THE REASON FOR LENS SUBLUXATION WAS PATIENT RELATED THE EXACT PATIENT ISSUE WAS NOT SPECIFIED BASED ON THIS INFORMATION THE REPORTED EVENT IS CONSIDERED UNRELATED TO THE DEVICE | Migration or Expulsion of Device | Device problem code
6978741 | 93828349 | N | ACCORDING TO THE INITIAL REPORTER THE REASON FOR LENS SUBLUXATION WAS PATIENT RELATED THE EXACT PATIENT ISSUE WAS NOT SPECIFIED BASED ON THIS INFORMATION THE REPORTED EVENT IS CONSIDERED UNRELATED TO THE DEVICE | Device Migration | Medical Evaluation result code
7092265 | 94212000 | D | IT WAS REPORTED THAT THE RIGHT VENTRICULAR RV LEAD EXHIBITED ELEVATED SENSING INTEGRITY COUNTER SIC THE LEAD WAS TESTED BUT NO PERMANENT PROGRAMMING WAS PERFORMED THE LEAD REMAINS IN USE NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
7092265 | 94212000 | D | IT WAS REPORTED THAT THE RIGHT VENTRICULAR RV LEAD EXHIBITED ELEVATED SENSING INTEGRITY COUNTER SIC THE LEAD WAS TESTED BUT NO PERMANENT PROGRAMMING WAS PERFORMED THE LEAD REMAINS IN USE NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT | Failure to Sense | Device Problem Code
6209426 | 95073081 | N | FINDINGS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Component or Accessory Incompatibility | Device problem code
6209426 | 95073081 | N | FINDINGS INSPECTED OPENEDUSED ENLITE SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Interoperability Problem Identified | Medical Evaluation result code
6215984 | 95073193 | N | INSPECTED ONE OPENED AND USED SENSOR WE PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Adverse Event Without Identified Device or Use Problem | Device Problem Code
6215984 | 95073193 | N | INSPECTED ONE OPENED AND USED SENSOR WE PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6215984 | 95073193 | N | INSPECTED ONE OPENED AND USED SENSOR WE PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6212978 | 95073234 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
6212978 | 95073234 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR FAILED PER SPECIFICATIONS | Interoperability Problem Identified | Medical Evaluation result code
6802628 | 95757577 | D | SUBSEQUENT TO THE INITIAL FILED MEDWATCH REPORT THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED THE PATIENT HAD EXPIRED THE SAME DAY B NO ADDITIONAL INFORMATION WAS PROVIDED | Death | Patient problem code
6802628 | 95757577 | D | SUBSEQUENT TO THE INITIAL FILED MEDWATCH REPORT THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED THE PATIENT HAD EXPIRED THE SAME DAY B NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
6583486 | 96989041 | D | IT WAS REPORTED THAT A NEW CATHETER WAS TORN DURING CT SCAN INJECTION MORE INFO TO BE OBTAINED NO PATIENT INJURY REPORTED | No Consequences Or Impact To Patient | Patient problem code
6583486 | 96989041 | D | IT WAS REPORTED THAT A NEW CATHETER WAS TORN DURING CT SCAN INJECTION MORE INFO TO BE OBTAINED NO PATIENT INJURY REPORTED | Material Split, Cut or Torn | Device Problem Code
6876774 | 99139904 | D | ADDITIONAL INFORMATION WAS RECEIVED IT WAS REPORTED THAT THE PNEUMOCEPHALUS WAS THOUGHT TO BE RELATED TO THE SURGERY AND ANESTHESIA BUT THE EXACT CAUSE WAS NOT INDICATED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
6876774 | 99139904 | D | ADDITIONAL INFORMATION WAS RECEIVED IT WAS REPORTED THAT THE PNEUMOCEPHALUS WAS THOUGHT TO BE RELATED TO THE SURGERY AND ANESTHESIA BUT THE EXACT CAUSE WAS NOT INDICATED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
6876309 | 99139977 | D | ADDITIONAL INFORMATION WAS RECEIVED IT WAS REPORTED THAT THE PNEUMOCEPHALUS WAS THOUGHT TO BE RELATED TO THE SURGERY AND ANESTHESIA BUT THE EXACT CAUSE WAS NOT INDICATED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
6876309 | 99139977 | D | ADDITIONAL INFORMATION WAS RECEIVED IT WAS REPORTED THAT THE PNEUMOCEPHALUS WAS THOUGHT TO BE RELATED TO THE SURGERY AND ANESTHESIA BUT THE EXACT CAUSE WAS NOT INDICATED NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
6883603 | 101032753 | D | ADDITIONAL INFORMATION RECEIVED FROM THE AUTHOR REPORTED THAT THE WEIGHT GAIN WAS NOT CONSIDERED RELATED TO THE DEVICE THE PATIENT ALREADY SUFFERED WITH OBESITY PRIOR TO THE SURGERY AND CONTINUED TO GAIN WEIGHT FOLLOWING THE SURGERY SHE WAS TREATED WITH GASTROPLASTY | Insufficient Information | Device Problem Code
6883603 | 101032753 | D | ADDITIONAL INFORMATION RECEIVED FROM THE AUTHOR REPORTED THAT THE WEIGHT GAIN WAS NOT CONSIDERED RELATED TO THE DEVICE THE PATIENT ALREADY SUFFERED WITH OBESITY PRIOR TO THE SURGERY AND CONTINUED TO GAIN WEIGHT FOLLOWING THE SURGERY SHE WAS TREATED WITH GASTROPLASTY | Weight Fluctuations | Patient problem code
6613209 | 101041194 | D | ADDITIONAL INFORMATION WAS RECEIVED THERE WAS NO FURTHER INFORMATION ABOUT THE IMPEDANCES NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
6613209 | 101041194 | D | ADDITIONAL INFORMATION WAS RECEIVED THERE WAS NO FURTHER INFORMATION ABOUT THE IMPEDANCES NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
6841407 | 101085616 | D | ADDITIONAL INFORMATION RECEIVED FROM THE PATIENTS FAMILY MEMBER REPORTED THAT THE PATIENTS DEVICE WAS EXPLANTED ON B A CULTURE THAT HAD BEEN PERFORMED HAD DETERMINED IT TO BE A STAPH INFECTION | Infection | Patient problem code
6841407 | 101085616 | D | ADDITIONAL INFORMATION RECEIVED FROM THE PATIENTS FAMILY MEMBER REPORTED THAT THE PATIENTS DEVICE WAS EXPLANTED ON B A CULTURE THAT HAD BEEN PERFORMED HAD DETERMINED IT TO BE A STAPH INFECTION | Insufficient Information | Device Problem Code
4471123 | 101522713 | D | ADDITIONAL INFORMATION WAS RECEIVED THE OUTCOME WAS UPDATED TO ONGOING THE PATIENT HAD A CLINIC VISIT B AND REPORTED THE SYMPTOMS NEVER RESOLVED AS WAS PREVIOUSLY REPORTED IN B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
4471123 | 101522713 | D | ADDITIONAL INFORMATION WAS RECEIVED THE OUTCOME WAS UPDATED TO ONGOING THE PATIENT HAD A CLINIC VISIT B AND REPORTED THE SYMPTOMS NEVER RESOLVED AS WAS PREVIOUSLY REPORTED IN B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
6263646 | 102901686 | D | ADDITIONAL INFORMATION FROM THE HEALTHCARE PROFESSIONAL OF THE CLINICAL STUDY REPORTED THAT NO ACTIONSINTERVENTIONS WERE TAKEN TO RESOLVE THE SHOCKING THE CAUSE OF THE SHOCKING WAS NOT DETERMINED | Shock | Patient problem code
6263646 | 102901686 | D | ADDITIONAL INFORMATION FROM THE HEALTHCARE PROFESSIONAL OF THE CLINICAL STUDY REPORTED THAT NO ACTIONSINTERVENTIONS WERE TAKEN TO RESOLVE THE SHOCKING THE CAUSE OF THE SHOCKING WAS NOT DETERMINED | Insufficient Information | Device Problem Code
5064681 | 103098359 | D | ADDITIONAL INFORMATION RECEIVED IT WAS REPORTED THE PATIENT EXPERIENCED MUSCLE TIGHTNESS IN THEIR FEET NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Muscular Rigidity | Patient problem code
5064681 | 103098359 | D | ADDITIONAL INFORMATION RECEIVED IT WAS REPORTED THE PATIENT EXPERIENCED MUSCLE TIGHTNESS IN THEIR FEET NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
4471115 | 103548705 | D | ADDITIONAL INFORMATION WAS RECEIVED THE OUTCOME WAS UPDATED TO ONGOING THE PATIENT HAD A CLINIC VISIT B AND REPORTED THE SYMPTOMS NEVER RESOLVED AS WAS PREVIOUSLY REPORTED IN B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
4471115 | 103548705 | D | ADDITIONAL INFORMATION WAS RECEIVED THE OUTCOME WAS UPDATED TO ONGOING THE PATIENT HAD A CLINIC VISIT B AND REPORTED THE SYMPTOMS NEVER RESOLVED AS WAS PREVIOUSLY REPORTED IN B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Insufficient Information | Device Problem Code
6987604 | 105567918 | N | FOLLOWUP DEVICE EVALUATION IN B THERE WERE INSTANCES OF VIBRATION FAILURE FOUND DURING INVESTIGATION THIS REPORT IS PROCESSED UNDER EXEMPTION NUMBER B THIS MDR REPORT IS PART OF THE PILOT PROGRAM | Vibration | Device problem code
6987604 | 105567918 | N | FOLLOWUP DEVICE EVALUATION IN B THERE WERE INSTANCES OF VIBRATION FAILURE FOUND DURING INVESTIGATION THIS REPORT IS PROCESSED UNDER EXEMPTION NUMBER B THIS MDR REPORT IS PART OF THE PILOT PROGRAM | No Device Problem Found | Medical Evaluation result code
6873336 | 108773924 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE X MM GRAFTMASTER STENT DELIVERY SYSTEM HAD BEEN IN AND OUT OF THE BODY SEVERAL TIMES NO ADDITIONAL INFORMATION WAS PROVIDED | No Information | Patient problem code
6873336 | 108773924 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE X MM GRAFTMASTER STENT DELIVERY SYSTEM HAD BEEN IN AND OUT OF THE BODY SEVERAL TIMES NO ADDITIONAL INFORMATION WAS PROVIDED | Component or Accessory Incompatibility | Device Problem Code
7561974 | 109826248 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE LEVELS BELOW MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT OVER MGDL AFTER BEING DISCHARGED AND MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN JUICE AND FOOD TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS UNRESPONSIVENESS THE CUSTOMER WAS WEARING THE INSULIN PUMP DURING THE INCIDENT THE CUSTOMER STATED THE PUMP DID NOT SUSPEND WHEN IT WAS SUPPOSED TO TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
7561974 | 109826248 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE LEVELS BELOW MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT OVER MGDL AFTER BEING DISCHARGED AND MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS GIVEN JUICE AND FOOD TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS UNRESPONSIVENESS THE CUSTOMER WAS WEARING THE INSULIN PUMP DURING THE INCIDENT THE CUSTOMER STATED THE PUMP DID NOT SUSPEND WHEN IT WAS SUPPOSED TO TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Pumping Problem | Device problem code
7563661 | 109882640 | N | THERE IS NO INDICATION THAT THE WAS EXPLANTED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON AND JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7563661 | 109882640 | N | THERE IS NO INDICATION THAT THE WAS EXPLANTED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON AND JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7562294 | 109888545 | N | B MANUFACTURING SITE EVALUATION ACCORDING TO THE AVAILABLE INFORMATION THERE WERE NO NEGATIVE CONSEQUENCES FOR THE PATIENT THE PIN CUTTER SHOWS NO OUTWARDLY DEFECTS THE CLAMP ARRIVED IN THREE PARTS PIN UPPER DISC AND LOWER DISC USED TEST AND ANALYSIS MICROSCOPE DIGITAL CAMERA A VISUAL INSPECTION OF THE LOWER DISC ESPECIALLY THE PIN AND ITS CUTTING EDGE THE LOWER DISC EXHIBITS A PIN WHICH WAS CUT SKEWED AND UNDERNEATH THE LAST RIMS SO THE DISC MUST CAME OFF FURTHERMORE PERFORMED SOME CUTTING TESTS WITH THE PIN CUTTER THE PIN CUTTER WORKED PROPERLY WITHOUT ANY PROBLEMS THE DETACHED IMPLANT AND THE PIN CUTTER IN THEIR INCOMING CONDITION THE MANUFACTURING DOCUMENTS HAVE CHECKED AND FOUND TO BE ACCORDING TO SPECIFICATION VALID DURING THE TIME OF PRODUCTION THE ROOT CAUSE FOR THE PROBLEM IS MOST PROBABLY USAGE RELATED THE PIN OF THE CLAMP WAS CUT UNDERNEATH THE LAST FIXATION RIM A CAPA IS NOT NECESSARY | Detachment of Device or device Component | Device problem code
7562294 | 109888545 | N | B MANUFACTURING SITE EVALUATION ACCORDING TO THE AVAILABLE INFORMATION THERE WERE NO NEGATIVE CONSEQUENCES FOR THE PATIENT THE PIN CUTTER SHOWS NO OUTWARDLY DEFECTS THE CLAMP ARRIVED IN THREE PARTS PIN UPPER DISC AND LOWER DISC USED TEST AND ANALYSIS MICROSCOPE DIGITAL CAMERA A VISUAL INSPECTION OF THE LOWER DISC ESPECIALLY THE PIN AND ITS CUTTING EDGE THE LOWER DISC EXHIBITS A PIN WHICH WAS CUT SKEWED AND UNDERNEATH THE LAST RIMS SO THE DISC MUST CAME OFF FURTHERMORE PERFORMED SOME CUTTING TESTS WITH THE PIN CUTTER THE PIN CUTTER WORKED PROPERLY WITHOUT ANY PROBLEMS THE DETACHED IMPLANT AND THE PIN CUTTER IN THEIR INCOMING CONDITION THE MANUFACTURING DOCUMENTS HAVE CHECKED AND FOUND TO BE ACCORDING TO SPECIFICATION VALID DURING THE TIME OF PRODUCTION THE ROOT CAUSE FOR THE PROBLEM IS MOST PROBABLY USAGE RELATED THE PIN OF THE CLAMP WAS CUT UNDERNEATH THE LAST FIXATION RIM A CAPA IS NOT NECESSARY | Interoperability Problem Identified | Medical Evaluation result code
7567521 | 110027410 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7567521 | 110027410 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7568441 | 110037185 | D | THE PATIENT WAS PREOPERATIVELY DIAGNOSED WITH PSEUDOARTHROSIS LUMBAR SPINE FAILED INSTRUMENTATION AND UNDERWENT THE FOLLOWING PROCEDURE REMOVAL OF INSTRUMENTATION TS REINSERTION OF PEDICLE SCREWS TL POSTERIOR LATERAL FUSION TS SEGMENTAL INSTRUMENTATION TS AND FLUOROSCOPY IN WHICH RHBMPACS WAS USED AS PER OPNOTES PEDICLE SCREWS HAD BEEN PLACED IN SEGMENTAL FASHION FROM TS PRIOR TO THE FUSION BILATERAL RODS WERE PLACED IN THE SCREWS AND SET SCREWS WERE USED TO SECURE THE RODS TO THE SCREWS AND FINAL TIGHTENING OCCURRED AFTER CONFIRMATORY IMAGING WAS OBTAINED INCLUDING RECONNECTION TO THE ILIAC BOLT A STACKED FOUR ROD TECHNIQUE WAS USED THE FUSION LEVELS INCLUDE TS TO PERFORM THE FUSION A HIGH SPEED FLUTED BURR WAS USED TO DECORTICATE THE LEVELS INDICATED FACETECTOMIES HAD BEEN PERFORMED AS WELL ONCE THE DECORTICATION PROCESS WAS PERFORMED AND BLEEDING BONE WAS IDENTIFIED AT THE FUSION LEVELS RODS WERE PLACED WITHIN THE PEDICLE SCREWS AND SET SCREWS APPLIED WITH FINAL TIGHTENING OCCURRING FOLLOWING THIS THE AUTOGENOUS LOCAL BONE GRAFT LISTED ABOVE WERE USED TO FACILITATE FUSION THE WOUND WAS COPIOUSLY IRRIGATED WITH NORMAL SALINE AND HEMOSTASIS WAS OBTAINED VANCOMYCIN GM FREE WAS PLACED IN THE WOUND PRIOR TO CLOSURE A DEEP DRAIN WAS THEN PLACED THE WOUND WAS THEN CLOSED IN LAYERS A STERILE DRESSING WAS APPLIED THE PATIENT WAS THEN PLACED IN THE PACU EXTUBATED THE ESTIMATED BLOOD LOSS AND COMPLICATIONS LOSS WERE AS LISTED ABOVE ON B THE PATIENT WAS DISCHARGED FROM THE FACILITY | Insufficient Information | Device problem code
7568441 | 110037185 | D | THE PATIENT WAS PREOPERATIVELY DIAGNOSED WITH PSEUDOARTHROSIS LUMBAR SPINE FAILED INSTRUMENTATION AND UNDERWENT THE FOLLOWING PROCEDURE REMOVAL OF INSTRUMENTATION TS REINSERTION OF PEDICLE SCREWS TL POSTERIOR LATERAL FUSION TS SEGMENTAL INSTRUMENTATION TS AND FLUOROSCOPY IN WHICH RHBMPACS WAS USED AS PER OPNOTES PEDICLE SCREWS HAD BEEN PLACED IN SEGMENTAL FASHION FROM TS PRIOR TO THE FUSION BILATERAL RODS WERE PLACED IN THE SCREWS AND SET SCREWS WERE USED TO SECURE THE RODS TO THE SCREWS AND FINAL TIGHTENING OCCURRED AFTER CONFIRMATORY IMAGING WAS OBTAINED INCLUDING RECONNECTION TO THE ILIAC BOLT A STACKED FOUR ROD TECHNIQUE WAS USED THE FUSION LEVELS INCLUDE TS TO PERFORM THE FUSION A HIGH SPEED FLUTED BURR WAS USED TO DECORTICATE THE LEVELS INDICATED FACETECTOMIES HAD BEEN PERFORMED AS WELL ONCE THE DECORTICATION PROCESS WAS PERFORMED AND BLEEDING BONE WAS IDENTIFIED AT THE FUSION LEVELS RODS WERE PLACED WITHIN THE PEDICLE SCREWS AND SET SCREWS APPLIED WITH FINAL TIGHTENING OCCURRING FOLLOWING THIS THE AUTOGENOUS LOCAL BONE GRAFT LISTED ABOVE WERE USED TO FACILITATE FUSION THE WOUND WAS COPIOUSLY IRRIGATED WITH NORMAL SALINE AND HEMOSTASIS WAS OBTAINED VANCOMYCIN GM FREE WAS PLACED IN THE WOUND PRIOR TO CLOSURE A DEEP DRAIN WAS THEN PLACED THE WOUND WAS THEN CLOSED IN LAYERS A STERILE DRESSING WAS APPLIED THE PATIENT WAS THEN PLACED IN THE PACU EXTUBATED THE ESTIMATED BLOOD LOSS AND COMPLICATIONS LOSS WERE AS LISTED ABOVE ON B THE PATIENT WAS DISCHARGED FROM THE FACILITY | Hemorrhage | Patient problem code
7568441 | 110037185 | D | THE PATIENT WAS PREOPERATIVELY DIAGNOSED WITH PSEUDOARTHROSIS LUMBAR SPINE FAILED INSTRUMENTATION AND UNDERWENT THE FOLLOWING PROCEDURE REMOVAL OF INSTRUMENTATION TS REINSERTION OF PEDICLE SCREWS TL POSTERIOR LATERAL FUSION TS SEGMENTAL INSTRUMENTATION TS AND FLUOROSCOPY IN WHICH RHBMPACS WAS USED AS PER OPNOTES PEDICLE SCREWS HAD BEEN PLACED IN SEGMENTAL FASHION FROM TS PRIOR TO THE FUSION BILATERAL RODS WERE PLACED IN THE SCREWS AND SET SCREWS WERE USED TO SECURE THE RODS TO THE SCREWS AND FINAL TIGHTENING OCCURRED AFTER CONFIRMATORY IMAGING WAS OBTAINED INCLUDING RECONNECTION TO THE ILIAC BOLT A STACKED FOUR ROD TECHNIQUE WAS USED THE FUSION LEVELS INCLUDE TS TO PERFORM THE FUSION A HIGH SPEED FLUTED BURR WAS USED TO DECORTICATE THE LEVELS INDICATED FACETECTOMIES HAD BEEN PERFORMED AS WELL ONCE THE DECORTICATION PROCESS WAS PERFORMED AND BLEEDING BONE WAS IDENTIFIED AT THE FUSION LEVELS RODS WERE PLACED WITHIN THE PEDICLE SCREWS AND SET SCREWS APPLIED WITH FINAL TIGHTENING OCCURRING FOLLOWING THIS THE AUTOGENOUS LOCAL BONE GRAFT LISTED ABOVE WERE USED TO FACILITATE FUSION THE WOUND WAS COPIOUSLY IRRIGATED WITH NORMAL SALINE AND HEMOSTASIS WAS OBTAINED VANCOMYCIN GM FREE WAS PLACED IN THE WOUND PRIOR TO CLOSURE A DEEP DRAIN WAS THEN PLACED THE WOUND WAS THEN CLOSED IN LAYERS A STERILE DRESSING WAS APPLIED THE PATIENT WAS THEN PLACED IN THE PACU EXTUBATED THE ESTIMATED BLOOD LOSS AND COMPLICATIONS LOSS WERE AS LISTED ABOVE ON B THE PATIENT WAS DISCHARGED FROM THE FACILITY | Blood Loss | Patient problem code
7568441 | 110037185 | D | THE PATIENT WAS PREOPERATIVELY DIAGNOSED WITH PSEUDOARTHROSIS LUMBAR SPINE FAILED INSTRUMENTATION AND UNDERWENT THE FOLLOWING PROCEDURE REMOVAL OF INSTRUMENTATION TS REINSERTION OF PEDICLE SCREWS TL POSTERIOR LATERAL FUSION TS SEGMENTAL INSTRUMENTATION TS AND FLUOROSCOPY IN WHICH RHBMPACS WAS USED AS PER OPNOTES PEDICLE SCREWS HAD BEEN PLACED IN SEGMENTAL FASHION FROM TS PRIOR TO THE FUSION BILATERAL RODS WERE PLACED IN THE SCREWS AND SET SCREWS WERE USED TO SECURE THE RODS TO THE SCREWS AND FINAL TIGHTENING OCCURRED AFTER CONFIRMATORY IMAGING WAS OBTAINED INCLUDING RECONNECTION TO THE ILIAC BOLT A STACKED FOUR ROD TECHNIQUE WAS USED THE FUSION LEVELS INCLUDE TS TO PERFORM THE FUSION A HIGH SPEED FLUTED BURR WAS USED TO DECORTICATE THE LEVELS INDICATED FACETECTOMIES HAD BEEN PERFORMED AS WELL ONCE THE DECORTICATION PROCESS WAS PERFORMED AND BLEEDING BONE WAS IDENTIFIED AT THE FUSION LEVELS RODS WERE PLACED WITHIN THE PEDICLE SCREWS AND SET SCREWS APPLIED WITH FINAL TIGHTENING OCCURRING FOLLOWING THIS THE AUTOGENOUS LOCAL BONE GRAFT LISTED ABOVE WERE USED TO FACILITATE FUSION THE WOUND WAS COPIOUSLY IRRIGATED WITH NORMAL SALINE AND HEMOSTASIS WAS OBTAINED VANCOMYCIN GM FREE WAS PLACED IN THE WOUND PRIOR TO CLOSURE A DEEP DRAIN WAS THEN PLACED THE WOUND WAS THEN CLOSED IN LAYERS A STERILE DRESSING WAS APPLIED THE PATIENT WAS THEN PLACED IN THE PACU EXTUBATED THE ESTIMATED BLOOD LOSS AND COMPLICATIONS LOSS WERE AS LISTED ABOVE ON B THE PATIENT WAS DISCHARGED FROM THE FACILITY | Hemostasis | Patient problem code
7570098 | 110111201 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Calibration Problem | Device problem code
7570098 | 110111201 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7570221 | 110111729 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Calibration Problem | Device problem code
7570221 | 110111729 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7570170 | 110111917 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Calibration Problem | Device problem code
7570170 | 110111917 | N | THE DEVICE WAS TESTED AND SERVICED BY A THIRD PARTY SERVICER THEY REPORTED TO MMDG THAT THE PUMP WAS INFUSING AT A VOLUMETRIC RATE THAT MMDG CONSIDERS REPORTABLE THE THIRD PARTY SERVICER ALSO ADVISED THAT THEY COMPLETED CALIBRATION TO CORRECT THE PUMP | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7575667 | 110421381 | D | IT WAS REPORTED THAT THE CATHETERS HAD NO HOLE ON THE END WERE PAINFUL WHEN INSERTED AND MADE THE PATIENT BLEED THE CATHETERS WERE ALLEGEDLY THINNER THAN THE USUAL ONES | Inadequacy of Device Shape and/or Size | Device problem code
7575667 | 110421381 | D | IT WAS REPORTED THAT THE CATHETERS HAD NO HOLE ON THE END WERE PAINFUL WHEN INSERTED AND MADE THE PATIENT BLEED THE CATHETERS WERE ALLEGEDLY THINNER THAN THE USUAL ONES | Pain | Patient problem code
7575667 | 110421381 | D | IT WAS REPORTED THAT THE CATHETERS HAD NO HOLE ON THE END WERE PAINFUL WHEN INSERTED AND MADE THE PATIENT BLEED THE CATHETERS WERE ALLEGEDLY THINNER THAN THE USUAL ONES | Hemorrhage | Patient problem code
7581320 | 110447401 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7581320 | 110447401 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7581323 | 110447720 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7581323 | 110447720 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7581336 | 110447750 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7581336 | 110447750 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7573390 | 110458597 | N | AFFECTED DEVICES WERE RETURNED DURING VISUAL INSPECTION OF RETURNED AFFECTED PRODUCT THE INVESTIGATOR WAS ABLE TO CONFIRM CUSTOMER CLAIM OF DEFLATED MASKS INVESTIGATOR DID PULL THE BATCH RECORD AND REVIEW NO DISCREPANCIES WERE FOUND IN THE BATCH RECORD | Deflation Problem | Device problem code
7573390 | 110458597 | N | AFFECTED DEVICES WERE RETURNED DURING VISUAL INSPECTION OF RETURNED AFFECTED PRODUCT THE INVESTIGATOR WAS ABLE TO CONFIRM CUSTOMER CLAIM OF DEFLATED MASKS INVESTIGATOR DID PULL THE BATCH RECORD AND REVIEW NO DISCREPANCIES WERE FOUND IN THE BATCH RECORD | No Device Problem Found | Medical Evaluation result code
7566868 | 110471383 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON MAY WHICH APPEARED VISUALLY WEAK POSITIVE THE CAMERA REPORT APPEARED OPTIMAL COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Connection Problem | Device problem code
7566868 | 110471383 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON MAY WHICH APPEARED VISUALLY WEAK POSITIVE THE CAMERA REPORT APPEARED OPTIMAL COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Interoperability Problem Identified | Medical Evaluation result code
7583221 | 110577940 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS NOT IMPACTED REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | No Consequences Or Impact To Patient | Patient problem code
7583221 | 110577940 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS NOT IMPACTED REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | Device Alarm System | Device problem code
7586174 | 110583319 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME BEFORE THE INCIDENT AND MGDL AT ADMISSION THE CUSTOMER HAD A LOW AND BEING DISORIENTED GAVE THEMSELVES BOLUSES OF UNITS IN A ROW THE CUSTOMER USED GLUCOSE TABLETS AND WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS HEADACHE A POSSIBLE SEIZURE AND DISORIENTATION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Seizures | Patient problem code
7586174 | 110583319 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME BEFORE THE INCIDENT AND MGDL AT ADMISSION THE CUSTOMER HAD A LOW AND BEING DISORIENTED GAVE THEMSELVES BOLUSES OF UNITS IN A ROW THE CUSTOMER USED GLUCOSE TABLETS AND WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS HEADACHE A POSSIBLE SEIZURE AND DISORIENTATION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Headache | Patient problem code
7586174 | 110583319 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME BEFORE THE INCIDENT AND MGDL AT ADMISSION THE CUSTOMER HAD A LOW AND BEING DISORIENTED GAVE THEMSELVES BOLUSES OF UNITS IN A ROW THE CUSTOMER USED GLUCOSE TABLETS AND WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS HEADACHE A POSSIBLE SEIZURE AND DISORIENTATION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
7586174 | 110583319 | D | CUSTOMER CALLED AND REPORTED THAT THEY RECEIVED EMERGENCY MEDICAL ASSISTANCE AND GOT HOSPITALIZED DUE TO LOW BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME BEFORE THE INCIDENT AND MGDL AT ADMISSION THE CUSTOMER HAD A LOW AND BEING DISORIENTED GAVE THEMSELVES BOLUSES OF UNITS IN A ROW THE CUSTOMER USED GLUCOSE TABLETS AND WAS GIVEN INTRAVENOUS FLUIDS TO TREAT THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS HEADACHE A POSSIBLE SEIZURE AND DISORIENTATION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN LESS THAN HOURS TROUBLESHOOTING WAS COMPLETED THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hypoglycemia | Patient Problem Code
7589681 | 110701916 | N | DATE OF EVENT UNKNOWN NOT PROVIDED BUT THE BEST ESTIMATE DATE IS BETWEEN B AND B B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7589681 | 110701916 | N | DATE OF EVENT UNKNOWN NOT PROVIDED BUT THE BEST ESTIMATE DATE IS BETWEEN B AND B B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7586404 | 110705386 | D | ENDURANT STENT GRAFT SYSTEMS WERE IMPLANTED DURING CGEVAR TECHNIQUE REPAIR OF JUXTARENAL ABDOMINAL AORTIC ANEURYSMS THE FOLLOWING MALFUNCTIONS WERE OBSERVED TYPE IA AND TYPE II ENDOLEAKS AND ILIAC LEG KINKING | Appropriate Term/Code Not Available | Device problem code
7586404 | 110705386 | D | ENDURANT STENT GRAFT SYSTEMS WERE IMPLANTED DURING CGEVAR TECHNIQUE REPAIR OF JUXTARENAL ABDOMINAL AORTIC ANEURYSMS THE FOLLOWING MALFUNCTIONS WERE OBSERVED TYPE IA AND TYPE II ENDOLEAKS AND ILIAC LEG KINKING | Hemorrhage | Patient problem code
7590307 | 110763041 | N | THE REPORTER CONTACTED ABBOTT DIABETES CARE ALLEGING THE METER WILL NOT TURN ON THIS COMPLAINT IS BEING REPORTED BECAUSE THE REPORTED ISSUE WAS NOT RESOLVED WITH TROUBLESHOOTING THERE WAS NO REPORT OF DEATH SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT | Failure to Power Up | Device problem code
7590307 | 110763041 | N | THE REPORTER CONTACTED ABBOTT DIABETES CARE ALLEGING THE METER WILL NOT TURN ON THIS COMPLAINT IS BEING REPORTED BECAUSE THE REPORTED ISSUE WAS NOT RESOLVED WITH TROUBLESHOOTING THERE WAS NO REPORT OF DEATH SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT | Loss of Power | Medical Evaluation result code
7591349 | 110862073 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7591349 | 110862073 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7581413 | 110895659 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7581413 | 110895659 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7590142 | 110965723 | N | PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | No Display / Image | Device problem code
7590142 | 110965723 | N | PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | Image Artifact | Medical Evaluation result code
7598484 | 111009925 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Calibration Problem | Device problem code
7598484 | 111009925 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Device Displays Incorrect Message | Device Problem Code
7598484 | 111009925 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7598490 | 111012397 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Calibration Problem | Device problem code
7598490 | 111012397 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Device Displays Incorrect Message | Device Problem Code
7598490 | 111012397 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7598453 | 111012807 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Device Displays Incorrect Message | Device Problem Code
7598453 | 111012807 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Calibration Problem | Device problem code
7598453 | 111012807 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7598481 | 111013322 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Calibration Problem | Device problem code
7598481 | 111013322 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Device Displays Incorrect Message | Device Problem Code
7598481 | 111013322 | N | PER THE MANUFACTURERS SUBSIDIARY UPON CALIBRATING THE BLOOD PARAMETER MONITOR BPM A CALIBRATOR ERROR IS DISPLAYED A SECOND CALIBRATION IS COMPLETED SUCCESSFULLY BUT THE SVO WAS BEING DISPLAYED AT POST CALIBRATION THE REPORTED ISSUE HAPPENS INCONSISTENTLY BUT ON A REGULAR BASIS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7600976 | 111054766 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7600976 | 111054766 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7600968 | 111054903 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7600968 | 111054903 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7600961 | 111055155 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7600961 | 111055155 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7602729 | 111111830 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7602729 | 111111830 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | Device Alarm System | Device problem code
7605120 | 111166376 | N | CLINICAL REVIEW THERE IS NO DOCUMENTATION TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENTS NAUSEA VOMITING AND CONFUSION WITH SUBSEQUENT DIAGNOSIS OF HOSPITAL ACQUIRED PERITONITIS WITH HOSPITALIZATION THE PATIENTS RECORDS DOCUMENT THAT HE WAS DISCHARGED B AND RETURNED B WHEN HE WAS DIAGNOSED WITH PERITONITIS AND SEPSIS THE PATIENT HAD BEEN COMPLETING CAPD DURING THE INITIAL HOSPITALIZATION AS THERE WAS NO ACCESS TO A CYCLER THERE IS NO REPORTED ALLEGATION OF A MALFUNCTION AGAINST THE LIBERTY CYCLER OR A LEAK WITH THE CYCLER SET IN REGARDS TO THIS EVENT IT IS UNKNOWN WHEN THE PATIENT LAST COMPLETED PD THERAPY ON THE LIBERTY CYCLER PRIOR TO DIAGNOSIS OF THE PERITONITIS A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE PLANTS INVESTIGATION B | Leak / Splash | Device problem code
7605120 | 111166376 | N | CLINICAL REVIEW THERE IS NO DOCUMENTATION TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENTS NAUSEA VOMITING AND CONFUSION WITH SUBSEQUENT DIAGNOSIS OF HOSPITAL ACQUIRED PERITONITIS WITH HOSPITALIZATION THE PATIENTS RECORDS DOCUMENT THAT HE WAS DISCHARGED B AND RETURNED B WHEN HE WAS DIAGNOSED WITH PERITONITIS AND SEPSIS THE PATIENT HAD BEEN COMPLETING CAPD DURING THE INITIAL HOSPITALIZATION AS THERE WAS NO ACCESS TO A CYCLER THERE IS NO REPORTED ALLEGATION OF A MALFUNCTION AGAINST THE LIBERTY CYCLER OR A LEAK WITH THE CYCLER SET IN REGARDS TO THIS EVENT IT IS UNKNOWN WHEN THE PATIENT LAST COMPLETED PD THERAPY ON THE LIBERTY CYCLER PRIOR TO DIAGNOSIS OF THE PERITONITIS A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE PLANTS INVESTIGATION B | Leakage/Seal | Medical Evaluation result code
7591362 | 111169826 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7591362 | 111169826 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7591375 | 111169995 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7591375 | 111169995 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7610590 | 111437097 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO BOLUSES DELIVERED FAILING TO LOWER BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
7610590 | 111437097 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO BOLUSES DELIVERED FAILING TO LOWER BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS | Positioning Problem | Device problem code
7613110 | 111455322 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7613110 | 111455322 | N | B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7618005 | 111637550 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618005 | 111637550 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618007 | 111637891 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618007 | 111637891 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618028 | 111638139 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618028 | 111638139 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618015 | 111638681 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618015 | 111638681 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618020 | 111639344 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618020 | 111639344 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618018 | 111639546 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618018 | 111639546 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618025 | 111639671 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618025 | 111639671 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7618016 | 111641253 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Break | Device problem code
7618016 | 111641253 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | Fracture Problem | Medical Evaluation result code
7620356 | 111709380 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7620356 | 111709380 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7620473 | 111710137 | D | MEDTRONIC RECEIVED INFORMATION THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7620473 | 111710137 | D | MEDTRONIC RECEIVED INFORMATION THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7602596 | 111843530 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES WERE FOUND TO BE OPTIMAL AND ALIGNED EVENT LOG SHOWED NO ERRORS DURING PROCESSING AND SPECIMEN IS ACCEPTABLE BASED ON COLOR CHECK IMAGES AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION ON MAY AND THE FSE FOUND NO HARDWARE ISSUES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Connection Problem | Device problem code
7602596 | 111843530 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES WERE FOUND TO BE OPTIMAL AND ALIGNED EVENT LOG SHOWED NO ERRORS DURING PROCESSING AND SPECIMEN IS ACCEPTABLE BASED ON COLOR CHECK IMAGES AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION ON MAY AND THE FSE FOUND NO HARDWARE ISSUES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Interoperability Problem Identified | Medical Evaluation result code
7624603 | 111848999 | N | IF IMPLANTED GIVE DATE UNKNOWN IF THE LENS WAS IMPLANTED IF EXPLANTED GIVE DATE UNKNOWN IF THE LENS WAS IMPLANTED AND THEREFORE UNKNOWN IF EXPLANTED ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON AND JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
7624603 | 111848999 | N | IF IMPLANTED GIVE DATE UNKNOWN IF THE LENS WAS IMPLANTED IF EXPLANTED GIVE DATE UNKNOWN IF THE LENS WAS IMPLANTED AND THEREFORE UNKNOWN IF EXPLANTED ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON AND JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7627734 | 111974559 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS WITH BLOOD GLUCOSE OF MGDL BLOOD GLUCOSE VALUE WHEN ADMITTED TO HOSPITAL MGLDL OTHER BLOOD GLUCOSE WAS MGDL THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS ABDOMINAL PAIN THE CUSTOMER WAS TREATED WITH MANUAL INJECTION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Pain, Abdominal | Patient problem code
7627734 | 111974559 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS WITH BLOOD GLUCOSE OF MGDL BLOOD GLUCOSE VALUE WHEN ADMITTED TO HOSPITAL MGLDL OTHER BLOOD GLUCOSE WAS MGDL THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS ABDOMINAL PAIN THE CUSTOMER WAS TREATED WITH MANUAL INJECTION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
7627734 | 111974559 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEY WERE HOSPITALIZED DUE TO HIGH BLOOD GLUCOSE AND DIABETIC KETOACIDOSIS WITH BLOOD GLUCOSE OF MGDL BLOOD GLUCOSE VALUE WHEN ADMITTED TO HOSPITAL MGLDL OTHER BLOOD GLUCOSE WAS MGDL THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS ABDOMINAL PAIN THE CUSTOMER WAS TREATED WITH MANUAL INJECTION THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE HOSPITALIZATION THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient problem code
7618499 | 112026835 | N | BOSTON SCIENTIFIC TECHNICAL SERVICES TS REVIEWED DATA SENT IN FOR ANALYSIS AND CONFIRMED THERE HAVE BEEN ABRUPT INCREASES IN THE RV PACING IMPEDANCE SINCE B THERE WAS NO EVIDENCE OF ANY STORED EPISODES INDICATING NOISE THE BIVENTRICULAR PACING PERCENT WAS CONSISTENT AT AROUND PERCENT TS DISCUSSED PERFORMING ADDITIONAL AGGRESSIVE PATIENT TESTING AT ROUTINE FOLLOWUP APPOINTMENTS REMOTE MONITORING WAS ALSO RECOMMENDED TO FURTHER WATCH THE ISSUE NO FURTHER ACTION IS PLANNED TO BE TAKEN AT THIS TIME | Impedance Problem | Device problem code
7618499 | 112026835 | N | BOSTON SCIENTIFIC TECHNICAL SERVICES TS REVIEWED DATA SENT IN FOR ANALYSIS AND CONFIRMED THERE HAVE BEEN ABRUPT INCREASES IN THE RV PACING IMPEDANCE SINCE B THERE WAS NO EVIDENCE OF ANY STORED EPISODES INDICATING NOISE THE BIVENTRICULAR PACING PERCENT WAS CONSISTENT AT AROUND PERCENT TS DISCUSSED PERFORMING ADDITIONAL AGGRESSIVE PATIENT TESTING AT ROUTINE FOLLOWUP APPOINTMENTS REMOTE MONITORING WAS ALSO RECOMMENDED TO FURTHER WATCH THE ISSUE NO FURTHER ACTION IS PLANNED TO BE TAKEN AT THIS TIME | Noise, Audible | Device problem code
7618499 | 112026835 | N | BOSTON SCIENTIFIC TECHNICAL SERVICES TS REVIEWED DATA SENT IN FOR ANALYSIS AND CONFIRMED THERE HAVE BEEN ABRUPT INCREASES IN THE RV PACING IMPEDANCE SINCE B THERE WAS NO EVIDENCE OF ANY STORED EPISODES INDICATING NOISE THE BIVENTRICULAR PACING PERCENT WAS CONSISTENT AT AROUND PERCENT TS DISCUSSED PERFORMING ADDITIONAL AGGRESSIVE PATIENT TESTING AT ROUTINE FOLLOWUP APPOINTMENTS REMOTE MONITORING WAS ALSO RECOMMENDED TO FURTHER WATCH THE ISSUE NO FURTHER ACTION IS PLANNED TO BE TAKEN AT THIS TIME | Impedance Problem Identified | Medical Evaluation result code
7617055 | 112040545 | N | MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER IN THIS INSTANCE THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I DRIVE A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED THE EXAM WAS REREAD AND CONFIRMED WITH THE CUSTOMER THAT THE REPORT WAS PRESENT NO FURTHER REVIEWINVESTIGATION WILL BE PERFORMED FOR CAUSE | Device Displays Incorrect Message | Device problem code
7617055 | 112040545 | N | MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER IN THIS INSTANCE THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I DRIVE A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED THE EXAM WAS REREAD AND CONFIRMED WITH THE CUSTOMER THAT THE REPORT WAS PRESENT NO FURTHER REVIEWINVESTIGATION WILL BE PERFORMED FOR CAUSE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7602309 | 112150863 | N | TESTING WAS PERFORMED AT ALERE B ON RETAINED KIT LOT WITH THE FOLLOWING INTERNAL WHOLE BLOOD AND SERUMPLASMA CONTROL SAMPLES B ALL TEST RESULTS WERE VALID AND PERFORMED AS EXPECTED ADDITIONALLY THE MANUFACTURING BATCH RECORDS FOR LOT WERE REVIEWED THIS LOT MET THE REQUIRED RELEASE SPECIFICATIONS A REVIEW OF THE COMPLAINTS REPORTED FALSE POSITIVE OR UNCONFIRMED FALSE POSITIVE RELATED TO LOT NUMBER SHOWED THAT THE COMPLAINT RATE IS B THE EVIDENCE AVAILABLE DOES NOT INDICATE THAT THE PRODUCT IS PERFORMING OUTSIDE LABEL CLAIMS ADDITIONAL INVESTIGATION REVEALED THAT THE CUSTOMER WAS NOT FOLLOWING LABEL INSTRUCTIONS FOR PERFORMING THE ALERE DETERMINE HIV AGAB COMBO TEST THERE WERE ISSUES IDENTIFIED IN THE SAMPLE COLLECTION AND APPLICATION AND ISSUES WITH THE ADDITION OF THE CHASE BUFFER WHICH CAN BE ASSOCIATED WITH FALSE RESULTS NEGATIVE AND POSITIVE THE AVAILABLE EVIDENCE SUGGESTS THAT THIS DEVICE LOT IS PERFORMING WITHIN LABELED CLAIMS | False Positive Result | Device problem code
7602309 | 112150863 | N | TESTING WAS PERFORMED AT ALERE B ON RETAINED KIT LOT WITH THE FOLLOWING INTERNAL WHOLE BLOOD AND SERUMPLASMA CONTROL SAMPLES B ALL TEST RESULTS WERE VALID AND PERFORMED AS EXPECTED ADDITIONALLY THE MANUFACTURING BATCH RECORDS FOR LOT WERE REVIEWED THIS LOT MET THE REQUIRED RELEASE SPECIFICATIONS A REVIEW OF THE COMPLAINTS REPORTED FALSE POSITIVE OR UNCONFIRMED FALSE POSITIVE RELATED TO LOT NUMBER SHOWED THAT THE COMPLAINT RATE IS B THE EVIDENCE AVAILABLE DOES NOT INDICATE THAT THE PRODUCT IS PERFORMING OUTSIDE LABEL CLAIMS ADDITIONAL INVESTIGATION REVEALED THAT THE CUSTOMER WAS NOT FOLLOWING LABEL INSTRUCTIONS FOR PERFORMING THE ALERE DETERMINE HIV AGAB COMBO TEST THERE WERE ISSUES IDENTIFIED IN THE SAMPLE COLLECTION AND APPLICATION AND ISSUES WITH THE ADDITION OF THE CHASE BUFFER WHICH CAN BE ASSOCIATED WITH FALSE RESULTS NEGATIVE AND POSITIVE THE AVAILABLE EVIDENCE SUGGESTS THAT THIS DEVICE LOT IS PERFORMING WITHIN LABELED CLAIMS | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7606916 | 112156439 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES ARE OPTIMAL AND ALIGNED EVENT LOG SHOWS NO ERRORS DURING PROCESSING AND SPECIMENS ARE ACCEPTABLE BASED ON COLOR CHECK IMAGES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Connection Problem | Device problem code
7606916 | 112156439 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES ARE OPTIMAL AND ALIGNED EVENT LOG SHOWS NO ERRORS DURING PROCESSING AND SPECIMENS ARE ACCEPTABLE BASED ON COLOR CHECK IMAGES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Interoperability Problem Identified | Medical Evaluation result code
7633883 | 112183260 | D | ENDURANT STENT GRAFTS WERE IMPLANTED IN PATIENTS FOR THE ENDOVASCULAR TREATMENT OF ABDOMINAL AORTIC ANEURYSMS AVERAGE DIAMETER OF MM FROM JANUARY TO DECEMBER HE FOLLOWING DEVICE MALFUNCTIONS WERE OBSERVED TYPE IA ENDOLEAK TYPE IB ENDOLEAK TYPE II ENDOLEAK TYPE III ENDOLEAK UNKNOWN SUBTYPE | Hemorrhage | Patient problem code
7633883 | 112183260 | D | ENDURANT STENT GRAFTS WERE IMPLANTED IN PATIENTS FOR THE ENDOVASCULAR TREATMENT OF ABDOMINAL AORTIC ANEURYSMS AVERAGE DIAMETER OF MM FROM JANUARY TO DECEMBER HE FOLLOWING DEVICE MALFUNCTIONS WERE OBSERVED TYPE IA ENDOLEAK TYPE IB ENDOLEAK TYPE II ENDOLEAK TYPE III ENDOLEAK UNKNOWN SUBTYPE | Adverse Event Without Identified Device or Use Problem | Device problem code
7624321 | 112203215 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Component Incompatible | Device problem code
7624321 | 112203215 | N | THE PROBLEM WAS DUE TO A COMPONENT FAILURE WE ARE UNAWARE ABOUT PATIENT INJURY | Interoperability Problem Identified | Medical Evaluation result code
7596959 | 112301363 | N | CORRECTION MEDTRONIC RECEIVED INFORMATION THAT THE AWARE DATE WAS B AFTER FURTHER REVIEW THE INITIAL MDR WAS SUBMITTED IN ERROR AND THE REPORTED EVENT SHOULD NOT HAVE BEEN CONSIDERED A REPORTABLE MALFUNCTION AS THE HIGH QUALITY D IMAGES WERE USED FOR THE PROCEDURE AND THERE WAS NO INDICATION THAT LOWER QUALITY FLUORO IMAGES WERE NEEDED FOR THE PROCEDURE THE REPORTED ISSUE WAS UNLIKELY TO CAUSE SERIOUS HARM | No Device Problem Found | Medical Evaluation result code
7615956 | 112312160 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES ARE OPTIMAL AND ALIGNED EVENT LOG SHOWS NO ERRORS DURING PROCESSING AND SPECIMENS ARE ACCEPTABLE BASED ON COLOR CHECK IMAGES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT THE INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA ALGORITHM TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE USER IS VISUALLY CONFIRMING ALL NEGATIVE CAPTURER READY ASSAY RESULTS REPORTED BY THE INSTRUMENT PER IMMUCOR TECHNICAL COMMUNICATION B THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Connection Problem | Device problem code
7615956 | 112312160 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO TRANSFER CAMERA IMAGES EVENT LOGS AND RESULT BATCHES CAMERA IMAGES ARE OPTIMAL AND ALIGNED EVENT LOG SHOWS NO ERRORS DURING PROCESSING AND SPECIMENS ARE ACCEPTABLE BASED ON COLOR CHECK IMAGES COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT THE INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA ALGORITHM TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE USER IS VISUALLY CONFIRMING ALL NEGATIVE CAPTURER READY ASSAY RESULTS REPORTED BY THE INSTRUMENT PER IMMUCOR TECHNICAL COMMUNICATION B THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Interoperability Problem Identified | Medical Evaluation result code
7640296 | 112424063 | N | THE CUSTOMERS REPORT OF A LEAK WAS CONFIRMED VISUAL INSPECTION SHOWED THAT ONE OF THE PVC TUBINGS ABOVE THE DRIP CHAMBER WAS DETACHED FROM ITS INLET PORT ON THE DRIP CHAMBER CAP FURTHER INSPECTION OF THE TUBINGS SEPARATED END SHOWED INSUFFICIENT SOLVENT THE ROOT CAUSE OF THE SEPARATION WAS IDENTIFIED AS INSUFFICIENT SOLVENT BEING APPLIED AT THE ENGAGEMENT OF THE TUBING | Material Separation | Device problem code
7640296 | 112424063 | N | THE CUSTOMERS REPORT OF A LEAK WAS CONFIRMED VISUAL INSPECTION SHOWED THAT ONE OF THE PVC TUBINGS ABOVE THE DRIP CHAMBER WAS DETACHED FROM ITS INLET PORT ON THE DRIP CHAMBER CAP FURTHER INSPECTION OF THE TUBINGS SEPARATED END SHOWED INSUFFICIENT SOLVENT THE ROOT CAUSE OF THE SEPARATION WAS IDENTIFIED AS INSUFFICIENT SOLVENT BEING APPLIED AT THE ENGAGEMENT OF THE TUBING | Leak / Splash | Device problem code
7640296 | 112424063 | N | THE CUSTOMERS REPORT OF A LEAK WAS CONFIRMED VISUAL INSPECTION SHOWED THAT ONE OF THE PVC TUBINGS ABOVE THE DRIP CHAMBER WAS DETACHED FROM ITS INLET PORT ON THE DRIP CHAMBER CAP FURTHER INSPECTION OF THE TUBINGS SEPARATED END SHOWED INSUFFICIENT SOLVENT THE ROOT CAUSE OF THE SEPARATION WAS IDENTIFIED AS INSUFFICIENT SOLVENT BEING APPLIED AT THE ENGAGEMENT OF THE TUBING | Detachment of Device or device Component | Device problem code
7640296 | 112424063 | N | THE CUSTOMERS REPORT OF A LEAK WAS CONFIRMED VISUAL INSPECTION SHOWED THAT ONE OF THE PVC TUBINGS ABOVE THE DRIP CHAMBER WAS DETACHED FROM ITS INLET PORT ON THE DRIP CHAMBER CAP FURTHER INSPECTION OF THE TUBINGS SEPARATED END SHOWED INSUFFICIENT SOLVENT THE ROOT CAUSE OF THE SEPARATION WAS IDENTIFIED AS INSUFFICIENT SOLVENT BEING APPLIED AT THE ENGAGEMENT OF THE TUBING | Leakage/Seal | Medical Evaluation result code
7640296 | 112424063 | N | THE CUSTOMERS REPORT OF A LEAK WAS CONFIRMED VISUAL INSPECTION SHOWED THAT ONE OF THE PVC TUBINGS ABOVE THE DRIP CHAMBER WAS DETACHED FROM ITS INLET PORT ON THE DRIP CHAMBER CAP FURTHER INSPECTION OF THE TUBINGS SEPARATED END SHOWED INSUFFICIENT SOLVENT THE ROOT CAUSE OF THE SEPARATION WAS IDENTIFIED AS INSUFFICIENT SOLVENT BEING APPLIED AT THE ENGAGEMENT OF THE TUBING | Fracture Problem | Medical Evaluation result code
7644599 | 112544116 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
7644599 | 112544116 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
7644632 | 112544283 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
7644632 | 112544283 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
7644615 | 112544348 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Component Incompatible | Device problem code
7644615 | 112544348 | N | IMPLANT LOSS IS KNOWN TO OCCUR CONSIDERED IN THE RISK MANAGEMENT FILE AND COMMUNICATED IN THE IFU THERE ARE NO INDICATIONS THAT THE OCCURRED IS A RESULT OF MANUFACTURING OR COMPONENT FAILURE THESE INCIDENTS DO NOT INVOLVE SERIOUS INJURY AND ARE NOT CONSIDERED AS SAFETY ISSUES BUT THEY ARE REPORTED DUE TO THE INTERVENTION REQUIRED FOR THE PATIENT TO BE ABLE TO CONTINUE USING THE BONE ANCHORED HEARING SYSTEM | Interoperability Problem Identified | Medical Evaluation result code
7626952 | 112551484 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON MAY WHICH APPEARED VISUALLY POSITIVE THE CAMERA REPORT APPEARED ACCEPTABLE COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Connection Problem | Device problem code
7626952 | 112551484 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON MAY WHICH APPEARED VISUALLY POSITIVE THE CAMERA REPORT APPEARED ACCEPTABLE COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Interoperability Problem Identified | Medical Evaluation result code
7635650 | 112554683 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION ON B WHICH APPEARED VISUALLY WEAK POSITIVE THE CAMERA REPORT APPEARED OPTIMAL COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Connection Problem | Device problem code
7654614 | 112865235 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT NOTED THE ADHESIVE WAS WET | Insufficient Information | Device problem code
7654614 | 112865235 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT NOTED THE ADHESIVE WAS WET | No Known Impact Or Consequence To Patient | Patient Problem Code
7656006 | 112942875 | D | INFORMATION WAS RECEIVED FROM THE MANUFACTURER REPRESENTATIVE REP REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR SPINAL PAIN IT WAS REPORTED THAT THE PATIENT HAD A FEVER AND INFECTION AT THE SURGICAL SITE NO FACTORS WERE REPORTED TO HAVE CONTRIBUTED TO THE ISSUE NO DIAGNOSTIC TROUBLESHOOTING WAS PERFORMED ANTIBIOTICS WERE GIVEN AND THE PATIENTS SYSTEM WAS EXPLANTED ON B THE INS AND LEADS WERE REMOVED AND WOULD NOT BE RETURNED FOR ANALYSIS AS THE CUSTOMER DISCARDED THEM IT WAS REPORTED THAT THE ISSUE WAS RESOLVED AT THE TIME OF THE EVENT THE DEVICE WAS NOT REPLACED AT THE TIME OF EXPLANT BUT IT WAS INDICATED THAT IT WOULD BE REPLACED IN THE FUTURE WITH A DEVICE FROM THE SAME MANUFACTURER THE CUSTOMER WAS NOT NOTIFIED THAT THE PRODUCT SHOULD BE RETURNED BECAUSE THERE WAS NO DEVICE ISSUE REPORTED THE EVENT OCCURRED POSTOPERATIVELY ON B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | Adverse Event Without Identified Device or Use Problem | Device problem code
7656006 | 112942875 | D | INFORMATION WAS RECEIVED FROM THE MANUFACTURER REPRESENTATIVE REP REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR INS FOR SPINAL PAIN IT WAS REPORTED THAT THE PATIENT HAD A FEVER AND INFECTION AT THE SURGICAL SITE NO FACTORS WERE REPORTED TO HAVE CONTRIBUTED TO THE ISSUE NO DIAGNOSTIC TROUBLESHOOTING WAS PERFORMED ANTIBIOTICS WERE GIVEN AND THE PATIENTS SYSTEM WAS EXPLANTED ON B THE INS AND LEADS WERE REMOVED AND WOULD NOT BE RETURNED FOR ANALYSIS AS THE CUSTOMER DISCARDED THEM IT WAS REPORTED THAT THE ISSUE WAS RESOLVED AT THE TIME OF THE EVENT THE DEVICE WAS NOT REPLACED AT THE TIME OF EXPLANT BUT IT WAS INDICATED THAT IT WOULD BE REPLACED IN THE FUTURE WITH A DEVICE FROM THE SAME MANUFACTURER THE CUSTOMER WAS NOT NOTIFIED THAT THE PRODUCT SHOULD BE RETURNED BECAUSE THERE WAS NO DEVICE ISSUE REPORTED THE EVENT OCCURRED POSTOPERATIVELY ON B NO FURTHER COMPLICATIONS WERE REPORTEDANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
7657849 | 112994211 | D | A HELIFX ENDOANCHORING SYSTEM WAS USED PROPHYLACTICALLY AS AN ACCESSORY DEVICE IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM ABDOMINAL AORTIC ANEURYSM A XXMM GORE EXCLUDER BIFURCATED STENT GRAFT AN ENDURANT XX CONTRALATERAL LIMB A XMM GORE EXCLUDER ILIAC LIMB EXTENSION A XX GORE TAG EXTENSION LIMB AND A GORE TAG THORACIC GRAFT USED AS A CUFF WERE ALSO IMPLANTED DURING THE PROCEDURE A TOTAL OF ENDOANCHORS WERE USED DURING THE PROCEDURE THE AORTIC DIAMETER AT MM BELOW THE LOWEST RENAL WAS MM AND THE AORTIC DIAMETER AT MM BELOW THE LOWEST RENAL WAS MM THE NECK LENGTH WAS REPORTED AS MM AND THE AORTIC DIAMETER AT THE DISTAL END OF THE PROXIMAL NECK WAS MM THE PROXIMAL NECK ANGLE WAS REPORTED AS DEGREES LOCALIZED THROMBUS WAS REPORTED AT THE SEAL ZONE WITH CIRCUMFERENCE COVERED BY THROMBUS WITH A MAXIMUM THICKNESS OF MM IT WAS REPORTED THE PATIENT PRESENTED AND HALF YEARS LATER WITH SUDDEN ONSET OF DIFFUSE LOWER ABDOMINAL PAIN THAT RADIATED INTO THE BACK ONE EPISODE OF EMESIS AND RECENT BLOOD DURING BOWEL MOVEMENTS IT WAS REPORTED THAT A TYPE IA ENDOLEAK WAS OBSERVED ON A CT SCAN AND ABDOMINAL RADIOGRAPH THE CAUSE OF THE EVENT IS UNKNOWN IT WAS NOTED TO NOT BE RELATED TO THE STUDY PROCEDURE AND TO HAVE A CAUSAL RELATIONSHIP TO THE ENDOANCHORS AN EMERGENCY INTERVENTION PROCEDURE WAS CARRIED OUT THE NEXT DAY TO TREAT THE TYPE IA ENDOLEAK AND AN ENDURANT BIFURCATE STENT GRAFT XXMM AND LIMB EXTENSION XXMM WERE IMPLANTED IN THE PATIENT THE EVENT WAS REPORTED TO BE RESOLVED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Insufficient Information | Device problem code
7657849 | 112994211 | D | A HELIFX ENDOANCHORING SYSTEM WAS USED PROPHYLACTICALLY AS AN ACCESSORY DEVICE IN A PATIENT FOR THE ENDOVASCULAR TREATMENT OF A MM ABDOMINAL AORTIC ANEURYSM A XXMM GORE EXCLUDER BIFURCATED STENT GRAFT AN ENDURANT XX CONTRALATERAL LIMB A XMM GORE EXCLUDER ILIAC LIMB EXTENSION A XX GORE TAG EXTENSION LIMB AND A GORE TAG THORACIC GRAFT USED AS A CUFF WERE ALSO IMPLANTED DURING THE PROCEDURE A TOTAL OF ENDOANCHORS WERE USED DURING THE PROCEDURE THE AORTIC DIAMETER AT MM BELOW THE LOWEST RENAL WAS MM AND THE AORTIC DIAMETER AT MM BELOW THE LOWEST RENAL WAS MM THE NECK LENGTH WAS REPORTED AS MM AND THE AORTIC DIAMETER AT THE DISTAL END OF THE PROXIMAL NECK WAS MM THE PROXIMAL NECK ANGLE WAS REPORTED AS DEGREES LOCALIZED THROMBUS WAS REPORTED AT THE SEAL ZONE WITH CIRCUMFERENCE COVERED BY THROMBUS WITH A MAXIMUM THICKNESS OF MM IT WAS REPORTED THE PATIENT PRESENTED AND HALF YEARS LATER WITH SUDDEN ONSET OF DIFFUSE LOWER ABDOMINAL PAIN THAT RADIATED INTO THE BACK ONE EPISODE OF EMESIS AND RECENT BLOOD DURING BOWEL MOVEMENTS IT WAS REPORTED THAT A TYPE IA ENDOLEAK WAS OBSERVED ON A CT SCAN AND ABDOMINAL RADIOGRAPH THE CAUSE OF THE EVENT IS UNKNOWN IT WAS NOTED TO NOT BE RELATED TO THE STUDY PROCEDURE AND TO HAVE A CAUSAL RELATIONSHIP TO THE ENDOANCHORS AN EMERGENCY INTERVENTION PROCEDURE WAS CARRIED OUT THE NEXT DAY TO TREAT THE TYPE IA ENDOLEAK AND AN ENDURANT BIFURCATE STENT GRAFT XXMM AND LIMB EXTENSION XXMM WERE IMPLANTED IN THE PATIENT THE EVENT WAS REPORTED TO BE RESOLVED NO ADDITIONAL CLINICAL SEQUELAE WERE REPORTED AND THE PATIENT IS FINE | Hemorrhage | Patient problem code
7657841 | 112998585 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS PATIENT TOOK OFF THE POD BECAUSE SHE WAS DUE TO CHANGE IT AND NOTICED THE CANNULA WASNT IN HER SKIN | No Known Impact Or Consequence To Patient | Patient Problem Code
7657841 | 112998585 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MGDL WHILE WEARING THE POD LONGER THAN HOURS PATIENT TOOK OFF THE POD BECAUSE SHE WAS DUE TO CHANGE IT AND NOTICED THE CANNULA WASNT IN HER SKIN | Insufficient Information | Device problem code
7658388 | 113014964 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | Positioning Problem | Device problem code
7658388 | 113014964 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
7658735 | 113026474 | D | PATIENT WAS INITIALLY IMPLANTED WITH A BIFURCATED SUPRARENAL AND LIMB GRAFT IN PATIENT PRESENTED RECENTLY WITH AN ENDOLEAK TYPE IIIB THE PHYSICIAN PLANS TO RELINE AND PROCEDURE IS SCHEDULED FOR B NO ADDITIONAL PATIENT SEQUALAE HAS BEEN REPORTED AT THIS TIME | Insufficient Information | Device problem code
7658735 | 113026474 | D | PATIENT WAS INITIALLY IMPLANTED WITH A BIFURCATED SUPRARENAL AND LIMB GRAFT IN PATIENT PRESENTED RECENTLY WITH AN ENDOLEAK TYPE IIIB THE PHYSICIAN PLANS TO RELINE AND PROCEDURE IS SCHEDULED FOR B NO ADDITIONAL PATIENT SEQUALAE HAS BEEN REPORTED AT THIS TIME | Hemorrhage | Patient problem code
7659192 | 113045591 | D | A PATIENT ON CONTINUOUS CYCLING PERITONEAL DIALYSIS CCPD FOR RENAL REPLACEMENT THERAPY RRT CONTACTED TECHNICAL SUPPORT FOR DRAINING SLOWLY THE PATIENT DIDNT FEEL WELL AND WAS NAUSEOUS THE PATIENT REPORTED HAVING HER CATHETER NOT A FRESENIUS PRODUCT REPOSITIONED EARLIER IN THE DAY AND HAD RECEIVED A REPLACEMENT CYCLER THE EVENING BEFORE THE PATIENT WAS IN DWELL OF AND TECHNICAL SUPPORT ASSISTED THE PATIENT WITH A STAT DRAIN THE PATIENT WAS STILL DRAINING SLOWLY BUT FELT A LITTLE BETTER UPON FOLLOWUP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE PDRN ON B IT WAS STATED THAT THE PATIENT UNDERWENT A CATHETER ASSESSMENT DATE UNKNOWN AND WAS ADVISED TO HAVE THE CATHETER REPLACED ON B THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR THE CATHETER REPLACEMENT SURGERY THE PATIENT WAS KEPT IN THE HOSPITAL FOR OBSERVATION AND UNDERWENT HEMODIALYSIS HD THERAPY AS THE PATIENT WAS FLUID OVERLOADED AND TO ALLOW FOR THE NEW CATHETER SITE TO HEAL THE PATIENT WAS DISCHARGED B AND ADVISED TO COMPLETE A FEW MORE HD TREATMENTS IN ORDER TO PULL OFF THE EXCESS FLUID THE PATIENT REFUSED HD AS THE TREATMENT SCARED HER THE PATIENT RESUMED PD THERAPY BUT WAS STILL EXPERIENCING DRAIN ISSUES THE PATIENTS WEIGHT INCREASED BY B DUE TO THE VOLUME OVERLOAD THE PATIENT DID IMPROVE WITH PD THERAPY AND THE CLINIC PDRN CONFIRMED THE PATIENT WAS DRAINING ADEQUATELY THE PATIENT UTILIZED DELFLEX TO PULL OFF THE EXCESS FLUID WHEN NORMALLY SHE UTILIZES THE DELFLEX | Nausea | Patient problem code
7659192 | 113045591 | D | A PATIENT ON CONTINUOUS CYCLING PERITONEAL DIALYSIS CCPD FOR RENAL REPLACEMENT THERAPY RRT CONTACTED TECHNICAL SUPPORT FOR DRAINING SLOWLY THE PATIENT DIDNT FEEL WELL AND WAS NAUSEOUS THE PATIENT REPORTED HAVING HER CATHETER NOT A FRESENIUS PRODUCT REPOSITIONED EARLIER IN THE DAY AND HAD RECEIVED A REPLACEMENT CYCLER THE EVENING BEFORE THE PATIENT WAS IN DWELL OF AND TECHNICAL SUPPORT ASSISTED THE PATIENT WITH A STAT DRAIN THE PATIENT WAS STILL DRAINING SLOWLY BUT FELT A LITTLE BETTER UPON FOLLOWUP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE PDRN ON B IT WAS STATED THAT THE PATIENT UNDERWENT A CATHETER ASSESSMENT DATE UNKNOWN AND WAS ADVISED TO HAVE THE CATHETER REPLACED ON B THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR THE CATHETER REPLACEMENT SURGERY THE PATIENT WAS KEPT IN THE HOSPITAL FOR OBSERVATION AND UNDERWENT HEMODIALYSIS HD THERAPY AS THE PATIENT WAS FLUID OVERLOADED AND TO ALLOW FOR THE NEW CATHETER SITE TO HEAL THE PATIENT WAS DISCHARGED B AND ADVISED TO COMPLETE A FEW MORE HD TREATMENTS IN ORDER TO PULL OFF THE EXCESS FLUID THE PATIENT REFUSED HD AS THE TREATMENT SCARED HER THE PATIENT RESUMED PD THERAPY BUT WAS STILL EXPERIENCING DRAIN ISSUES THE PATIENTS WEIGHT INCREASED BY B DUE TO THE VOLUME OVERLOAD THE PATIENT DID IMPROVE WITH PD THERAPY AND THE CLINIC PDRN CONFIRMED THE PATIENT WAS DRAINING ADEQUATELY THE PATIENT UTILIZED DELFLEX TO PULL OFF THE EXCESS FLUID WHEN NORMALLY SHE UTILIZES THE DELFLEX | Insufficient Information | Device problem code
7660170 | 113152244 | D | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN IN THE FEMORAL AREA AND BURNING SENSATION IN LOWER BACK ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Pain | Patient problem code
7660170 | 113152244 | D | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN IN THE FEMORAL AREA AND BURNING SENSATION IN LOWER BACK ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Insufficient Information | Device problem code
7660170 | 113152244 | D | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN IN THE FEMORAL AREA AND BURNING SENSATION IN LOWER BACK ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE | Burning Sensation | Patient problem code
7659669 | 113178357 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP REGARDING AN IMPLANTABLE NEUROSTIMULATOR INS FOR PARKINSONS DUAL MOVEMENT DISORDERS IT WAS REPORTED THAT THE PATIENTS BATTERY DISCHARGED THEIR SYMPTOMS RETURNED WHICH WERE TREMOR RIGIDITY AND BRADYKINESIA THE REP THEN INTERROGATED WITH THE DEVICE RECHARGED IT AND TURNED THE DEVICE ON THE ISSUES WERE RESOLVED THERE WERE NO FURTHER COMPLICATIONS REPORTED | Adverse Event Without Identified Device or Use Problem | Device problem code
7659669 | 113178357 | D | INFORMATION WAS RECEIVED BY A MANUFACTURE REPRESENTATIVE REP REGARDING AN IMPLANTABLE NEUROSTIMULATOR INS FOR PARKINSONS DUAL MOVEMENT DISORDERS IT WAS REPORTED THAT THE PATIENTS BATTERY DISCHARGED THEIR SYMPTOMS RETURNED WHICH WERE TREMOR RIGIDITY AND BRADYKINESIA THE REP THEN INTERROGATED WITH THE DEVICE RECHARGED IT AND TURNED THE DEVICE ON THE ISSUES WERE RESOLVED THERE WERE NO FURTHER COMPLICATIONS REPORTED | Muscular Rigidity | Patient problem code
7663583 | 113179711 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | No Known Impact Or Consequence To Patient | Patient Problem Code
7663583 | 113179711 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | Positioning Problem | Device problem code
7658948 | 113259977 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | No Known Impact Or Consequence To Patient | Patient Problem Code
7658948 | 113259977 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | Positioning Problem | Device problem code
7658815 | 113287862 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS | Insufficient Information | Device problem code
7658815 | 113287862 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS | No Known Impact Or Consequence To Patient | Patient Problem Code
7658830 | 113287909 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED CANNULA DID NOT GO ALL THE WAY INTO THE SKIN AND DID NOT KNOW THIS UNTIL THE POD WAS TAKEN OFF THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM AND FELT THE BACK OF THE ARM WAS WET AS TREATMENT A NEW POD WAS APPLIED | No Known Impact Or Consequence To Patient | Patient Problem Code
7658830 | 113287909 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED CANNULA DID NOT GO ALL THE WAY INTO THE SKIN AND DID NOT KNOW THIS UNTIL THE POD WAS TAKEN OFF THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM AND FELT THE BACK OF THE ARM WAS WET AS TREATMENT A NEW POD WAS APPLIED | Positioning Problem | Device problem code
7666140 | 113305456 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS AND ADMITTED TO THE HOSPITAL TREATMENT INCLUDED INTRAVENOUS DELIVERY OF FLUIDS SALINE AND INSULIN | Insufficient Information | Device problem code
7666140 | 113305456 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN PATIENT WAS DIAGNOSED WITH DIABETIC KETOACIDOSIS AND ADMITTED TO THE HOSPITAL TREATMENT INCLUDED INTRAVENOUS DELIVERY OF FLUIDS SALINE AND INSULIN | Diabetic Ketoacidosis | Patient problem code
7666317 | 113310961 | D | OUS MDR IT WAS REPORTED THAT MONTHS AFTER THE IMPLANT EXIT BLOCK OCCURRED WITH STABLE SENSING VALUES INTRAOPERATIVELY THE LEAD DID NOT SHOW ANY PECULIARITIES THE GENERATOR WAS EXPLANTED | Exit Block | Patient problem code
7666317 | 113310961 | D | OUS MDR IT WAS REPORTED THAT MONTHS AFTER THE IMPLANT EXIT BLOCK OCCURRED WITH STABLE SENSING VALUES INTRAOPERATIVELY THE LEAD DID NOT SHOW ANY PECULIARITIES THE GENERATOR WAS EXPLANTED | Insufficient Information | Device problem code
7658919 | 113397320 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | Positioning Problem | Device problem code
7658919 | 113397320 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | No Known Impact Or Consequence To Patient | Patient Problem Code
7659601 | 113397322 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR MORE THAN HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM PATIENT IS ALSO REPORTING THERE WAS BLEEDING AT THE INSERTION SITE AND A BLOOD CLOT WAS FOUND AS WELL PATIENT REPORTEDLY DRANK A LARGE SUM OF WATER AND HER BG LEVELS WENT DOWN A NEW POD WAS APPLIED | Positioning Problem | Device problem code
7659601 | 113397322 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR MORE THAN HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM PATIENT IS ALSO REPORTING THERE WAS BLEEDING AT THE INSERTION SITE AND A BLOOD CLOT WAS FOUND AS WELL PATIENT REPORTEDLY DRANK A LARGE SUM OF WATER AND HER BG LEVELS WENT DOWN A NEW POD WAS APPLIED | Hemorrhage | Patient problem code
7659601 | 113397322 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR MORE THAN HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM PATIENT IS ALSO REPORTING THERE WAS BLEEDING AT THE INSERTION SITE AND A BLOOD CLOT WAS FOUND AS WELL PATIENT REPORTEDLY DRANK A LARGE SUM OF WATER AND HER BG LEVELS WENT DOWN A NEW POD WAS APPLIED | Hematoma | Patient problem code
7658968 | 113397365 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | Positioning Problem | Device problem code
7658968 | 113397365 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | No Known Impact Or Consequence To Patient | Patient Problem Code
7659016 | 113397537 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS AND NOT FEELING CANNULA PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN AS TREATMENT INSULIN WAS DELIVERED VIA INSULIN PEN PATIENT NO LONGER HAS DEVICE THEREFORE UNABLE TO REQUEST RETURN | No Known Impact Or Consequence To Patient | Patient Problem Code
7659016 | 113397537 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS AND NOT FEELING CANNULA PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN AS TREATMENT INSULIN WAS DELIVERED VIA INSULIN PEN PATIENT NO LONGER HAS DEVICE THEREFORE UNABLE TO REQUEST RETURN | Positioning Problem | Device problem code
7658884 | 113397748 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | Positioning Problem | Device problem code
7658884 | 113397748 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT FELT INSULIN ON THE SKIN AND THAT THE CANNULA WAS NOT INSERTED AS IT SHOULD THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT PATIENT DID CORRECTIONS BUT THAT DID NOT LOWER THE BG LEVELS | No Known Impact Or Consequence To Patient | Patient Problem Code
7670256 | 113510954 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS ARM FOR LESS THAN AN HOUR IT WAS ALSO STATED THAT THE CANNULA WAS NOT IN AND IT HURT THAT PATIENTS SITE | No Code Available | Patient problem code
7670256 | 113510954 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS ARM FOR LESS THAN AN HOUR IT WAS ALSO STATED THAT THE CANNULA WAS NOT IN AND IT HURT THAT PATIENTS SITE | Positioning Problem | Device problem code
7670459 | 113512859 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED APPROXIMATELY MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT REPORTED THE POD FELL OFF A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Device Fell | Device problem code
7670459 | 113512859 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED APPROXIMATELY MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM THE PATIENT REPORTED THE POD FELL OFF A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
7655527 | 113520229 | D | NEW INFORMATION NOTED THAT THE PATIENT WOULD ONLY BE DISCHARGED HOME ON B IF THE PNEUMOTHORAX HAD RESOLVED OR IF THE PATIENT WAS STABLE AS THE PATIENT WAS DISCHARGED ON B THE PATIENT WAS IN STABLE CONDITION | No Consequences Or Impact To Patient | Patient problem code
7655527 | 113520229 | D | NEW INFORMATION NOTED THAT THE PATIENT WOULD ONLY BE DISCHARGED HOME ON B IF THE PNEUMOTHORAX HAD RESOLVED OR IF THE PATIENT WAS STABLE AS THE PATIENT WAS DISCHARGED ON B THE PATIENT WAS IN STABLE CONDITION | Insufficient Information | Device problem code
7673743 | 113536691 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD FOR MORE THAN HOURS ON THE ABDOMEN NO FURTHER INFORMATION WAS PROVIDED | Hyperglycemia | Patient problem code
7673743 | 113536691 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD FOR MORE THAN HOURS ON THE ABDOMEN NO FURTHER INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7675876 | 113620887 | D | ONE PATIENT B STATED THAT HE EXPERIENCED A BURNING SENSATION IN THE EYES AFTER USING THE MEDICAL DEVICE PUBLIX MULTIPURPOSE SOLUTION THE PATIENT MEDICAL RECORD IS NOT AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
7675876 | 113620887 | D | ONE PATIENT B STATED THAT HE EXPERIENCED A BURNING SENSATION IN THE EYES AFTER USING THE MEDICAL DEVICE PUBLIX MULTIPURPOSE SOLUTION THE PATIENT MEDICAL RECORD IS NOT AVAILABLE AT THIS TIME | Burning Sensation | Patient problem code
7674396 | 113630280 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN NO FURTHER INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7674396 | 113630280 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN NO FURTHER INFORMATION WAS PROVIDED | No Known Impact Or Consequence To Patient | Patient Problem Code
7674179 | 113632571 | D | THE PATIENT REPORTED BELIEVING THAT THE CANNULA WAS NOT INSERTED PROPERLY DUE TO SMELLING AND FEELING INSULIN ON HER SKIN HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS | Positioning Problem | Device problem code
7674179 | 113632571 | D | THE PATIENT REPORTED BELIEVING THAT THE CANNULA WAS NOT INSERTED PROPERLY DUE TO SMELLING AND FEELING INSULIN ON HER SKIN HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS | No Known Impact Or Consequence To Patient | Patient Problem Code
7674174 | 113634020 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM AS TREATMENT A NEW POD WAS APPLIED | Insufficient Information | Device problem code
7674174 | 113634020 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM AS TREATMENT A NEW POD WAS APPLIED | No Known Impact Or Consequence To Patient | Patient Problem Code
7678724 | 113710251 | D | THE PATIENT WITH THIS DEVICE IMPLANTED HAD AN MRI WITHOUT THE MRI PROGRAM ON THE DEVICE ACTIVATED AS A RESULT THE BATTERY DISCHARGED AND THE ICD WENT INTO EOS MODE | Defective Device | Device problem code
7678724 | 113710251 | D | THE PATIENT WITH THIS DEVICE IMPLANTED HAD AN MRI WITHOUT THE MRI PROGRAM ON THE DEVICE ACTIVATED AS A RESULT THE BATTERY DISCHARGED AND THE ICD WENT INTO EOS MODE | No Known Impact Or Consequence To Patient | Patient problem code
7670033 | 113754951 | D | OUR CUSTOMER COMPLAINED ABOUT THE INFUSION LINEWHICH IS ONLY PARTIAL THROUGH THAT THE EYE COLLAPSED DURING VITRECTOMY | Collapse | Patient problem code
7670033 | 113754951 | D | OUR CUSTOMER COMPLAINED ABOUT THE INFUSION LINEWHICH IS ONLY PARTIAL THROUGH THAT THE EYE COLLAPSED DURING VITRECTOMY | Insufficient Information | Device problem code
7672224 | 113785049 | D | NEW INFORMATION RECEIVED NOTED THAT THE LEAD WAS CAPPED ON B DURING DEVICE EXPLANT THE PATIENT WAS STABLE PERIOPERATIVELY AND WAS DISCHARGED ON THE SAME DAY OF PROCEDURE | Insufficient Information | Device problem code
7672224 | 113785049 | D | NEW INFORMATION RECEIVED NOTED THAT THE LEAD WAS CAPPED ON B DURING DEVICE EXPLANT THE PATIENT WAS STABLE PERIOPERATIVELY AND WAS DISCHARGED ON THE SAME DAY OF PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
7684128 | 113866977 | D | IT WAS REPORTED THAT THE PATIENTS PUMP WAS ERODING THROUGH THE SKIN TO PREVENT AN INFECTION THE PUMP WAS SUBSEQUENTLY EXPLANTED | Skin Erosion | Patient problem code
7684128 | 113866977 | D | IT WAS REPORTED THAT THE PATIENTS PUMP WAS ERODING THROUGH THE SKIN TO PREVENT AN INFECTION THE PUMP WAS SUBSEQUENTLY EXPLANTED | Material Erosion | Device problem code
7684086 | 113874203 | D | IT WAS REPORTED THAT THE CUSTOMER HAD AN OCCLUSION ISSUE WITH THEIR INFUSION SET BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE INCIDENT THE CUSTOMER WAS ADVISED TO CHANGE THEIR INFUSION SET NO TROUBLESHOOTING WAS PERFORMED NOTHING WILL BE REPLACED OR RETURNED | Infusion or Flow Problem | Device problem code
7684086 | 113874203 | D | IT WAS REPORTED THAT THE CUSTOMER HAD AN OCCLUSION ISSUE WITH THEIR INFUSION SET BLOOD GLUCOSE VALUE WAS UNKNOWN AT THE TIME OF THE INCIDENT THE CUSTOMER WAS ADVISED TO CHANGE THEIR INFUSION SET NO TROUBLESHOOTING WAS PERFORMED NOTHING WILL BE REPLACED OR RETURNED | No Known Impact Or Consequence To Patient | Patient Problem Code
7688386 | 114001628 | D | IT WAS REPORTED THE PATIENT COMPLAINED OF A BURNING SENSATION AT THE SCS IPG SITE EVEN WHEN THE STIMULATION WAS OFF SUBSEQUENTLY SURGICAL INTERVENTION WAS TAKEN AND THE PATIENTS SCS IPG WAS REPLACED | Appropriate Term/Code Not Available | Device problem code
7688386 | 114001628 | D | IT WAS REPORTED THE PATIENT COMPLAINED OF A BURNING SENSATION AT THE SCS IPG SITE EVEN WHEN THE STIMULATION WAS OFF SUBSEQUENTLY SURGICAL INTERVENTION WAS TAKEN AND THE PATIENTS SCS IPG WAS REPLACED | Burning Sensation | Patient problem code
7681513 | 114033001 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM NO FURTHER INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7681513 | 114033001 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM NO FURTHER INFORMATION WAS PROVIDED | No Known Impact Or Consequence To Patient | Patient Problem Code
7682448 | 114040991 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR MORE THAN HOURS THE NEEDLE MECHANISM DID RETRACT AS INTENDED DURING ACTIVATION POD WAS DISPOSED OFF NO FURTHER INFORMATION IS AVAILABLE | No Known Impact Or Consequence To Patient | Patient Problem Code
7682448 | 114040991 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD FOR MORE THAN HOURS THE NEEDLE MECHANISM DID RETRACT AS INTENDED DURING ACTIVATION POD WAS DISPOSED OFF NO FURTHER INFORMATION IS AVAILABLE | Adverse Event Without Identified Device or Use Problem | Device problem code
7681616 | 114041309 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM AS TREATMENT A NEW POD WAS APPLIED | Insufficient Information | Device problem code
7681616 | 114041309 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM AS TREATMENT A NEW POD WAS APPLIED | No Known Impact Or Consequence To Patient | Patient Problem Code
7681639 | 114047262 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | Insufficient Information | Device problem code
7681639 | 114047262 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
7690525 | 114087095 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN LONGER THAN HOURS AFTER SMELLINGSEEING INSULIN AT SITE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
7690525 | 114087095 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN LONGER THAN HOURS AFTER SMELLINGSEEING INSULIN AT SITE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
7692840 | 114139752 | D | IT WAS REPORTED THAT A NEEDLE STICK INJURY OCCURRED FROM A UNSPECIFIED BD PEN NEEDLE TO A WORKER AT A HOSPITAL IT IS UNKNOWN IF THE NEEDLE WAS CLEAN OR DIRTY NO MEDICAL INTERVENTION WAS REPORTED | Physical Resistance / Sticking | Device problem code
7692840 | 114139752 | D | IT WAS REPORTED THAT A NEEDLE STICK INJURY OCCURRED FROM A UNSPECIFIED BD PEN NEEDLE TO A WORKER AT A HOSPITAL IT IS UNKNOWN IF THE NEEDLE WAS CLEAN OR DIRTY NO MEDICAL INTERVENTION WAS REPORTED | Injury | Patient problem code
7691240 | 114145949 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS WHEN REMOVED FROM THE INFUSION SITE LEG THE PODS CANNULA WAS FOUND CURVED TO THE RIGHT AS TREATMENT SHE GAVE CORRECTIVE BOLUSES AND A NEW POD WAS APPLIED | No Known Impact Or Consequence To Patient | Patient Problem Code
7691240 | 114145949 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS WHEN REMOVED FROM THE INFUSION SITE LEG THE PODS CANNULA WAS FOUND CURVED TO THE RIGHT AS TREATMENT SHE GAVE CORRECTIVE BOLUSES AND A NEW POD WAS APPLIED | Insufficient Information | Device problem code
7693469 | 114146256 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A PARASTOMAL HERNIA REPAIR SURGERY ON B AND MESH WAS IMPLANTED IT WAS REPORTED THE PATIENT EXPERIENCED CHRONIC PAIN BOWEL OBSTRUCTIONS AND HERNIA RECURRENCE NO ADDITIONAL INFORMATION IS REPORTED | Insufficient Information | Device problem code
7693469 | 114146256 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A PARASTOMAL HERNIA REPAIR SURGERY ON B AND MESH WAS IMPLANTED IT WAS REPORTED THE PATIENT EXPERIENCED CHRONIC PAIN BOWEL OBSTRUCTIONS AND HERNIA RECURRENCE NO ADDITIONAL INFORMATION IS REPORTED | Pain | Patient problem code
7693469 | 114146256 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A PARASTOMAL HERNIA REPAIR SURGERY ON B AND MESH WAS IMPLANTED IT WAS REPORTED THE PATIENT EXPERIENCED CHRONIC PAIN BOWEL OBSTRUCTIONS AND HERNIA RECURRENCE NO ADDITIONAL INFORMATION IS REPORTED | Hernia | Patient problem code
7693469 | 114146256 | D | IT WAS REPORTED BY AN ATTORNEY THAT THE PATIENT UNDERWENT A PARASTOMAL HERNIA REPAIR SURGERY ON B AND MESH WAS IMPLANTED IT WAS REPORTED THE PATIENT EXPERIENCED CHRONIC PAIN BOWEL OBSTRUCTIONS AND HERNIA RECURRENCE NO ADDITIONAL INFORMATION IS REPORTED | Obstruction | Patient problem code
7686745 | 114172742 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7686745 | 114172742 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER HAD MANUAL INJECTIONS AVAILABLE FOR INSULIN THERAPY | Device Alarm System | Device problem code
7691092 | 114174079 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS TREATMENT INCLUDED CORRECTING HER BLOOD GLUCOSE WITH A NEW POD THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B THE DEVICE WAS RETURNED FOR EVALUATION | Insufficient Information | Device problem code
7691092 | 114174079 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS TREATMENT INCLUDED CORRECTING HER BLOOD GLUCOSE WITH A NEW POD THE PATIENTS BLOOD GLUCOSE HISTORY IS AS FOLLOWS B THE DEVICE WAS RETURNED FOR EVALUATION | No Known Impact Or Consequence To Patient | Patient Problem Code
7694128 | 114178164 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | Positioning Problem | Device problem code
7694128 | 114178164 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ARM A NEW POD WAS APPLIED IMMEDIATELY AFTER THE EVENT | No Known Impact Or Consequence To Patient | Patient Problem Code
7694297 | 114180805 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED ON AN UNKNOWN DATE DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7694297 | 114180805 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED ON AN UNKNOWN DATE DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7695380 | 114214434 | D | IT WAS REPORTED THAT THE BD VACUTAINER SAFETYLOK BLOOD COLLECTION SETS SAFETY DID NOT REMAIN CLOSED RESULTING IN A NEEDLE STICK INJURY THERE WAS NO REPORT OF MEDICAL INTERVENTION NOTED | Physical Resistance / Sticking | Device problem code
7695380 | 114214434 | D | IT WAS REPORTED THAT THE BD VACUTAINER SAFETYLOK BLOOD COLLECTION SETS SAFETY DID NOT REMAIN CLOSED RESULTING IN A NEEDLE STICK INJURY THERE WAS NO REPORT OF MEDICAL INTERVENTION NOTED | No Consequences Or Impact To Patient | Patient problem code
7696939 | 114277335 | D | ON B THE PATIENT UNDERWENT SURGICAL DEBRANCHING OF THE LEFT SUBCLAVIAN ARTERY THE LEFT COMMON CAROTID ARTERY AND THE INNOMINATE ARTERY PRIOR TO ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM TWO CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES PROXIMAL TGUJB DISTAL TGUJB WERE SUCCESSFULLY DEPLOYED ALTHOUGH A BIRD BEAK CONFIGURATION WAS OBSERVED IN THE PROXIMAL PART OF TGUJB THE PROCEDURE WAS COMPLETED WITHOUT FURTHER TREATMENT SINCE NO ENDOLEAK AND NO VESSEL RUPTUREDISSECTION WAS OBSERVED THE PATIENT TOLERATED THE PROCEDURE ON B A FOLLOWUP COMPUTED TOMOGRAPHY CT REVEALED A RETROGRADE THORACIC AORTIC DISSECTION ORIGINATING AROUND THE PROXIMAL END OF TGUJB DEVICE AND LEADING TO A CORONARY ARTERY AN OPEN SURGERY WAS PERFORMED TO TREAT THE DISSECTION THE CLINICAL SPECIALIST VISITED THE MEDICAL FACILITY ON B AND FOLLOWING INFORMATION WAS PROVIDED BY THE MEDICAL FACILITY EXCESSIVE BALLOONING WAS PERFORMED DURING THE PROCEDURE AND THE TIP OF THE BALLOON WAS THOUGHT TO BE POSITIONED OUTSIDE OF THE PROXIMAL PART OF CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES THEREFORE THE DISSECTION WAS CAUSED BY THE BALLOON INSTEAD OF THE CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES ALTHOUGH THE DISSECTION WAS NOT NOTICED DURING THE PROCEDURE AFTER THE DISSECTION WAS OBSERVED WITH THE FOLLOW UP CT THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY THEREFORE IT IS UNKNOWN WHAT KIND OF TREATMENT WAS PERFORMED FOR THE DISSECTION IT IS ALSO UNKNOWN THE CURRENT PATIENT CONDITION | Hemorrhage | Patient problem code
7696939 | 114277335 | D | ON B THE PATIENT UNDERWENT SURGICAL DEBRANCHING OF THE LEFT SUBCLAVIAN ARTERY THE LEFT COMMON CAROTID ARTERY AND THE INNOMINATE ARTERY PRIOR TO ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM TWO CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES PROXIMAL TGUJB DISTAL TGUJB WERE SUCCESSFULLY DEPLOYED ALTHOUGH A BIRD BEAK CONFIGURATION WAS OBSERVED IN THE PROXIMAL PART OF TGUJB THE PROCEDURE WAS COMPLETED WITHOUT FURTHER TREATMENT SINCE NO ENDOLEAK AND NO VESSEL RUPTUREDISSECTION WAS OBSERVED THE PATIENT TOLERATED THE PROCEDURE ON B A FOLLOWUP COMPUTED TOMOGRAPHY CT REVEALED A RETROGRADE THORACIC AORTIC DISSECTION ORIGINATING AROUND THE PROXIMAL END OF TGUJB DEVICE AND LEADING TO A CORONARY ARTERY AN OPEN SURGERY WAS PERFORMED TO TREAT THE DISSECTION THE CLINICAL SPECIALIST VISITED THE MEDICAL FACILITY ON B AND FOLLOWING INFORMATION WAS PROVIDED BY THE MEDICAL FACILITY EXCESSIVE BALLOONING WAS PERFORMED DURING THE PROCEDURE AND THE TIP OF THE BALLOON WAS THOUGHT TO BE POSITIONED OUTSIDE OF THE PROXIMAL PART OF CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES THEREFORE THE DISSECTION WAS CAUSED BY THE BALLOON INSTEAD OF THE CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES ALTHOUGH THE DISSECTION WAS NOT NOTICED DURING THE PROCEDURE AFTER THE DISSECTION WAS OBSERVED WITH THE FOLLOW UP CT THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY THEREFORE IT IS UNKNOWN WHAT KIND OF TREATMENT WAS PERFORMED FOR THE DISSECTION IT IS ALSO UNKNOWN THE CURRENT PATIENT CONDITION | Aneurysm | Patient problem code
7696939 | 114277335 | D | ON B THE PATIENT UNDERWENT SURGICAL DEBRANCHING OF THE LEFT SUBCLAVIAN ARTERY THE LEFT COMMON CAROTID ARTERY AND THE INNOMINATE ARTERY PRIOR TO ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM TWO CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES PROXIMAL TGUJB DISTAL TGUJB WERE SUCCESSFULLY DEPLOYED ALTHOUGH A BIRD BEAK CONFIGURATION WAS OBSERVED IN THE PROXIMAL PART OF TGUJB THE PROCEDURE WAS COMPLETED WITHOUT FURTHER TREATMENT SINCE NO ENDOLEAK AND NO VESSEL RUPTUREDISSECTION WAS OBSERVED THE PATIENT TOLERATED THE PROCEDURE ON B A FOLLOWUP COMPUTED TOMOGRAPHY CT REVEALED A RETROGRADE THORACIC AORTIC DISSECTION ORIGINATING AROUND THE PROXIMAL END OF TGUJB DEVICE AND LEADING TO A CORONARY ARTERY AN OPEN SURGERY WAS PERFORMED TO TREAT THE DISSECTION THE CLINICAL SPECIALIST VISITED THE MEDICAL FACILITY ON B AND FOLLOWING INFORMATION WAS PROVIDED BY THE MEDICAL FACILITY EXCESSIVE BALLOONING WAS PERFORMED DURING THE PROCEDURE AND THE TIP OF THE BALLOON WAS THOUGHT TO BE POSITIONED OUTSIDE OF THE PROXIMAL PART OF CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES THEREFORE THE DISSECTION WAS CAUSED BY THE BALLOON INSTEAD OF THE CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES ALTHOUGH THE DISSECTION WAS NOT NOTICED DURING THE PROCEDURE AFTER THE DISSECTION WAS OBSERVED WITH THE FOLLOW UP CT THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY THEREFORE IT IS UNKNOWN WHAT KIND OF TREATMENT WAS PERFORMED FOR THE DISSECTION IT IS ALSO UNKNOWN THE CURRENT PATIENT CONDITION | Aortic Dissection | Patient problem code
7696939 | 114277335 | D | ON B THE PATIENT UNDERWENT SURGICAL DEBRANCHING OF THE LEFT SUBCLAVIAN ARTERY THE LEFT COMMON CAROTID ARTERY AND THE INNOMINATE ARTERY PRIOR TO ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM TWO CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES PROXIMAL TGUJB DISTAL TGUJB WERE SUCCESSFULLY DEPLOYED ALTHOUGH A BIRD BEAK CONFIGURATION WAS OBSERVED IN THE PROXIMAL PART OF TGUJB THE PROCEDURE WAS COMPLETED WITHOUT FURTHER TREATMENT SINCE NO ENDOLEAK AND NO VESSEL RUPTUREDISSECTION WAS OBSERVED THE PATIENT TOLERATED THE PROCEDURE ON B A FOLLOWUP COMPUTED TOMOGRAPHY CT REVEALED A RETROGRADE THORACIC AORTIC DISSECTION ORIGINATING AROUND THE PROXIMAL END OF TGUJB DEVICE AND LEADING TO A CORONARY ARTERY AN OPEN SURGERY WAS PERFORMED TO TREAT THE DISSECTION THE CLINICAL SPECIALIST VISITED THE MEDICAL FACILITY ON B AND FOLLOWING INFORMATION WAS PROVIDED BY THE MEDICAL FACILITY EXCESSIVE BALLOONING WAS PERFORMED DURING THE PROCEDURE AND THE TIP OF THE BALLOON WAS THOUGHT TO BE POSITIONED OUTSIDE OF THE PROXIMAL PART OF CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES THEREFORE THE DISSECTION WAS CAUSED BY THE BALLOON INSTEAD OF THE CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES ALTHOUGH THE DISSECTION WAS NOT NOTICED DURING THE PROCEDURE AFTER THE DISSECTION WAS OBSERVED WITH THE FOLLOW UP CT THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY THEREFORE IT IS UNKNOWN WHAT KIND OF TREATMENT WAS PERFORMED FOR THE DISSECTION IT IS ALSO UNKNOWN THE CURRENT PATIENT CONDITION | Dissection | Patient problem code
7696939 | 114277335 | D | ON B THE PATIENT UNDERWENT SURGICAL DEBRANCHING OF THE LEFT SUBCLAVIAN ARTERY THE LEFT COMMON CAROTID ARTERY AND THE INNOMINATE ARTERY PRIOR TO ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM ON B THE PATIENT UNDERWENT ENDOVASCULAR REPAIR OF A THORACIC AORTIC ANEURYSM TWO CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES PROXIMAL TGUJB DISTAL TGUJB WERE SUCCESSFULLY DEPLOYED ALTHOUGH A BIRD BEAK CONFIGURATION WAS OBSERVED IN THE PROXIMAL PART OF TGUJB THE PROCEDURE WAS COMPLETED WITHOUT FURTHER TREATMENT SINCE NO ENDOLEAK AND NO VESSEL RUPTUREDISSECTION WAS OBSERVED THE PATIENT TOLERATED THE PROCEDURE ON B A FOLLOWUP COMPUTED TOMOGRAPHY CT REVEALED A RETROGRADE THORACIC AORTIC DISSECTION ORIGINATING AROUND THE PROXIMAL END OF TGUJB DEVICE AND LEADING TO A CORONARY ARTERY AN OPEN SURGERY WAS PERFORMED TO TREAT THE DISSECTION THE CLINICAL SPECIALIST VISITED THE MEDICAL FACILITY ON B AND FOLLOWING INFORMATION WAS PROVIDED BY THE MEDICAL FACILITY EXCESSIVE BALLOONING WAS PERFORMED DURING THE PROCEDURE AND THE TIP OF THE BALLOON WAS THOUGHT TO BE POSITIONED OUTSIDE OF THE PROXIMAL PART OF CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES THEREFORE THE DISSECTION WAS CAUSED BY THE BALLOON INSTEAD OF THE CONFORMABLE GORE TAG THORACIC ENDOPROSTHESES ALTHOUGH THE DISSECTION WAS NOT NOTICED DURING THE PROCEDURE AFTER THE DISSECTION WAS OBSERVED WITH THE FOLLOW UP CT THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY THEREFORE IT IS UNKNOWN WHAT KIND OF TREATMENT WAS PERFORMED FOR THE DISSECTION IT IS ALSO UNKNOWN THE CURRENT PATIENT CONDITION | Insufficient Information | Device problem code
7696151 | 114278020 | D | IT WAS REPORTED THAT AFTER USE WITH A BD ULTRAFINE INSULIN SYRINGE HAD NEEDLE STICK INJURY TO THE NURSE WHEN TRIED TO RECAP THERE WAS NO REPORT OF MEDICAL INTERVENTION | Physical Resistance / Sticking | Device problem code
7696151 | 114278020 | D | IT WAS REPORTED THAT AFTER USE WITH A BD ULTRAFINE INSULIN SYRINGE HAD NEEDLE STICK INJURY TO THE NURSE WHEN TRIED TO RECAP THERE WAS NO REPORT OF MEDICAL INTERVENTION | No Consequences Or Impact To Patient | Patient problem code
7698899 | 114330600 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA THE PATIENTS POD DOUBLE BEEPED BUT THEN IT DEACTIVATED AND THE PDM STATES THE POD EXPIRED BUT THE PATIENT NEVER GOT ANY ALARM OR ALERT TO LET THE PATIENT KNOW IT WAS SET TO TWELVE HOURS PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEFT THIGH PATIENT BOLUS BEFORE GOING TO THE HOSPITAL THE PATIENT ENDED UP IN THE HOSPITAL ON MONDAY B PATIENT HAD THE STOMACH FLU BUT THE POD FAILURE INCREASED THE SYMPTOMS THIS HAPPENED AFTER THE POD FAILED AND IT SENT THE PATIENT INTO DKA SYMPTOMS REPORTED INCLUDE STOMACH FLU SYMPTOMS PRIOR TO THE EVENT FOR ABOUT A WEEK AND PATIENT WAS NOT ABLE TO EAT OR DRINK MUCH DURING THAT TIME THEN SYMPTOMS HEIGHTEN WITH THE DKA AND PATIENT WAS VOMITING HAD NAUSEA AND CONSTIPATION THE PATIENT WAS TREATED WITH HEAVY FLUID TO TRY AND REVERSE IT AND RAN LABS BECAUSE THE PATIENTS NUMBERS WERE ALARMING CONSIDERING FLUIDS WERE GIVEN FOR FIVE HOURS AND NOTHING WAS REALLY CHANGED PATIENT WAS GIVEN ZOFRAN FOR NAUSEA AND LACATIVES FOR CONSTIPATION AN HOUR BEFORE THE POD DEACTIVATED THE DOCTOR MAY HAVE SUSPENDED INSULIN FOR MINUTES THEN RESUMED IT PATIENT DOES NOT HAVE THE POD ANYMORE SO DEVICE IS NOT AVAILABLE FOR RETURN | Nausea | Patient problem code
7698899 | 114330600 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA THE PATIENTS POD DOUBLE BEEPED BUT THEN IT DEACTIVATED AND THE PDM STATES THE POD EXPIRED BUT THE PATIENT NEVER GOT ANY ALARM OR ALERT TO LET THE PATIENT KNOW IT WAS SET TO TWELVE HOURS PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEFT THIGH PATIENT BOLUS BEFORE GOING TO THE HOSPITAL THE PATIENT ENDED UP IN THE HOSPITAL ON MONDAY B PATIENT HAD THE STOMACH FLU BUT THE POD FAILURE INCREASED THE SYMPTOMS THIS HAPPENED AFTER THE POD FAILED AND IT SENT THE PATIENT INTO DKA SYMPTOMS REPORTED INCLUDE STOMACH FLU SYMPTOMS PRIOR TO THE EVENT FOR ABOUT A WEEK AND PATIENT WAS NOT ABLE TO EAT OR DRINK MUCH DURING THAT TIME THEN SYMPTOMS HEIGHTEN WITH THE DKA AND PATIENT WAS VOMITING HAD NAUSEA AND CONSTIPATION THE PATIENT WAS TREATED WITH HEAVY FLUID TO TRY AND REVERSE IT AND RAN LABS BECAUSE THE PATIENTS NUMBERS WERE ALARMING CONSIDERING FLUIDS WERE GIVEN FOR FIVE HOURS AND NOTHING WAS REALLY CHANGED PATIENT WAS GIVEN ZOFRAN FOR NAUSEA AND LACATIVES FOR CONSTIPATION AN HOUR BEFORE THE POD DEACTIVATED THE DOCTOR MAY HAVE SUSPENDED INSULIN FOR MINUTES THEN RESUMED IT PATIENT DOES NOT HAVE THE POD ANYMORE SO DEVICE IS NOT AVAILABLE FOR RETURN | Diabetic Ketoacidosis | Patient problem code
7698899 | 114330600 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA THE PATIENTS POD DOUBLE BEEPED BUT THEN IT DEACTIVATED AND THE PDM STATES THE POD EXPIRED BUT THE PATIENT NEVER GOT ANY ALARM OR ALERT TO LET THE PATIENT KNOW IT WAS SET TO TWELVE HOURS PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEFT THIGH PATIENT BOLUS BEFORE GOING TO THE HOSPITAL THE PATIENT ENDED UP IN THE HOSPITAL ON MONDAY B PATIENT HAD THE STOMACH FLU BUT THE POD FAILURE INCREASED THE SYMPTOMS THIS HAPPENED AFTER THE POD FAILED AND IT SENT THE PATIENT INTO DKA SYMPTOMS REPORTED INCLUDE STOMACH FLU SYMPTOMS PRIOR TO THE EVENT FOR ABOUT A WEEK AND PATIENT WAS NOT ABLE TO EAT OR DRINK MUCH DURING THAT TIME THEN SYMPTOMS HEIGHTEN WITH THE DKA AND PATIENT WAS VOMITING HAD NAUSEA AND CONSTIPATION THE PATIENT WAS TREATED WITH HEAVY FLUID TO TRY AND REVERSE IT AND RAN LABS BECAUSE THE PATIENTS NUMBERS WERE ALARMING CONSIDERING FLUIDS WERE GIVEN FOR FIVE HOURS AND NOTHING WAS REALLY CHANGED PATIENT WAS GIVEN ZOFRAN FOR NAUSEA AND LACATIVES FOR CONSTIPATION AN HOUR BEFORE THE POD DEACTIVATED THE DOCTOR MAY HAVE SUSPENDED INSULIN FOR MINUTES THEN RESUMED IT PATIENT DOES NOT HAVE THE POD ANYMORE SO DEVICE IS NOT AVAILABLE FOR RETURN | Constipation | Patient problem code
7698899 | 114330600 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA THE PATIENTS POD DOUBLE BEEPED BUT THEN IT DEACTIVATED AND THE PDM STATES THE POD EXPIRED BUT THE PATIENT NEVER GOT ANY ALARM OR ALERT TO LET THE PATIENT KNOW IT WAS SET TO TWELVE HOURS PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEFT THIGH PATIENT BOLUS BEFORE GOING TO THE HOSPITAL THE PATIENT ENDED UP IN THE HOSPITAL ON MONDAY B PATIENT HAD THE STOMACH FLU BUT THE POD FAILURE INCREASED THE SYMPTOMS THIS HAPPENED AFTER THE POD FAILED AND IT SENT THE PATIENT INTO DKA SYMPTOMS REPORTED INCLUDE STOMACH FLU SYMPTOMS PRIOR TO THE EVENT FOR ABOUT A WEEK AND PATIENT WAS NOT ABLE TO EAT OR DRINK MUCH DURING THAT TIME THEN SYMPTOMS HEIGHTEN WITH THE DKA AND PATIENT WAS VOMITING HAD NAUSEA AND CONSTIPATION THE PATIENT WAS TREATED WITH HEAVY FLUID TO TRY AND REVERSE IT AND RAN LABS BECAUSE THE PATIENTS NUMBERS WERE ALARMING CONSIDERING FLUIDS WERE GIVEN FOR FIVE HOURS AND NOTHING WAS REALLY CHANGED PATIENT WAS GIVEN ZOFRAN FOR NAUSEA AND LACATIVES FOR CONSTIPATION AN HOUR BEFORE THE POD DEACTIVATED THE DOCTOR MAY HAVE SUSPENDED INSULIN FOR MINUTES THEN RESUMED IT PATIENT DOES NOT HAVE THE POD ANYMORE SO DEVICE IS NOT AVAILABLE FOR RETURN | Insufficient Information | Device problem code
7698899 | 114330600 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA THE PATIENTS POD DOUBLE BEEPED BUT THEN IT DEACTIVATED AND THE PDM STATES THE POD EXPIRED BUT THE PATIENT NEVER GOT ANY ALARM OR ALERT TO LET THE PATIENT KNOW IT WAS SET TO TWELVE HOURS PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEFT THIGH PATIENT BOLUS BEFORE GOING TO THE HOSPITAL THE PATIENT ENDED UP IN THE HOSPITAL ON MONDAY B PATIENT HAD THE STOMACH FLU BUT THE POD FAILURE INCREASED THE SYMPTOMS THIS HAPPENED AFTER THE POD FAILED AND IT SENT THE PATIENT INTO DKA SYMPTOMS REPORTED INCLUDE STOMACH FLU SYMPTOMS PRIOR TO THE EVENT FOR ABOUT A WEEK AND PATIENT WAS NOT ABLE TO EAT OR DRINK MUCH DURING THAT TIME THEN SYMPTOMS HEIGHTEN WITH THE DKA AND PATIENT WAS VOMITING HAD NAUSEA AND CONSTIPATION THE PATIENT WAS TREATED WITH HEAVY FLUID TO TRY AND REVERSE IT AND RAN LABS BECAUSE THE PATIENTS NUMBERS WERE ALARMING CONSIDERING FLUIDS WERE GIVEN FOR FIVE HOURS AND NOTHING WAS REALLY CHANGED PATIENT WAS GIVEN ZOFRAN FOR NAUSEA AND LACATIVES FOR CONSTIPATION AN HOUR BEFORE THE POD DEACTIVATED THE DOCTOR MAY HAVE SUSPENDED INSULIN FOR MINUTES THEN RESUMED IT PATIENT DOES NOT HAVE THE POD ANYMORE SO DEVICE IS NOT AVAILABLE FOR RETURN | Vomiting | Patient problem code
7692103 | 114333871 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7692103 | 114333871 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | False Alarm | Device problem code
7699356 | 114356226 | D | IT WAS REPORTED BY A CUSTOMER SERVICE MANAGER THAT A PATIENT DEVELOPED AN INFECTION AT THE PODS INSERTION SITE WHILE WEARING THE POD AND WAS IN THE HOSPITAL | Insufficient Information | Device problem code
7699356 | 114356226 | D | IT WAS REPORTED BY A CUSTOMER SERVICE MANAGER THAT A PATIENT DEVELOPED AN INFECTION AT THE PODS INSERTION SITE WHILE WEARING THE POD AND WAS IN THE HOSPITAL | Infection | Patient problem code
7699747 | 114356386 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7699747 | 114356386 | D | MEDTRONIC RECEIVED INFORMATION THAT ONE DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7699523 | 114357862 | D | MEDTRONIC RECEIVED INFORMATION THAT FOUR DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7699523 | 114357862 | D | MEDTRONIC RECEIVED INFORMATION THAT FOUR DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient Problem Code
7702723 | 114434356 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS MOTHER SAID THE CANNULA WAS NOT IN AND IT WAS CRIMPED OVER THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS AS TREATMENT PATIENTS MOTHER APPLIED A NEW POD AND GAVE A MANUAL INJECTION | Positioning Problem | Device problem code
7702723 | 114434356 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS MOTHER SAID THE CANNULA WAS NOT IN AND IT WAS CRIMPED OVER THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS AS TREATMENT PATIENTS MOTHER APPLIED A NEW POD AND GAVE A MANUAL INJECTION | No Known Impact Or Consequence To Patient | Patient Problem Code
7703478 | 114467497 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE ELECTROCARDIOGRAM ECG INDICATED COMPLETE HEART BLOCK CHB SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7703478 | 114467497 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE ELECTROCARDIOGRAM ECG INDICATED COMPLETE HEART BLOCK CHB SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7704981 | 114533849 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Shock | Patient problem code
7704981 | 114533849 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Insufficient Information | Device problem code
7704981 | 114533849 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Burning Sensation | Patient problem code
7704983 | 114535837 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Shock | Patient problem code
7704983 | 114535837 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Insufficient Information | Device problem code
7704983 | 114535837 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED THE PATIENT TRIAL PROCEDURE WAS ABANDONED DURING THE TRIAL IMPLANT THE PATIENT EXPERIENCED A SHOCKING AND BURNING SENSATION IN THEIR LEFT LEG | Burning Sensation | Patient problem code
7704710 | 114540225 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED AN AIR DETECTION GOING OFF WITH NO BUBBLES IN SET THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | Insufficient Information | Device problem code
7704710 | 114540225 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED AN AIR DETECTION GOING OFF WITH NO BUBBLES IN SET THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO PATIENT INVOLVEMENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Patient Involvement | Patient problem code
7705858 | 114547775 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM | No Known Impact Or Consequence To Patient | Patient Problem Code
7705858 | 114547775 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD LONGER THAN HOURS ON THE ARM | Insufficient Information | Device problem code
7701667 | 114547820 | D | VALIANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Dissection | Patient problem code
7701667 | 114547820 | D | VALIANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Hemorrhage | Patient problem code
7701667 | 114547820 | D | VALIANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Insufficient Information | Device problem code
7701679 | 114549103 | D | TALENT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Dissection | Patient problem code
7701679 | 114549103 | D | TALENT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Hemorrhage | Patient problem code
7701679 | 114549103 | D | TALENT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN JANUARY AND JANUARY THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Insufficient Information | Device problem code
7699699 | 114550356 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS AS TREATMENT A NEW POD WAS APPLIED AND BG IMPROVED | No Known Impact Or Consequence To Patient | Patient Problem Code
7699699 | 114550356 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPBUTTOCKS AS TREATMENT A NEW POD WAS APPLIED AND BG IMPROVED | Positioning Problem | Device problem code
7698289 | 114551561 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED BETWEEN AND MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN HYPERGLYCEMIA TREATED BY PATIENT WITH A MANUAL INJECTION OF UNITS OF INSULIN WITH AN INSULIN PEN | Hyperglycemia | Patient problem code
7698289 | 114551561 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED BETWEEN AND MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN HYPERGLYCEMIA TREATED BY PATIENT WITH A MANUAL INJECTION OF UNITS OF INSULIN WITH AN INSULIN PEN | Insufficient Information | Device problem code
7706225 | 114567526 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK WAS REPORTED TWO DAYS FOLLOWING A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7706225 | 114567526 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK WAS REPORTED TWO DAYS FOLLOWING A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7706552 | 114581111 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED B THE PATIENT WAS HOSPITALIZED FOR INCREASING DYSPHAGIA AND DYSARTHRIA SYMPTOMS REPORTEDLY THE PATIENT WAS DISCHARGED WITH AN OUTPATIENT TREATMENT RECOMMENDATION NO FURTHER INFORMATION WAS AVAILABLE AT THIS TIME | Dysphagia | Patient problem code
7706552 | 114581111 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED B THE PATIENT WAS HOSPITALIZED FOR INCREASING DYSPHAGIA AND DYSARTHRIA SYMPTOMS REPORTEDLY THE PATIENT WAS DISCHARGED WITH AN OUTPATIENT TREATMENT RECOMMENDATION NO FURTHER INFORMATION WAS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
7706612 | 114582935 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED POLAND THE PATIENT WAS HOSPITALIZED FOR INCREASING DYSPHAGIA AND DYSARTHRIA SYMPTOMS REPORTEDLY THE PATIENT WAS DISCHARGED WITH AN OUTPATIENT TREATMENT RECOMMENDATION NO FURTHER INFORMATION WAS AVAILABLE AT THIS TIME | Insufficient Information | Device problem code
7706612 | 114582935 | D | DEVICE OF REFERENCE MFR REPORT IT WAS REPORTED POLAND THE PATIENT WAS HOSPITALIZED FOR INCREASING DYSPHAGIA AND DYSARTHRIA SYMPTOMS REPORTEDLY THE PATIENT WAS DISCHARGED WITH AN OUTPATIENT TREATMENT RECOMMENDATION NO FURTHER INFORMATION WAS AVAILABLE AT THIS TIME | Dysphagia | Patient problem code
7706758 | 114588703 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Insufficient Information | Device problem code
7706758 | 114588703 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Pulmonary Edema | Patient problem code
7706753 | 114588723 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Insufficient Information | Device problem code
7706753 | 114588723 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Pulmonary Edema | Patient problem code
7706772 | 114589054 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Pulmonary Edema | Patient problem code
7706772 | 114589054 | D | IT WAS REPORTED THAT A PATIENT WAS HOSPITALIZED DUE TO A PULMONARY EDEMA THAT DEVELOPED POST PROCEDURE THE ISSUE WAS RESOLVED AND PATIENT WAS DISCHARGED | Insufficient Information | Device problem code
7707403 | 114608567 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SIX YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Aneurysm | Patient problem code
7707403 | 114608567 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SIX YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Insufficient Information | Device problem code
7707403 | 114608567 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SIX YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Hemorrhage | Patient problem code
7707429 | 114609286 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SEVEN YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Insufficient Information | Device problem code
7707429 | 114609286 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SEVEN YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Hemorrhage | Patient problem code
7707429 | 114609286 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY SEVEN YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH A TYPE III ENDOLEAK IT IS UNKNOWN IF REINTERVENTION HAS BEEN PLANNED AND THE PATIENTS CURRENT STATUS IS UNKNOWN | Aneurysm | Patient problem code
7698511 | 114614965 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD ON THE ARM THE PATIENT WAS ABLE TO ACTIVATE A NEW POD AND HER BG LEVELS LOWER TO MMOLL MGDL | No Known Impact Or Consequence To Patient | Patient Problem Code
7698511 | 114614965 | D | THE PATIENT REPORTED HER BLOOD GLUCOSE REACHED MMOLL MGDL AFTER WEARING THE POD ON THE ARM THE PATIENT WAS ABLE TO ACTIVATE A NEW POD AND HER BG LEVELS LOWER TO MMOLL MGDL | Insufficient Information | Device problem code
7707959 | 114671927 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED AIR DETECTION ALARM WENT OFF WITH NO BUBBLES IN THE SET THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | No Known Impact Or Consequence To Patient | Patient problem code
7707959 | 114671927 | D | IT WAS REPORTED THAT A SPECTRUM PUMP EXPERIENCED AIR DETECTION ALARM WENT OFF WITH NO BUBBLES IN THE SET THERE WAS NO KNOWN PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT NO ADDITIONAL INFORMATION IS AVAILABLE | Insufficient Information | Device problem code
7697557 | 114679747 | D | ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN PATIENTS IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Hemorrhage | Patient problem code
7697557 | 114679747 | D | ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED IN PATIENTS IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Insufficient Information | Device problem code
7698206 | 114686514 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS IN THE INFUSION SITE ABDOMEN PATIENT HAD MODERATE KETONES AS TREATMENT PATIENT CONTINUED TO ADMINISTER INSULIN REMOVED THE POD AND APPLIED A NEW POD | Insufficient Information | Device problem code
7698206 | 114686514 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD FOR LONGER THAN HOURS IN THE INFUSION SITE ABDOMEN PATIENT HAD MODERATE KETONES AS TREATMENT PATIENT CONTINUED TO ADMINISTER INSULIN REMOVED THE POD AND APPLIED A NEW POD | No Known Impact Or Consequence To Patient | Patient Problem Code
7697564 | 114689851 | D | TALENT STENT GRAFT SYSTEMS WERE USED IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Hemorrhage | Patient problem code
7697564 | 114689851 | D | TALENT STENT GRAFT SYSTEMS WERE USED IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Insufficient Information | Device problem code
7706097 | 114693769 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE RISING BG PATIENT DELIVERED A CORRECTION WHEN CHECKING THE CANNULA PATIENT FOUND INSULIN ON ADHESIVE AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | No Known Impact Or Consequence To Patient | Patient Problem Code
7706097 | 114693769 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE RISING BG PATIENT DELIVERED A CORRECTION WHEN CHECKING THE CANNULA PATIENT FOUND INSULIN ON ADHESIVE AND WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE | Positioning Problem | Device problem code
7698573 | 114695056 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS | No Known Impact Or Consequence To Patient | Patient Problem Code
7698573 | 114695056 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE HIPBUTTOCKS | Insufficient Information | Device problem code
7699878 | 114698024 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED OVER MGDL WHILE WEARING THE POD ON HIS HIPBUTTOCKS FOR BETWEEN AND HOURS TREATMENT INCLUDED REPLACING THE POD AND WORKING OUT THE DEVICE WAS RETURNED FOR EVALUATION | Insufficient Information | Device problem code
7699878 | 114698024 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED OVER MGDL WHILE WEARING THE POD ON HIS HIPBUTTOCKS FOR BETWEEN AND HOURS TREATMENT INCLUDED REPLACING THE POD AND WORKING OUT THE DEVICE WAS RETURNED FOR EVALUATION | No Known Impact Or Consequence To Patient | Patient Problem Code
7706224 | 114702551 | D | THE DEVICES HANDPIECE HAD ALLEGEDLY DISCHARGED ENERGY AFTER THE TRIGGER BUTTON WAS RELEASED | Component or Accessory Incompatibility | Device problem code
7706224 | 114702551 | D | THE DEVICES HANDPIECE HAD ALLEGEDLY DISCHARGED ENERGY AFTER THE TRIGGER BUTTON WAS RELEASED | No Information | Patient problem code
7710339 | 114706081 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7710339 | 114706081 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7701672 | 114708970 | D | ENDURANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN B AND B THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Insufficient Information | Device problem code
7701672 | 114708970 | D | ENDURANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN B AND B THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Dissection | Patient problem code
7701672 | 114708970 | D | ENDURANT STENT GRAFTS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS BETWEEN B AND B THE FOLLOWING EVENTS WERE OBSERVED MALFUNCTION UNKNOWN ENDOLEAKS | Hemorrhage | Patient problem code
7697561 | 114713301 | D | ENDURANT IIS STENT GRAFT SYSTEMS WERE USED IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Insufficient Information | Device problem code
7697561 | 114713301 | D | ENDURANT IIS STENT GRAFT SYSTEMS WERE USED IN AORTIC AND CHIMNEY DEVICE COMBINATIONS IN PATIENTS FOR THE TREATMENT OF PARARENAL AORTIC DISEASE BETWEEN AND THE FOLLOWING ADVERSE EVENTS WERE OBSERVED TYPE IA ENDOLEAKS | Hemorrhage | Patient problem code
7704560 | 114719297 | D | IT WAS REPORTED THAT A RIDGE DEVELOPED ON THE FOLEY CATHETER BALLOON DURING REMOVAL OF THE CATHETER THE PATIENT REPORTEDLY EXPERIENCED PAIN A BURNING SENSATION AND BLEEDING DURING REMOVAL OF THE CATHETER NO MEDICAL INTERVENTION WAS REPORTED | Burning Sensation | Patient problem code
7704560 | 114719297 | D | IT WAS REPORTED THAT A RIDGE DEVELOPED ON THE FOLEY CATHETER BALLOON DURING REMOVAL OF THE CATHETER THE PATIENT REPORTEDLY EXPERIENCED PAIN A BURNING SENSATION AND BLEEDING DURING REMOVAL OF THE CATHETER NO MEDICAL INTERVENTION WAS REPORTED | Pain | Patient problem code
7704560 | 114719297 | D | IT WAS REPORTED THAT A RIDGE DEVELOPED ON THE FOLEY CATHETER BALLOON DURING REMOVAL OF THE CATHETER THE PATIENT REPORTEDLY EXPERIENCED PAIN A BURNING SENSATION AND BLEEDING DURING REMOVAL OF THE CATHETER NO MEDICAL INTERVENTION WAS REPORTED | Insufficient Information | Device problem code
7704560 | 114719297 | D | IT WAS REPORTED THAT A RIDGE DEVELOPED ON THE FOLEY CATHETER BALLOON DURING REMOVAL OF THE CATHETER THE PATIENT REPORTEDLY EXPERIENCED PAIN A BURNING SENSATION AND BLEEDING DURING REMOVAL OF THE CATHETER NO MEDICAL INTERVENTION WAS REPORTED | Hemorrhage | Patient problem code
7710719 | 114745504 | D | IT WAS REPORTED THAT THE PUMP DECLARED A TEMPERATURE ALARM THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER WAS ABLE TO CLEAR THE ALARM AND CONTINUED USING THE PUMP FOR INSULIN THERAPY | False Alarm | Device problem code
7710719 | 114745504 | D | IT WAS REPORTED THAT THE PUMP DECLARED A TEMPERATURE ALARM THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER WAS ABLE TO CLEAR THE ALARM AND CONTINUED USING THE PUMP FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7711875 | 114771877 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR BETWEEN AND HOURS TREATMENT INCLUDED DOING A MANUAL INJECTION USING A PEN THE DEVICE WAS RETURNED FOR EVALUATION | Insufficient Information | Device problem code
7711875 | 114771877 | D | THE PATIENT REPORTED THAT HER BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ABDOMEN FOR BETWEEN AND HOURS TREATMENT INCLUDED DOING A MANUAL INJECTION USING A PEN THE DEVICE WAS RETURNED FOR EVALUATION | No Known Impact Or Consequence To Patient | Patient Problem Code
7712483 | 114799070 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED SIX DAYS LATER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7712483 | 114799070 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED SIX DAYS LATER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7713724 | 114832361 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED TWO DAYS LATER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7713724 | 114832361 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED TWO DAYS LATER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7713947 | 114834248 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7713947 | 114834248 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE COMPLETE HEART BLOCK CHB WAS REPORTED AND A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7714782 | 114859945 | D | IT WAS REPORTED THAT DURING DEVICE UPGRADE PROCEDURE THE PATIENT EXPERIENCED A CARDIOVASCULAR COLLAPSE WHEN THE PHYSICIAN WAS IMPLANTING THE RIGHT VENTRICULAR LEAD EMERGENT CHEST TUBE WAS PLACED THE LEAD WAS EXTRACTED THE PATIENT WAS STABLE AFTER THE PROCEDURE | Collapse | Patient problem code
7714782 | 114859945 | D | IT WAS REPORTED THAT DURING DEVICE UPGRADE PROCEDURE THE PATIENT EXPERIENCED A CARDIOVASCULAR COLLAPSE WHEN THE PHYSICIAN WAS IMPLANTING THE RIGHT VENTRICULAR LEAD EMERGENT CHEST TUBE WAS PLACED THE LEAD WAS EXTRACTED THE PATIENT WAS STABLE AFTER THE PROCEDURE | Insufficient Information | Device problem code
7701303 | 114866176 | D | NEW INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS HOSPITALIZED AND THE DEVICE WAS REPROGRAMMED THE PATIENTS CONDITION WAS GOOD AND THEY WERE DISCHARGED FROM THE HOSPITAL | No Consequences Or Impact To Patient | Patient problem code
7701303 | 114866176 | D | NEW INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS HOSPITALIZED AND THE DEVICE WAS REPROGRAMMED THE PATIENTS CONDITION WAS GOOD AND THEY WERE DISCHARGED FROM THE HOSPITAL | Insufficient Information | Device problem code
7715776 | 114885452 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RASH UNDER THE THERAPY PADS AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS GIVEN A CREAM FROM HER DOCTOR THE PATIENTS IRRITATION IMPROVED | Irritation | Patient problem code
7715776 | 114885452 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RASH UNDER THE THERAPY PADS AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS GIVEN A CREAM FROM HER DOCTOR THE PATIENTS IRRITATION IMPROVED | Insufficient Information | Device problem code
7715776 | 114885452 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A RASH UNDER THE THERAPY PADS AND THAT IT HAD GOTTEN WORSE AFTER WEARING THE LIFEVEST THE PATIENT WAS GIVEN A CREAM FROM HER DOCTOR THE PATIENTS IRRITATION IMPROVED | Rash | Patient problem code
7717260 | 114930270 | D | OUS MDR AN EXIT BLOCK OF THE ATRIAL LEAD WAS REPORTED DAY AFTER THE IMPLANTATION THE PATIENT WAS HOSPITALIZED FOR ATRIAL LEAD EXTRACTION AND IMPLANTATION OF A NEW LEAD | Exit Block | Patient problem code
7717260 | 114930270 | D | OUS MDR AN EXIT BLOCK OF THE ATRIAL LEAD WAS REPORTED DAY AFTER THE IMPLANTATION THE PATIENT WAS HOSPITALIZED FOR ATRIAL LEAD EXTRACTION AND IMPLANTATION OF A NEW LEAD | Insufficient Information | Device problem code
7717282 | 114932251 | D | CUSTOMER CALLED AND REPORTED THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS HIGH DUE TO BEING OFF THE PUMP THE CUSTOMER EXPERIENCED HIGH SYMPTOMS SUCH AS NOT FEELING WELL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE CUSTOMER ALSO HAD A LOW THAT REQUIRED EMERGENCY MEDICAL ASSISTANCE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Hyperglycemia | Patient Problem Code
7717282 | 114932251 | D | CUSTOMER CALLED AND REPORTED THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE ON B WITH BLOOD GLUCOSE OF MGDL AT THE TIME OF THE INCIDENT THE CUSTOMER WAS AT MGDL AT THE TIME OF THE CALL THE CUSTOMER WAS HIGH DUE TO BEING OFF THE PUMP THE CUSTOMER EXPERIENCED HIGH SYMPTOMS SUCH AS NOT FEELING WELL THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP DURING THE INCIDENT WITHIN HOURS TROUBLESHOOTING WAS NOT COMPLETED AS THE CUSTOMER DECLINED THE CUSTOMER ALSO HAD A LOW THAT REQUIRED EMERGENCY MEDICAL ASSISTANCE THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS | Insufficient Information | Device problem code
7716956 | 114939333 | D | AS REPORTED AFTER SHUTTLING DOWN BACK A MYNXGRIP VASCULAR CLOSURE DEVICE VCD THE WHITE ADVANCER TUBE WAS NOT EXPOSED TO ADVANCE THE SEALANT MANUAL COMPRESSION WAS APPLIED FOR AN UNKNOWN DURATION TO ACHIEVE HEMOSTASIS THE PATIENTS HOSPITALIZATION WAS NOT EXTENDED | Insufficient Information | Device problem code
7716956 | 114939333 | D | AS REPORTED AFTER SHUTTLING DOWN BACK A MYNXGRIP VASCULAR CLOSURE DEVICE VCD THE WHITE ADVANCER TUBE WAS NOT EXPOSED TO ADVANCE THE SEALANT MANUAL COMPRESSION WAS APPLIED FOR AN UNKNOWN DURATION TO ACHIEVE HEMOSTASIS THE PATIENTS HOSPITALIZATION WAS NOT EXTENDED | Hemostasis | Patient problem code
7713298 | 114945397 | D | ENDURANT SENT GRAFT SYSTEMS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT FOR ABDOMINAL AORTIC ANEURYSM AAA ON DATES BETWEEN JANUARY AND DECEMBER THE FOLLOWING ADVERSE EVENTS WERE REPORTED TYPE IA ENDOLEAKS AND ANEURYSM ENLARGEMENT | Aneurysm | Patient problem code
7713298 | 114945397 | D | ENDURANT SENT GRAFT SYSTEMS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT FOR ABDOMINAL AORTIC ANEURYSM AAA ON DATES BETWEEN JANUARY AND DECEMBER THE FOLLOWING ADVERSE EVENTS WERE REPORTED TYPE IA ENDOLEAKS AND ANEURYSM ENLARGEMENT | Insufficient Information | Device problem code
7713298 | 114945397 | D | ENDURANT SENT GRAFT SYSTEMS WERE IMPLANTED INTO PATIENTS FOR THE ENDOVASCULAR TREATMENT FOR ABDOMINAL AORTIC ANEURYSM AAA ON DATES BETWEEN JANUARY AND DECEMBER THE FOLLOWING ADVERSE EVENTS WERE REPORTED TYPE IA ENDOLEAKS AND ANEURYSM ENLARGEMENT | Hemorrhage | Patient problem code
7718835 | 114987284 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY FIVE YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH AN ENDOLEAK AT AN INDETERMINATE ORIGIN PER CT IT IS UNKNOWN IF THE PHYSICIAN PLANS TO TREAT THE PATIENT | Aneurysm | Patient problem code
7718835 | 114987284 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY FIVE YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH AN ENDOLEAK AT AN INDETERMINATE ORIGIN PER CT IT IS UNKNOWN IF THE PHYSICIAN PLANS TO TREAT THE PATIENT | Hemorrhage | Patient problem code
7718835 | 114987284 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY FIVE YEARS POST INITIAL PROCEDURE THE PATIENT PRESENTED WITH AN ENDOLEAK AT AN INDETERMINATE ORIGIN PER CT IT IS UNKNOWN IF THE PHYSICIAN PLANS TO TREAT THE PATIENT | Insufficient Information | Device problem code
7718937 | 114989373 | D | THE PATIENT WAS INITIALLY TREATED WITH THE OVATION ABDOMINAL STENT GRAFT SYSTEM POST INITIAL IMPLANT PROCEDURE CT CONFIRMED TYPE IA AND IIIB ENDOLEAKS THERE IS NO REPORT OF PATIENT SEQUELAE | Insufficient Information | Device problem code
7718937 | 114989373 | D | THE PATIENT WAS INITIALLY TREATED WITH THE OVATION ABDOMINAL STENT GRAFT SYSTEM POST INITIAL IMPLANT PROCEDURE CT CONFIRMED TYPE IA AND IIIB ENDOLEAKS THERE IS NO REPORT OF PATIENT SEQUELAE | Hemorrhage | Patient problem code
7720434 | 115057809 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE AN ELECTROCARDIOGRAM ECG INDICATED A COMPLETE HEART BLOCK CHB SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7720434 | 115057809 | D | MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE AN ELECTROCARDIOGRAM ECG INDICATED A COMPLETE HEART BLOCK CHB SUBSEQUENTLY A PERMANENT PACEMAKER WAS IMPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7722441 | 115101542 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS MGDL AND A BOLUS WAS DELIVERED CUSTOMER HAD RECENTLY CONSUMED A MEAL AND WAS THOUGHT TO BE THE CAUSE OF ELEVATED BG CUSTOMER SYSTEM CHECK WAS PERFORMED WITH TANDEM TECHNICAL SUPPORT AND NO DEVICE ISSUE WAS IDENTIFIED BG LOWERED TO MGDL AT TIME OF REPORT RECOMMENDATION WAS MADE FOR CUSTOMER TO CONSULT WITH A HEALTHCARE PROVIDER | Adverse Event Without Identified Device or Use Problem | Device problem code
7722441 | 115101542 | D | IT WAS REPORTED THAT THE CUSTOMERS BLOOD GLUCOSE BG WAS MGDL AND A BOLUS WAS DELIVERED CUSTOMER HAD RECENTLY CONSUMED A MEAL AND WAS THOUGHT TO BE THE CAUSE OF ELEVATED BG CUSTOMER SYSTEM CHECK WAS PERFORMED WITH TANDEM TECHNICAL SUPPORT AND NO DEVICE ISSUE WAS IDENTIFIED BG LOWERED TO MGDL AT TIME OF REPORT RECOMMENDATION WAS MADE FOR CUSTOMER TO CONSULT WITH A HEALTHCARE PROVIDER | No Known Impact Or Consequence To Patient | Patient Problem Code
7671542 | 115111549 | D | UPON SECONDARY REVIEW IT WAS DETERMINED THAT THE REPORTED DEVICE HAD NO INDICATION OF A PRODUCT MALFUNCTION | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7671542 | 115111549 | D | UPON SECONDARY REVIEW IT WAS DETERMINED THAT THE REPORTED DEVICE HAD NO INDICATION OF A PRODUCT MALFUNCTION | No Information | Patient Problem Code
7723020 | 115121717 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7723020 | 115121717 | D | MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A PERMANENT PACEMAKER WAS IMPLANTED DUE TO COMPLETE HEART BLOCK CHB NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7715600 | 115228931 | D | NEW INFORMATION REPORTED ON NOTED THAT THE PHYSICIAN ELECTED NOT TO PERFORM PROGRAMMING CHANGES A SUCCESSFUL DEFIBRILLATION THRESHOLD TEST WAS CONDUCTED AND THE PATIENT WAS DISCHARGED THE PATIENT WAS ASYMPTOMATIC | No Consequences Or Impact To Patient | Patient problem code
7715600 | 115228931 | D | NEW INFORMATION REPORTED ON NOTED THAT THE PHYSICIAN ELECTED NOT TO PERFORM PROGRAMMING CHANGES A SUCCESSFUL DEFIBRILLATION THRESHOLD TEST WAS CONDUCTED AND THE PATIENT WAS DISCHARGED THE PATIENT WAS ASYMPTOMATIC | Insufficient Information | Device problem code
7729535 | 115356697 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | Insufficient Information | Device problem code
7729535 | 115356697 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE ABDOMEN | No Known Impact Or Consequence To Patient | Patient Problem Code
7727065 | 115357748 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS SHE STATED THAT THE CANNULA DID NOT LOOK NORMAL WAS NOT STRAIGHT WHEN THE POD WAS REMOVED TREATMENT INCLUDED INJECTIONS AND CHANGING THE POD | No Known Impact Or Consequence To Patient | Patient Problem Code
7727065 | 115357748 | D | THE MOTHER REPORTED THAT THE PATIENTS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HER ARM FOR BETWEEN AND HOURS SHE STATED THAT THE CANNULA DID NOT LOOK NORMAL WAS NOT STRAIGHT WHEN THE POD WAS REMOVED TREATMENT INCLUDED INJECTIONS AND CHANGING THE POD | Insufficient Information | Device problem code
7729041 | 115357916 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BG LEVELS DECREASED TO MMOLL MGDL | Hyperglycemia | Patient problem code
7729041 | 115357916 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA AND THE PATIENTS BG LEVELS DECREASED TO MMOLL MGDL | Insufficient Information | Device problem code
7731229 | 115371327 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY ONE YEAR POST INITIAL PROCEDURE A TYPE III ENDOLEAK WAS DETECTED THE PHYSICIAN ELECTED TO TREAT THE PATIENT BY IMPLANTING A CTAG DEVICE NONENDOLOGIX ON B | Insufficient Information | Device problem code
7731229 | 115371327 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY ONE YEAR POST INITIAL PROCEDURE A TYPE III ENDOLEAK WAS DETECTED THE PHYSICIAN ELECTED TO TREAT THE PATIENT BY IMPLANTING A CTAG DEVICE NONENDOLOGIX ON B | Aneurysm | Patient problem code
7731229 | 115371327 | D | THE PATIENT WAS INITIALLY TREATED WITH THE AFX ABDOMINAL AORTIC ANEURYSM STENT APPROXIMATELY ONE YEAR POST INITIAL PROCEDURE A TYPE III ENDOLEAK WAS DETECTED THE PHYSICIAN ELECTED TO TREAT THE PATIENT BY IMPLANTING A CTAG DEVICE NONENDOLOGIX ON B | Hemorrhage | Patient problem code
7731353 | 115376142 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG FOR TREATMENT PATIENT GAVE HIMSELF A MANUAL SHOT OF UNITS OF INSULIN TO BRING BG LEVELS DOWN | Positioning Problem | Device problem code
7731353 | 115376142 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG FOR TREATMENT PATIENT GAVE HIMSELF A MANUAL SHOT OF UNITS OF INSULIN TO BRING BG LEVELS DOWN | No Known Impact Or Consequence To Patient | Patient Problem Code
7731553 | 115395831 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA AND DEHYDRATION PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEG THE PATIENT WAS TREATED WITH AN INSULIN DRIP THEN PATIENT WAS TAKEN OFF INSULIN DRIP AND PLACED A POD ON BUT BG LEVELS BEGAN TO RISE PATIENTS MEDICATION INCLUDES IMLODAPEN AND LEXAPRO THE PATIENT WAS IN THE INTENSIVE CARE UNIT ICU AND WAS BEING TREATED ON A SLIDING SCALE OF INSULIN PATIENT STATED INSULIN COULD HAVE BEEN SUSPENDED BUT DID NOT KNOW | Diabetic Ketoacidosis | Patient problem code
7731553 | 115395831 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA AND DEHYDRATION PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEG THE PATIENT WAS TREATED WITH AN INSULIN DRIP THEN PATIENT WAS TAKEN OFF INSULIN DRIP AND PLACED A POD ON BUT BG LEVELS BEGAN TO RISE PATIENTS MEDICATION INCLUDES IMLODAPEN AND LEXAPRO THE PATIENT WAS IN THE INTENSIVE CARE UNIT ICU AND WAS BEING TREATED ON A SLIDING SCALE OF INSULIN PATIENT STATED INSULIN COULD HAVE BEEN SUSPENDED BUT DID NOT KNOW | Dehydrated | Patient problem code
7731553 | 115395831 | D | IT WAS REPORTED THAT A PATIENT HAD BEEN HOSPITALIZED WITH DIABETIC KETOACIDOSIS DKA AND DEHYDRATION PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE INFUSION SITE LEG THE PATIENT WAS TREATED WITH AN INSULIN DRIP THEN PATIENT WAS TAKEN OFF INSULIN DRIP AND PLACED A POD ON BUT BG LEVELS BEGAN TO RISE PATIENTS MEDICATION INCLUDES IMLODAPEN AND LEXAPRO THE PATIENT WAS IN THE INTENSIVE CARE UNIT ICU AND WAS BEING TREATED ON A SLIDING SCALE OF INSULIN PATIENT STATED INSULIN COULD HAVE BEEN SUSPENDED BUT DID NOT KNOW | Insufficient Information | Device problem code
7731653 | 115481868 | D | IT WAS REPORTED THAT THE CUSTOMER MISCALCULATED A FOOD BOLUS AND INADVERTENTLY ENTERED MORE GRAMS THAN WHAT SHE INTENDED TO ENTER SUBSEQUENTLY CUSTOMER OVERBOLUSED BLOOD GLUCOSE WAS MGDL CANDYJUICE WERE CONSUMED TO ADDRESS BG TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT FOUND NO DEVICE ISSUE UPON FOLLOWUP BG WAS MGDL | No Known Impact Or Consequence To Patient | Patient Problem Code
7731653 | 115481868 | D | IT WAS REPORTED THAT THE CUSTOMER MISCALCULATED A FOOD BOLUS AND INADVERTENTLY ENTERED MORE GRAMS THAN WHAT SHE INTENDED TO ENTER SUBSEQUENTLY CUSTOMER OVERBOLUSED BLOOD GLUCOSE WAS MGDL CANDYJUICE WERE CONSUMED TO ADDRESS BG TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT FOUND NO DEVICE ISSUE UPON FOLLOWUP BG WAS MGDL | Adverse Event Without Identified Device or Use Problem | Device problem code
7722986 | 115481962 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7722986 | 115481962 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | False Alarm | Device problem code
7733237 | 115491002 | D | USER HAD DIFFICULTY UNSCREWING PEN NEEDLE FROM HUMALOG RESULTED IN AN ACCIDENTAL NEEDLE STICK | Needle Stick/Puncture | Patient problem code
7733237 | 115491002 | D | USER HAD DIFFICULTY UNSCREWING PEN NEEDLE FROM HUMALOG RESULTED IN AN ACCIDENTAL NEEDLE STICK | Insufficient Information | Device problem code
7735505 | 115568151 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
7735505 | 115568151 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Positioning Problem | Device problem code
7735516 | 115568205 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA BLOOD WAS NOTED IN THE CANNULA UPON REMOVAL | Hyperglycemia | Patient problem code
7735516 | 115568205 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS REACHED MMOLL MGDL WHILE WEARING THE POD ON THE ABDOMEN THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA BLOOD WAS NOTED IN THE CANNULA UPON REMOVAL | Positioning Problem | Device problem code
7735519 | 115569501 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA BLOOD WAS NOTED IN THE CANNULA UPON REMOVAL | Positioning Problem | Device problem code
7735519 | 115569501 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA BLOOD WAS NOTED IN THE CANNULA UPON REMOVAL | Hyperglycemia | Patient problem code
7735533 | 115570093 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Positioning Problem | Device problem code
7735533 | 115570093 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD THE PATIENT WAS UNSURE THE CANNULA WAS IN AND A NEW POD WAS APPLIED TO TREAT THE HYPERGLYCEMIA | Hyperglycemia | Patient problem code
7737483 | 115646681 | D | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BURNING SENSATION ALL OVER INSIDE THE CHEEKS WHILE HE WAS WEARING COMFORT SPLINT DEVICE | Burning Sensation | Patient problem code
7737483 | 115646681 | D | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BURNING SENSATION ALL OVER INSIDE THE CHEEKS WHILE HE WAS WEARING COMFORT SPLINT DEVICE | Insufficient Information | Device problem code
7740960 | 115771406 | D | INFORMATION WAS RECEIVED THAT A SMITHS MEDICAL CATHETER IS DIFFICULT TO PERFORM VENIPUNCTURE WITH | Insufficient Information | Device problem code
7740960 | 115771406 | D | INFORMATION WAS RECEIVED THAT A SMITHS MEDICAL CATHETER IS DIFFICULT TO PERFORM VENIPUNCTURE WITH | Venipuncture | Patient problem code
7651114 | 115845142 | N | DATE SENT TO FDA PATIENT CODE SURGICAL INTERVENTION MESH MIGRATION ADDITIONAL NARRATIVE IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY ON B DUE TO MESH MIGRATION IT WAS ALSO REPORTED THAT FOLLOWING THE PROCEDURE THE PATIENT EXPERIENCED PAIN AND ADHESIONS NO ADDITIONAL INFORMATION WAS PROVIDED | Migration or Expulsion of Device | Device problem code
7651114 | 115845142 | N | DATE SENT TO FDA PATIENT CODE SURGICAL INTERVENTION MESH MIGRATION ADDITIONAL NARRATIVE IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY ON B DUE TO MESH MIGRATION IT WAS ALSO REPORTED THAT FOLLOWING THE PROCEDURE THE PATIENT EXPERIENCED PAIN AND ADHESIONS NO ADDITIONAL INFORMATION WAS PROVIDED | Device Migration | Medical Evaluation result code
7695567 | 115848077 | N | DATE SENT TO FDA ADDITIONAL NARRATIVE IT WAS REPORTED THAT THE PATIENT EXPERIENCED ABDOMINAL PAIN NAUSEA AND OBSTRUCTION FOLLOWING THE PROCEDURE IT WAS REPORTED THAT THE PATIENT UNDERWENT REMOVAL OF MESH ON B NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7746010 | 115895324 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS B ALLEGING THAT HER ONETOUCH ULTRA METER READ INACCURATELY HIGH COMPARED TO FEELINGSNORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT DETAILED THAT THE ALLEGED ISSUE BEGAN ON B AROUND PM THE PATIENT REPORTED THAT SHE OBTAINED AN ALLEGED READING ON THE SUBJECT METER OF MMOLL ON THE SUBJECT METER METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES WITH INSULIN SELFADJUSTER AND REPORTED THAT ON B PM SHE TOOK UNITS OF NOVORAPID INSULIN IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED PRODUCT ISSUE BEGAN SHE BECAME DISORIENTATED AND WAS TAKEN TO HOSPITAL BY HER HUSBAND AT AROUND PM THE PATIENT STATED THAT WHILST IN HOSPITAL SHE WAS MONITORED BUT DID NOT RECEIVE TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE THE TEST STRIPS WERE STORED CORRECTLY AND HAD NOT EXPIRED THE PATIENT DID NOT HAVE CONTROL SOLUTION TO PERFORM A TEST REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
7746010 | 115895324 | D | ON B THE LAY USERPATIENT CONTACTED LIFESCAN LFS B ALLEGING THAT HER ONETOUCH ULTRA METER READ INACCURATELY HIGH COMPARED TO FEELINGSNORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED ON CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT DETAILED THAT THE ALLEGED ISSUE BEGAN ON B AROUND PM THE PATIENT REPORTED THAT SHE OBTAINED AN ALLEGED READING ON THE SUBJECT METER OF MMOLL ON THE SUBJECT METER METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES WITH INSULIN SELFADJUSTER AND REPORTED THAT ON B PM SHE TOOK UNITS OF NOVORAPID INSULIN IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT REPORTED THAT MINUTES AFTER THE ALLEGED PRODUCT ISSUE BEGAN SHE BECAME DISORIENTATED AND WAS TAKEN TO HOSPITAL BY HER HUSBAND AT AROUND PM THE PATIENT STATED THAT WHILST IN HOSPITAL SHE WAS MONITORED BUT DID NOT RECEIVE TREATMENT DURING TROUBLESHOOTING THE CCA NOTED THAT THE METER WAS SET TO THE CORRECT UNIT OF MEASURE THE TEST STRIPS WERE STORED CORRECTLY AND HAD NOT EXPIRED THE PATIENT DID NOT HAVE CONTROL SOLUTION TO PERFORM A TEST REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Insufficient Information | Device problem code
7722210 | 115961999 | N | THIS REPORT IS BEING SUBMITTED AS FOLLOW UP NO TO PROVIDE ADDITIONAL INVESTIGATION RESULTS BASED ON THE COMPLAINT DESCRIPTION THE PHYSICIAN USED A FR PROCEDURAL SHEATH WAS USED THIS IS AGAINST THE INDICATION OF THE DEVICE WHICH LIKELY RESULTED IN THE BLEEDING THAT WAS OBSERVED THE IFU STATES THE ANGIOSEAL DEVICE IS INDICATED FOR USE IN CLOSING AND REDUCING TIME TO HEMOSTASIS AT THE FEMORAL ARTERIAL PUNCTURE SITE IN PATIENTS WHO HAVE UNDERGONE OR INTERVENTIONAL PROCEDURES USING AN FRENCH OR SMALLER PROCEDURAL SHEATH FOR THE F ANGIOSEAL DEVICE | No Device Problem Found | Medical Evaluation result code
7678275 | 115966789 | N | PUMP WAS RECEIVED WITH NO BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH INSPECTED LCD CONNECTOR AND NO UNLOCKED CONNECTOR NOTED | Component or Accessory Incompatibility | Device problem code
7678275 | 115966789 | N | PUMP WAS RECEIVED WITH NO BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH INSPECTED LCD CONNECTOR AND NO UNLOCKED CONNECTOR NOTED | Interoperability Problem Identified | Medical Evaluation result code
7664818 | 115977018 | N | CORRECTION FOR INFORMATION MISSING FROM INITIAL MDR | Component Missing | Device problem code
7664818 | 115977018 | N | CORRECTION FOR INFORMATION MISSING FROM INITIAL MDR | Missing or Inadequate Safety Measures | Medical Evaluation result code
7702006 | 115980475 | N | THE CUSTOMERS METER WAS RETURNED FOR INVESTIGATION THE RETURNED METER AND MASTER LOT STRIPS WERE TESTED IN COMPARISON TO A RETENTION METER AND MASTER LOT STRIPS HUMAN BLOOD SAMPLES FROM WARFARIN DONORS WERE USED DONOR INR INR DONOR INR INR DONOR HCT DONOR HCT TESTING RESULTS DONOR RETENTION METER WITH MASTER LOT STRIPS INR CUSTOMER METER WITH MASTER LOT STRIPS INR DONOR RETENTION METER WITH MASTER LOT STRIPS INR CUSTOMER METER WITH MASTER LOT STRIPS INR ALL INR VALUES WERE WITHIN THE SPECIFIED MAXIMUM DIFFERENCE BETWEEN MEASUREMENTS NO ERROR MESSAGES OCCURRED THE RETURNED AND THE RETENTION MATERIAL MEET THE SPECIFICATIONS | No Device Problem Found | Medical Evaluation result code
7671505 | 116111870 | N | THE KEYBOARD TOUCH PAD WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT ALL KEYS AND THE MOUSE PAD WERE FUNCTIONING NORMALLY EXCEPT THE LEFT CLICK BUTTON UNDER THE MOUSE PAD WAS INTERMITTENT KEYBOARD WAS MISSING TWO BUTTONS ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Component Missing | Device problem code
7671505 | 116111870 | N | THE KEYBOARD TOUCH PAD WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT ALL KEYS AND THE MOUSE PAD WERE FUNCTIONING NORMALLY EXCEPT THE LEFT CLICK BUTTON UNDER THE MOUSE PAD WAS INTERMITTENT KEYBOARD WAS MISSING TWO BUTTONS ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Electrical /Electronic Property Problem | Device problem code
7671505 | 116111870 | N | THE KEYBOARD TOUCH PAD WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT ALL KEYS AND THE MOUSE PAD WERE FUNCTIONING NORMALLY EXCEPT THE LEFT CLICK BUTTON UNDER THE MOUSE PAD WAS INTERMITTENT KEYBOARD WAS MISSING TWO BUTTONS ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Noise, Audible | Device problem code
7671505 | 116111870 | N | THE KEYBOARD TOUCH PAD WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT ALL KEYS AND THE MOUSE PAD WERE FUNCTIONING NORMALLY EXCEPT THE LEFT CLICK BUTTON UNDER THE MOUSE PAD WAS INTERMITTENT KEYBOARD WAS MISSING TWO BUTTONS ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Missing or Inadequate Safety Measures | Medical Evaluation result code
7669544 | 116126077 | N | THE DEVICE WAS RECEIVED IN NORMAL WORKING CONDITIONS WITH BATTERY NEAR BEGINNING OF LIFE BOL LEVEL ELECTRICAL AND MECHANICAL ANALYSIS WAS PERFORMED INCLUDING SENSITIVITY AND CAPTURE TESTING NO ANOMALIES WERE DETECTED VISUAL INSPECTION OF THE HEADER AND CONNECTORS REVEALED NO FOREIGN MATERIAL OR CONTAMINANTS THAT COULD CONTRIBUTE TO THE REPORTED COMPLAINT A BATTERY ASSESSMENT WAS PERFORMED AND THE BATTERY WAS WITHIN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE LONGEVITY REMAINING THE REPORTED COMPLAINTS OF A SENSING AND CAPTURE ANOMALY WERE NOT CONFIRMED | Device Contamination with Chemical or Other Material | Device problem code
7669544 | 116126077 | N | THE DEVICE WAS RECEIVED IN NORMAL WORKING CONDITIONS WITH BATTERY NEAR BEGINNING OF LIFE BOL LEVEL ELECTRICAL AND MECHANICAL ANALYSIS WAS PERFORMED INCLUDING SENSITIVITY AND CAPTURE TESTING NO ANOMALIES WERE DETECTED VISUAL INSPECTION OF THE HEADER AND CONNECTORS REVEALED NO FOREIGN MATERIAL OR CONTAMINANTS THAT COULD CONTRIBUTE TO THE REPORTED COMPLAINT A BATTERY ASSESSMENT WAS PERFORMED AND THE BATTERY WAS WITHIN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE LONGEVITY REMAINING THE REPORTED COMPLAINTS OF A SENSING AND CAPTURE ANOMALY WERE NOT CONFIRMED | Capturing Problem | Device problem code
7669544 | 116126077 | N | THE DEVICE WAS RECEIVED IN NORMAL WORKING CONDITIONS WITH BATTERY NEAR BEGINNING OF LIFE BOL LEVEL ELECTRICAL AND MECHANICAL ANALYSIS WAS PERFORMED INCLUDING SENSITIVITY AND CAPTURE TESTING NO ANOMALIES WERE DETECTED VISUAL INSPECTION OF THE HEADER AND CONNECTORS REVEALED NO FOREIGN MATERIAL OR CONTAMINANTS THAT COULD CONTRIBUTE TO THE REPORTED COMPLAINT A BATTERY ASSESSMENT WAS PERFORMED AND THE BATTERY WAS WITHIN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE LONGEVITY REMAINING THE REPORTED COMPLAINTS OF A SENSING AND CAPTURE ANOMALY WERE NOT CONFIRMED | Contamination of Environment by Device | Medical Evaluation result code
7753985 | 116203547 | N | A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICES COULD NOT BE PERFORMED AS ITEM AND LOTSERIAL NUMBERS WERE NOT AVAILABLE PER THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE ADVERSE EVENTS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK | Insufficient Information | Device problem code
7753985 | 116203547 | N | A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICES COULD NOT BE PERFORMED AS ITEM AND LOTSERIAL NUMBERS WERE NOT AVAILABLE PER THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE ADVERSE EVENTS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK | No Device Problem Found | Medical Evaluation result code
7755151 | 116215654 | N | ACCORDING TO THE IFU ADVERSE EVENTS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK AND ANEURYSM ENLARGEMENT USERS ARE MADE AWARE OF THE RISKS ASSOCIATED WITH TYPE II ENDOLEAKS IN THE IFU AND ARE INSTRUCTED TO CONSIDER THE RISKS AND BENEFITS DISCUSSED IN THE IFU FOR EACH PATIENT BEFORE USING THE DEVICES | Insufficient Information | Device problem code
7723192 | 116224530 | N | ONE HEARTSTART MRX LITHIUM ION BATTERY PACK WAS RETURNED FOR FAILURE ANALYSIS THE REPORTED PROBLEM WAS VERIFIEDDUPLICATED IN THE LAB AND TRACED TO AN EXTERNAL POWER MODULE LOCKOUT AC LOCKOUT | No Device Problem Found | Medical Evaluation result code
7672179 | 116245889 | N | THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
7757300 | 116248865 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Insufficient Information | Device problem code
7757300 | 116248865 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Fall | Patient problem code
7731327 | 116258337 | N | THE PUMP WAS RETURNED TO ANIMAS AND INVESTIGATED ON JUL WITH THE FOLLOWING FINDINGS THE PISTON ROD WOULD NOT RETRACT UPON INVESTIGATION OF THE INTERNAL COMPONENTS A FOREIGN OBJECT WAS FOUND IN THE CARTRIDGE COMPARTMENT ONCE THE OBJECT WAS REMOVED THE PISTON ROD FUNCTIONED CORRECTLY | Retraction Problem | Device problem code
7731327 | 116258337 | N | THE PUMP WAS RETURNED TO ANIMAS AND INVESTIGATED ON JUL WITH THE FOLLOWING FINDINGS THE PISTON ROD WOULD NOT RETRACT UPON INVESTIGATION OF THE INTERNAL COMPONENTS A FOREIGN OBJECT WAS FOUND IN THE CARTRIDGE COMPARTMENT ONCE THE OBJECT WAS REMOVED THE PISTON ROD FUNCTIONED CORRECTLY | No Device Problem Found | Medical Evaluation result code
7674406 | 116274003 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY MOTOR POSITION ENCODER ERROR ALARM DUE TO MOTOR ERROR ALARMS | Component or Accessory Incompatibility | Device problem code
7674406 | 116274003 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY MOTOR POSITION ENCODER ERROR ALARM DUE TO MOTOR ERROR ALARMS | Failure to Deliver | Device problem code
7674406 | 116274003 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY MOTOR POSITION ENCODER ERROR ALARM DUE TO MOTOR ERROR ALARMS | Inaccurate Delivery | Device Problem Code
7674406 | 116274003 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY MOTOR POSITION ENCODER ERROR ALARM DUE TO MOTOR ERROR ALARMS | No Device Problem Found | Medical Evaluation result code
7672669 | 116279892 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO EXCESSIVE NO DELIVERY ALARMS NOTED | No Device Problem Found | Medical Evaluation result code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Loss of Power | Device problem code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Audible Alarm | Device problem code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Display / Image | Device Problem Code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Disconnection | Device problem code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
7676574 | 116280914 | N | PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS PUMP MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO THE ELECTRICAL BOARD DURING VISUAL INSPECTION HOWEVER PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ON THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
7701108 | 116373917 | N | ADDITIONAL INFORMATION DEVICE MANUFACTURE DATE PROVIDED A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT TESTING REVEALED THAT THE REPORTED ISSUE COULD BE REPLICATED AS THE PROBE WOULD NOT MOVE IN THE NEW TRAJECTORY AS DESCRIBED IN THE REPORTED ISSUE | Mechanical Jam | Device problem code
7701108 | 116373917 | N | ADDITIONAL INFORMATION DEVICE MANUFACTURE DATE PROVIDED A MEDTRONIC REPRESENTATIVE WENT TO THE SITE TO TEST THE EQUIPMENT TESTING REVEALED THAT THE REPORTED ISSUE COULD BE REPLICATED AS THE PROBE WOULD NOT MOVE IN THE NEW TRAJECTORY AS DESCRIBED IN THE REPORTED ISSUE | Mechanical Problem Identified | Medical Evaluation result code
7762407 | 116418620 | D | IT WAS REPORTED THAT THE PATIENT HAD TO BE HOSPITALIZED WITH BLOOD GLUCOSE BG READING AT MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE BACK AT THE HOSPITAL SHE WAS TREATED WITH AN INFUSION AND HER BG LOWERED TO MMOLL MGDL | Insufficient Information | Device problem code
7762407 | 116418620 | D | IT WAS REPORTED THAT THE PATIENT HAD TO BE HOSPITALIZED WITH BLOOD GLUCOSE BG READING AT MMOLL MGDL AFTER WEARING THE POD BETWEEN AND HOURS ON THE BACK AT THE HOSPITAL SHE WAS TREATED WITH AN INFUSION AND HER BG LOWERED TO MMOLL MGDL | No Known Impact Or Consequence To Patient | Patient Problem Code
7693278 | 116471708 | N | ADDITIONAL INFORMATION WAS RECEIVED WHICH WAS NOT INCLUDED WITH THE INITIAL REPORT THE ADDITIONAL INFORMATION HAS BEEN PROVIDED WITH THIS REPORT THE INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | Electrical Power Problem | Device problem code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | Inaccurate Delivery | Device Problem Code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | Component or Accessory Incompatibility | Device problem code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | Inappropriate or Unexpected Reset | Device Problem Code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | Failure to Deliver | Device problem code
7662554 | 116636992 | N | PUMP PASSED THE DISPLACEMENT TEST UNIT RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED OR FLUSH DRIVE SUPPORT DISK UNABLE TO PERFORM THE UNEXPECTED RESTART ERROR TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY POSSIBLE OVER DELIVERY AND CARELINK DUE TO MOTOR ERROR ALARM PUMP RECEIVED WITH NORMAL OPERATING CURRENT PUMP PASSED SELF TEST PUMP PASSED OFF NO POWER TEST | No Device Problem Found | Medical Evaluation result code
7657696 | 116639828 | N | NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
7663258 | 116639854 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM AND EXCESSIVE NO DELIVERY ALARMS NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Failure to Deliver | Device problem code
7663258 | 116639854 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM AND EXCESSIVE NO DELIVERY ALARMS NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Inaccurate Delivery | Device Problem Code
7663258 | 116639854 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM AND EXCESSIVE NO DELIVERY ALARMS NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Device Alarm System | Device problem code
7663258 | 116639854 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM AND EXCESSIVE NO DELIVERY ALARMS NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Component or Accessory Incompatibility | Device problem code
7663258 | 116639854 | N | PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM AND EXCESSIVE NO DELIVERY ALARMS NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | No Device Problem Found | Medical Evaluation result code
7733564 | 116643623 | N | NONREPORTABLE RATIONALE BASED ON ADDITIONAL INFORMATION RECEIVED FROM DR B ON AUGUST IT WAS CONFIRMED THAT THE WINGSPAN STENT PERFORMED AS INTENDED AND WAS NOT DIRECTLY RELATED TO THE VESSEL PERFORATION AND SUBSEQUENT ADVERSE EVENTS BUT THAT THE VESSEL PERFORATION WAS DUE TO A GUIDEWIRE UNKNOWN MANUFACTURER PLACED INTO A SMALL PERFORATING VESSEL THEREFORE THERE IS NO ALLEGATION AGAINST THE WINGSPAN STENT THE MANUFACTURER HAS REVIEWED ALL INFORMATION AND DETERMINED THIS EVENT NO LONGER MEETS THE REQUIREMENT OF THE REPORTABLE EVENT FOR THE DEVICE IN QUESTION | Material Perforation | Device problem code
7733564 | 116643623 | N | NONREPORTABLE RATIONALE BASED ON ADDITIONAL INFORMATION RECEIVED FROM DR B ON AUGUST IT WAS CONFIRMED THAT THE WINGSPAN STENT PERFORMED AS INTENDED AND WAS NOT DIRECTLY RELATED TO THE VESSEL PERFORATION AND SUBSEQUENT ADVERSE EVENTS BUT THAT THE VESSEL PERFORATION WAS DUE TO A GUIDEWIRE UNKNOWN MANUFACTURER PLACED INTO A SMALL PERFORATING VESSEL THEREFORE THERE IS NO ALLEGATION AGAINST THE WINGSPAN STENT THE MANUFACTURER HAS REVIEWED ALL INFORMATION AND DETERMINED THIS EVENT NO LONGER MEETS THE REQUIREMENT OF THE REPORTABLE EVENT FOR THE DEVICE IN QUESTION | Deformation Problem | Medical Evaluation result code
7772758 | 116754760 | N | ACCORDING TO THE GORE EXCLUDER AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE IFU COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE EXCLUDER AAA ENDOPROSTHESIS THAT MAY OCCUR ANDOR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ENDOLEAK USERS ARE MADE AWARE OF THE RISKS ASSOCIATED WITH TYPE II ENDOLEAKS IN THE IFU AND ARE INSTRUCTED TO CONSIDER THE RISKS AND BENEFITS DISCUSSED IN THE IFU FOR EACH PATIENT BEFORE USING THE DEVICES | Insufficient Information | Device problem code
7772758 | 116754761 | D | ON B THE PATIENT UNDERWENT TREATMENT OF A MM ABDOMINAL AORTIC ANEURYSM USING TWO GORE EXCLUDER AAA ENDOPROSTHESES RLT PLCON B A TYPE II ENDOLEAK WAS IDENTIFIED ON B EMBOLIZATION OF THE LUMBAR ARTERY ANEURYSM SAC AND RIGHT RENAL ACCESSORY ARTERY WAS PERFORMED | Aneurysm | Patient problem code
7772758 | 116754761 | D | ON B THE PATIENT UNDERWENT TREATMENT OF A MM ABDOMINAL AORTIC ANEURYSM USING TWO GORE EXCLUDER AAA ENDOPROSTHESES RLT PLCON B A TYPE II ENDOLEAK WAS IDENTIFIED ON B EMBOLIZATION OF THE LUMBAR ARTERY ANEURYSM SAC AND RIGHT RENAL ACCESSORY ARTERY WAS PERFORMED | Hemorrhage | Patient problem code
7772758 | 116754761 | D | ON B THE PATIENT UNDERWENT TREATMENT OF A MM ABDOMINAL AORTIC ANEURYSM USING TWO GORE EXCLUDER AAA ENDOPROSTHESES RLT PLCON B A TYPE II ENDOLEAK WAS IDENTIFIED ON B EMBOLIZATION OF THE LUMBAR ARTERY ANEURYSM SAC AND RIGHT RENAL ACCESSORY ARTERY WAS PERFORMED | Insufficient Information | Device problem code
7728009 | 116785566 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | Material Erosion | Device problem code
7728009 | 116785566 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | No Device Problem Found | Medical Evaluation result code
7712477 | 116787821 | N | THE FIELD COMPLAINT OF THE TRANSMITTER BEING FRIED FROM A THUNDERSTORM WAS NOT VERIFIED HOWEVER A NO POWER ISSUE WAS CONFIRMED FINAL ANALYSIS FOUND AN AC POWER ADAPTER ANOMALY WHICH CONTRIBUTED TO THE POWER UP ANOMALY | Electrical Power Problem | Device problem code
7712477 | 116787821 | N | THE FIELD COMPLAINT OF THE TRANSMITTER BEING FRIED FROM A THUNDERSTORM WAS NOT VERIFIED HOWEVER A NO POWER ISSUE WAS CONFIRMED FINAL ANALYSIS FOUND AN AC POWER ADAPTER ANOMALY WHICH CONTRIBUTED TO THE POWER UP ANOMALY | Energy Storage System Problem | Medical Evaluation result code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | Detachment of Device or device Component | Device problem code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | Material Protrusion / Extrusion | Device problem code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | Inaccurate Delivery | Device Problem Code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | Failure to Deliver | Device problem code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | Component or Accessory Incompatibility | Device Problem Code
7675622 | 116789539 | N | DEVICE WAS RECEIVED WITH DETACHED END CAP AND LOOSE OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO DETACHED END CAP | No Device Problem Found | Medical Evaluation result code
7672991 | 116790344 | N | ONE FR ANGIOSEAL VIP WAS RETURNED FOR EVALUATION THE CARRIER TUBE ASSEMBLY WAS RETURNED SEPARATE ALONG WITH THE ARTERIOTOMY LOCATOR AND HEMOSTASIS SHEATH THE RETURNED DEVICE WAS SUBJECTED TO VISUAL ANALYSIS THERE WAS NO ANCHOR COLLAGEN OR TAMPER TUBE RETURNED NO SUTURE WAS VISIBLE OUTSIDE THE DISTAL END OF THE CARRIER TUBE TIP THE CARRIER TUBE HUB WAS NOT MATED WITH THE HEMOSTASIS SHEATH AND THE DEPLOYMENT SLEEVE WAS IN THE REAR POSITION THE HUB OF THE CARRIER TUBE WAS REMOVED AND ALL TURNS OF THE SUTURE WERE NOTED WOUND ON THE SPOOL THE CARRIER TUBE WAS CUT THROUGH LONGITUDINALLY AND THE SHEARED END OF THE SUTURE WAS EXAMINED IT WAS CONFIRMED THAT THE SUTURE HAD SEVERED WITHIN THE TUBE THERE IS NO EVIDENCE THAT THIS EVENT WAS RELATED TO A DEVICE DEFECT OR MALFUNCTION FROM THE POSITION OF THE DEPLOYMENT SLEEVE IT IS POSSIBLE THAT NOT POSITIONING THE DEPLOYMENT SLEEVE APPROPRIATELY BEFORE WITHDRAWAL MAY HAVE PLACED ON THE SUTURE WHICH WHEN COMBINED WITH OFF AXIS FORCES COULD HAVE LED TO THIS EVENT THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DEFINITIVELY DETERMINED BASED ON THE AVAILABLE INFORMATION | Defective Device | Device problem code
7672991 | 116790344 | N | ONE FR ANGIOSEAL VIP WAS RETURNED FOR EVALUATION THE CARRIER TUBE ASSEMBLY WAS RETURNED SEPARATE ALONG WITH THE ARTERIOTOMY LOCATOR AND HEMOSTASIS SHEATH THE RETURNED DEVICE WAS SUBJECTED TO VISUAL ANALYSIS THERE WAS NO ANCHOR COLLAGEN OR TAMPER TUBE RETURNED NO SUTURE WAS VISIBLE OUTSIDE THE DISTAL END OF THE CARRIER TUBE TIP THE CARRIER TUBE HUB WAS NOT MATED WITH THE HEMOSTASIS SHEATH AND THE DEPLOYMENT SLEEVE WAS IN THE REAR POSITION THE HUB OF THE CARRIER TUBE WAS REMOVED AND ALL TURNS OF THE SUTURE WERE NOTED WOUND ON THE SPOOL THE CARRIER TUBE WAS CUT THROUGH LONGITUDINALLY AND THE SHEARED END OF THE SUTURE WAS EXAMINED IT WAS CONFIRMED THAT THE SUTURE HAD SEVERED WITHIN THE TUBE THERE IS NO EVIDENCE THAT THIS EVENT WAS RELATED TO A DEVICE DEFECT OR MALFUNCTION FROM THE POSITION OF THE DEPLOYMENT SLEEVE IT IS POSSIBLE THAT NOT POSITIONING THE DEPLOYMENT SLEEVE APPROPRIATELY BEFORE WITHDRAWAL MAY HAVE PLACED ON THE SUTURE WHICH WHEN COMBINED WITH OFF AXIS FORCES COULD HAVE LED TO THIS EVENT THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DEFINITIVELY DETERMINED BASED ON THE AVAILABLE INFORMATION | Activation, Positioning or Separation Problem | Device problem code
7672991 | 116790344 | N | ONE FR ANGIOSEAL VIP WAS RETURNED FOR EVALUATION THE CARRIER TUBE ASSEMBLY WAS RETURNED SEPARATE ALONG WITH THE ARTERIOTOMY LOCATOR AND HEMOSTASIS SHEATH THE RETURNED DEVICE WAS SUBJECTED TO VISUAL ANALYSIS THERE WAS NO ANCHOR COLLAGEN OR TAMPER TUBE RETURNED NO SUTURE WAS VISIBLE OUTSIDE THE DISTAL END OF THE CARRIER TUBE TIP THE CARRIER TUBE HUB WAS NOT MATED WITH THE HEMOSTASIS SHEATH AND THE DEPLOYMENT SLEEVE WAS IN THE REAR POSITION THE HUB OF THE CARRIER TUBE WAS REMOVED AND ALL TURNS OF THE SUTURE WERE NOTED WOUND ON THE SPOOL THE CARRIER TUBE WAS CUT THROUGH LONGITUDINALLY AND THE SHEARED END OF THE SUTURE WAS EXAMINED IT WAS CONFIRMED THAT THE SUTURE HAD SEVERED WITHIN THE TUBE THERE IS NO EVIDENCE THAT THIS EVENT WAS RELATED TO A DEVICE DEFECT OR MALFUNCTION FROM THE POSITION OF THE DEPLOYMENT SLEEVE IT IS POSSIBLE THAT NOT POSITIONING THE DEPLOYMENT SLEEVE APPROPRIATELY BEFORE WITHDRAWAL MAY HAVE PLACED ON THE SUTURE WHICH WHEN COMBINED WITH OFF AXIS FORCES COULD HAVE LED TO THIS EVENT THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DEFINITIVELY DETERMINED BASED ON THE AVAILABLE INFORMATION | Break | Device problem code
7672991 | 116790344 | N | ONE FR ANGIOSEAL VIP WAS RETURNED FOR EVALUATION THE CARRIER TUBE ASSEMBLY WAS RETURNED SEPARATE ALONG WITH THE ARTERIOTOMY LOCATOR AND HEMOSTASIS SHEATH THE RETURNED DEVICE WAS SUBJECTED TO VISUAL ANALYSIS THERE WAS NO ANCHOR COLLAGEN OR TAMPER TUBE RETURNED NO SUTURE WAS VISIBLE OUTSIDE THE DISTAL END OF THE CARRIER TUBE TIP THE CARRIER TUBE HUB WAS NOT MATED WITH THE HEMOSTASIS SHEATH AND THE DEPLOYMENT SLEEVE WAS IN THE REAR POSITION THE HUB OF THE CARRIER TUBE WAS REMOVED AND ALL TURNS OF THE SUTURE WERE NOTED WOUND ON THE SPOOL THE CARRIER TUBE WAS CUT THROUGH LONGITUDINALLY AND THE SHEARED END OF THE SUTURE WAS EXAMINED IT WAS CONFIRMED THAT THE SUTURE HAD SEVERED WITHIN THE TUBE THERE IS NO EVIDENCE THAT THIS EVENT WAS RELATED TO A DEVICE DEFECT OR MALFUNCTION FROM THE POSITION OF THE DEPLOYMENT SLEEVE IT IS POSSIBLE THAT NOT POSITIONING THE DEPLOYMENT SLEEVE APPROPRIATELY BEFORE WITHDRAWAL MAY HAVE PLACED ON THE SUTURE WHICH WHEN COMBINED WITH OFF AXIS FORCES COULD HAVE LED TO THIS EVENT THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DEFINITIVELY DETERMINED BASED ON THE AVAILABLE INFORMATION | Fracture Problem | Medical Evaluation result code
7621956 | 116809627 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | No Device Problem Found | Medical Evaluation result code
7693116 | 116811100 | N | NO FURTHER INVESTIGATION IS WARRANTED BASED ON THE INFORMATION THAT THE PATIENT ANNULUS ANATOMY DID NOT ALLOW A PROPER VALVE POSITIONING LEADING TO CORONARY OBSTRUCTION THE EVENT IS RELATED TO THE PATIENTS SPECIFIC ANATOMY AND THE PROCEDURE AND IS NOT RELATED TO THE DEVICE | Obstruction of Flow | Device problem code
7693116 | 116811100 | N | NO FURTHER INVESTIGATION IS WARRANTED BASED ON THE INFORMATION THAT THE PATIENT ANNULUS ANATOMY DID NOT ALLOW A PROPER VALVE POSITIONING LEADING TO CORONARY OBSTRUCTION THE EVENT IS RELATED TO THE PATIENTS SPECIFIC ANATOMY AND THE PROCEDURE AND IS NOT RELATED TO THE DEVICE | Operational Problem Identified | Medical Evaluation result code
7706670 | 116861205 | N | ADDITIONAL INFORMATION UNIQUE DEVICE IDENTIFICATION UDI AND DEVICE MANUFACTURE DATE PROVIDED A SOFTWARE INVESTIGATION ANALYSIS WAS INITIATED TO DETERMINE THE PROBABLE CAUSE OF THE ISSUE THROUGH REVIEW OF THE REPORTED ISSUE ANALYSIS FOUND THAT THE WORKFLOW BY THE SITE AND MOVEMENT PATHS RESULTED IN THE IMPRECISION ANALYSIS FOUND THAT THE SOFTWARE FUNCTIONED AS DESIGNED | Imprecision | Device problem code
7706670 | 116861205 | N | ADDITIONAL INFORMATION UNIQUE DEVICE IDENTIFICATION UDI AND DEVICE MANUFACTURE DATE PROVIDED A SOFTWARE INVESTIGATION ANALYSIS WAS INITIATED TO DETERMINE THE PROBABLE CAUSE OF THE ISSUE THROUGH REVIEW OF THE REPORTED ISSUE ANALYSIS FOUND THAT THE WORKFLOW BY THE SITE AND MOVEMENT PATHS RESULTED IN THE IMPRECISION ANALYSIS FOUND THAT THE SOFTWARE FUNCTIONED AS DESIGNED | Non-Functional Defect | Medical Evaluation result code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Inaccurate Delivery | Device Problem Code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Failure to Deliver | Device problem code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Mechanical Jam | Device problem code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Device Alarm System | Device problem code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Component or Accessory Incompatibility | Device problem code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | Interoperability Problem Identified | Medical Evaluation result code
7681963 | 116926530 | N | INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING INSULIN PUMP PASSED DISPLACEMENT TEST SELFTEST REWIND TEST BASIC OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | No Device Problem Found | Medical Evaluation result code
7661830 | 116927802 | N | DEVICE PASSED THE DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST | No Device Problem Found | Medical Evaluation result code
7664519 | 116931969 | N | DEVICE PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED NO DRIVE OR MOTOR ANOMALY NOTED DEVICE PROGRAMMED WITH MULTIPLE BOLUS DELIVERIES AND ALL BOLUS DELIVERED PROPERLY THEIR INDICATED AMOUNTS AND WERE PROPERLY RECORDED IN THE DAILY HISTORY NOTED | No Device Problem Found | Medical Evaluation result code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Component or Accessory Incompatibility | Device problem code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Failure to Deliver | Device problem code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Mechanical Jam | Device problem code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Device Alarm System | Device problem code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Inaccurate Delivery | Device Problem Code
7661909 | 116932959 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED DOME SWITCH ON BOLUS ESC ACT UP ARROW AND DOWN ARROW BUTTON CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD NO BUTTON ERROR ALARM DURING TESTING DEVICE PASSED DISPLACEMENT TEST SELF TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | No Device Problem Found | Medical Evaluation result code
7684221 | 116933219 | N | DEVICE PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST | No Device Problem Found | Medical Evaluation result code
7664295 | 116933314 | N | DEVICE PASSED BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST NO MOTOR ERROR ALARM OR EXCESSIVE NO DELIVERY ALARM NOTED MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7779340 | 116993610 | D | OUS MDR IT WAS REPORTED THAT MONTHS AFTER THE IMPLANTATION AN EXIT BLOCK OCCURRED THUS A REVISION WAS PERFORMED DURING THE PROCEDURE THE EXIT BLOCK WAS ALSO MEASURED WITH AN EXTERNAL DEVICE THE LEAD WAS EXPLANTED AND REPLACED | Exit Block | Patient problem code
7779340 | 116993610 | D | OUS MDR IT WAS REPORTED THAT MONTHS AFTER THE IMPLANTATION AN EXIT BLOCK OCCURRED THUS A REVISION WAS PERFORMED DURING THE PROCEDURE THE EXIT BLOCK WAS ALSO MEASURED WITH AN EXTERNAL DEVICE THE LEAD WAS EXPLANTED AND REPLACED | Insufficient Information | Device problem code
7723256 | 116999499 | N | ADDITIONAL INFORMATION MEDTRONIC REPRESENTATIVE REPORTED THAT THERE WAS A LARGE PIECE OF METAL IN THE FILED OF VIEW WHICH CAUSED THE ARTIFACT | Signal Artifact | Device problem code
7723256 | 116999499 | N | ADDITIONAL INFORMATION MEDTRONIC REPRESENTATIVE REPORTED THAT THERE WAS A LARGE PIECE OF METAL IN THE FILED OF VIEW WHICH CAUSED THE ARTIFACT | Image Artifact | Medical Evaluation result code
7717889 | 117019634 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Calibration Problem | Device problem code
7717889 | 117019634 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Failure to Calibrate | Device problem code
7717889 | 117019634 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7717884 | 117019639 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Failure to Calibrate | Device problem code
7717884 | 117019639 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Calibration Problem | Device problem code
7717884 | 117019639 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7717897 | 117019851 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Failure to Calibrate | Device problem code
7717897 | 117019851 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Calibration Problem | Device problem code
7717897 | 117019851 | N | THE MANUFACTURER HAS COMPLETED THE EVALUATION FOR A VENTILATOR THAT WAS RETURNED WITH AN ALLEGATION THE DEVICE FAILED A TEST STEP DURING TESTING THE DEVICE WAS EVALUATED AT THE MANUFACTURERS SERVICE CENTER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7713393 | 117020609 | N | ADDITIONAL INFORMATION WAS RECEIVED STATING THERE WAS AN ISSUE WITH THE ATRIAL PORT ON THIS DEVICE THE PHYSICIAN REPORTED THE DEVICE DID NOT RECOGNIZE THAT A LEAD WAS INSERTED IN THE PORT AS NO SENSING OR PACING IMPEDANCE WAS OBSERVED EVEN AFTER THE SETSCREWS HAD BEEN SECURED | Compatibility Problem | Device Problem Code
7713393 | 117020609 | N | ADDITIONAL INFORMATION WAS RECEIVED STATING THERE WAS AN ISSUE WITH THE ATRIAL PORT ON THIS DEVICE THE PHYSICIAN REPORTED THE DEVICE DID NOT RECOGNIZE THAT A LEAD WAS INSERTED IN THE PORT AS NO SENSING OR PACING IMPEDANCE WAS OBSERVED EVEN AFTER THE SETSCREWS HAD BEEN SECURED | Impedance Problem | Device problem code
7713393 | 117020609 | N | ADDITIONAL INFORMATION WAS RECEIVED STATING THERE WAS AN ISSUE WITH THE ATRIAL PORT ON THIS DEVICE THE PHYSICIAN REPORTED THE DEVICE DID NOT RECOGNIZE THAT A LEAD WAS INSERTED IN THE PORT AS NO SENSING OR PACING IMPEDANCE WAS OBSERVED EVEN AFTER THE SETSCREWS HAD BEEN SECURED | Interoperability Problem Identified | Medical Evaluation result code
7668258 | 117022888 | N | CORRECTION OUTCOMES ATTRIBUTED TO ADVERSE WERE LEFT BLANK THE CORRECT OUTCOMES ATTRIBUTED TO ADVERSE IS REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENTDAMAGE DEVICES B | Break | Device problem code
7668258 | 117022888 | N | CORRECTION OUTCOMES ATTRIBUTED TO ADVERSE WERE LEFT BLANK THE CORRECT OUTCOMES ATTRIBUTED TO ADVERSE IS REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENTDAMAGE DEVICES B | Fracture Problem | Medical Evaluation result code
7696705 | 117023151 | N | ON B IT WAS REPORTED TO MENTOR THAT THE PATIENT FELT LUMP ON THE LEFT BREAST IMPLANT THERE WAS FREE SILICONE AND CALCIFICATIONS ON THE LEFT BREAST IMPLANT AND LEFT AXILLARY NODES THERE WAS A PALPABLE MASS ON THE LEFT BREAST THE PATIENT WAS NOT HOSPITALIZED B | Insufficient Information | Device problem code
7678060 | 117023260 | N | B ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DEMOGRAPHICS KNOWN AT THE TIME OF INDEX PROCEDURE AGE B BMI B CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE URINARY INCONTINENCE NEVER CURED SINCE PHYSICIANS OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT LIKELY CAUSES ISSUES WITH TECHNIQUE OF INSERTION MISSED AT CHECK CYSTOSCOPYURETHROSCOPY EROSION ANY CONCURRENT PROCEDUREDEVICE IMPLANTATION NONE WERE THERE ANY INTRAOPERATIVE COMPLICATIONS NONE WHEN WAS THE MESH EXPOSURE FIRST NOTED BY A PHYSICIAN B MESH EXPOSURE SYMPTOMS AND DIAGNOSTIC CONFIRMATION ONGOING STRESS INCONTINENCE MESH SEEN IN URETHRAL LUMEN AT CYSTOSCOPY DESCRIBE ANY MEDICALSURGICAL INTERVENTION FOR EXPOSURE INCLUDING DATES AND SURGICAL FINDINGS REFERRED TO TERTIARY CENTRE FOR REMOVAL OF MESH THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE OTHER RELEVANT PATIENT HISTORYCONCOMITANT MEDICATIONS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE | Material Erosion | Device problem code
7678060 | 117023260 | N | B ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED DEMOGRAPHICS KNOWN AT THE TIME OF INDEX PROCEDURE AGE B BMI B CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE URINARY INCONTINENCE NEVER CURED SINCE PHYSICIANS OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT LIKELY CAUSES ISSUES WITH TECHNIQUE OF INSERTION MISSED AT CHECK CYSTOSCOPYURETHROSCOPY EROSION ANY CONCURRENT PROCEDUREDEVICE IMPLANTATION NONE WERE THERE ANY INTRAOPERATIVE COMPLICATIONS NONE WHEN WAS THE MESH EXPOSURE FIRST NOTED BY A PHYSICIAN B MESH EXPOSURE SYMPTOMS AND DIAGNOSTIC CONFIRMATION ONGOING STRESS INCONTINENCE MESH SEEN IN URETHRAL LUMEN AT CYSTOSCOPY DESCRIBE ANY MEDICALSURGICAL INTERVENTION FOR EXPOSURE INCLUDING DATES AND SURGICAL FINDINGS REFERRED TO TERTIARY CENTRE FOR REMOVAL OF MESH THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE OTHER RELEVANT PATIENT HISTORYCONCOMITANT MEDICATIONS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE | No Device Problem Found | Medical Evaluation result code
7587413 | 117025433 | N | THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THAT THIS DEVICE MET ALL MANUFACTURING SPECIFICATION FOR PRODUCT RELEASED FOR DISTRIBUTION NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT ONE PERIPHERAL INTRAVENOUS CATHETER WAS RECEIVED FOR DEVICE EVALUATION VISUAL INSPECTION UNDER A MICROSCOPE DID NOT REVEAL ANY ANOMALIES OR DISCREPANCIES THE NEEDLE WAS FOUND TO BE SHARP AND THE BEVEL SHOWED NO DEFECT THE WALL THICKNESS AND THE TRIM LENGTH WERE FOUND TO BE WITHIN MANUFACTURING SPECIFICATION THE COMPLAINT COULD NOT BE CONFIRMED WHILE NO DEFINITIVE ROOT CAUSE TO THE REPORTED ISSUE COULD BE DETERMINED THIS INVESTIGATION REVEALED NO INTRINSIC EVIDENCE TO SUGGEST A ROOT CAUSE RELATED TO MANUFACTURING | No Device Problem Found | Medical Evaluation result code
7781671 | 117121434 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Insufficient Information | Device problem code
7781671 | 117121434 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Fall | Patient problem code
7781666 | 117123586 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Insufficient Information | Device problem code
7781666 | 117123586 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Fall | Patient problem code
7690410 | 117177237 | N | THE INVESTIGATION HAS BEEN COMPLETED BASED ON THE ANALYSIS THE ALLEGED WAKE BUTTON ISSUE WAS VERIFIED THE ALLEGED TOUCHSCREEN WAS NOT VERIFIED AND THE ALLEGED TEMPERATURE ALARM ISSUE WAS UNABLE TO BE VERIFIED DUE TO A DIFFERENT ISSUE BEING IDENTIFIED | No Device Problem Found | Medical Evaluation result code
7786139 | 117202310 | D | IT WAS REPORTED THAT A LOSS OF ASPIRATION OCCURRED AND THE PATIENT EXPERIENCED BURNING PAIN TWO ANGIOJET SOLENT OMNI CATHETERS WERE SELECTED FOR A DEEP VEIN THROMBOSIS THROMBECTOMY PROCEDURE IN THE LOWER LIMBS WITH SLIGHT VENOUS COMPRESSION SYNDROME A REGULAR ML SALINE BAG WITH UNITS OF HEPARIN SODIUM WAS USED THE FIRST ANGIOJET SOLENT OMNI WAS SET UP THE TOP OF THE SALINE BAG WAS PIERCED WITH MULTIPLE SMALL HOLES FOR ELIMINATING NEGATIVE PRESSURE THE NEEDLE WAS INSERTED INTO THE ML SALINE BAG WITH THOUSAND UNITS OF UROKINASE AND AFTER SECONDS OF FLUSHING THE PHYSICIAN RAN POWER PULSE AFTER THE POWER PULSE RAN FOR ABOUT SECONDS THE GASBAG WAS MAKING RHYTHMICAL SPUTTER SOUNDS THE PHYSICIAN NOTICED THAT THE GASBAG WAS NOT INFLATING REGULARLY THE GASBAG COULD NOT SHRINK REGULARLY WHEN PRIMING AND COULD NOT INHALE TO INFLATE TO PROVIDE ENOUGH AIR PRESSURE THE PHYSICIAN CHANGED TO MANUAL POWER PULSE AND REPLACED THE ANGIOJET CATHETER THE SECOND CATHETER WAS FLUSHED FOR SECONDS AND THROMBECTOMY MODE WAS STARTED AN APPROACH TO THE PROXIMAL PART FROM THE POPLITEAL VEIN AT CMS WAS MADE WHEN THE SCREEN SHOWED CC THE GASBAG MADE BANGING SOUNDS AND ASPIRATION WAS LOWING DOWN WITH LESS THAN CC OF LIQUID EXCHANGED PER SECOND AND THE PATIENT WAS FEELING A BURNING PAIN IN THE POPLITEAL VEIN THE PHYSICIAN STOPPED USING THE ANGIOJET CATHETER AND USED AN UNSPECIFIED F GUIDING INSTEAD THE PATIENTS STATUS IS STABLE | No Consequences Or Impact To Patient | Patient problem code
7786139 | 117202310 | D | IT WAS REPORTED THAT A LOSS OF ASPIRATION OCCURRED AND THE PATIENT EXPERIENCED BURNING PAIN TWO ANGIOJET SOLENT OMNI CATHETERS WERE SELECTED FOR A DEEP VEIN THROMBOSIS THROMBECTOMY PROCEDURE IN THE LOWER LIMBS WITH SLIGHT VENOUS COMPRESSION SYNDROME A REGULAR ML SALINE BAG WITH UNITS OF HEPARIN SODIUM WAS USED THE FIRST ANGIOJET SOLENT OMNI WAS SET UP THE TOP OF THE SALINE BAG WAS PIERCED WITH MULTIPLE SMALL HOLES FOR ELIMINATING NEGATIVE PRESSURE THE NEEDLE WAS INSERTED INTO THE ML SALINE BAG WITH THOUSAND UNITS OF UROKINASE AND AFTER SECONDS OF FLUSHING THE PHYSICIAN RAN POWER PULSE AFTER THE POWER PULSE RAN FOR ABOUT SECONDS THE GASBAG WAS MAKING RHYTHMICAL SPUTTER SOUNDS THE PHYSICIAN NOTICED THAT THE GASBAG WAS NOT INFLATING REGULARLY THE GASBAG COULD NOT SHRINK REGULARLY WHEN PRIMING AND COULD NOT INHALE TO INFLATE TO PROVIDE ENOUGH AIR PRESSURE THE PHYSICIAN CHANGED TO MANUAL POWER PULSE AND REPLACED THE ANGIOJET CATHETER THE SECOND CATHETER WAS FLUSHED FOR SECONDS AND THROMBECTOMY MODE WAS STARTED AN APPROACH TO THE PROXIMAL PART FROM THE POPLITEAL VEIN AT CMS WAS MADE WHEN THE SCREEN SHOWED CC THE GASBAG MADE BANGING SOUNDS AND ASPIRATION WAS LOWING DOWN WITH LESS THAN CC OF LIQUID EXCHANGED PER SECOND AND THE PATIENT WAS FEELING A BURNING PAIN IN THE POPLITEAL VEIN THE PHYSICIAN STOPPED USING THE ANGIOJET CATHETER AND USED AN UNSPECIFIED F GUIDING INSTEAD THE PATIENTS STATUS IS STABLE | Insufficient Information | Device problem code
7605949 | 117305318 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | Material Erosion | Device problem code
7605949 | 117305318 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | No Device Problem Found | Medical Evaluation result code
7676188 | 117306355 | N | DEVICE WAS RECEIVED WITH BLANK DISPLAY DUE TO POWER CONNECTOR NOT PLUGGED IN PROPERLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7676188 | 117306355 | N | DEVICE WAS RECEIVED WITH BLANK DISPLAY DUE TO POWER CONNECTOR NOT PLUGGED IN PROPERLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7676188 | 117306355 | N | DEVICE WAS RECEIVED WITH BLANK DISPLAY DUE TO POWER CONNECTOR NOT PLUGGED IN PROPERLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | No Display / Image | Device Problem Code
7676188 | 117306355 | N | DEVICE WAS RECEIVED WITH BLANK DISPLAY DUE TO POWER CONNECTOR NOT PLUGGED IN PROPERLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Connection Problem | Device problem code
7676188 | 117306355 | N | DEVICE WAS RECEIVED WITH BLANK DISPLAY DUE TO POWER CONNECTOR NOT PLUGGED IN PROPERLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7672677 | 117307541 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | Component or Accessory Incompatibility | Device Problem Code
7672677 | 117307541 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7679680 | 117307682 | N | UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST OCCLUSION TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT PASSED REWIND PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Mechanical Jam | Device problem code
7679680 | 117307682 | N | UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST OCCLUSION TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT PASSED REWIND PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Missing or Inadequate Safety Measures | Medical Evaluation result code
7679680 | 117307682 | N | UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST OCCLUSION TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT PASSED REWIND PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Component Missing | Device problem code
7670592 | 117312882 | N | COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | Component or Accessory Incompatibility | Device Problem Code
7670592 | 117312882 | N | COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | No Device Problem Found | Medical Evaluation result code
7661779 | 117313155 | N | INSULIN PUMP WAS MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH DETACHED END CAP UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIMECOMPROMISED FORCE SENSOR SYSTEM AND EXCESSIVE NO DELIVERY TEST OR VERIFY WATCHDOG TIMEOUT ALARM DUE TO DETACHED END CAP | No Display / Image | Device Problem Code
7661779 | 117313155 | N | INSULIN PUMP WAS MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH DETACHED END CAP UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIMECOMPROMISED FORCE SENSOR SYSTEM AND EXCESSIVE NO DELIVERY TEST OR VERIFY WATCHDOG TIMEOUT ALARM DUE TO DETACHED END CAP | Inaccurate Delivery | Device Problem Code
7661779 | 117313155 | N | INSULIN PUMP WAS MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH DETACHED END CAP UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIMECOMPROMISED FORCE SENSOR SYSTEM AND EXCESSIVE NO DELIVERY TEST OR VERIFY WATCHDOG TIMEOUT ALARM DUE TO DETACHED END CAP | Detachment of Device or device Component | Device problem code
7661779 | 117313155 | N | INSULIN PUMP WAS MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH DETACHED END CAP UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIMECOMPROMISED FORCE SENSOR SYSTEM AND EXCESSIVE NO DELIVERY TEST OR VERIFY WATCHDOG TIMEOUT ALARM DUE TO DETACHED END CAP | Failure to Deliver | Device problem code
7661779 | 117313155 | N | INSULIN PUMP WAS MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH DETACHED END CAP UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIMECOMPROMISED FORCE SENSOR SYSTEM AND EXCESSIVE NO DELIVERY TEST OR VERIFY WATCHDOG TIMEOUT ALARM DUE TO DETACHED END CAP | Operational Problem Identified | Medical Evaluation result code
7672257 | 117313243 | N | THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
7672257 | 117313243 | N | THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Mechanical Jam | Device problem code
7672257 | 117313243 | N | THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Operational Problem Identified | Medical Evaluation result code
7672257 | 117313243 | N | THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
7666537 | 117315678 | N | UNIT ALARMED MOTOR ERROR DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM OPERATING BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARM | Inaccurate Delivery | Device Problem Code
7666537 | 117315678 | N | UNIT ALARMED MOTOR ERROR DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM OPERATING BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARM | Failure to Deliver | Device problem code
7666537 | 117315678 | N | UNIT ALARMED MOTOR ERROR DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM OPERATING BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARM | Component or Accessory Incompatibility | Device Problem Code
7666537 | 117315678 | N | UNIT ALARMED MOTOR ERROR DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM OPERATING BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARM | Interoperability Problem Identified | Medical Evaluation result code
7666537 | 117315678 | N | UNIT ALARMED MOTOR ERROR DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM OPERATING BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARM | Operational Problem Identified | Medical Evaluation result code
7628480 | 117328178 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY FAULT ISOLATED TO ELECTRICAL BOARD WE WERE UNABLE TO PERFORM THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST OR VERIFY MISSING SEGMENTS AND PARTIAL DISPLAY DUE TO BLANK DISPLAY | Display or Visual Feedback Problem | Device problem code
7628480 | 117328178 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY FAULT ISOLATED TO ELECTRICAL BOARD WE WERE UNABLE TO PERFORM THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST OR VERIFY MISSING SEGMENTS AND PARTIAL DISPLAY DUE TO BLANK DISPLAY | Component Missing | Device problem code
7628480 | 117328178 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY FAULT ISOLATED TO ELECTRICAL BOARD WE WERE UNABLE TO PERFORM THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST OR VERIFY MISSING SEGMENTS AND PARTIAL DISPLAY DUE TO BLANK DISPLAY | No Display / Image | Device Problem Code
7628480 | 117328178 | N | THE INSULIN PUMP WAS RECEIVED WITH A BLANK DISPLAY FAULT ISOLATED TO ELECTRICAL BOARD WE WERE UNABLE TO PERFORM THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST OR VERIFY MISSING SEGMENTS AND PARTIAL DISPLAY DUE TO BLANK DISPLAY | Missing or Inadequate Safety Measures | Medical Evaluation result code
7672791 | 117329061 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPEC UNIT PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Component Missing | Device problem code
7672791 | 117329061 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPEC UNIT PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Mechanical Jam | Device problem code
7672791 | 117329061 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPEC UNIT PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Missing or Inadequate Safety Measures | Medical Evaluation result code
7672791 | 117329061 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPEC UNIT PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST | Mechanical Problem Identified | Medical Evaluation result code
7679703 | 117336873 | N | UNIT PASSED THE DISPLACEMENT TEST AND SELFTEST NO PUMP ERROR NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Inappropriate or Unexpected Reset | Device Problem Code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Inaccurate Delivery | Device Problem Code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Failure to Deliver | Device problem code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Failure of Device to Self-Test | Device problem code
7665421 | 117354165 | N | UNIT WAS RECEIVED WITH A RF PIC FAILED SELFTEST DURING THE SELFTEST PROBLEM ISOLATED TO THE RADIO FREQUENCY BOARD HOWEVER THE PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS UNIT PASSED THE UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Operational Problem Identified | Medical Evaluation result code
7663165 | 117459873 | N | UNIT RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEFLUSHED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7663165 | 117459873 | N | UNIT RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEFLUSHED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Component or Accessory Incompatibility | Device Problem Code
7663165 | 117459873 | N | UNIT RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEFLUSHED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7663165 | 117459873 | N | UNIT RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEFLUSHED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Operational Problem Identified | Medical Evaluation result code
7663464 | 117475357 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY BATTERY OUT LIMIT ALARM DUE TO BLANK DISPLAY | Inaccurate Delivery | Device Problem Code
7663464 | 117475357 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY BATTERY OUT LIMIT ALARM DUE TO BLANK DISPLAY | Inappropriate or Unexpected Reset | Device Problem Code
7663464 | 117475357 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY BATTERY OUT LIMIT ALARM DUE TO BLANK DISPLAY | Failure to Deliver | Device problem code
7663464 | 117475357 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY BATTERY OUT LIMIT ALARM DUE TO BLANK DISPLAY | No Display / Image | Device problem code
7663464 | 117475357 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY BATTERY OUT LIMIT ALARM DUE TO BLANK DISPLAY | Operational Problem Identified | Medical Evaluation result code
7668811 | 117476978 | N | IT IS NOW REPORTED THAT THE PATIENT EXPERIENCED A SCABBED SKIN WOUND OVER THE RECEIVER STIMULATOR AREA AND SUBSEQUENTLY THE SCABBED AREA WAS REMOVED ON B THE PATIENT ALSO EXPERIENCED PURULENT DRAINAGE ON B THE DEVICE WAS EXPLANTED WITH A SKIN GRAFT AND HOSPITAL STAY AND THE PATIENT WAS ADMINISTERED IV ANTIBIOTICS THE ISSUES HAVE REPORTEDLY SINCE RESOLVED | Leak / Splash | Device problem code
7668811 | 117476978 | N | IT IS NOW REPORTED THAT THE PATIENT EXPERIENCED A SCABBED SKIN WOUND OVER THE RECEIVER STIMULATOR AREA AND SUBSEQUENTLY THE SCABBED AREA WAS REMOVED ON B THE PATIENT ALSO EXPERIENCED PURULENT DRAINAGE ON B THE DEVICE WAS EXPLANTED WITH A SKIN GRAFT AND HOSPITAL STAY AND THE PATIENT WAS ADMINISTERED IV ANTIBIOTICS THE ISSUES HAVE REPORTEDLY SINCE RESOLVED | Leakage/Seal | Medical Evaluation result code
7724162 | 117478184 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX RETRACTED AND CLOGGED WITH BLOOD AND TISSUE XRAY EXAMINATION SHOWED THE INNER COIL WAS OVER TORQUED AT THE CONNECTOR REGION CONSISTENT WITH THE ATTEMPTED IMPLANT PROCEDURE AFTER THE HELIX HOUSING WAS CLEANED THE HELIX FUNCTIONED NORMALLY THE REPORTED FIELD EVENT OF THE INABILITY TO EXTEND THE HELIX OF THE TENDRIL STS LEAD WAS CONFIRMED | Complete Blockage | Device problem code
7724162 | 117478184 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX RETRACTED AND CLOGGED WITH BLOOD AND TISSUE XRAY EXAMINATION SHOWED THE INNER COIL WAS OVER TORQUED AT THE CONNECTOR REGION CONSISTENT WITH THE ATTEMPTED IMPLANT PROCEDURE AFTER THE HELIX HOUSING WAS CLEANED THE HELIX FUNCTIONED NORMALLY THE REPORTED FIELD EVENT OF THE INABILITY TO EXTEND THE HELIX OF THE TENDRIL STS LEAD WAS CONFIRMED | Retraction Problem | Device problem code
7724162 | 117478184 | N | AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX RETRACTED AND CLOGGED WITH BLOOD AND TISSUE XRAY EXAMINATION SHOWED THE INNER COIL WAS OVER TORQUED AT THE CONNECTOR REGION CONSISTENT WITH THE ATTEMPTED IMPLANT PROCEDURE AFTER THE HELIX HOUSING WAS CLEANED THE HELIX FUNCTIONED NORMALLY THE REPORTED FIELD EVENT OF THE INABILITY TO EXTEND THE HELIX OF THE TENDRIL STS LEAD WAS CONFIRMED | Operational Problem Identified | Medical Evaluation result code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Positioning Failure | Device problem code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Component Missing | Device problem code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Mechanical Jam | Device problem code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Component or Accessory Incompatibility | Device problem code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Mechanical Problem Identified | Medical Evaluation result code
7662594 | 117478367 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR ALARMS WERE NOTED UNIT RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER UNIT ALARMED PUMP ERROR DURING REWIND DUE TO CORRODE MOTOR HOME SWITCH UNABLE TO PERFORM REWIND PRIMESEATING TEST BASIC OCCLUSION TEST OR FORCE SENSOR TEST DUE TO PUMP ERROR UNIT RECEIVED WITH OPERATING CURRENTS WITHIN SPECS AND PASSED SELF TEST | Missing or Inadequate Safety Measures | Medical Evaluation result code
7669856 | 117478622 | N | CONCLUSION DEVICE INVESTIGATIONS DID NOT REVEAL ANY DEVICE DEFECT OR PROBLEM WHICH IS EXPECTED TO HAVE BEEN PRESENT WHILST IMPLANTED THIS FINDING WAS EXPECTED BECAUSE ACCORDING TO THE PATIENT REPORT THE DEVICE WAS EXPLANTED DUE TO AN EPISODE OF MENINGITIS A REVIEW OF THE DEVICES STERILIZATION RECORDS SHOWS THAT THE DEVICE HAS BEEN SUBJECT TO A VALID STERILIZATION PROCESS NO INFORMATION AVAILABLE POINTS TO THE IMPLANT BEING THE SOURCE OF INFECTION THE PATIENT IS RECOVERING WELL FOLLOWING EXPLANTATION AND ANTIBIOTICS TREATMENT THIS IS A FINAL REPORT | Defective Device | Device problem code
7669856 | 117478622 | N | CONCLUSION DEVICE INVESTIGATIONS DID NOT REVEAL ANY DEVICE DEFECT OR PROBLEM WHICH IS EXPECTED TO HAVE BEEN PRESENT WHILST IMPLANTED THIS FINDING WAS EXPECTED BECAUSE ACCORDING TO THE PATIENT REPORT THE DEVICE WAS EXPLANTED DUE TO AN EPISODE OF MENINGITIS A REVIEW OF THE DEVICES STERILIZATION RECORDS SHOWS THAT THE DEVICE HAS BEEN SUBJECT TO A VALID STERILIZATION PROCESS NO INFORMATION AVAILABLE POINTS TO THE IMPLANT BEING THE SOURCE OF INFECTION THE PATIENT IS RECOVERING WELL FOLLOWING EXPLANTATION AND ANTIBIOTICS TREATMENT THIS IS A FINAL REPORT | No Device Problem Found | Medical Evaluation result code
7672786 | 117478632 | N | UNIT PASSED REWIND SEATING BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO PUMP ERROR WERE NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7662863 | 117479004 | N | UNIT WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST NO UNEXPECTED LOW BATTERY BATTERY OUT LIMIT AND FAILED BATTERY TEST ALARMS DURING TESTING ALSO UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST NO UNEXPECTED PRIMEFILL ANOMALY NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7665261 | 117479079 | N | UNIT RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST SELF TEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST PUMP PASSED THE DELIVERY ACCURACY TEST AT INCHES NO EXCESSIVE NO DELIVERY ALARMS NOTED | No Device Problem Found | Medical Evaluation result code
7665441 | 117479103 | N | NO UNEXPECTED NO DELIVERY ALARM OR AUDIOBEEP LOUD NOISE ANOMALY NOTED DURING TESTING UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST NO UNEXPECTED PRIMEFILL ANOMALY NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7668904 | 117479156 | N | NO MOTOR ERROR ALARM OR UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7663133 | 117479361 | N | NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST NO UNEXPECTED PRIMEFILL ANOMALY NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Inappropriate or Unexpected Reset | Device Problem Code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Inaccurate Delivery | Device Problem Code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Leak / Splash | Device problem code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Failure to Deliver | Device problem code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Operational Problem Identified | Medical Evaluation result code
7669036 | 117479669 | N | UNIT ALARMED UNEXPECTED RESTART AFTER BATTERY CHANGE AND RESETS TIMEDATE TO FACTORY DEFAULT DUE TO CONNECTOR VOLTAGE LEAK AT INTERFACE BOARD HOWEVER UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMECOMPROMISED FORCE SENSOR SYSTEM TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Leakage/Seal | Medical Evaluation result code
7670523 | 117479748 | N | INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST INSULIN PUMP WAS TESTED WITH NO PRIMEFILL ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7674946 | 117479803 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Inappropriate or Unexpected Reset | Device Problem Code
7674946 | 117479803 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7674946 | 117479803 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Battery Problem | Device Problem Code
7674946 | 117479803 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Energy Storage System Problem | Medical Evaluation result code
7674987 | 117479865 | N | INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO EXCESSIVE NO DELIVERY ALARMS NOTED | No Device Problem Found | Medical Evaluation result code
6559317 | 117479928 | N | THE ASTRAL DEVICE WAS RETURNED TO RESMED FOR AN INVESTIGATION REVIEW OF THE DEVICE DATA LOGS CONFIRMED THE REPORTED SYSTEM FAULT BASED ON ALL AVAILABLE EVIDENCE AND COMPLAINT INVESTIGATIONS OF A SIMILAR NATURE THE INVESTIGATION DETERMINED THAT THE ROOT CAUSE WAS DEVICE OPERATION IN A TEMPERATURE OUTSIDE OF THE DEVICE SPECIFICATION RESMEDS RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE B | Defective Device | Device problem code
6559317 | 117479928 | N | THE ASTRAL DEVICE WAS RETURNED TO RESMED FOR AN INVESTIGATION REVIEW OF THE DEVICE DATA LOGS CONFIRMED THE REPORTED SYSTEM FAULT BASED ON ALL AVAILABLE EVIDENCE AND COMPLAINT INVESTIGATIONS OF A SIMILAR NATURE THE INVESTIGATION DETERMINED THAT THE ROOT CAUSE WAS DEVICE OPERATION IN A TEMPERATURE OUTSIDE OF THE DEVICE SPECIFICATION RESMEDS RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE B | Device Difficult to Operate | Medical Evaluation result code
7661845 | 117480048 | N | THE PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST A WATER FILLED RESERVOIR WAS USED DURING TESTING SEVERAL BOLUS DELIVERIES WERE PROGRAMMED AND DELIVERED OBSERVED LIQUID EXIT THE TUBING DURING THE BOLUS DELIVERIES ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN NO DELIVERY ANOMALIES NOTED DURING TESTING | Failure to Deliver | Device problem code
7661845 | 117480048 | N | THE PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST A WATER FILLED RESERVOIR WAS USED DURING TESTING SEVERAL BOLUS DELIVERIES WERE PROGRAMMED AND DELIVERED OBSERVED LIQUID EXIT THE TUBING DURING THE BOLUS DELIVERIES ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN NO DELIVERY ANOMALIES NOTED DURING TESTING | Inaccurate Delivery | Device Problem Code
7669188 | 117480448 | N | UNIT RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED RESTART ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST ALSO UNIT PASSED THE DELIVERY ACCURACY TEST AT INCHES NO POSSIBLE OVERUNDER DELIVERY ANOMALY NOTED OR NO RESERVOIR ALARM NOTED DURING TESTING | Failure to Deliver | Device problem code
7669188 | 117480448 | N | UNIT RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED RESTART ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST ALSO UNIT PASSED THE DELIVERY ACCURACY TEST AT INCHES NO POSSIBLE OVERUNDER DELIVERY ANOMALY NOTED OR NO RESERVOIR ALARM NOTED DURING TESTING | Inaccurate Delivery | Device Problem Code
7669188 | 117480448 | N | UNIT RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED RESTART ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST ALSO UNIT PASSED THE DELIVERY ACCURACY TEST AT INCHES NO POSSIBLE OVERUNDER DELIVERY ANOMALY NOTED OR NO RESERVOIR ALARM NOTED DURING TESTING | Inappropriate or Unexpected Reset | Device Problem Code
7669188 | 117480448 | N | UNIT RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED RESTART ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST ALSO UNIT PASSED THE DELIVERY ACCURACY TEST AT INCHES NO POSSIBLE OVERUNDER DELIVERY ANOMALY NOTED OR NO RESERVOIR ALARM NOTED DURING TESTING | Operational Problem Identified | Medical Evaluation result code
7690516 | 117623779 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING OR IN THE INSULIN PUMP TRACE DOWNLOAD DEVICE MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTED BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN TEN MINUTES AND NO UNEXPECTED PUMP ERROR ALARM NOTED | Component or Accessory Incompatibility | Device problem code
7690516 | 117623779 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING OR IN THE INSULIN PUMP TRACE DOWNLOAD DEVICE MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTED BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN TEN MINUTES AND NO UNEXPECTED PUMP ERROR ALARM NOTED | Loss of Power | Device problem code
7690516 | 117623779 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING OR IN THE INSULIN PUMP TRACE DOWNLOAD DEVICE MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTED BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN TEN MINUTES AND NO UNEXPECTED PUMP ERROR ALARM NOTED | No Audible Alarm | Device problem code
7690516 | 117623779 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING OR IN THE INSULIN PUMP TRACE DOWNLOAD DEVICE MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTED BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN TEN MINUTES AND NO UNEXPECTED PUMP ERROR ALARM NOTED | Operational Problem Identified | Medical Evaluation result code
7671150 | 117624331 | N | CORRECTION IN FOLLOWUP NO OF MANUFACTURER REPORT NUMBER NEW INFORMATION RECEIVED NOTES THAT THE LEAD DISLODGEMENTS WERE DISCOVERED DURING ROUTINE FOLLOWUP IN CLINIC WHEN INCREASED CAPTURE THRESHOLD WAS OBSERVED ON BOTH LEADS CHEST XRAY REVEALED THAT THE ONE OF THE LEAD UNSPECIFIED WAS PERFORATING THE APEX ADDITIONAL INTERVENTION TO TREAT THE PERFORATION WAS NOT REQUIRED AFTER LEADS EXPLANT PATIENT WAS STABLE BEFORE DURING AND AFTER THE PROCEDURE | Material Perforation | Device problem code
7671150 | 117624331 | N | CORRECTION IN FOLLOWUP NO OF MANUFACTURER REPORT NUMBER NEW INFORMATION RECEIVED NOTES THAT THE LEAD DISLODGEMENTS WERE DISCOVERED DURING ROUTINE FOLLOWUP IN CLINIC WHEN INCREASED CAPTURE THRESHOLD WAS OBSERVED ON BOTH LEADS CHEST XRAY REVEALED THAT THE ONE OF THE LEAD UNSPECIFIED WAS PERFORATING THE APEX ADDITIONAL INTERVENTION TO TREAT THE PERFORATION WAS NOT REQUIRED AFTER LEADS EXPLANT PATIENT WAS STABLE BEFORE DURING AND AFTER THE PROCEDURE | Deformation Problem | Medical Evaluation result code
7661720 | 117626715 | N | UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT NO PUMP ERROR ALARMS NOTED DURING TESTING UNIT COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND THE GUARDIAN LINK TRANSMITTER THE CORRECT TEST BLOOD GLUCOSE VALUE WAS DISPLAYED ON THE SENSOR GLUCOSE GRAPH SCREEN NO UNEXPECTED SENSOR UPDATING CHANGE SENSOR AND CHANGED EARLY MESSAGES NOTED DURING TESTING THE AUDIO WAS TOGGLED ON AND OFF WITH NO AUDIOBEEP ANOMALIES OR SPEAKER ANOMALIES NOTED DURING TESTING UNIT RECEIVED WITH PUMP ERROR DURING SELFTEST PROBLEM ISOLATED TO KEYPAD ASSEMBLY | Device Displays Incorrect Message | Device Problem Code
7661720 | 117626715 | N | UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT NO PUMP ERROR ALARMS NOTED DURING TESTING UNIT COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND THE GUARDIAN LINK TRANSMITTER THE CORRECT TEST BLOOD GLUCOSE VALUE WAS DISPLAYED ON THE SENSOR GLUCOSE GRAPH SCREEN NO UNEXPECTED SENSOR UPDATING CHANGE SENSOR AND CHANGED EARLY MESSAGES NOTED DURING TESTING THE AUDIO WAS TOGGLED ON AND OFF WITH NO AUDIOBEEP ANOMALIES OR SPEAKER ANOMALIES NOTED DURING TESTING UNIT RECEIVED WITH PUMP ERROR DURING SELFTEST PROBLEM ISOLATED TO KEYPAD ASSEMBLY | Component or Accessory Incompatibility | Device problem code
7661720 | 117626715 | N | UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT NO PUMP ERROR ALARMS NOTED DURING TESTING UNIT COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND THE GUARDIAN LINK TRANSMITTER THE CORRECT TEST BLOOD GLUCOSE VALUE WAS DISPLAYED ON THE SENSOR GLUCOSE GRAPH SCREEN NO UNEXPECTED SENSOR UPDATING CHANGE SENSOR AND CHANGED EARLY MESSAGES NOTED DURING TESTING THE AUDIO WAS TOGGLED ON AND OFF WITH NO AUDIOBEEP ANOMALIES OR SPEAKER ANOMALIES NOTED DURING TESTING UNIT RECEIVED WITH PUMP ERROR DURING SELFTEST PROBLEM ISOLATED TO KEYPAD ASSEMBLY | Interoperability Problem Identified | Medical Evaluation result code
7710871 | 117630646 | N | THE FAILURE INVESTIGATION HAS BEEN COMPLETED AND THE ALLEGED MALFUNCTION ISSUE WAS VERIFIED HOWEVER THE ALLEGED MINIMUM FILL ISSUE COULD NOT BE VERIFIED DUE TO THE MALFUNCTION ISSUE | Filling Problem | Device Problem Code
7710871 | 117630646 | N | THE FAILURE INVESTIGATION HAS BEEN COMPLETED AND THE ALLEGED MALFUNCTION ISSUE WAS VERIFIED HOWEVER THE ALLEGED MINIMUM FILL ISSUE COULD NOT BE VERIFIED DUE TO THE MALFUNCTION ISSUE | Operational Problem Identified | Medical Evaluation result code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Component Missing | Device problem code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Failure to Deliver | Device problem code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Display or Visual Feedback Problem | Device problem code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Inappropriate or Unexpected Reset | Device Problem Code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Inaccurate Delivery | Device Problem Code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Component or Accessory Incompatibility | Device problem code
7660750 | 117630654 | N | DEVICE RECEIVED WITH ALL BUTTON NO RESPONSE DUE TO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING TEST UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY MISSING SEGMENTS OR PARTIAL DISPLAY ANOMALY DUE TO BUTTONS NOT RESPONDING | Interoperability Problem Identified | Medical Evaluation result code
7588964 | 117632448 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Defective Device | Device problem code
7588964 | 117632448 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Device Difficult to Operate | Medical Evaluation result code
7663768 | 117693723 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO PRIME ANOMALY OR UNEXPECTED MAX FILLED REACHED ALARM ANOMALY NOTED MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7657914 | 117694094 | N | INSULIN PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | No Display / Image | Device Problem Code
7657914 | 117694094 | N | INSULIN PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | Image Artifact | Medical Evaluation result code
7664657 | 117695625 | N | INSULIN PUMP PASSED ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARM NO FAILED BATTERY TEST ALERT AND NO LOW BATTERY TEST ALARM NOTED DURING TEST MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7673222 | 117696611 | N | INSULIN PUMP PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST INSULIN PUMP WAS MONITORED AND NO CONSTANT TONE ANOMALY NOTED NO BATTERY POWER LOSS ALARM NOTED | No Device Problem Found | Medical Evaluation result code
7668201 | 117696689 | N | INSULIN PUMP PASSED UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED RESTART ALARM NOTED | No Device Problem Found | Medical Evaluation result code
7672031 | 117697028 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7673421 | 117697801 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7663496 | 117697825 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7717266 | 117728586 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | Material Erosion | Device problem code
7717266 | 117728586 | N | AN EROSION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE THE STERILIZATION PROCESS WAS INVESTIGATED THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS SUCH AS GAS CONCENTRATION TEMPERATURE HUMIDITY ETC ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE ADDITIONALLY AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT IN SUMMARY THE EROSION WAS NOT DEVICE RELATED | Deformation Problem | Medical Evaluation result code
7662576 | 117728862 | N | INSULIN PUMP PASSED SELF TEST DISPLACEMENT TEST AND REWIND TEST NO MISSING SEGMENTSPARTIAL DISPLAY OR DISPLAY ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7689439 | 117731152 | N | ADDITIONAL INFORMATION SOFTWARE ANALYSIS A SOFTWARE INVESTIGATION ANALYSIS WAS INITIATED TO DETERMINE THE PROBABLE CAUSE OF THE ISSUE THROUGH IMAGE ANALYSIS ANALYSIS FOUND THAT THERE WAS NONUNIFORMITY AND NOISE IN THE INTENSITIES ESPECIALLY IN THE SOFT TISSUE THIS CAN CAUSE THE CREATION OF A POOR INITIAL MODEL THERE WAS NO INDICATION OF A SOFTWARE ANOMALY OR MALFUNCTION | No Device Problem Found | Medical Evaluation result code
7707904 | 117731516 | N | ADDITIONAL INFORMATION MANUFACTURE DATE AND DEVICE EVALUATION THE SURGEON MONITOR AND LCD CABLE WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT THE LCD CABLE FOR THE MONITOR WAS WORKING WELL WITH NO FAILURE THE MONITOR WAS PLUGGED IN FOR ANALYSIS AND DISPLAYED A GOOD IMAGE SHORTLY AFTERWARDS THE SCREEN WENT BLACK IT WAS DETERMINED TO BE A LCD PLUG MALFUNCTION | No Display / Image | Device problem code
7707904 | 117731516 | N | ADDITIONAL INFORMATION MANUFACTURE DATE AND DEVICE EVALUATION THE SURGEON MONITOR AND LCD CABLE WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT THE LCD CABLE FOR THE MONITOR WAS WORKING WELL WITH NO FAILURE THE MONITOR WAS PLUGGED IN FOR ANALYSIS AND DISPLAYED A GOOD IMAGE SHORTLY AFTERWARDS THE SCREEN WENT BLACK IT WAS DETERMINED TO BE A LCD PLUG MALFUNCTION | Image Artifact | Medical Evaluation result code
7656434 | 117731517 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR ALARM | No Device Problem Found | Medical Evaluation result code
7664858 | 117732977 | N | ADDITIONAL INFORMATION LOG ANALYSIS A SOFTWARE ANALYSIS WAS INITIATED TO ANALYZE RECEIVED LOGS MICROCAL LOGS WERE INSPECTED AND SHOWED NO INDICATION OF A SOFTWARE MALFUNCTION THERE WAS NOT ENOUGH INFORMATION AVAILABLE TO DETERMINE ROOT CAUSE OF THE REPORTED ISSUE | No Device Problem Found | Medical Evaluation result code
7721252 | 117732978 | N | HEALTH PROFESSIONAL PREVIOUSLY STATED NO INFORMATION THE CORRECT RESPONSE IS YES OCCUPATION SHOULD HAVE STATED NURSE RATHER THAN OTHER HEALTH CARE PROFESSIONAL | Insufficient Information | Device problem code
7705707 | 117734338 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST AT INCHES | Insufficient Information | Device problem code
7717080 | 117736845 | N | DEVISE PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST AT INCHES | Insufficient Information | Device problem code
7707950 | 117745502 | N | CORRECTION PLEASE REMOVE ADVERSE EVENT AS THE ISSUE WAS NOTED DURING THE IMPLANT PROCEDURE PLEASE REMOVE REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENTDAMAGE AS THE ISSUE WAS NOTED DURING THE IMPLANT PROCEDURE TYPE OF REPORTABLE EVENT SHOULD HAVE BEEN MALFUNCTION INSTEAD OF SERIOUS INJURY | Break | Device problem code
7707950 | 117745502 | N | CORRECTION PLEASE REMOVE ADVERSE EVENT AS THE ISSUE WAS NOTED DURING THE IMPLANT PROCEDURE PLEASE REMOVE REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENTDAMAGE AS THE ISSUE WAS NOTED DURING THE IMPLANT PROCEDURE TYPE OF REPORTABLE EVENT SHOULD HAVE BEEN MALFUNCTION INSTEAD OF SERIOUS INJURY | Fracture Problem | Medical Evaluation result code
7682918 | 117747608 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7682918 | 117747608 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7682918 | 117747608 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7682918 | 117747608 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7679816 | 117753257 | N | UNIT RECEIVED WITH BLANK DISPLAY DUE TO SHIFTED BATTERY TUBE ASSEMBLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY RATTLING OF PUMP DUE TO DETACHED END CAP | No Display / Image | Device problem code
7679816 | 117753257 | N | UNIT RECEIVED WITH BLANK DISPLAY DUE TO SHIFTED BATTERY TUBE ASSEMBLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY RATTLING OF PUMP DUE TO DETACHED END CAP | Detachment of Device or device Component | Device problem code
7679816 | 117753257 | N | UNIT RECEIVED WITH BLANK DISPLAY DUE TO SHIFTED BATTERY TUBE ASSEMBLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY RATTLING OF PUMP DUE TO DETACHED END CAP | Noise, Audible | Device problem code
7679816 | 117753257 | N | UNIT RECEIVED WITH BLANK DISPLAY DUE TO SHIFTED BATTERY TUBE ASSEMBLY UNABLE TO PERFORM DISPLACEMENT TEST DUE TO BLANK DISPLAY RATTLING OF PUMP DUE TO DETACHED END CAP | Interoperability Problem Identified | Medical Evaluation result code
7681911 | 117753772 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7681911 | 117753772 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7681911 | 117753772 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7681911 | 117753772 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7657500 | 117753844 | N | DEVICE PASSED THE FUNCTIONAL TESTING INCLUDING THE OPERATING CURRENTS MEASUREMENT SELF TEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO LED ANOMALY NOTED NO DUAL OR SQUARE BOLUS ANOMALY OR BOLUS WIZARD ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7662128 | 117755355 | N | DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED THE SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARMS UNEXPECTED EXCESSIVE NO DELIVERY ALARMS OR DISPLACEMENT ANOMALIES NOTED DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST AT B INCHES NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | Inaccurate Delivery | Device Problem Code
7662128 | 117755355 | N | DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED THE SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARMS UNEXPECTED EXCESSIVE NO DELIVERY ALARMS OR DISPLACEMENT ANOMALIES NOTED DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST AT B INCHES NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | Component or Accessory Incompatibility | Device problem code
7662128 | 117755355 | N | DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED THE SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARMS UNEXPECTED EXCESSIVE NO DELIVERY ALARMS OR DISPLACEMENT ANOMALIES NOTED DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST AT B INCHES NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | Failure to Deliver | Device problem code
7662128 | 117755355 | N | DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED THE SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARMS UNEXPECTED EXCESSIVE NO DELIVERY ALARMS OR DISPLACEMENT ANOMALIES NOTED DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST AT B INCHES NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | Inappropriate or Unexpected Reset | Device Problem Code
7662128 | 117755355 | N | DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED THE SELFTEST UNEXPECTED RESTART ERROR TEST DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARMS UNEXPECTED EXCESSIVE NO DELIVERY ALARMS OR DISPLACEMENT ANOMALIES NOTED DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST AT B INCHES NOTED THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | Operational Problem Identified | Medical Evaluation result code
7661851 | 117755787 | N | THE DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST AT INCHES THE DEVICE WAS PROGRAMMED WITH MULTIPLE BOLUSES AND MONITORED ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN THE DEVICE THEN WAS PROGRAMMED WITH MULTIPLE BASAL PROFILES AND MONITORED ALL BASAL PROFILES DELIVERED THEIR INDICATED AMOUNTS AND WERE VERIFIED IN THE DAILY AND SUMMARY HISTORY SCREENS THE UNITS LEFT AT THE INSULIN PUMP DISPLAY MATCHED PROPERLY THE UNITS LEFT AT TEST RESERVOIR NO DELIVERY ANOMALY BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING THE AUTO MODE FEATURE IS OPERATING PROPERLY | Failure to Deliver | Device problem code
7661851 | 117755787 | N | THE DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST AT INCHES THE DEVICE WAS PROGRAMMED WITH MULTIPLE BOLUSES AND MONITORED ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN THE DEVICE THEN WAS PROGRAMMED WITH MULTIPLE BASAL PROFILES AND MONITORED ALL BASAL PROFILES DELIVERED THEIR INDICATED AMOUNTS AND WERE VERIFIED IN THE DAILY AND SUMMARY HISTORY SCREENS THE UNITS LEFT AT THE INSULIN PUMP DISPLAY MATCHED PROPERLY THE UNITS LEFT AT TEST RESERVOIR NO DELIVERY ANOMALY BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING THE AUTO MODE FEATURE IS OPERATING PROPERLY | Inaccurate Delivery | Device Problem Code
7661851 | 117755787 | N | THE DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST AT INCHES THE DEVICE WAS PROGRAMMED WITH MULTIPLE BOLUSES AND MONITORED ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN THE DEVICE THEN WAS PROGRAMMED WITH MULTIPLE BASAL PROFILES AND MONITORED ALL BASAL PROFILES DELIVERED THEIR INDICATED AMOUNTS AND WERE VERIFIED IN THE DAILY AND SUMMARY HISTORY SCREENS THE UNITS LEFT AT THE INSULIN PUMP DISPLAY MATCHED PROPERLY THE UNITS LEFT AT TEST RESERVOIR NO DELIVERY ANOMALY BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING THE AUTO MODE FEATURE IS OPERATING PROPERLY | Operational Problem Identified | Medical Evaluation result code
7664139 | 117766318 | N | DEVICE PASSED BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST NO MOTOR ERROR ALARM NOTED MOTOR PASSED MOTOR TEST NO MOTOR POSITION ENCODER ERROR ALARM IN HISTORY FILE NOTED | No Device Problem Found | Medical Evaluation result code
7712047 | 117834444 | D | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP STATING THE CAUSE OF THE BURNING SENSATION WAS NOT DETERMINED THE SURGEON WANTED TO SCHEDULE THE PATIENT FOR REVISION | Insufficient Information | Device problem code
7712047 | 117834444 | D | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP STATING THE CAUSE OF THE BURNING SENSATION WAS NOT DETERMINED THE SURGEON WANTED TO SCHEDULE THE PATIENT FOR REVISION | Burning Sensation | Patient problem code
7798782 | 117856485 | N | NO PATIENT SAMPLE OR CUSTOMER PRODUCT WAS RETURNED TO IMMUCOR FOR IMMUCOR INVESTIGATION IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE TEST WELL IMAGE IN QUESTION WHICH APPEARED AS VISUALLY POSITIVE THE CAMERA REPORT WAS OPTIMAL THERE WERE NO ASSOCIATED ERRORS OR WARNINGS ON THE EVENT LOG AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION ON JUL THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Connection Problem | Device problem code
7798782 | 117856485 | N | NO PATIENT SAMPLE OR CUSTOMER PRODUCT WAS RETURNED TO IMMUCOR FOR IMMUCOR INVESTIGATION IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON JUL TO ASSESS THE TEST WELL IMAGE IN QUESTION WHICH APPEARED AS VISUALLY POSITIVE THE CAMERA REPORT WAS OPTIMAL THERE WERE NO ASSOCIATED ERRORS OR WARNINGS ON THE EVENT LOG AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE TO ASSESS THE TESTING INSTRUMENT IN QUESTION ON JUL THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Interoperability Problem Identified | Medical Evaluation result code
7802308 | 117857936 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED AS WEAK POSITIVE THE CAMERA IMAGES WERE ALIGNED AND FOCUSED AND THE INSTRUMENT EVENT LOG SHOWED NO ERRORS DURING TESTING AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON JUL TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Connection Problem | Device problem code
7802308 | 117857936 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION WHICH APPEARED AS WEAK POSITIVE THE CAMERA IMAGES WERE ALIGNED AND FOCUSED AND THE INSTRUMENT EVENT LOG SHOWED NO ERRORS DURING TESTING AN IMMUCOR FIELD SERVICE ENGINEER FSE VISITED THE CUSTOMER SITE ON JUL TO ASSESS THE TESTING INSTRUMENT IN QUESTION THE ECHO INSTRUMENT IS PERFORMING AS EXPECTED WITH THE EXCEPTION OF THE CAMERA MODULE ABILITY TO APPROPRIATELY IDENTIFY SOME TRUE POSITIVE RESULTS COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT THE IMMUCOR INTERNAL REPORT RECORD NUMBER FOR THIS REPORT IS B | Interoperability Problem Identified | Medical Evaluation result code
7696952 | 117862910 | N | CORRECTION DEVICE MANUFACTURE DATE UPDATED TO PROPER VALUE THE SUSPECTED POWER RELAY WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT THE RETURNED HARDWARE FAILED BENCH TESTING ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Electrical /Electronic Property Problem | Device problem code
7696952 | 117862910 | N | CORRECTION DEVICE MANUFACTURE DATE UPDATED TO PROPER VALUE THE SUSPECTED POWER RELAY WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT THE RETURNED HARDWARE FAILED BENCH TESTING ANALYSIS FOUND THAT THE REPORTED EVENT WAS RELATED TO A ELECTRICAL ISSUE | Energy Storage System Problem | Medical Evaluation result code
7663061 | 117872009 | N | NO FROZEN SCREEN OR BUTTON ERROR ALARM NOTED UNIT TIMECLOCK MOVED DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED CREASED ACT AND UP BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED | Component or Accessory Incompatibility | Device Problem Code
7663061 | 117872009 | N | NO FROZEN SCREEN OR BUTTON ERROR ALARM NOTED UNIT TIMECLOCK MOVED DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED CREASED ACT AND UP BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED | No Display / Image | Device problem code
7663061 | 117872009 | N | NO FROZEN SCREEN OR BUTTON ERROR ALARM NOTED UNIT TIMECLOCK MOVED DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED CREASED ACT AND UP BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED | Interoperability Problem Identified | Medical Evaluation result code
7664121 | 117872700 | N | DEVICE PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND REWIND TEST NO MOTOR ERROR ALARM NOTED MOTOR PASSED MOTOR TEST NO BATTERY POWER LOSS ALARM NOTED THE BATTERY ICONS FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
7673064 | 117874206 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST HOWEVER INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR INSULIN PUMP RECEIVED WITH SERIAL NUMBER LABEL FADING | Electrical Power Problem | Device problem code
7673064 | 117874206 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST HOWEVER INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR INSULIN PUMP RECEIVED WITH SERIAL NUMBER LABEL FADING | Computer Software Problem | Device problem code
7673064 | 117874206 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST HOWEVER INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR INSULIN PUMP RECEIVED WITH SERIAL NUMBER LABEL FADING | Software Problem Identified | Medical Evaluation result code
7700029 | 117877894 | N | ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS EXPERIENCING PAIN AS THE BATTERY PACK FROM HER SCS WAS STICKING OUT AND WAS CATCHING ON THINGS THE PATIENT UNDERWENT A REVISION PROCEDURE AND WAS DOING WELL POSTOPERATIVELY | No Device Problem Found | Medical Evaluation result code
7720955 | 117879249 | N | THE RECIPIENT WAS REPORTEDLY REIMPLANTED THE RECIPIENT IS DOING WELL THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS SEVERED NEAR THE FANTAIL PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED AN ELECTRICAL TEST FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Jam | Device problem code
7720955 | 117879249 | N | THE RECIPIENT WAS REPORTEDLY REIMPLANTED THE RECIPIENT IS DOING WELL THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS SEVERED NEAR THE FANTAIL PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED AN ELECTRICAL TEST FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Problem Identified | Medical Evaluation result code
7649785 | 117880994 | N | B ONE EMPTY ML WATER BOTTLE WAS RECEIVED FOR EVALUATION IT CAME USED WITH THE HUMIDIFIER ADAPTOR ATTACHED TO THE BOTTLE UPON VISUAL INSPECTION IT WAS OBSERVED THERE IS A CHANNEL FROM THE BOTTOM OF THE ADAPTOR PORT TO THE TOP OF THE BOTTLE THE WATER BOTTLE WAS FILLED FULL WITH WATER THE WATER BOTTLEADAPTOR ASSEMBLY WAS APPLIED TO A FLOW METER THE FLOW METER WAS SET TO LPM AND NO BUBBLES WERE OBSERVED IN THE BOTTLE UPON INCREASING THE AIR FLOW BUBBLES APPEARED IN THE BOTTLE THE COMPLAINT IS CONFIRMED A NONCONFORMANCE WAS OPENED TO ADDRESS THIS ISSUE | No Device Problem Found | Medical Evaluation result code
7710768 | 117997281 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE TO THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Disconnection | Device problem code
7710768 | 117997281 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE TO THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
7710768 | 117997281 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE TO THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Loss of Power | Device problem code
7710768 | 117997281 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE TO THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Audible Alarm | Device problem code
7710768 | 117997281 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE TO THE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON THE ELECTRICAL BOARD THE DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Loss of Power | Device Problem Code
7657473 | 117997308 | N | PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD NO AUDIOBEEP ANOMALY NOTED UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | No Display / Image | Device Problem Code
7657473 | 117997308 | N | PUMP RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD NO AUDIOBEEP ANOMALY NOTED UNABLE TO PERFORM ANY TESTS DUE TO BLANK DISPLAY | Image Artifact | Medical Evaluation result code
7659726 | 118001029 | N | PUMP PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST NO BATTERY POWER LOSS ALARM NOTED | No Audible Alarm | Device problem code
7659726 | 118001029 | N | PUMP PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST NO BATTERY POWER LOSS ALARM NOTED | Electrical Power Problem | Device problem code
7659726 | 118001029 | N | PUMP PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST NO BATTERY POWER LOSS ALARM NOTED | Loss of Power | Device problem code
7659726 | 118001029 | N | PUMP PASSED IDLE CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST NO BATTERY POWER LOSS ALARM NOTED | Loss of Power | Device Problem Code
7670978 | 118001213 | N | DEVICE RECEIVED WITH INTERMITTENT KEYPAD UNRESPONSIVE DUE TO PEELING KEYPAD OVERLAY UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO BUTTON UNRESPONSIVE DEVICE RECEIVED WITH FADED SERIAL NUMBER LABEL | Component or Accessory Incompatibility | Device problem code
7670978 | 118001213 | N | DEVICE RECEIVED WITH INTERMITTENT KEYPAD UNRESPONSIVE DUE TO PEELING KEYPAD OVERLAY UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO BUTTON UNRESPONSIVE DEVICE RECEIVED WITH FADED SERIAL NUMBER LABEL | Peeled / Delaminated | Device problem code
7670978 | 118001213 | N | DEVICE RECEIVED WITH INTERMITTENT KEYPAD UNRESPONSIVE DUE TO PEELING KEYPAD OVERLAY UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO BUTTON UNRESPONSIVE DEVICE RECEIVED WITH FADED SERIAL NUMBER LABEL | Interoperability Problem Identified | Medical Evaluation result code
7663153 | 118003960 | N | NO UNEXPECTED INSULIN FLOW BLOCKED ALARMS NOTED DURING TESTING UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELFTEST AND DISPLACEMENT ACCURACY TEST PUMP PASSED THE DELIVERY ACCURACY TEST NO UNEXPECTED NO CALIBRATION OCCURRED MESSAGES OR BLOOD GLUCOSE REQUIRED MESSAGES NOTED DURING TESTING THE CORRECT TEST BLOOD GLUCOSE VALUE WAS SENT FROM GLUCOSE METER AND RECEIVED BY UNIT UNDER TEST UNIT COMMUNICATED PROPERLY WITH BLOOD GLUCOSE METER MONITORED WITH THE UNIT COMMUNICATING PROPERLY WITH SENSOR TESTER AND THE SENSOR TRANSMITTER THE AUTO MODE FEATURE ACTIVATED PROPERLY NO AUTO MODE ANOMALIES NOTED DURING TESTING | Partial Blockage | Device problem code
7663153 | 118003960 | N | NO UNEXPECTED INSULIN FLOW BLOCKED ALARMS NOTED DURING TESTING UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELFTEST AND DISPLACEMENT ACCURACY TEST PUMP PASSED THE DELIVERY ACCURACY TEST NO UNEXPECTED NO CALIBRATION OCCURRED MESSAGES OR BLOOD GLUCOSE REQUIRED MESSAGES NOTED DURING TESTING THE CORRECT TEST BLOOD GLUCOSE VALUE WAS SENT FROM GLUCOSE METER AND RECEIVED BY UNIT UNDER TEST UNIT COMMUNICATED PROPERLY WITH BLOOD GLUCOSE METER MONITORED WITH THE UNIT COMMUNICATING PROPERLY WITH SENSOR TESTER AND THE SENSOR TRANSMITTER THE AUTO MODE FEATURE ACTIVATED PROPERLY NO AUTO MODE ANOMALIES NOTED DURING TESTING | Calibration Problem | Device problem code
7663153 | 118003960 | N | NO UNEXPECTED INSULIN FLOW BLOCKED ALARMS NOTED DURING TESTING UNIT PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELFTEST AND DISPLACEMENT ACCURACY TEST PUMP PASSED THE DELIVERY ACCURACY TEST NO UNEXPECTED NO CALIBRATION OCCURRED MESSAGES OR BLOOD GLUCOSE REQUIRED MESSAGES NOTED DURING TESTING THE CORRECT TEST BLOOD GLUCOSE VALUE WAS SENT FROM GLUCOSE METER AND RECEIVED BY UNIT UNDER TEST UNIT COMMUNICATED PROPERLY WITH BLOOD GLUCOSE METER MONITORED WITH THE UNIT COMMUNICATING PROPERLY WITH SENSOR TESTER AND THE SENSOR TRANSMITTER THE AUTO MODE FEATURE ACTIVATED PROPERLY NO AUTO MODE ANOMALIES NOTED DURING TESTING | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7757300 | 118005104 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Fall | Patient problem code
7757300 | 118005104 | D | INDEX SURGERY B NONREVISION OF OPTETRAK LOGIC KNEE COMPONENTS DUE TO PATIENT FALL THE CASE REPORT FORM INDICATES THESE EVENTS ARE DEFINITELY NOT RELATED TO DEVICES AND TO PROCEDURE THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES | Insufficient Information | Device problem code
7709571 | 118010344 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT TEST ACTIVE CURRENT MEASUREMENT TEST AND SELF TEST HOWEVER INSULIN PUMP ALARMED INSULIN FLOW BLOCKED DURING THE SEAT TEST DUE TO OUT OF SPECIFICATION FORCE SENSOR ZERO OFFSET MV UNABLE TO VERIFY DRIVEMOTOR ANOMALY AND PERFORM THE DISPLACEMENT REWIND BASIC OCCLUSION FORCE SENSOR AND OCCLUSION TEST DUE TO UNEXPECTED INSULIN FLOW BLOCKED ALARM THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7668951 | 118011825 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT TIME MOVED NO FROZEN SCREEN BLANK DISPLAY OR AUDIOBEEP ANOMALY NOTED ALL BUTTONS FUNCTIONING PROPERLY NO MOISTUREDAMAGEUNLOCKED LCD KEYPAD CONNECTOR DURING VISUAL INSPECTION NOTED UNIT WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED RESTART ERROR ALARM OR RESET TIMEDATE ANOMALY AFTER CHANGE BATTERY NOTED DURING TESTING | No Display / Image | Device Problem Code
7668951 | 118011825 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT TIME MOVED NO FROZEN SCREEN BLANK DISPLAY OR AUDIOBEEP ANOMALY NOTED ALL BUTTONS FUNCTIONING PROPERLY NO MOISTUREDAMAGEUNLOCKED LCD KEYPAD CONNECTOR DURING VISUAL INSPECTION NOTED UNIT WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED RESTART ERROR ALARM OR RESET TIMEDATE ANOMALY AFTER CHANGE BATTERY NOTED DURING TESTING | Inappropriate or Unexpected Reset | Device Problem Code
7668951 | 118011825 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION UNIT TIME MOVED NO FROZEN SCREEN BLANK DISPLAY OR AUDIOBEEP ANOMALY NOTED ALL BUTTONS FUNCTIONING PROPERLY NO MOISTUREDAMAGEUNLOCKED LCD KEYPAD CONNECTOR DURING VISUAL INSPECTION NOTED UNIT WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED RESTART ERROR ALARM OR RESET TIMEDATE ANOMALY AFTER CHANGE BATTERY NOTED DURING TESTING | Incorrect Dimension | Medical Evaluation result code
7670291 | 118011858 | N | THE PUMP ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK PUMP PASSED THE DISPLACEMENT TEST UNABLE TO CONFIRM MOTOR ERROR ALARM OR PERFORM THE PRIME TEST OCCLUSION TEST AND NO DELIVERY TEST DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7670291 | 118011858 | N | THE PUMP ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK PUMP PASSED THE DISPLACEMENT TEST UNABLE TO CONFIRM MOTOR ERROR ALARM OR PERFORM THE PRIME TEST OCCLUSION TEST AND NO DELIVERY TEST DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7670291 | 118011858 | N | THE PUMP ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK PUMP PASSED THE DISPLACEMENT TEST UNABLE TO CONFIRM MOTOR ERROR ALARM OR PERFORM THE PRIME TEST OCCLUSION TEST AND NO DELIVERY TEST DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Component or Accessory Incompatibility | Device Problem Code
7670291 | 118011858 | N | THE PUMP ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK PUMP PASSED THE DISPLACEMENT TEST UNABLE TO CONFIRM MOTOR ERROR ALARM OR PERFORM THE PRIME TEST OCCLUSION TEST AND NO DELIVERY TEST DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Interoperability Problem Identified | Medical Evaluation result code
7719511 | 118033526 | N | REPORTER ADDRESS WAS CORRECTED METHOD IN ADDITION TO METHOD AND RETAINED SAMPLES FROM THE SAME LOT WERE ALSO INSPECTED AND NO DEFECT WAS FOUND ALL INDICATIONS SHOW THE M ELECTROSURGICAL PATIENT PLATE IN QUESTION MET ALL PRODUCT SPECIFICATIONS WITH NO DEFECTS END OF REPORT | No Device Problem Found | Medical Evaluation result code
7812056 | 118090226 | D | ON B THE LAY USER PATIENT CONTACTED LIFESCAN LFS ALLEGING HER ONETOUCH ULTRA METER WAS READING INACCURATELY HIGH IN COMPARISON TO FEELINGSNORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT STATED THAT THE ALLEGED PRODUCT ISSUE BEGAN ON B SHE REPORTED THAT SHE OBTAINED AN ALLEGED INACCURATELY HIGH RESULT OF AND MGDL ON THE SUBJECT METER COMPARED TO FEELINGSNORMAL RESULTS METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES WITH ORAL MEDICATION GLIPIZIDE MG IN THE MORNING AND METFORMIN IN THE EVENING DIET AND EXERCISE AND REPORTED THAT SHE EXERCISED IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT STATED THAT ON AN UNSPECIFIED TIME AFTER THE ALLEGED PRODUCT ISSUE BEGAN SHE DEVELOPED THE SYMPTOMS OF FELT DIZZY COULD HARDLY WALK AND FELT HER SUGAR WAS LOW THE PATIENT REPORTED THAT SHE SELFTREATED WITH FOOD AND ORDRINK ATE SOMETHING WITH SUGAR AT THE TIME OF TROUBLE SHOOTING THE CCA CONFIRMED THAT THE SUBJECT METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS THAT WERE USED WERE STORED CORRECTLY AND HAD NOT EXPIRED THE PATIENT DID NOT HAVE CONTROL SOLUTION TO PERFORM A TEST REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Insufficient Information | Device problem code
7812056 | 118090226 | D | ON B THE LAY USER PATIENT CONTACTED LIFESCAN LFS ALLEGING HER ONETOUCH ULTRA METER WAS READING INACCURATELY HIGH IN COMPARISON TO FEELINGSNORMAL RESULTS THE COMPLAINT WAS CLASSIFIED BASED CUSTOMER CARE ADVOCATE CCA DOCUMENTATION THE PATIENT STATED THAT THE ALLEGED PRODUCT ISSUE BEGAN ON B SHE REPORTED THAT SHE OBTAINED AN ALLEGED INACCURATELY HIGH RESULT OF AND MGDL ON THE SUBJECT METER COMPARED TO FEELINGSNORMAL RESULTS METER TO FEELINGSNORMAL RESULTS COMPARISONS DO NOT MEET THE CRITERIA NECESSARY FOR LFS TO DETERMINE THE POSSIBILITY OF INACCURACY THE PATIENT MANAGES HER DIABETES WITH ORAL MEDICATION GLIPIZIDE MG IN THE MORNING AND METFORMIN IN THE EVENING DIET AND EXERCISE AND REPORTED THAT SHE EXERCISED IN RESPONSE TO THE ALLEGED PRODUCT ISSUE THE PATIENT STATED THAT ON AN UNSPECIFIED TIME AFTER THE ALLEGED PRODUCT ISSUE BEGAN SHE DEVELOPED THE SYMPTOMS OF FELT DIZZY COULD HARDLY WALK AND FELT HER SUGAR WAS LOW THE PATIENT REPORTED THAT SHE SELFTREATED WITH FOOD AND ORDRINK ATE SOMETHING WITH SUGAR AT THE TIME OF TROUBLE SHOOTING THE CCA CONFIRMED THAT THE SUBJECT METER WAS SET TO THE CORRECT UNIT OF MEASURE AND THE TEST STRIPS THAT WERE USED WERE STORED CORRECTLY AND HAD NOT EXPIRED THE PATIENT DID NOT HAVE CONTROL SOLUTION TO PERFORM A TEST REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT THIS COMPLAINT IS BEING REPORTED BECAUSE THE PATIENT REPORTEDLY DEVELOPED SIGNS ANDOR SYMPTOMS THAT THAT MEET LFS CRITERIA FOR A SERIOUS INJURY ADVERSE EVENT AFTER THE ALLEGED PRODUCT ISSUE BEGAN | Injury | Patient problem code
7667372 | 118102501 | N | UNIT PASSED THE UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED RESTART ERROR ALARM OR RESET TIMEDATE ANOMALY AFTER CHANGE BATTERY NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7681934 | 118128400 | N | VOLUNTARY MW AND USER FACILITY REPORT NUMBER WAS RECEIVED B THE FORM STATED THAT THE PATIENTS CARDIOMEMS PILLOW WAS SMOKING HOWEVER IT HAS BEEN CONFIRMED BY THE PATIENT THAT THE SMOKE WAS COMING FROM THE UNIT ITSELF | Smoking | Device problem code
7681934 | 118128400 | N | VOLUNTARY MW AND USER FACILITY REPORT NUMBER WAS RECEIVED B THE FORM STATED THAT THE PATIENTS CARDIOMEMS PILLOW WAS SMOKING HOWEVER IT HAS BEEN CONFIRMED BY THE PATIENT THAT THE SMOKE WAS COMING FROM THE UNIT ITSELF | Non-Functional Defect | Medical Evaluation result code
7662239 | 118130030 | N | INSULIN PUMP PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENTPARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED ANY ISSUED ON SCREEN WERE NOTED | Component Missing | Device problem code
7662239 | 118130030 | N | INSULIN PUMP PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENTPARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED ANY ISSUED ON SCREEN WERE NOTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7675990 | 118131136 | N | DEVICE PASSED SELF TEST AND DISPLACEMENT TEST SOFTWARE ERROR DETECTED NOTED IN THE FORMATTED HISTORY FILE DUE TO SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
7663801 | 118132880 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO UNEXPECTED RESTART ERROR ALARM NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7670691 | 118134453 | N | DEVICE ALARMED MOTOR ERROR DURING BASIC OCCLUSION TEST DUE TO FAULTY FORCED SENSOR READING GOLD UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST OR VERIFY COMPROMISED FORCE SENSOR SYSTEM ALARM OR PRIMEFILL ANOMALY DUE TO MOTOR ERROR ALARM HOWEVER DEVICE PASSED REWIND TEST AND DISPLACEMENT TEST MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7692468 | 118135243 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED UNIT PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST UNIT HAD MISSING END CAP STICKER | No Device Problem Found | Medical Evaluation result code
7670365 | 118135471 | N | PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO PRIMEFILL ANOMALY AND NO MOTOR ERROR ALARM NOTED DURING TEST MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7659874 | 118136350 | N | PUMP PASSED THE OPERATING CURRENTS MEASUREMENT SELF TEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED BATTERY OUT LIMIT ALARMPOWER OFF ALARM ANOMALY OR SPI TO MOTOR PIC COMMUNICATION ERROR ALARM ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7682399 | 118137342 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Positioning Failure | Device problem code
7682399 | 118137342 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Mechanical Jam | Device problem code
7682399 | 118137342 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Component Missing | Device problem code
7682399 | 118137342 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Missing or Inadequate Safety Measures | Medical Evaluation result code
7682399 | 118137342 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Mechanical Problem Identified | Medical Evaluation result code
7670155 | 118137956 | N | DEVICE PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN THE ALARM HISTORY DUE TO HISTORY FILE OVERWRITTEN NO WEAK BATTERY ALARM NOTED | Battery Problem | Device problem code
7670155 | 118137956 | N | DEVICE PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN THE ALARM HISTORY DUE TO HISTORY FILE OVERWRITTEN NO WEAK BATTERY ALARM NOTED | Mechanical Jam | Device Problem Code
7670155 | 118137956 | N | DEVICE PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN THE ALARM HISTORY DUE TO HISTORY FILE OVERWRITTEN NO WEAK BATTERY ALARM NOTED | Component or Accessory Incompatibility | Device Problem Code
7670155 | 118137956 | N | DEVICE PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING UNABLE TO VERIFY MOTOR POSITION ENCODER ERROR ALARM IN THE ALARM HISTORY DUE TO HISTORY FILE OVERWRITTEN NO WEAK BATTERY ALARM NOTED | Interoperability Problem Identified | Medical Evaluation result code
7675837 | 118138917 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Component Missing | Device problem code
7675837 | 118138917 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Mechanical Jam | Device problem code
7675837 | 118138917 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O RING | Mechanical Problem Identified | Medical Evaluation result code
7657975 | 118138928 | N | DEVICE PASSED THE DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY OR OCCLUSION ALARM NOTED NO MOTOR ERROR ALARM NOTED MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7672035 | 118139355 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST HARDWARE LOW LEVEL FAILURES ERROR WAS PRESENT IN THE DEVICE TRACE DOWNLOAD PROBLEM ISOLATED TO KEYPAD ASSEMBLY TOGGLED ON AND OFF THE AUDIO FEATURE WITH THE AUDIO FEATURE FUNCTIONING PROPERLY DURING TESTING | No Device Problem Found | Medical Evaluation result code
7706219 | 118140168 | N | DEVICE RECEIVED WITH MISSING RETAINER RESERVOIR WAS UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7706219 | 118140168 | N | DEVICE RECEIVED WITH MISSING RETAINER RESERVOIR WAS UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7680713 | 118140368 | N | DEVICE RECEIVED WITH MISSING RETAINER AND RESERVOIR TUBE ORING DEVICE WAS NOT TESTED FOR DISPLACEMENT TEST DUE TO MISSING RETAINER RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7680713 | 118140368 | N | DEVICE RECEIVED WITH MISSING RETAINER AND RESERVOIR TUBE ORING DEVICE WAS NOT TESTED FOR DISPLACEMENT TEST DUE TO MISSING RETAINER RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7676110 | 118143410 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7676110 | 118143410 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7676110 | 118143410 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7668910 | 118143445 | N | UNIT POWER UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED UNIT WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED OFF NO POWER LOW BATTERY WEAK BATTERY BATTERY OUT LIMIT AND FAILED BATTERY TEST ALARMS NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7657542 | 118143850 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NO UNLOCKED LCD KEYPAD CONNECTOR THE DEVICE PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENT OR PARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED FLASHING ANOMALIES NOTED | Component or Accessory Incompatibility | Device problem code
7657542 | 118143850 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NO UNLOCKED LCD KEYPAD CONNECTOR THE DEVICE PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENT OR PARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED FLASHING ANOMALIES NOTED | Component Missing | Device problem code
7657542 | 118143850 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NO UNLOCKED LCD KEYPAD CONNECTOR THE DEVICE PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENT OR PARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED FLASHING ANOMALIES NOTED | Display or Visual Feedback Problem | Device Problem Code
7657542 | 118143850 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH NO UNLOCKED LCD KEYPAD CONNECTOR THE DEVICE PASSED THE DISPLACEMENT AND SELF TEST NO UNEXPECTED MISSING SEGMENT OR PARTIAL DISPLAY ANOMALIES NOTED NO UNEXPECTED FLASHING ANOMALIES NOTED | Display or Visual Feedback Problem | Device problem code
7686498 | 118144487 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7686498 | 118144487 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7686498 | 118144487 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7686498 | 118144487 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7662930 | 118145048 | N | PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY PRIMEFILL ANOMALY DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Component or Accessory Incompatibility | Device Problem Code
7662930 | 118145048 | N | PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY PRIMEFILL ANOMALY DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7662930 | 118145048 | N | PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY PRIMEFILL ANOMALY DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7662930 | 118145048 | N | PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY PRIMEFILL ANOMALY DUE TO THE COMPROMISED FORCE SENSOR SYSTEM ALARM | Operational Problem Identified | Medical Evaluation result code
7670251 | 118145709 | N | PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM DURING TEST | No Device Problem Found | Medical Evaluation result code
7813284 | 118146212 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENTS MOTHER NOTED THE POD FELL OF WHILE HER SON WAS SWIMMING | Device Fell | Device problem code
7813284 | 118146212 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE LEVELS EXCEEDED MMOLL MGDL WHILE WEARING THE POD BETWEEN AND HOURS ON THE LEG THE PATIENTS MOTHER NOTED THE POD FELL OF WHILE HER SON WAS SWIMMING | No Known Impact Or Consequence To Patient | Patient Problem Code
7813378 | 118146632 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DESPITE MULTIPLE BOLUSES OF UNITS AND UNITS RESPECTIVELY BG DID NOT IMPROVE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | No Known Impact Or Consequence To Patient | Patient Problem Code
7813378 | 118146632 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DESPITE MULTIPLE BOLUSES OF UNITS AND UNITS RESPECTIVELY BG DID NOT IMPROVE PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE LEG | Positioning Problem | Device problem code
7636637 | 118148405 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Component or Accessory Incompatibility | Device Problem Code
7636637 | 118148405 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7636637 | 118148405 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7636637 | 118148405 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM | No Device Problem Found | Medical Evaluation result code
7675371 | 118152986 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP AND RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE ORING | Mechanical Jam | Device problem code
7675371 | 118152986 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP AND RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
7675371 | 118152986 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP AND RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE ORING | Positioning Failure | Device problem code
7675371 | 118152986 | N | UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE PCAP AND RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MISSING RESERVOIR TUBE ORING | Mechanical Problem Identified | Medical Evaluation result code
7668255 | 118153552 | N | PUMP RECEIVED WITH NO BUTTON RESPONSE AT UP ARROW BUTTON DUE TO UNLOCKED CONNECTOR ON THE LCD BOARD NO BUTTON ERROR ALARM FROZEN SCREEN AND TIME CLOCK ANOMALY NOTED DURING TESTING | Device Alarm System | Device problem code
7668255 | 118153552 | N | PUMP RECEIVED WITH NO BUTTON RESPONSE AT UP ARROW BUTTON DUE TO UNLOCKED CONNECTOR ON THE LCD BOARD NO BUTTON ERROR ALARM FROZEN SCREEN AND TIME CLOCK ANOMALY NOTED DURING TESTING | No Display / Image | Device problem code
7668255 | 118153552 | N | PUMP RECEIVED WITH NO BUTTON RESPONSE AT UP ARROW BUTTON DUE TO UNLOCKED CONNECTOR ON THE LCD BOARD NO BUTTON ERROR ALARM FROZEN SCREEN AND TIME CLOCK ANOMALY NOTED DURING TESTING | Component or Accessory Incompatibility | Device Problem Code
7668255 | 118153552 | N | PUMP RECEIVED WITH NO BUTTON RESPONSE AT UP ARROW BUTTON DUE TO UNLOCKED CONNECTOR ON THE LCD BOARD NO BUTTON ERROR ALARM FROZEN SCREEN AND TIME CLOCK ANOMALY NOTED DURING TESTING | Interoperability Problem Identified | Medical Evaluation result code
7668905 | 118153763 | N | PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7664241 | 118154403 | N | PUMP PASSED THE DISPLACEMENT TEST HOWEVER ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE PRIME TEST AND MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7812106 | 118157757 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT INSULIN PUMP DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM BLOOD GLUCOSE WAS UNKNOWN TROUBLESHOOTING WAS PERFORMED THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY ALERT OR REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING ADVISED THE INSULIN PUMP WILL NEED TO BE REPLACED ADVISED TO DISCONTINUE USE OF THE PUMP AND REVERT TO A BACKUP PLAN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | Battery Problem | Device problem code
7812106 | 118157757 | D | THE CUSTOMER REPORTED VIA PHONE CALL THAT INSULIN PUMP DID NOT RECEIVE A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM BLOOD GLUCOSE WAS UNKNOWN TROUBLESHOOTING WAS PERFORMED THIS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY ALERT OR REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING ADVISED THE INSULIN PUMP WILL NEED TO BE REPLACED ADVISED TO DISCONTINUE USE OF THE PUMP AND REVERT TO A BACKUP PLAN THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS | No Known Impact Or Consequence To Patient | Patient problem code
7814047 | 118159360 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A HEAT RASH ON HER BACK THE PATIENT WAS PRESCRIBED CORTISONE CREAM BY HER PHYSICIAN THE PATIENT REPORTED THAT THE RASH IMPROVED | Insufficient Information | Device problem code
7814047 | 118159360 | D | A US DISTRIBUTOR REPORTED THAT THE PATIENT HAD DEVELOPED A HEAT RASH ON HER BACK THE PATIENT WAS PRESCRIBED CORTISONE CREAM BY HER PHYSICIAN THE PATIENT REPORTED THAT THE RASH IMPROVED | Rash | Patient problem code
7814761 | 118200376 | N | B LABELING INDICATES INSERTING THE SENSOR AND WEARING THE ADHESIVE PATCH MIGHT CAUSE INFECTION BLEEDING PAIN OR SKIN IRRITATIONS EG REDNESS SWELLING BRUISING ITCHING SCARRING OR SKIN DISCOLORATION | Insufficient Information | Device problem code
7814574 | 118200935 | N | ON B THE EVALUATION WAS COMPLETED FOR PARTTTNSE LOT SERIAL B THE EVALUATION SHOWED THAT PRIOR TO THE EVENT THE TIP HAD BEEN USED FOR TREATMENTS THE TIP PASSED FLOW AND THERMISTOR TESTING NO LEAK OR FUNCTIONAL TEST WAS PERFORMED AND VISUAL TESTING FAILED DUE TO DIELECTRIC BREAKDOWN BEING PRESENT ADDITIONAL INVESTIGATION RESULTS ARE PENDING | Results Pending Completion of Investigation | Medical Evaluation result code
7814574 | 118200935 | N | ON B THE EVALUATION WAS COMPLETED FOR PARTTTNSE LOT SERIAL B THE EVALUATION SHOWED THAT PRIOR TO THE EVENT THE TIP HAD BEEN USED FOR TREATMENTS THE TIP PASSED FLOW AND THERMISTOR TESTING NO LEAK OR FUNCTIONAL TEST WAS PERFORMED AND VISUAL TESTING FAILED DUE TO DIELECTRIC BREAKDOWN BEING PRESENT ADDITIONAL INVESTIGATION RESULTS ARE PENDING | Leakage/Seal | Medical Evaluation result code
7806886 | 118254839 | N | A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THAT THE SCREWS WHICH SECURE THE FUNNEL TO THE PIPE BECAME LOOSE SUBSEQUENTLY ALLOWING WATER TO LEAK OUT ONTO THE FLOOR THE TECHNICIAN RESECURED THE SCREWS TESTED THE STERILIZER AND CONFIRMED IT TO BE OPERATING ACCORDING TO SPECIFICATIONS THE STERILIZER WAS MANUFACTURED IN AND IS NOT UNDER STERIS SERVICE AGREEMENT FOR MAINTENANCE ACTIVITIES THE USER FACILITY IS RESPONSIBLE FOR MAINTENANCE ACTIVITIES THE TECHNICIAN COUNSELED THE USER FACILITY ON PROPERLY MAINTAINING THEIR STERILIZER AND NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Leak / Splash | Device problem code
7806886 | 118254839 | N | A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THAT THE SCREWS WHICH SECURE THE FUNNEL TO THE PIPE BECAME LOOSE SUBSEQUENTLY ALLOWING WATER TO LEAK OUT ONTO THE FLOOR THE TECHNICIAN RESECURED THE SCREWS TESTED THE STERILIZER AND CONFIRMED IT TO BE OPERATING ACCORDING TO SPECIFICATIONS THE STERILIZER WAS MANUFACTURED IN AND IS NOT UNDER STERIS SERVICE AGREEMENT FOR MAINTENANCE ACTIVITIES THE USER FACILITY IS RESPONSIBLE FOR MAINTENANCE ACTIVITIES THE TECHNICIAN COUNSELED THE USER FACILITY ON PROPERLY MAINTAINING THEIR STERILIZER AND NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Detachment of Device or device Component | Device Problem Code
7806886 | 118254839 | N | A STERIS SERVICE TECHNICIAN ARRIVED ONSITE TO INSPECT THE STERILIZER AND IDENTIFIED THAT THE SCREWS WHICH SECURE THE FUNNEL TO THE PIPE BECAME LOOSE SUBSEQUENTLY ALLOWING WATER TO LEAK OUT ONTO THE FLOOR THE TECHNICIAN RESECURED THE SCREWS TESTED THE STERILIZER AND CONFIRMED IT TO BE OPERATING ACCORDING TO SPECIFICATIONS THE STERILIZER WAS MANUFACTURED IN AND IS NOT UNDER STERIS SERVICE AGREEMENT FOR MAINTENANCE ACTIVITIES THE USER FACILITY IS RESPONSIBLE FOR MAINTENANCE ACTIVITIES THE TECHNICIAN COUNSELED THE USER FACILITY ON PROPERLY MAINTAINING THEIR STERILIZER AND NO ADDITIONAL ISSUES HAVE BEEN REPORTED | Leakage/Seal | Medical Evaluation result code
7657061 | 118281131 | N | DEVICE PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER DEVICE ALSO HAD A MISSING RESERVOIR TUBE ORING | Mechanical Jam | Device problem code
7657061 | 118281131 | N | DEVICE PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER DEVICE ALSO HAD A MISSING RESERVOIR TUBE ORING | Positioning Failure | Device problem code
7657061 | 118281131 | N | DEVICE PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER DEVICE ALSO HAD A MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
7657061 | 118281131 | N | DEVICE PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER DEVICE ALSO HAD A MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
7657061 | 118281131 | N | DEVICE PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER DEVICE ALSO HAD A MISSING RESERVOIR TUBE ORING | Mechanical Problem Identified | Medical Evaluation result code
7690369 | 118281979 | N | ADDITIONAL INFORMATION THE POE INJECTOR WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT DURING ALL TESTING AND FINAL RPI TESTING UNIT STABLE FULLY FUNCTIONAL NO INTERMITTENT FAULTS OCCURRED UNIT RAN ON TESTER OVERNIGHT AND NOW GREEN POWER LED IS RED POWER WAS UNPLUGGED AND REPLUGGED AND LED STAYED RED ELECTRICAL ISSUE CONFIRMED THIS ISSUE WAS DOCUMENTED IN A MEDTRONIC NAVIGATION HARDWARE ANOMALY TRACKING DATABASE | Electrical /Electronic Property Problem | Device problem code
7690369 | 118281979 | N | ADDITIONAL INFORMATION THE POE INJECTOR WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ANALYSIS FOUND THAT DURING ALL TESTING AND FINAL RPI TESTING UNIT STABLE FULLY FUNCTIONAL NO INTERMITTENT FAULTS OCCURRED UNIT RAN ON TESTER OVERNIGHT AND NOW GREEN POWER LED IS RED POWER WAS UNPLUGGED AND REPLUGGED AND LED STAYED RED ELECTRICAL ISSUE CONFIRMED THIS ISSUE WAS DOCUMENTED IN A MEDTRONIC NAVIGATION HARDWARE ANOMALY TRACKING DATABASE | Energy Storage System Problem | Medical Evaluation result code
7672792 | 118288097 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7672792 | 118288097 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7672792 | 118288097 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7815374 | 118288747 | N | B INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR PASSED PER SPECIFICATION ALSO PERFORMED BICARBONATE BUFFER TEST SENSOR PASSED PER SPECIFICATIONS WITH ACCURATE READINGS ALSO PERFORMED VISUAL INSPECTION AND CHECKED INSERTION NEEDLE FOR BURRSHOOKS AND DULL NEEDLE TIP DEFECTS AND NONE WAS FOUND INSERTION NEEDLE PASSED PER SPECIFICATIONS | Dull, Blunt | Device problem code
7815374 | 118288747 | N | B INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST SENSOR PASSED PER SPECIFICATION ALSO PERFORMED BICARBONATE BUFFER TEST SENSOR PASSED PER SPECIFICATIONS WITH ACCURATE READINGS ALSO PERFORMED VISUAL INSPECTION AND CHECKED INSERTION NEEDLE FOR BURRSHOOKS AND DULL NEEDLE TIP DEFECTS AND NONE WAS FOUND INSERTION NEEDLE PASSED PER SPECIFICATIONS | Stress Problem Identified | Medical Evaluation result code
7815398 | 118289091 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE NO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY AUDIOVIBRATE ANOMALY DUE TO BUTTONS NOT RESPONDING | Failure to Deliver | Device problem code
7815398 | 118289091 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE NO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY AUDIOVIBRATE ANOMALY DUE TO BUTTONS NOT RESPONDING | Component or Accessory Incompatibility | Device problem code
7815398 | 118289091 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE NO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY AUDIOVIBRATE ANOMALY DUE TO BUTTONS NOT RESPONDING | Inappropriate or Unexpected Reset | Device Problem Code
7815398 | 118289091 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE NO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY AUDIOVIBRATE ANOMALY DUE TO BUTTONS NOT RESPONDING | Inaccurate Delivery | Device Problem Code
7815398 | 118289091 | N | THE INSULIN PUMP WAS RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE NO UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY AUDIOVIBRATE ANOMALY DUE TO BUTTONS NOT RESPONDING | Interoperability Problem Identified | Medical Evaluation result code
7676302 | 118290802 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7676302 | 118290802 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7676302 | 118290802 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7676302 | 118290802 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7674687 | 118291974 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Disconnection | Device problem code
7674687 | 118291974 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
7674687 | 118291974 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN V FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
7723139 | 118292264 | N | USER FACILITY ONLY ONE PROCESSOR PCA WAS RETURNED FOR FAILURE ANALYSIS THE REPORTED PROBLEM WAS DUPLICATED DURING TESTING IN THE LAB J PINS WERE FOUND TO BE LIFTED AND COMPLETELY DETACHED FROM THE PCA RESULTING TO DEFIB TEST FAILURE DURING OPERATIONAL CHECK | Detachment of Device or device Component | Device problem code
7723139 | 118292264 | N | USER FACILITY ONLY ONE PROCESSOR PCA WAS RETURNED FOR FAILURE ANALYSIS THE REPORTED PROBLEM WAS DUPLICATED DURING TESTING IN THE LAB J PINS WERE FOUND TO BE LIFTED AND COMPLETELY DETACHED FROM THE PCA RESULTING TO DEFIB TEST FAILURE DURING OPERATIONAL CHECK | Interoperability Problem Identified | Medical Evaluation result code
7669880 | 118292986 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORINGRESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINERUNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7669880 | 118292986 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORINGRESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINERUNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7669880 | 118292986 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORINGRESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINERUNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7669880 | 118292986 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORINGRESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINERUNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7672668 | 118294917 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7672668 | 118294917 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7672668 | 118294917 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7687120 | 118299090 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7687120 | 118299090 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7687120 | 118299090 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7676305 | 118306969 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7676305 | 118306969 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7676305 | 118306969 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7684219 | 118309644 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7684219 | 118309644 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7675884 | 118309694 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7675884 | 118309694 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7675884 | 118309694 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7675884 | 118309694 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Aspiration | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Disability | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Fever | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Infection | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Injury | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Pain | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Purulent Discharge | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Swelling | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Toxicity | Patient problem code
7218108 | 118309747 | D | PFS AND MEDICAL RECORDS RECEIVED PFS ALLEGES INJURY HIP PAIN DISABILITY AND INFECTION AFTER REVIEW OF MEDICAL RECORDS THE PATIENT WAS REVISED TO ADDRESS INFECTION AND IMPLANTATION OF ANTIBIOTIC SPACER WITH STAGED REVISION AT A LATER TIME PATIENT UNDERWENT AN ASPIRATION AND SINCE THEN EXPERIENCED FEVER PAIN AND SWELLING REVISION NOTES REPORTED A LARGE AMOUNT OF PURULENT FLUID WHICH WAS POSITIVE FOR GRAMPOSITIVE COCCI AND HAD NUMEROUS WHITE BLOOD CELLS THERE WERE LARGE AMOUNT OF METALLOSIS DEBRIS ENCOUNTERED AN EXTENDED TROCHANTERIC OSTEOTOMY WAS NEEDED TO REMOVE THE STEM METAL ION LEVELS ARE ABOVE PPB ADDED CUP SINCE IT WAS REVISED INFECTION WAS CURED AND NEW HIP PROSTHESIS WAS IMPLANTED ON B DOI B DOR B LEFT HIP | Appropriate Term/Code Not Available | Device Problem Code
7693430 | 118311219 | N | INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST INSULIN PUMP REWINDS PROPERLY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED HOWEVER PUMP UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER AND RESERVOIR TUBE ORING | Component Missing | Device problem code
7693430 | 118311219 | N | INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST INSULIN PUMP REWINDS PROPERLY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED HOWEVER PUMP UNABLE TO PERFORM THE DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER AND RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
7680769 | 118311430 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7680769 | 118311430 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7680769 | 118311430 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7675985 | 118311544 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7675985 | 118311544 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7675985 | 118311544 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7680659 | 118312015 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7680659 | 118312015 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7680659 | 118312015 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7680989 | 118313970 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7680989 | 118313970 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7680989 | 118313970 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7692997 | 118314300 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7692997 | 118314300 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7692997 | 118314300 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7680946 | 118315988 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7680946 | 118315988 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7680946 | 118315988 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7676918 | 118316183 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7676918 | 118316183 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7676918 | 118316183 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7676918 | 118316183 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7680663 | 118317506 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7680663 | 118317506 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7680663 | 118317506 | N | DEVICE HAD A MISSING RETAINER AND RESERVOIR TUBE O RING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7693814 | 118317917 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR O RING RESERVOIR RETAINER UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7693814 | 118317917 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR O RING RESERVOIR RETAINER UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7693814 | 118317917 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR O RING RESERVOIR RETAINER UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7693814 | 118317917 | N | DEVICE ARRIVED WITH MISSING RETAINER AND RESERVOIR O RING RESERVOIR RETAINER UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7679865 | 118318967 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7679865 | 118318967 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7679865 | 118318967 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7679865 | 118318967 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR UNABLE TO LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7694228 | 118319625 | N | ONE PROCESSOR PCA WAS RETURNED FOR FAILURE ANALYSIS THE REPORTED PROBLEM WAS DUPLICATED DURING TESTING IN THE LAB J PINS WERE FOUND TO BE LIFTED AND COMPLETELY DETACHED FROM THE PCA RESULTING IN DEFIB TEST FAILURE DURING OPERATIONAL CHECK | Detachment of Device or device Component | Device problem code
7694228 | 118319625 | N | ONE PROCESSOR PCA WAS RETURNED FOR FAILURE ANALYSIS THE REPORTED PROBLEM WAS DUPLICATED DURING TESTING IN THE LAB J PINS WERE FOUND TO BE LIFTED AND COMPLETELY DETACHED FROM THE PCA RESULTING IN DEFIB TEST FAILURE DURING OPERATIONAL CHECK | Interoperability Problem Identified | Medical Evaluation result code
7822043 | 118396886 | D | INDEX SURGERY REVISION DUE TO PATIENT FALL | Fall | Patient problem code
7822043 | 118396886 | D | INDEX SURGERY REVISION DUE TO PATIENT FALL | Insufficient Information | Device problem code
7822044 | 118397380 | D | INDEX SURGERY REVISION DUE TO PATIENT FALL | Insufficient Information | Device problem code
7822044 | 118397380 | D | INDEX SURGERY REVISION DUE TO PATIENT FALL | Fall | Patient problem code
7814340 | 118403601 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT A MANUAL INJECTION WAS DELIVERED AND NEW POD WAS APPLIED THE POD WAS DISCARDED | No Known Impact Or Consequence To Patient | Patient Problem Code
7814340 | 118403601 | D | IT WAS REPORTED THAT PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE AS TREATMENT A MANUAL INJECTION WAS DELIVERED AND NEW POD WAS APPLIED THE POD WAS DISCARDED | Positioning Problem | Device problem code
7703697 | 118601870 | N | THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT ALLEGEDLY LIT A CIGARETTE WHILE USING AN EVERFLO OXYGEN CONCENTRATOR THE PATIENT RECEIVED FIRST DEGREE BURNS TO HER CHEST AND ABDOMEN THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR FIRST DEGREE BURNS AND COPD EXACERBATION RELATED TO SMOKE INHALATION THE DURABLE MEDICAL EQUIPMENT DME CONFIRMED THAT THE DEVICE WILL NOT BE RETURNING FOR INVESTIGATION | Smoking | Device problem code
7703697 | 118601870 | N | THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT ALLEGEDLY LIT A CIGARETTE WHILE USING AN EVERFLO OXYGEN CONCENTRATOR THE PATIENT RECEIVED FIRST DEGREE BURNS TO HER CHEST AND ABDOMEN THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR FIRST DEGREE BURNS AND COPD EXACERBATION RELATED TO SMOKE INHALATION THE DURABLE MEDICAL EQUIPMENT DME CONFIRMED THAT THE DEVICE WILL NOT BE RETURNING FOR INVESTIGATION | No Device Problem Found | Medical Evaluation result code
7836760 | 118886464 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS MOTHER FELT AROUND THE POD AND SHE FELT THE CANNULA POKE HER THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN AND THE ADHESIVE WAS PULLING AWAY THE PATIENT HAD MODERATE KETONES AND NOW THEY ARE BACK TO NORMAL AS TREATMENT A NEW POD WAS APPLIED AND A SHOT OF INSULIN WAS ADMINISTERED | No Known Impact Or Consequence To Patient | Patient Problem Code
7836760 | 118886464 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENTS MOTHER FELT AROUND THE POD AND SHE FELT THE CANNULA POKE HER THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE ABDOMEN AND THE ADHESIVE WAS PULLING AWAY THE PATIENT HAD MODERATE KETONES AND NOW THEY ARE BACK TO NORMAL AS TREATMENT A NEW POD WAS APPLIED AND A SHOT OF INSULIN WAS ADMINISTERED | Positioning Problem | Device problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Ambulation Difficulties | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Arthralgia | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Distress | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Pain | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Scar Tissue | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Swelling | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Toxicity | Patient problem code
6740868 | 118924535 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Insufficient Information | Device Problem Code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Ambulation Difficulties | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Arthralgia | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Distress | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Pain | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Scar Tissue | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Swelling | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Toxicity | Patient problem code
6740767 | 118924874 | D | ASR LITIGATION RECORDS RECEIVED LITIGATION ALLEGES JOINT PAIN SWELLING DIFFICULTY WALKING PULLED MUSCLE SCAR TISSUE INABILITY TO HEAL WITHOUT SECOND SURGERY EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN THE BLOOD AND EMOTIONAL DISTRESS | Insufficient Information | Device Problem Code
7658422 | 119094942 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7658422 | 119094942 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7659601 | 119232682 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND ALARM NOTED IN DATA A SAFETY FEATURE NOT CONSIDERED A MALFUNCTION | Component or Accessory Incompatibility | Device problem code
7659601 | 119232682 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND ALARM NOTED IN DATA A SAFETY FEATURE NOT CONSIDERED A MALFUNCTION | Incompatible Component/ Accessory | Medical Evaluation result code
7846616 | 119237740 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD MORE THAN HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPSBUTTOCKS FOR TREATMENT PATIENT GAVE A BOLUS AND INCREASED BASAL RATE | No Known Impact Or Consequence To Patient | Patient Problem Code
7846616 | 119237740 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD MORE THAN HOURS DUE TO ELEVATED BG LEVELS PATIENT WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE HIPSBUTTOCKS FOR TREATMENT PATIENT GAVE A BOLUS AND INCREASED BASAL RATE | Positioning Problem | Device problem code
7670373 | 119243369 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND THE ADHESIVE PAD WAS RETURNED ON THE DEVICE AND NO PROBLEMS WERE FOUND THAT COULD CAUSE THE ADHESIVE TO NOT STICK | Component or Accessory Incompatibility | Device problem code
7670373 | 119243369 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND THE ADHESIVE PAD WAS RETURNED ON THE DEVICE AND NO PROBLEMS WERE FOUND THAT COULD CAUSE THE ADHESIVE TO NOT STICK | Incompatible Component/ Accessory | Medical Evaluation result code
7848578 | 119318665 | N | MISSING LOT INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT A FOLLOW UP REPORT WILL BE SUBMITTED IF THE INFORMATION BECOMES AVAILABLE B | Component Missing | Device problem code
7848578 | 119318665 | N | MISSING LOT INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT A FOLLOW UP REPORT WILL BE SUBMITTED IF THE INFORMATION BECOMES AVAILABLE B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7848056 | 119324795 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED PLEASE EXPLAIN WHAT IS MEANT BY TISSUE WAS CUT ALL THE WAY THROUGH BEFORE STAPLING WAS COMPLETED WAS THE STAPLE LINE AND CUT LINE EQUAL IN LENGTH THE STAPLE LINE WAS NORMAL LENGTH PAST CUT LINE WERE ANY ISSUES NOTE WITH STAPLE FORM NO WHAT WAS THE PRODUCT CODE OF RELOAD USED WAS THE DEVICE DIFFICULT TO CLOSE OR FIRE IT SEEMED TO BE A BIT EASIER ON THE FIRST STROKE OUT OF ON THE FIRST RELOAD FIRED AFTER THAT THE FOLLOWING STROKES SEEMED NORMAL WHEN WAS THE LEAK IDENTIFIED WHAT WAS THE LOCATION OF THE ST STAPLE LINE ON SIGMOID COLON WAS ANY TISSUE MOVEMENT NOTICED DURING FIRING NO WAS THE PROCEDURE COMPLETED SUCCESSFULLY WITH SAME DEVICE YES BUT THE PATIENT WAS BROUGHT BACK LATER DOES THE SURGEON WAIT SECONDS PRIOR TO FIRING YES HOW LONG AFTER INITIAL PROCEDURE WAS LEAK DETECTED NOT POSITIVE BUT I THINK IT WAS A COUPLE OF DAYS WHY DOES SURGEON BELIEVE POSTOP LEAK WAS RELATED TO ISSUE OBSERVED DURING INITIAL SURGERY BECAUSE THE LEAK OCCURRED AT THE LOCATION OF THE FIRST STROKE OF THE FIRST FIRING WHAT IS THE CURRENT PATIENT STATUS UNKNOWN BUT I MET WITH SURGEON LAST WEEK AND I GOT THE IMPRESSION THE PATIENT WAS RECOVERING FINE | Difficult to Open or Close | Device problem code
7848056 | 119324795 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED PLEASE EXPLAIN WHAT IS MEANT BY TISSUE WAS CUT ALL THE WAY THROUGH BEFORE STAPLING WAS COMPLETED WAS THE STAPLE LINE AND CUT LINE EQUAL IN LENGTH THE STAPLE LINE WAS NORMAL LENGTH PAST CUT LINE WERE ANY ISSUES NOTE WITH STAPLE FORM NO WHAT WAS THE PRODUCT CODE OF RELOAD USED WAS THE DEVICE DIFFICULT TO CLOSE OR FIRE IT SEEMED TO BE A BIT EASIER ON THE FIRST STROKE OUT OF ON THE FIRST RELOAD FIRED AFTER THAT THE FOLLOWING STROKES SEEMED NORMAL WHEN WAS THE LEAK IDENTIFIED WHAT WAS THE LOCATION OF THE ST STAPLE LINE ON SIGMOID COLON WAS ANY TISSUE MOVEMENT NOTICED DURING FIRING NO WAS THE PROCEDURE COMPLETED SUCCESSFULLY WITH SAME DEVICE YES BUT THE PATIENT WAS BROUGHT BACK LATER DOES THE SURGEON WAIT SECONDS PRIOR TO FIRING YES HOW LONG AFTER INITIAL PROCEDURE WAS LEAK DETECTED NOT POSITIVE BUT I THINK IT WAS A COUPLE OF DAYS WHY DOES SURGEON BELIEVE POSTOP LEAK WAS RELATED TO ISSUE OBSERVED DURING INITIAL SURGERY BECAUSE THE LEAK OCCURRED AT THE LOCATION OF THE FIRST STROKE OF THE FIRST FIRING WHAT IS THE CURRENT PATIENT STATUS UNKNOWN BUT I MET WITH SURGEON LAST WEEK AND I GOT THE IMPRESSION THE PATIENT WAS RECOVERING FINE | Leak / Splash | Device problem code
7848056 | 119324795 | N | B BATCH UNK WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT THEREFORE WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NONCONFORMANCE ADDITIONAL INFORMATION REQUESTED AND RECEIVED PLEASE EXPLAIN WHAT IS MEANT BY TISSUE WAS CUT ALL THE WAY THROUGH BEFORE STAPLING WAS COMPLETED WAS THE STAPLE LINE AND CUT LINE EQUAL IN LENGTH THE STAPLE LINE WAS NORMAL LENGTH PAST CUT LINE WERE ANY ISSUES NOTE WITH STAPLE FORM NO WHAT WAS THE PRODUCT CODE OF RELOAD USED WAS THE DEVICE DIFFICULT TO CLOSE OR FIRE IT SEEMED TO BE A BIT EASIER ON THE FIRST STROKE OUT OF ON THE FIRST RELOAD FIRED AFTER THAT THE FOLLOWING STROKES SEEMED NORMAL WHEN WAS THE LEAK IDENTIFIED WHAT WAS THE LOCATION OF THE ST STAPLE LINE ON SIGMOID COLON WAS ANY TISSUE MOVEMENT NOTICED DURING FIRING NO WAS THE PROCEDURE COMPLETED SUCCESSFULLY WITH SAME DEVICE YES BUT THE PATIENT WAS BROUGHT BACK LATER DOES THE SURGEON WAIT SECONDS PRIOR TO FIRING YES HOW LONG AFTER INITIAL PROCEDURE WAS LEAK DETECTED NOT POSITIVE BUT I THINK IT WAS A COUPLE OF DAYS WHY DOES SURGEON BELIEVE POSTOP LEAK WAS RELATED TO ISSUE OBSERVED DURING INITIAL SURGERY BECAUSE THE LEAK OCCURRED AT THE LOCATION OF THE FIRST STROKE OF THE FIRST FIRING WHAT IS THE CURRENT PATIENT STATUS UNKNOWN BUT I MET WITH SURGEON LAST WEEK AND I GOT THE IMPRESSION THE PATIENT WAS RECOVERING FINE | Leakage/Seal | Medical Evaluation result code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Difficult to Remove | Device Problem Code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Misfire | Device problem code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Difficult to Open or Close | Device problem code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Fire | Device problem code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Leak / Splash | Device problem code
7850331 | 119345484 | N | B BATCH UNK MONTH AND DAY UNKNOWN ONLY YEAR IS KNOWN THE LOTBATCH WAS NOT PROVIDED THEREFORE THE MANUFACTURING RECORDS COULD NOT BE REVIEWED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED WHAT WERE THE INDICATIONS FOR SURGERY DID THE PATIENT RECEIVE ANY PREOP CHEMOTHERAPY OR RADIATION WHAT PURSE STRING TECHNIQUE WAS UTILIZED WAS THE DEVICE DIFFICULT TO CLOSE WAS THE DEVICE DIFFICULT TO FIRE DID THE HEALTHCARE PROFESSIONAL WAIT SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE TACTILE FEEDBACK WHEN FIRING THE DEVICE WAS THERE DIFFICULTY IN REMOVING THE DEVICE WERE THE DONUTS INSPECTED IF SO PLEASE DESCRIBE WAS A COMPLETE WASHER TRANSECTION CONFIRMED WAS A LEAK TEST PERFORMED AT THE END OF THE CASE WHAT DOES THE SURGEON BELIEVE CAUSED THE LEAK WHAT IS THE RELATIONSHIP TO THE STAPLER HOW MANY DAYS POSTOP DID THE LEAK OCCUR HOW WAS THE LEAK IDENTIFIED HOW WAS THE LEAK ADDRESSED WERE ANY STAPLE FORMATION ISSUES IDENTIFIED AT ANY POINT IN PATIENT CARE WHAT IS THE CURRENT PATIENT STATUS RESPONSE I DO NOT HAVE ANY MORE INFORMATION ABOUT THIS COMPLAINT | Leakage/Seal | Medical Evaluation result code
7850742 | 119373667 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF LUMP REDNESS AND SWELLING IN THE LEFT CHEEK RIGHT CHEEK AND CHIN WERE ALSO GETTING RED AND SWOLLEN AND NODULE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TABLE AND TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY AND DURATION OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Migration or Expulsion of Device | Device problem code
7850742 | 119373667 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF LUMP REDNESS AND SWELLING IN THE LEFT CHEEK RIGHT CHEEK AND CHIN WERE ALSO GETTING RED AND SWOLLEN AND NODULE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TABLE AND TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY AND DURATION OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leak / Splash | Device problem code
7850742 | 119373667 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF LUMP REDNESS AND SWELLING IN THE LEFT CHEEK RIGHT CHEEK AND CHIN WERE ALSO GETTING RED AND SWOLLEN AND NODULE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TABLE AND TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY AND DURATION OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Material Discolored | Device problem code
7850742 | 119373667 | N | FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME THE EVENTS OF LUMP REDNESS AND SWELLING IN THE LEFT CHEEK RIGHT CHEEK AND CHIN WERE ALSO GETTING RED AND SWOLLEN AND NODULE ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS DEVICE LABELING WARNINGS TREATMENT SITE REACTIONS CONSIST MAINLY OF SHORTTERM INFLAMMATORY SYMPTOMS AND GENERALLY RESOLVE WITHIN TO WEEKS REFER TO THE ADVERSE EVENTS SECTION FOR DETAILS PRECAUTIONS AS WITH ALL TRANSCUTANEOUS PROCEDURES DERMAL FILLER IMPLANTATION CARRIES A RISK OF INFECTION STANDARD PRECAUTIONS ASSOCIATED WITH INJECTABLE MATERIALS SHOULD BE FOLLOWED PATIENTS MAY EXPERIENCE LATE ONSET NODULES WITH USE OF DERMAL FILLERS INCLUDING JUVÉDERM VOLUMA XC REFER TO ADVERSE EVENTS SECTION FOR DETAILS ADVERSE EVENTS TABLE AND TREATMENT SITE RESPONSES BY MAXIMUM SEVERITY AND DURATION OCCURRING IN OF SUBJECTS AFTER INITIAL TREATMENT N POSSIBLE TREATMENT SITE RESPONSES INCLUDE TENDERNESS SWELLING FIRMNESS LUMPSBUMPS BRUISING PAIN REDNESS DISCOLORATION AND ITCHING C OTHER SAFETY DATA POSTMARKET SURVEILLANCE JUVÉDERM VOLUMA WITHOUT LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND JUVÉDERM VOLUMA XC ALSO KNOWN AS JUVÉDERM VOLUMA WITH LIDOCAINE HAS BEEN MARKETED OUTSIDE THE US SINCE AND IN THE US SINCE THE FOLLOWING AES WERE RECEIVED FROM POSTMARKET SURVEILLANCE FOR JUVÉDERM VOLUMA WITH AND WITHOUT LIDOCAINE WITH A FREQUENCY OF EVENTS OR MORE AND WERE NOT OBSERVED IN THE CLINICAL STUDY THIS INCLUDES REPORTS RECEIVED GLOBALLY FROM ALL SOURCES INCLUDING SCIENTIFIC JOURNALS AND VOLUNTARY REPORTS ALL AES OBTAINED THROUGH POSTMARKET SURVEILLANCE ARE LISTED IN ORDER OF NUMBER OF REPORTS RECEIVED INFLAMMATORY REACTION LACK OF CORRECTION INFECTION MIGRATION ALLERGIC REACTION ABSCESS PARESTHESIA VASCULAR OCCLUSION DRAINAGE NECROSIS VISION ABNORMALITIES MALAISE SCARRING NAUSEA GRANULOMA DEEPER WRINKLE AND DYSPNEA REPORTED TREATMENTS INCLUDE ANTIBIOTICS STEROIDS ANTISEPTIC CREAMS HYALURONIDASE ANTIINFLAMMATORIES ANTIHISTAMINES NEEDLE ASPIRATION EYE DROPS RADIO FREQUENCY THERAPY HYPERBARIC OXYGEN TREATMENT LASER TREATMENT ICE MASSAGE WARM COMPRESS ANALGESICS ANTIVIRALS ULTRASOUND THERAPY EXCISION DRAINAGE AND SURGERY | Leakage/Seal | Medical Evaluation result code
7848578 | 119390054 | N | MISSING LOT INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT A FOLLOW UP REPORT WILL BE SUBMITTED IF THE INFORMATION BECOMES AVAILABLE B | Component Missing | Device problem code
7848578 | 119390054 | N | MISSING LOT INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT A FOLLOW UP REPORT WILL BE SUBMITTED IF THE INFORMATION BECOMES AVAILABLE B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7859395 | 119643923 | N | THE PHYSICIAN REPORTED THAT HE DID NOT FEEL THAT THE DEVICE LED TO THE INFECTION AND THAT IT WAS THE INCISION ITSELF THAT CAUSED THE INFECTION HOWEVER BECAUSE THE DEVICE MAY HAVE CONTRIBUTED TO THE INFECTION THE MANUFACTURER IS REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE MANUFACTURER WAS NOT MADE AWARE IN A TIMELY MANNER NO ANALYSIS OR TESTING CAN BE DONE CURRENT DEVICE LABELING ADDRESSES THE REPORTED EVENT AS FOLLOWS POSSIBLE COMPLICATIONS INCLUDE INFECTION SEROMA PAIN SCAFFOLD MIGRATION WOUND DEHISCENCE HEMORRHAGE ADHESIONS HEMATOMA INFLAMMATION EXTRUSION AND RECURRENCE OF THE SOFT TISSUE DEFECT | No Device Problem Found | Medical Evaluation result code
7859395 | 119643923 | N | THE PHYSICIAN REPORTED THAT HE DID NOT FEEL THAT THE DEVICE LED TO THE INFECTION AND THAT IT WAS THE INCISION ITSELF THAT CAUSED THE INFECTION HOWEVER BECAUSE THE DEVICE MAY HAVE CONTRIBUTED TO THE INFECTION THE MANUFACTURER IS REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE MANUFACTURER WAS NOT MADE AWARE IN A TIMELY MANNER NO ANALYSIS OR TESTING CAN BE DONE CURRENT DEVICE LABELING ADDRESSES THE REPORTED EVENT AS FOLLOWS POSSIBLE COMPLICATIONS INCLUDE INFECTION SEROMA PAIN SCAFFOLD MIGRATION WOUND DEHISCENCE HEMORRHAGE ADHESIONS HEMATOMA INFLAMMATION EXTRUSION AND RECURRENCE OF THE SOFT TISSUE DEFECT | No Device Problem Found | Medical Evaluation result code
7683038 | 119709448 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | No Device Problem Found | Medical Evaluation result code
7674352 | 119731830 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA DISLODGING FROM THE INFUSION SITE OR CAUSE HIGH BLOOD GLUCOSE LEVELS WAS FOUND | Defective Device | Device problem code
7674352 | 119731830 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA DISLODGING FROM THE INFUSION SITE OR CAUSE HIGH BLOOD GLUCOSE LEVELS WAS FOUND | Device Difficult to Operate | Medical Evaluation result code
7679978 | 119845656 | N | B DEXCOM WAS MADE AWARE ON B THAT ON B THE PATIENT EXPERIENCED A LOSS OF CONNECTION | Connection Problem | Device problem code
7679978 | 119845656 | N | B DEXCOM WAS MADE AWARE ON B THAT ON B THE PATIENT EXPERIENCED A LOSS OF CONNECTION | Interoperability Problem Identified | Medical Evaluation result code
7649275 | 119868659 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7649275 | 119868659 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7649275 | 119868659 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7649275 | 119868659 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7670923 | 119976366 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7670923 | 119976366 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Interoperability Problem Identified | Medical Evaluation result code
7694278 | 119976647 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7694278 | 119976647 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Interoperability Problem Identified | Medical Evaluation result code
7644699 | 119985489 | N | COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | Device Alarm System | Device problem code
7644699 | 119985489 | N | COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | Mechanical Problem Identified | Medical Evaluation result code
7870614 | 120019367 | N | THE PHYSICIAN REPORTED THAT HE DID NOT FEEL THAT THE DEVICE LED TO THE INFECTION AND THAT IT WAS THE INCISION ITSELF THAT CAUSED THE INFECTION HOWEVER BECAUSE THE DEVICE MAY HAVE CONTRIBUTED TO THE INFECTION THE MANUFACTURER IS REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE MANUFACTURER WAS NOT MADE AWARE IN A TIMELY MANNER NO ANALYSIS OR TESTING CAN BE DONE CURRENT DEVICE LABELING ADDRESSES THE REPORTED EVENT AS FOLLOWS POSSIBLE COMPLICATIONS INCLUDE INFECTION SEROMA PAIN SCAFFOLD MIGRATION WOUND DEHISCENCE HEMORRHAGE ADHESIONS HEMATOMA INFLAMMATION EXTRUSION AND RECURRENCE OF THE SOFT TISSUE DEFECT | Defective Device | Device problem code
7870614 | 120019367 | N | THE PHYSICIAN REPORTED THAT HE DID NOT FEEL THAT THE DEVICE LED TO THE INFECTION AND THAT IT WAS THE INCISION ITSELF THAT CAUSED THE INFECTION HOWEVER BECAUSE THE DEVICE MAY HAVE CONTRIBUTED TO THE INFECTION THE MANUFACTURER IS REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE MANUFACTURER WAS NOT MADE AWARE IN A TIMELY MANNER NO ANALYSIS OR TESTING CAN BE DONE CURRENT DEVICE LABELING ADDRESSES THE REPORTED EVENT AS FOLLOWS POSSIBLE COMPLICATIONS INCLUDE INFECTION SEROMA PAIN SCAFFOLD MIGRATION WOUND DEHISCENCE HEMORRHAGE ADHESIONS HEMATOMA INFLAMMATION EXTRUSION AND RECURRENCE OF THE SOFT TISSUE DEFECT | Migration or Expulsion of Device | Device problem code
7870614 | 120019367 | N | THE PHYSICIAN REPORTED THAT HE DID NOT FEEL THAT THE DEVICE LED TO THE INFECTION AND THAT IT WAS THE INCISION ITSELF THAT CAUSED THE INFECTION HOWEVER BECAUSE THE DEVICE MAY HAVE CONTRIBUTED TO THE INFECTION THE MANUFACTURER IS REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE MANUFACTURER WAS NOT MADE AWARE IN A TIMELY MANNER NO ANALYSIS OR TESTING CAN BE DONE CURRENT DEVICE LABELING ADDRESSES THE REPORTED EVENT AS FOLLOWS POSSIBLE COMPLICATIONS INCLUDE INFECTION SEROMA PAIN SCAFFOLD MIGRATION WOUND DEHISCENCE HEMORRHAGE ADHESIONS HEMATOMA INFLAMMATION EXTRUSION AND RECURRENCE OF THE SOFT TISSUE DEFECT | Device Migration | Medical Evaluation result code
7703522 | 120139990 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7703522 | 120139990 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA FAILING TO INSERT CORRECTLY OR THE PUMP FAILING TO DELIVER INSULIN WAS FOUND | Interoperability Problem Identified | Medical Evaluation result code
7705823 | 120140031 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7705823 | 120140031 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Interoperability Problem Identified | Medical Evaluation result code
7699699 | 120140095 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7699699 | 120140095 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Interoperability Problem Identified | Medical Evaluation result code
7659276 | 120141686 | N | DATE SENT TO FDA B PATIENT CODE B IT WAS REPORTED THAT FOLLOWING INSERTION PATIENT EXPERIENCED SMALL BOWEL OBSTRUCTION AND UMBILICAL HERNIA NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7729565 | 120147607 | N | DEVICE PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST AT INCHES | No Device Problem Found | Medical Evaluation result code
7874416 | 120283904 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE INTERMITTENT TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | No Known Impact Or Consequence To Patient | Patient problem code
7874416 | 120283904 | D | IT WAS REPORTED THAT THE PUMP DECLARED MULTIPLE INTERMITTENT TEMPERATURE ALARMS THE PUMP WAS NOT EXPOSED TO ABNORMAL TEMPERATURE CONDITIONS THE CUSTOMERS BLOOD GLUCOSE LEVEL WAS MGDL REPORTEDLY THE CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY | Device Alarm System | Device problem code
7877868 | 120321751 | N | DEVICE EVALUATION OF MONITOR SN B AND BELT SN B HAS BEEN COMPLETED THE REPORTED PROBLEM PATIENT DEATH WAS INVESTIGATED AS RECEIVED THE MONITOR AND BELT PASSED INCOMING FUNCTIONAL TESTING DURING THE INCOMING FUNCTIONAL TESTING A HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES FOLLOWED BY A HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM DURING THE TRANSITION TO THE HZ SIGNAL THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM AUDIO MESSAGING AND SIREN ALARMS AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY J BIPHASIC PULSES THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY ECG ACQUISITION DETECTION ALGORITHM PERFORMANCE AND PULSE DELIVERY FUNCTIONALITY | No Device Problem Found | Medical Evaluation result code
7607745 | 120326627 | N | DEVICE PASSED THE DISPLACEMENT UNEXPECTED RESTART ERROR TEST AND SELF TEST NO UNEXPECTED RESTART ERROR ALARM NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | No Display / Image | Device Problem Code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | Failure to Deliver | Device problem code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | Scratched Material | Device problem code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | Inaccurate Delivery | Device Problem Code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7879418 | 120375566 | N | AFTER BATTERY INSTALLATION THE INSULIN PUMP DISPLAY START UP AND STOP FROZEN AT NUMBER FOLLOWED BY AN UNEXPECTED CONSTANT AUDIO TONE ALARM PROBLEM WAS ISOLATED TO MB UNABLE TO VERIFY OPERATIONAL CURRENTS OR PERFORM FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST DAT TEST AND DISPLACEMENT TEST DUE TO FROZEN SCREEN UNIT RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE THE INSULIN PUMP RECEIVED WITH STAINED END CAP STICKER AND BACK ADDRESS SERIAL NUMBER LABEL | Stress Problem Identified | Medical Evaluation result code
7876334 | 120395395 | D | TYPE III ENDOLEAK FROM STENT GRAFT ABDOMINAL EXPLORATION WITH PRIMARY CLOSURE OF AAA PT DECOMPENSATED WITH ACTIVE BLEEDING ABDOMEN OPENED AND AAA SAC INSPECTED FOUND TO HAVE BLOOD COMING THROUGH ENDOLOGIX FABRIC | Hemorrhage | Patient problem code
7876334 | 120395395 | D | TYPE III ENDOLEAK FROM STENT GRAFT ABDOMINAL EXPLORATION WITH PRIMARY CLOSURE OF AAA PT DECOMPENSATED WITH ACTIVE BLEEDING ABDOMEN OPENED AND AAA SAC INSPECTED FOUND TO HAVE BLOOD COMING THROUGH ENDOLOGIX FABRIC | Insufficient Information | Device problem code
7699818 | 120415507 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7699818 | 120415507 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7699240 | 120417804 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7699240 | 120417804 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7698926 | 120419840 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7698926 | 120419840 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7698401 | 120449567 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7698401 | 120449567 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7642161 | 120455069 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7642161 | 120455069 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7674913 | 120455096 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7674913 | 120455096 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7699169 | 120470784 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7699169 | 120470784 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7710179 | 120471024 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED INSPECTION OF THE DEVICE DID NOT FIND ANY ISSUES THAT WOULD RESULT IN THE BLEEDING OR BRUISING AT THE INFUSION SITE NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7710179 | 120471024 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED INSPECTION OF THE DEVICE DID NOT FIND ANY ISSUES THAT WOULD RESULT IN THE BLEEDING OR BRUISING AT THE INFUSION SITE NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Shock | Patient problem code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Shock, Cardiogenic | Patient problem code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Insufficient Information | Device problem code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Diarrhea | Patient problem code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Pain | Patient problem code
7883031 | 120478012 | D | IT WAS REPORTED THAT THE PATIENT HAD GASTROINTESTINAL GI BLEED AT TIME OF ADMISSION ON B FOR CARDIOGENIC SHOCK AND MULTISYSTEM ORGAN FAILURE OF UNKNOWN ETIOLOGY THE PATIENT FELT LIGHTHEADED HAD EPIGASTRIC PAIN AND NOTED A POOR APPETITE FOR SEVERAL DAYS ONE EPISODE OF VERY DARK EMESIS SIMILAR TO COFFEE GROUNDS OCCURRED THE MORNING OF B INR WAS AT THE OUTSIDE HOSPITAL THE PATIENT FIRST PRESENTED TO AND AT THE IMPLANTING CENTER AT THE TIME OF THE EVENT THE PATIENT WAS TAKING THE ANTICOAGULANT WARFARIN WHILE IN CARDIAC INTENSIVE CARE UNIT CICU THE PATIENT HAD MULTIPLE EPISODES OF MAROON COLORED HEMOCCULT POSITIVE STOOLS THE PATIENT WAS TRANSFUSED UNITS PACKED RED BLOOD CELLS PRBCS PRIOR TO TRANSFER OUT OF THE INTENSIVE CARE UNIT ICU THE PATIENT WAS ALSO TRANSFUSED U FRESH FROZEN PLASMA FFP AT AN UNSPECIFIED TIME AN ESOPHAGOGASTRODUODENOSCOPY EGD WAS NOT PERFORMED DUE TO THE PATIENTS INTUBATION FOR CRITICAL COMORBIDITIES AND THE POTENTIAL NEED FOR EXTENSIVE LAVAGE ON B THE PATIENT HAD DIARRHEA AND WAS CLOSTRIDIUM DIFFICILE C DIFF POSITIVE ON B THE PATIENT WAS DISCHARGED TO ACUTE INPATIENT REHAB WITH NO CURRENT SIGNS OF GI BLEEDING HEMOGLOBINHEMATOCRIT HH REMAINED STABLE THE EVENT REPORTEDLY RESOLVED ON B AND THE PATIENT WAS DISCHARGED FROM INPATIENT REHAB ON B IT WAS REPORTED THAT THE PATIENT HAD NO FURTHER BLEEDING | Hemorrhage | Patient problem code
7880749 | 120529080 | D | MARLEX MESH IMPLANTED FOR HERNIA REPAIR PT CURRENTLY HAS PAIN SWOLLEN LEFT TESTICLE BURNING SENSATION AND PRESSURE TO THE LOWER EXTREMITIES BLOATING AND TIGHTNESS | No Code Available | Patient problem code
7880749 | 120529080 | D | MARLEX MESH IMPLANTED FOR HERNIA REPAIR PT CURRENTLY HAS PAIN SWOLLEN LEFT TESTICLE BURNING SENSATION AND PRESSURE TO THE LOWER EXTREMITIES BLOATING AND TIGHTNESS | Insufficient Information | Device problem code
7880749 | 120529080 | D | MARLEX MESH IMPLANTED FOR HERNIA REPAIR PT CURRENTLY HAS PAIN SWOLLEN LEFT TESTICLE BURNING SENSATION AND PRESSURE TO THE LOWER EXTREMITIES BLOATING AND TIGHTNESS | Pain | Patient problem code
7880749 | 120529080 | D | MARLEX MESH IMPLANTED FOR HERNIA REPAIR PT CURRENTLY HAS PAIN SWOLLEN LEFT TESTICLE BURNING SENSATION AND PRESSURE TO THE LOWER EXTREMITIES BLOATING AND TIGHTNESS | Burning Sensation | Patient problem code
7880749 | 120529080 | D | MARLEX MESH IMPLANTED FOR HERNIA REPAIR PT CURRENTLY HAS PAIN SWOLLEN LEFT TESTICLE BURNING SENSATION AND PRESSURE TO THE LOWER EXTREMITIES BLOATING AND TIGHTNESS | Swelling | Patient problem code
7888899 | 120709830 | N | ANALYSIS OF THE DATA LOGS WERE PERFORMED IT WAS NOTED THAT THE ESTOP WAS NOT DISENGAGED FOR MINUTES WAS CAUSING THE ERROR MESSAGE THE SYSTEM WAS PERFORMING AS INTENDED | Device Displays Incorrect Message | Device problem code
7888899 | 120709830 | N | ANALYSIS OF THE DATA LOGS WERE PERFORMED IT WAS NOTED THAT THE ESTOP WAS NOT DISENGAGED FOR MINUTES WAS CAUSING THE ERROR MESSAGE THE SYSTEM WAS PERFORMING AS INTENDED | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7889045 | 120714292 | D | OUS MDR IT WAS REPORTED THAT THE LEAD HAD TO BE EXPLANTED DAYS AFTER THE IMPLANTATION DUE TO AN EXIT BLOCK | Insufficient Information | Device problem code
7889045 | 120714292 | D | OUS MDR IT WAS REPORTED THAT THE LEAD HAD TO BE EXPLANTED DAYS AFTER THE IMPLANTATION DUE TO AN EXIT BLOCK | Exit Block | Patient problem code
7896204 | 120964126 | N | THE EVENT OF PATIENT HAS SUFFERED BODILY INJURY AND RESULTING PAIN AND SUFFERING MENTAL ANGUISH DISABILITY DISFIGUREMENT AND LOSS OF THE CAPACITY FOR THE ENJOYMENT OF LIFE THE LOSSES ARE PERMANENT OR CONTINUING IN NATURE AND PATIENT WILL SUFFER THEM IN THE FUTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Deflation Problem | Device problem code
7896204 | 120964126 | N | THE EVENT OF PATIENT HAS SUFFERED BODILY INJURY AND RESULTING PAIN AND SUFFERING MENTAL ANGUISH DISABILITY DISFIGUREMENT AND LOSS OF THE CAPACITY FOR THE ENJOYMENT OF LIFE THE LOSSES ARE PERMANENT OR CONTINUING IN NATURE AND PATIENT WILL SUFFER THEM IN THE FUTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Calcified | Device problem code
7896204 | 120964126 | N | THE EVENT OF PATIENT HAS SUFFERED BODILY INJURY AND RESULTING PAIN AND SUFFERING MENTAL ANGUISH DISABILITY DISFIGUREMENT AND LOSS OF THE CAPACITY FOR THE ENJOYMENT OF LIFE THE LOSSES ARE PERMANENT OR CONTINUING IN NATURE AND PATIENT WILL SUFFER THEM IN THE FUTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT PRODUCT OR PATIENT DETAILS HAS BEEN REQUESTED NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DEVICE LABELING POTENTIAL ADVERSE EVENTS THAT MAY OCCUR WITH SALINEFILLED BREAST IMPLANT SURGERY INCLUDE REOPERATION PAIN WRINKLING ASYMMETRY IMPLANT PALPABILITYVISIBILITY IMPLANT REMOVAL CAPSULAR CONTRACTURE CHANGES IN NIPPLE AND BREAST SENSATION IMPLANT DISPLACEMENTMIGRATION IMPLANT DEFLATION SCARRING INFECTION HEMATOMASEROMA BREASTFEEDING COMPLICATIONS IMPLANT EXTRUSION NECROSIS DELAYED WOUND HEALING BREAST TISSUE ATROPHYCHEST WALL DEFORMITY CALCIUM DEPOSITS AND LYMPHADENOPATHY | Non-Functional Defect | Medical Evaluation result code
7897176 | 120979781 | D | THE NEEDLES ARE BENDING USER STATES THAT THE NEEDLE BEND IS HIGH UP NEAR WERE THE PLASTIC MEETS THE NEEDLE ALSO HE IS GETTING A BURNING SENSATION AFTER INJECTION | Material Twisted / Bent | Device problem code
7897176 | 120979781 | D | THE NEEDLES ARE BENDING USER STATES THAT THE NEEDLE BEND IS HIGH UP NEAR WERE THE PLASTIC MEETS THE NEEDLE ALSO HE IS GETTING A BURNING SENSATION AFTER INJECTION | Burning Sensation | Patient problem code
7813248 | 121035427 | N | ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEAK RATE WAS LESS THAN LPM AND IS THEREFORE NOT A SUFFICIENT LEAK TO REQUIRE MEDICAL INTERVENTION TO PREVENT PATIENT HARM OR AN INABILITY TO MECHANICALLY VENTILATE THIS WAS NOT A REPORTABLE MALFUNCTION ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEAK RATE WAS LESS THAN LPM AND IS THEREFORE NOT A SUFFICIENT LEAK TO REQUIRE MEDICAL INTERVENTION TO PREVENT PATIENT HARM OR AN INABILITY TO MECHANICALLY VENTILATE THIS WAS NOT A REPORTABLE MALFUNCTION | Leak / Splash | Device problem code
7813248 | 121035427 | N | ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEAK RATE WAS LESS THAN LPM AND IS THEREFORE NOT A SUFFICIENT LEAK TO REQUIRE MEDICAL INTERVENTION TO PREVENT PATIENT HARM OR AN INABILITY TO MECHANICALLY VENTILATE THIS WAS NOT A REPORTABLE MALFUNCTION ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEAK RATE WAS LESS THAN LPM AND IS THEREFORE NOT A SUFFICIENT LEAK TO REQUIRE MEDICAL INTERVENTION TO PREVENT PATIENT HARM OR AN INABILITY TO MECHANICALLY VENTILATE THIS WAS NOT A REPORTABLE MALFUNCTION | Leakage/Seal | Medical Evaluation result code
7900009 | 121431521 | N | THE INSULIN PUMP WAS RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST AND DISPLACEMENT TEST PUMP PASSED THE DAT TEST AT INCHES THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | Scratched Material | Device problem code
7900009 | 121431521 | N | THE INSULIN PUMP WAS RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST SELF TEST A ERROR TEST AND DISPLACEMENT TEST PUMP PASSED THE DAT TEST AT INCHES THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW AND CASE | No Device Problem Found | Medical Evaluation result code
7901650 | 121459121 | D | DURING FOLLOWUP EXIT BLOCK WAS OBSERVED ON THE RIGHT VENTRICULAR LEAD THE LEAD WAS EXPLANTED AND REPLACED AND THE PATIENT WAS IN STABLE CONDITION | Exit Block | Patient problem code
7901650 | 121459121 | D | DURING FOLLOWUP EXIT BLOCK WAS OBSERVED ON THE RIGHT VENTRICULAR LEAD THE LEAD WAS EXPLANTED AND REPLACED AND THE PATIENT WAS IN STABLE CONDITION | Insufficient Information | Device problem code
7677749 | 121463050 | N | REVIEW OF THIS EVENT ANDOR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO EVIDENCE TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT OR A SIMILAR DEVICE WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE MALFUNCTION WERE TO RECUR THEREFORE THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN CFR | No Device Problem Found | Medical Evaluation result code
7901939 | 121468875 | D | THE PATIENT REPORTED THAT HIS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS ABDOMEN FOR BETWEEN AND HOURS HE STATED THAT THE CANNULA APPEARED AT ABOUT A DEGREE ANGLE TO THE SIDE WHEN THE POD WAS REMOVED TREATMENT INCLUDED A BOLUS AND CHANGING THE POD B | No Known Impact Or Consequence To Patient | Patient Problem Code
7901939 | 121468875 | D | THE PATIENT REPORTED THAT HIS BLOOD GLUCOSE REACHED MGDL WHILE WEARING THE POD ON HIS ABDOMEN FOR BETWEEN AND HOURS HE STATED THAT THE CANNULA APPEARED AT ABOUT A DEGREE ANGLE TO THE SIDE WHEN THE POD WAS REMOVED TREATMENT INCLUDED A BOLUS AND CHANGING THE POD B | Insufficient Information | Device problem code
7687958 | 121474403 | N | THE INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT THE INSULIN PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY ALARMS NOTED DURING TESTING THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7687958 | 121474403 | N | THE INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT THE INSULIN PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY ALARMS NOTED DURING TESTING THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7687958 | 121474403 | N | THE INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT THE INSULIN PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY ALARMS NOTED DURING TESTING THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7687958 | 121474403 | N | THE INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT THE INSULIN PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED LOW BATTERY ALARMS NOTED DURING TESTING THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7597041 | 121639147 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND SELF TEST NO UNEXPECTED RESTART ERROR ALARM NOTED DURING TEST | Inappropriate or Unexpected Reset | Device Problem Code
7631055 | 121755600 | N | DATE SENT TO FDA B ADDITIONAL NARRATIVE IT WAS REPORTED THAT FOLLOWING PROCEDURE THE PATIENT EXPERIENCED BURNING SENSATION AND PAIN NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
7735602 | 121775150 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN OR TO THE IRRITATION AT THE INFUSION SITE WAS FOUND | Defective Device | Device problem code
7735602 | 121775150 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN OR TO THE IRRITATION AT THE INFUSION SITE WAS FOUND | Device Difficult to Operate | Medical Evaluation result code
7911306 | 121859180 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR LESS THAN AN HOUR IN THE INFUSION SITE PATIENT DID NOT FEEL THE INSERTION AND WHEN THE POD WAS REMOVED THERE WAS NO INSERTION MARK WHICH INDICATES A FAILURE OF THE NEEDLE MECHANISM TO DEPLOY AS INTENDED DURING ACTIVATION | Mechanical Problem | Device problem code
7911306 | 121859180 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS REACHED MGDL WHILE WEARING THE POD FOR LESS THAN AN HOUR IN THE INFUSION SITE PATIENT DID NOT FEEL THE INSERTION AND WHEN THE POD WAS REMOVED THERE WAS NO INSERTION MARK WHICH INDICATES A FAILURE OF THE NEEDLE MECHANISM TO DEPLOY AS INTENDED DURING ACTIVATION | No Known Impact Or Consequence To Patient | Patient Problem Code
7916501 | 121988549 | D | IT WAS REPORTED THAT THE PATIENTS VNS WAS REMOVED SOMETIME IN DUE TO THE LEAD ERODING THROUGH THE SKIN WHICH RESULTED IN AN INFECTION NO ADDITIONAL OR RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE | Material Erosion | Device problem code
7916501 | 121988549 | D | IT WAS REPORTED THAT THE PATIENTS VNS WAS REMOVED SOMETIME IN DUE TO THE LEAD ERODING THROUGH THE SKIN WHICH RESULTED IN AN INFECTION NO ADDITIONAL OR RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE | Infection | Patient problem code
7701505 | 122074684 | N | INSULIN PUMP PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST | Insufficient Information | Device problem code
7701505 | 122074684 | N | INSULIN PUMP PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
7693636 | 122076809 | N | THE INSULIN PUMP WAS RECEIVED WITH NO MOTOR MOVEMENT DURING REWIND DUE TO MOTOR GEARBOX JAMMED UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST AND OCCLUSION TEST DUE TO NO MOTOR MOVEMENT | Mechanical Jam | Device problem code
7693636 | 122076809 | N | THE INSULIN PUMP WAS RECEIVED WITH NO MOTOR MOVEMENT DURING REWIND DUE TO MOTOR GEARBOX JAMMED UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST AND OCCLUSION TEST DUE TO NO MOTOR MOVEMENT | Mechanical Problem Identified | Medical Evaluation result code
7672596 | 122080988 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES GRAY SCREEN ANOMALIES OR DISPLAY ANOMALIES NOTED DURING TESTING DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST DEVICE RECEIVED WITH A MINOR DEEP SCRATCH ON LCD WINDOW | Scratched Material | Device problem code
7672596 | 122080988 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES GRAY SCREEN ANOMALIES OR DISPLAY ANOMALIES NOTED DURING TESTING DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST DEVICE RECEIVED WITH A MINOR DEEP SCRATCH ON LCD WINDOW | Component Missing | Device problem code
7672596 | 122080988 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES GRAY SCREEN ANOMALIES OR DISPLAY ANOMALIES NOTED DURING TESTING DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST DEVICE RECEIVED WITH A MINOR DEEP SCRATCH ON LCD WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
7672596 | 122080988 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES GRAY SCREEN ANOMALIES OR DISPLAY ANOMALIES NOTED DURING TESTING DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST DEVICE RECEIVED WITH A MINOR DEEP SCRATCH ON LCD WINDOW | Stress Problem Identified | Medical Evaluation result code
7705754 | 122088031 | N | DEVICE PASSED REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7705754 | 122088031 | N | DEVICE PASSED REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7705754 | 122088031 | N | DEVICE PASSED REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND OCCLUSION TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7919526 | 122155405 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT ADMINISTERED BOLUS AND NOTICED LEVELS NOT DOING DOWN SO DID SEVEN MORE BOLUSES AND LEVELS WERE STILL RISING THE PATIENT SMELT AND FELT INSULIN AROUND THE POD AND DOES NOT THINK CANNULA WAS IN THE BODY THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE PATIENT ADMINISTERED ANOTHER SHOT AND BG LEVELS IS MGDL | No Known Impact Or Consequence To Patient | Patient Problem Code
7919526 | 122155405 | D | IT WAS REPORTED THAT A PATIENTS BLOOD GLUCOSE BG LEVELS EXCEEDED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT ADMINISTERED BOLUS AND NOTICED LEVELS NOT DOING DOWN SO DID SEVEN MORE BOLUSES AND LEVELS WERE STILL RISING THE PATIENT SMELT AND FELT INSULIN AROUND THE POD AND DOES NOT THINK CANNULA WAS IN THE BODY THUS WAS UNSURE IF CANNULA WAS PROPERLY SEATED IN THE INFUSION SITE PATIENT ADMINISTERED ANOTHER SHOT AND BG LEVELS IS MGDL | Positioning Problem | Device problem code
7701619 | 122198361 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7701619 | 122198361 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7701619 | 122198361 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7708923 | 122202219 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST B | Display or Visual Feedback Problem | Device problem code
7708923 | 122202219 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST B | No Device Problem Found | Medical Evaluation result code
7713076 | 122204453 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | Display or Visual Feedback Problem | Device problem code
7713076 | 122204453 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | No Device Problem Found | Medical Evaluation result code
7713076 | 122204453 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | No Display / Image | Device Problem Code
7713076 | 122204453 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | Image Artifact | Medical Evaluation result code
7662457 | 122207271 | N | DEVICE PASSED THE FUNCTIONAL TESTING INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST B | Insufficient Information | Device Problem Code
7662457 | 122207271 | N | DEVICE PASSED THE FUNCTIONAL TESTING INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST B | No Device Problem Found | Medical Evaluation result code
7692305 | 122211482 | N | INSULIN PUMP ERROR ALARM FOUND IN THE PUMP HISTORY DUE TO SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
7665099 | 122214848 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED VISUAL INSPECTION REVEALED A CUT THROUGH THE LEAD INSULATION THAT WAS INDUCED DURING THE IMPLANTEXPLANT PROCEDURE TESTING WAS COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE AND INNER INSULATION INTEGRITY MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS LABORATORY ANALYSIS DID NOT IDENTIFY ANY LEAD CHARACTERISTICS THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS | Insufficient Information | Device Problem Code
7665099 | 122214848 | N | UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED VISUAL INSPECTION REVEALED A CUT THROUGH THE LEAD INSULATION THAT WAS INDUCED DURING THE IMPLANTEXPLANT PROCEDURE TESTING WAS COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE AND INNER INSULATION INTEGRITY MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS LABORATORY ANALYSIS DID NOT IDENTIFY ANY LEAD CHARACTERISTICS THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS | No Device Problem Found | Medical Evaluation result code
7699430 | 122224382 | N | B THE INSULIN PUMP PASSED THE SELF TEST DISPLACEMENT TEST AND THE DISPLAY FUNCTIONED PROPERLY THERE WERE NO MISSING SEGMENTS NOR PARTIAL DISPLAY ANOMALY DURING TESTING | Display or Visual Feedback Problem | Device problem code
7699430 | 122224382 | N | B THE INSULIN PUMP PASSED THE SELF TEST DISPLACEMENT TEST AND THE DISPLAY FUNCTIONED PROPERLY THERE WERE NO MISSING SEGMENTS NOR PARTIAL DISPLAY ANOMALY DURING TESTING | Component Missing | Device problem code
7699430 | 122224382 | N | B THE INSULIN PUMP PASSED THE SELF TEST DISPLACEMENT TEST AND THE DISPLAY FUNCTIONED PROPERLY THERE WERE NO MISSING SEGMENTS NOR PARTIAL DISPLAY ANOMALY DURING TESTING | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699545 | 122224877 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699545 | 122224877 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699545 | 122224877 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699629 | 122226052 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699629 | 122226052 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699629 | 122226052 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699600 | 122226082 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699600 | 122226082 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7699600 | 122226082 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699651 | 122226098 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699651 | 122226098 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699651 | 122226098 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699558 | 122226127 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699558 | 122226127 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699558 | 122226127 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7699558 | 122226127 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699516 | 122226133 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699516 | 122226133 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699516 | 122226133 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699516 | 122226133 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7599671 | 122227432 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES NOTED DURING TESTING UNIT PASSED DISPLACEMENT TEST AND SELF TEST | Component Missing | Device problem code
7599671 | 122227432 | N | NO MISSING SEGMENTSPARTIAL DISPLAY ANOMALIES NOTED DURING TESTING UNIT PASSED DISPLACEMENT TEST AND SELF TEST | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699692 | 122227685 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7699692 | 122227685 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7699692 | 122227685 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7699692 | 122227685 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7702917 | 122227946 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7702917 | 122227946 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7702917 | 122227946 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER AND WAS UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7918715 | 122304380 | N | WE HAVE RECEIVED THE DEVICE FOR EVALUATION THE DEVICE DID NOT OPERATE AT ALL WHEN WE TESTED THE DEVICE DURING OUR EVALUATION FURTHER INSPECTION FOUND THAT A WIRE HAD DETACHED FROM THE PCB AND A SEAL HOUSING WAS FOUND TO BE OUT OF ALIGNMENT WITH THE HEATSINK TANK THE SEAL HOUSING WAS NOT TIGHTLY COUPLED TO THE MOTOR WHICH ALLOWED THE SEAL HOUSINGMOTOR TO ROTATE ENOUGH TO CAUSE THE WIRE TO BREAK WE BELIEVE THIS WAS AN ASSEMBLY ERROR ISOLATED TO JUST THIS HANDPIECE THERE WAS NO INJURY TO THE PATIENT AS THE RESULT OF THIS INCIDENT HOWEVER SURGEON HAD TO ABANDON THE PHLEBECTOMY SURGERY AS A RESULT OF THIS INCIDENT | Detachment of Device or device Component | Device problem code
7918715 | 122304380 | N | WE HAVE RECEIVED THE DEVICE FOR EVALUATION THE DEVICE DID NOT OPERATE AT ALL WHEN WE TESTED THE DEVICE DURING OUR EVALUATION FURTHER INSPECTION FOUND THAT A WIRE HAD DETACHED FROM THE PCB AND A SEAL HOUSING WAS FOUND TO BE OUT OF ALIGNMENT WITH THE HEATSINK TANK THE SEAL HOUSING WAS NOT TIGHTLY COUPLED TO THE MOTOR WHICH ALLOWED THE SEAL HOUSINGMOTOR TO ROTATE ENOUGH TO CAUSE THE WIRE TO BREAK WE BELIEVE THIS WAS AN ASSEMBLY ERROR ISOLATED TO JUST THIS HANDPIECE THERE WAS NO INJURY TO THE PATIENT AS THE RESULT OF THIS INCIDENT HOWEVER SURGEON HAD TO ABANDON THE PHLEBECTOMY SURGERY AS A RESULT OF THIS INCIDENT | Break | Device problem code
7918715 | 122304380 | N | WE HAVE RECEIVED THE DEVICE FOR EVALUATION THE DEVICE DID NOT OPERATE AT ALL WHEN WE TESTED THE DEVICE DURING OUR EVALUATION FURTHER INSPECTION FOUND THAT A WIRE HAD DETACHED FROM THE PCB AND A SEAL HOUSING WAS FOUND TO BE OUT OF ALIGNMENT WITH THE HEATSINK TANK THE SEAL HOUSING WAS NOT TIGHTLY COUPLED TO THE MOTOR WHICH ALLOWED THE SEAL HOUSINGMOTOR TO ROTATE ENOUGH TO CAUSE THE WIRE TO BREAK WE BELIEVE THIS WAS AN ASSEMBLY ERROR ISOLATED TO JUST THIS HANDPIECE THERE WAS NO INJURY TO THE PATIENT AS THE RESULT OF THIS INCIDENT HOWEVER SURGEON HAD TO ABANDON THE PHLEBECTOMY SURGERY AS A RESULT OF THIS INCIDENT | Fracture Problem | Medical Evaluation result code
7701561 | 122327794 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7701561 | 122327794 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7701561 | 122327794 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7701561 | 122327794 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7697344 | 122329366 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7697344 | 122329366 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7701540 | 122329396 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7701540 | 122329396 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7701540 | 122329396 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7701540 | 122329396 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7725241 | 122336412 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7725241 | 122336412 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7725241 | 122336412 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7725241 | 122336412 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Positioning Failure | Device problem code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Component Missing | Device problem code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Scratched Material | Device problem code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Mechanical Jam | Device problem code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Mechanical Problem Identified | Medical Evaluation result code
7717288 | 122337201 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER DEVICE RECEIVED WITH MINOR SCRATCHED DISPLAY WINDOW CASE AND KEYPAD OVERLAY | Stress Problem Identified | Medical Evaluation result code
7687415 | 122339634 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Inappropriate or Unexpected Reset | Device Problem Code
7687415 | 122339634 | N | INSULIN PUMP PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION RANGE PUMP WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED | Energy Storage System Problem | Medical Evaluation result code
7705575 | 122340777 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7705575 | 122340777 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7705575 | 122340777 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7693072 | 122341891 | N | THE DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED THE DEVICE WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED OFF NO POWER LOW BATTERY WEAK BATTERY BATTERY OUT LIMIT AND FAILED BATTERY TEST ALARMS NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7684196 | 122345514 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST B | Failure to Deliver | Device problem code
7684196 | 122345514 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST B | Inaccurate Delivery | Device Problem Code
7684196 | 122345514 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST B | Operational Problem Identified | Medical Evaluation result code
7695996 | 122346738 | N | THE INSULIN PUMP WAS RECEIVED WITH NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING THE PUMP WAS RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Failure to Deliver | Device problem code
7695996 | 122346738 | N | THE INSULIN PUMP WAS RECEIVED WITH NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING THE PUMP WAS RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Inaccurate Delivery | Device Problem Code
7695996 | 122346738 | N | THE INSULIN PUMP WAS RECEIVED WITH NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING THE PUMP WAS RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND TEST SELFTEST UNEXPECTED ERROR TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | Operational Problem Identified | Medical Evaluation result code
7562005 | 122351370 | N | THE DEVICE WAS RECEIVED WITH A CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO THE ELECTRICAL BOARD UNABLE TO PERFORM THE SELF TEST DISPLACEMENT TEST REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT TEST AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR ALARM | Component or Accessory Incompatibility | Device problem code
7562005 | 122351370 | N | THE DEVICE WAS RECEIVED WITH A CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO THE ELECTRICAL BOARD UNABLE TO PERFORM THE SELF TEST DISPLACEMENT TEST REWIND TEST PRIME AND SEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT TEST AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR ALARM | Interoperability Problem Identified | Medical Evaluation result code
7590414 | 122353148 | N | THE INSULIN PUMP WAS RECEIVED WITH A CRITICAL PUMP ERROR ALARM AND PUMP ERROR ALARM DUE TO THE FORCE SENSOR FLEX CABLE INCOMPLETELY PLUGGED IN UNABLE TO PERFORM THE SELF TEST DISPLACEMENT REWIND PRIME OR SEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR ALARM | Component or Accessory Incompatibility | Device problem code
7590414 | 122353148 | N | THE INSULIN PUMP WAS RECEIVED WITH A CRITICAL PUMP ERROR ALARM AND PUMP ERROR ALARM DUE TO THE FORCE SENSOR FLEX CABLE INCOMPLETELY PLUGGED IN UNABLE TO PERFORM THE SELF TEST DISPLACEMENT REWIND PRIME OR SEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR ALARM | Interoperability Problem Identified | Medical Evaluation result code
7721448 | 122360033 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Component or Accessory Incompatibility | Device problem code
7721448 | 122360033 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7721448 | 122360033 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7689105 | 122360176 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARM NOTED MOTOR PASSED MOTOR TEST B | No Device Problem Found | Medical Evaluation result code
7632932 | 122362018 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7632932 | 122362018 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7632932 | 122362018 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Operational Problem Identified | Medical Evaluation result code
7674117 | 122362752 | N | UNIT PASSED THE FUNCTIONAL TEST INCLUDING THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED PUMP ERROR NOTED DURING TESTING NO LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING UNIT COMMUNICATED PROPERLY WITH THE GUARDIAN LINK AND THE TEST PLUG NO INSULIN PUMPSENSOR COMMUNICATION ANOMALY SENSOR ANOMALY CHANGE SENSOR ALERT LOST SENSOR SIGNAL ALERT OR CALIBRATION ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7693004 | 122366403 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7693004 | 122366403 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7693004 | 122366403 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7706102 | 122366454 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP WAS RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARMS NOTED DURING TESTING NO UNEXPECTED REPLACE BATTERY NOW ALARMS OR LOW BATTERY ALARMS NOTED DURING TESTING INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7717556 | 122376735 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7717556 | 122376735 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7713672 | 122384853 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7713672 | 122384853 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7713672 | 122384853 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7713672 | 122384853 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7725491 | 122385593 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7725491 | 122385593 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7725491 | 122385593 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7725491 | 122385593 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7920348 | 122479984 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Apparent Adverse Event | Device problem code
7920348 | 122479984 | N | B PENTAX VIDEO COLONOSCOPE MODEL ECFI IS NOT DISTRIBUTED IN THE USA THEREFORE THE PMAK NUMBER IS NOT APPLICABLE | No Device Problem Found | Medical Evaluation result code
7696937 | 122517857 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSER OR PROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | No Device Problem Found | Medical Evaluation result code
7693083 | 122519880 | N | NO MISSING RETAINER OR MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
7661033 | 122520316 | N | PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO FAILED BATTERY TEST ALARMS NOTED DURING TESTING PUMP ERROR ALARM VARIABLE WAS CONFIRMED IN THE FORMATTED HISTORY FILE DUE TO A COLD SOLDER JOINT FOUND ON THE AMBIENT LIGHT SENSOR NO UNEXPECTED PUMP ERROR ALARM DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED PUMP ERROR DUE TO A SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
7717409 | 122523513 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7717409 | 122523513 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7717409 | 122523513 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7712218 | 122523781 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY ANY DEFICIENCY ALLEGATION ANOMALY DUE TO BUTTONS NOT RESPONDING | Failure to Deliver | Device problem code
7712218 | 122523781 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY ANY DEFICIENCY ALLEGATION ANOMALY DUE TO BUTTONS NOT RESPONDING | Inappropriate or Unexpected Reset | Device Problem Code
7712218 | 122523781 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY ANY DEFICIENCY ALLEGATION ANOMALY DUE TO BUTTONS NOT RESPONDING | Inaccurate Delivery | Device Problem Code
7712218 | 122523781 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY ANY DEFICIENCY ALLEGATION ANOMALY DUE TO BUTTONS NOT RESPONDING | Component or Accessory Incompatibility | Device problem code
7712218 | 122523781 | N | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST OR VERIFY ANY DEFICIENCY ALLEGATION ANOMALY DUE TO BUTTONS NOT RESPONDING | Interoperability Problem Identified | Medical Evaluation result code
7692978 | 122523877 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7692978 | 122523877 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7692978 | 122523877 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7692978 | 122523877 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7720467 | 122524000 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST B | Inaccurate Delivery | Device Problem Code
7720467 | 122524000 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO EXCESSIVE NO DELIVERY ALARM NOTED DURING TEST B | Failure to Deliver | Device problem code
7693218 | 122524002 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7693218 | 122524002 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7693218 | 122524002 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7693218 | 122524002 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7700185 | 122524732 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST DRIVE SUPPORT CAP WAS INSPECTED AND NO ANOMALY WAS NOTED | Mechanical Problem Identified | Medical Evaluation result code
7691290 | 122525220 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7691290 | 122525220 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7691290 | 122525220 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7691290 | 122525220 | N | THE DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7680817 | 122526695 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST NO UNEXPECTED MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING B | Component Missing | Device problem code
7680817 | 122526695 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST NO UNEXPECTED MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING B | Display or Visual Feedback Problem | Device problem code
7680817 | 122526695 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST NO UNEXPECTED MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7657592 | 122527196 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW NOTED | Display or Visual Feedback Problem | Device problem code
7657592 | 122527196 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW NOTED | No Display / Image | Device Problem Code
7657592 | 122527196 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW NOTED | Scratched Material | Device problem code
7657592 | 122527196 | N | INSULIN PUMP PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW NOTED | Stress Problem Identified | Medical Evaluation result code
7709171 | 122541199 | N | THE PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7709171 | 122541199 | N | THE PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7709171 | 122541199 | N | THE PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7628230 | 122550457 | N | B THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST THERE WAS NO CRITICAL PUMP ERROR DURING TESTING | No Device Problem Found | Medical Evaluation result code
7692202 | 122658322 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7692202 | 122658322 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Component or Accessory Incompatibility | Device problem code
7692202 | 122658322 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Interoperability Problem Identified | Medical Evaluation result code
7627705 | 122681754 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO LCD BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | No Display / Image | Device problem code
7627705 | 122681754 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO LCD BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Failure to Deliver | Device problem code
7627705 | 122681754 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO LCD BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Inappropriate or Unexpected Reset | Device Problem Code
7627705 | 122681754 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO LCD BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Inaccurate Delivery | Device Problem Code
7627705 | 122681754 | N | UNIT RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO LCD BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Operational Problem Identified | Medical Evaluation result code
7702439 | 122683728 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7702439 | 122683728 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7702439 | 122683728 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7702439 | 122683728 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7697358 | 122689930 | N | THE DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7697358 | 122689930 | N | THE DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7697358 | 122689930 | N | THE DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7697358 | 122689930 | N | THE DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7697879 | 122690745 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7697879 | 122690745 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7697879 | 122690745 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7697879 | 122690745 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7701609 | 122692111 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7701609 | 122692111 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7701609 | 122692111 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7701609 | 122692111 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7702411 | 122692340 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7702411 | 122692340 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7702411 | 122692340 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7701753 | 122692414 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7701753 | 122692414 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7701753 | 122692414 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7702445 | 122692810 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7702445 | 122692810 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7702445 | 122692810 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7702445 | 122692810 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7705542 | 122694145 | N | DEVICE RECEIVED WITH UNRESPONSIVE KEYPAD ASSEMBLY DUE TO KEYPAD DEFECT USED TEST KEYPAD UNIT TO GET SOFTWARE VERSION DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7705542 | 122694145 | N | DEVICE RECEIVED WITH UNRESPONSIVE KEYPAD ASSEMBLY DUE TO KEYPAD DEFECT USED TEST KEYPAD UNIT TO GET SOFTWARE VERSION DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7705542 | 122694145 | N | DEVICE RECEIVED WITH UNRESPONSIVE KEYPAD ASSEMBLY DUE TO KEYPAD DEFECT USED TEST KEYPAD UNIT TO GET SOFTWARE VERSION DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component or Accessory Incompatibility | Device problem code
7705542 | 122694145 | N | DEVICE RECEIVED WITH UNRESPONSIVE KEYPAD ASSEMBLY DUE TO KEYPAD DEFECT USED TEST KEYPAD UNIT TO GET SOFTWARE VERSION DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Defective Device | Device problem code
7705542 | 122694145 | N | DEVICE RECEIVED WITH UNRESPONSIVE KEYPAD ASSEMBLY DUE TO KEYPAD DEFECT USED TEST KEYPAD UNIT TO GET SOFTWARE VERSION DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7704066 | 122694527 | N | DEVICE PASSED DISPLACEMENT TEST AND SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY B | Component Missing | Device problem code
7704066 | 122694527 | N | DEVICE PASSED DISPLACEMENT TEST AND SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY B | Display or Visual Feedback Problem | Device problem code
7704066 | 122694527 | N | DEVICE PASSED DISPLACEMENT TEST AND SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY B | Component Missing | Device Problem Code
7705316 | 122694878 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7705316 | 122694878 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7705316 | 122694878 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7717581 | 122698039 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7717581 | 122698039 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7717581 | 122698039 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7697536 | 122698463 | N | DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7697536 | 122698463 | N | DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7697536 | 122698463 | N | DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7697536 | 122698463 | N | DEVICE WAS RECEIVED WITH MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE O RING AND MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7563870 | 122700248 | N | UNIT HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ALL BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY PRIMEFILL ANOMALY DUE TO UNRESPONSIVE BUTTON | Failure to Deliver | Device problem code
7563870 | 122700248 | N | UNIT HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ALL BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY PRIMEFILL ANOMALY DUE TO UNRESPONSIVE BUTTON | Component or Accessory Incompatibility | Device problem code
7563870 | 122700248 | N | UNIT HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ALL BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY PRIMEFILL ANOMALY DUE TO UNRESPONSIVE BUTTON | Inaccurate Delivery | Device Problem Code
7563870 | 122700248 | N | UNIT HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ALL BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY PRIMEFILL ANOMALY DUE TO UNRESPONSIVE BUTTON | Operational Problem Identified | Medical Evaluation result code
7563870 | 122700248 | N | UNIT HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ALL BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIMEDELIVERY TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY PRIMEFILL ANOMALY DUE TO UNRESPONSIVE BUTTON | Interoperability Problem Identified | Medical Evaluation result code
7722731 | 122700766 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7722731 | 122700766 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7722731 | 122700766 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7722731 | 122700766 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7663204 | 122702399 | N | THE INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE PUMP WAS UNABLE TO CONFIRM MOTOR POSITION ENCODER ERROR ALARM DUE TO OVERWRITTEN HISTORY FILE | No Device Problem Found | Medical Evaluation result code
7662268 | 122703009 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE INSULIN PUMP WAS MONITORED FOR A DAY AND NO MISSING SEGMENTS OR DISPLAY ANOMALY NOTED | Display or Visual Feedback Problem | Device problem code
7662268 | 122703009 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE INSULIN PUMP WAS MONITORED FOR A DAY AND NO MISSING SEGMENTS OR DISPLAY ANOMALY NOTED | Component Missing | Device problem code
7662268 | 122703009 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE INSULIN PUMP WAS MONITORED FOR A DAY AND NO MISSING SEGMENTS OR DISPLAY ANOMALY NOTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7700746 | 122716416 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7700746 | 122716416 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7700746 | 122716416 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7700746 | 122716416 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7937390 | 122805952 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
7937392 | 122806264 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
7693040 | 123376853 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7693040 | 123376853 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7693040 | 123376853 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7940299 | 123414014 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICES DEFIB OUTPUT WAS OUT OF SPECIFICATION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | No Consequences Or Impact To Patient | Patient problem code
7940299 | 123414014 | D | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT AGE AND GENDER UNKNOWN THE DEVICES DEFIB OUTPUT WAS OUT OF SPECIFICATION COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION | Insufficient Information | Device Problem Code
7730405 | 123420760 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | No Display / Image | Device problem code
7730405 | 123420760 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | Display or Visual Feedback Problem | Device problem code
7730405 | 123420760 | N | DEVICE PASSED THE OPERATING CURRENTS SELF TEST AND DISPLACEMENT TEST NO BLANK DISPLAY ANOMALY NOTED DURING TEST | Image Artifact | Medical Evaluation result code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Component Missing | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Loss of Power | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Positioning Failure | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Disconnection | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Electrical Power Problem | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | No Audible Alarm | Device problem code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7689393 | 123425053 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR LOW BATTERY POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY INSULIN PUMP HAD A MISSING RETAINER AND A MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER THE TEST PCAP AND RESERVOIR WILL NOT LOCK IN PLACE IN THE RESERVOIR COMPARTMENT DUE TO MISSING RETAINER | Energy Storage System Problem | Medical Evaluation result code
7957323 | 123425073 | D | MEDTRONIC RECEIVED INFORMATION THAT AN UNKNOWN DURATION POST IMPLANT OF THIS BIOPROSTHETIC VALVE THE PATIENT DEVELOPED COMPLETE HEART BLOCK WHICH REQUIRED THE IMPLANT OF A PERMANENT PACEMAKER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7957323 | 123425073 | D | MEDTRONIC RECEIVED INFORMATION THAT AN UNKNOWN DURATION POST IMPLANT OF THIS BIOPROSTHETIC VALVE THE PATIENT DEVELOPED COMPLETE HEART BLOCK WHICH REQUIRED THE IMPLANT OF A PERMANENT PACEMAKER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7914760 | 123500847 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION FOR ID WHICH APPEARED AS VISUALLY WEAK POSITIVE COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Connection Problem | Device problem code
7914760 | 123500847 | N | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD ON SEP TO ASSESS THE INSTRUMENT TEST WELL IMAGES IN QUESTION FOR ID WHICH APPEARED AS VISUALLY WEAK POSITIVE COMPLAINTS OF POSITIVE REACTIONS BEING INTERPRETED AS NEGATIVE BY THE GALILEO ECHO CAMERA ALGORITHM ARE BEING CORRECTED UNDER DESIGN CONTROL PROJECT B THE IMMUCOR TECHNICAL COMMUNICATION B IS BEING USED BY THE LAB TO VISUALLY CONFIRM ALL NEGATIVE ASSAY EVENTS REPORTED BY THE ECHO INSTRUMENT B | Interoperability Problem Identified | Medical Evaluation result code
7729241 | 123520934 | N | DEVICE PASSED SELF TEST AND DISPLACEMENT TEST NO FUZZY DISPLAY MISSING SEGMENTS PARTIAL DISPLAY OR LINES ON DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7697409 | 123545494 | N | INSULIN PUMP PASSED DISPLACEMENT TEST SELF TEST OFF NO POWER TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPEC RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO BLANK DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7734467 | 123695141 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7734467 | 123695141 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7734467 | 123695141 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7735133 | 123700554 | N | DEVICE RECEIVED WITH CRITICAL PUMP ERROR AND MOTOR SAFETY CIRCUIT FAILED TEST DUE TO SOFTWARE ERROR | Component or Accessory Incompatibility | Device problem code
7735133 | 123700554 | N | DEVICE RECEIVED WITH CRITICAL PUMP ERROR AND MOTOR SAFETY CIRCUIT FAILED TEST DUE TO SOFTWARE ERROR | Circuit Failure | Device problem code
7735133 | 123700554 | N | DEVICE RECEIVED WITH CRITICAL PUMP ERROR AND MOTOR SAFETY CIRCUIT FAILED TEST DUE TO SOFTWARE ERROR | Computer Software Problem | Device problem code
7735133 | 123700554 | N | DEVICE RECEIVED WITH CRITICAL PUMP ERROR AND MOTOR SAFETY CIRCUIT FAILED TEST DUE TO SOFTWARE ERROR | Mechanical Problem Identified | Medical Evaluation result code
7735133 | 123700554 | N | DEVICE RECEIVED WITH CRITICAL PUMP ERROR AND MOTOR SAFETY CIRCUIT FAILED TEST DUE TO SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7684768 | 123705172 | N | INSULIN UMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED | Loss of Power | Device problem code
7684768 | 123705172 | N | INSULIN UMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED | No Audible Alarm | Device problem code
7684768 | 123705172 | N | INSULIN UMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED | Disconnection | Device problem code
7684768 | 123705172 | N | INSULIN UMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED | Electrical Power Problem | Device problem code
7684768 | 123705172 | N | INSULIN UMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY AND POWER LOSS ALARM WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE LOADED VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED REPLACE BATTERY NOW ALARMS NOTED | Energy Storage System Problem | Medical Evaluation result code
6583601 | 123725425 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED CRITICAL PUMP ERROR OPEN BOOK DURING TESTING | Component or Accessory Incompatibility | Device problem code
6583601 | 123725425 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELFTEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED CRITICAL PUMP ERROR OPEN BOOK DURING TESTING | Interoperability Problem Identified | Medical Evaluation result code
6550461 | 123725632 | N | INSULIN PUMP PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO PUMP ERROR ALARM NOTED | Component or Accessory Incompatibility | Device problem code
6550461 | 123725632 | N | INSULIN PUMP PASSED THE REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO PUMP ERROR ALARM NOTED | Mechanical Problem Identified | Medical Evaluation result code
7721784 | 123726615 | N | COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | Loose or Intermittent Connection | Device problem code
7721784 | 123726615 | N | COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | Interoperability Problem Identified | Medical Evaluation result code
7728153 | 123729379 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7728153 | 123729379 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7728153 | 123729379 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7730289 | 123731497 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7730289 | 123731497 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7730289 | 123731497 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7730289 | 123731497 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7603964 | 123733141 | N | THE DEVICE WAS RECEIVED WITH A CRITICAL PUMP ERROR OPEN BOOK IMAGE ALARM AND PUMP ERROR ALARM DUE TO THE FORCE SENSOR FLEX CABLE INCOMPLETELY PLUGGED IN UNABLE TO PERFORM THE SELF TEST DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT TEST AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR OPEN BOOK IMAGE ALARM | Component or Accessory Incompatibility | Device problem code
7603964 | 123733141 | N | THE DEVICE WAS RECEIVED WITH A CRITICAL PUMP ERROR OPEN BOOK IMAGE ALARM AND PUMP ERROR ALARM DUE TO THE FORCE SENSOR FLEX CABLE INCOMPLETELY PLUGGED IN UNABLE TO PERFORM THE SELF TEST DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT TEST AND ACTIVE CURRENT MEASUREMENT TEST DUE TO CRITICAL PUMP ERROR OPEN BOOK IMAGE ALARM | Interoperability Problem Identified | Medical Evaluation result code
7734818 | 123734285 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7734818 | 123734285 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7734818 | 123734285 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7734691 | 123734434 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7734691 | 123734434 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7734691 | 123734434 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7920150 | 123844040 | N | ON AUG THE FIELD SERVICE ENGINEER FSE WAS AT THE CUSTOMER SITE AND FOUND THAT THE INSTRUMENT HAD BEEN GENERATING VOTE OUTS AND SYSTEM MESSAGES SINCE THE DAY BEFORE B THE CUSTOMER HAD OBSERVED THE THREE ABNORMALLY HIGH PLT RESULTS AND STOPPED THE INSTRUMENT DAILY CHECKS WERE RUN AND FAILED FOR RBC AND PLT BACKGROUND THE FSE THEN FOUND THAT THE RBC APERTURES WERE PARTIALLY CLOGGED THE FSE CLEANED THE APERTURES TO RESOLVE THE ISSUE THE ASSIGNABLE CAUSE FOR THE REPORTED EVENT WAS A PARTIALLY CLOGGED RBC APERTURE THE FSE ALSO COMPLETED B TO UPDATE THE SOFTWARE TO VERSION THE FSE VERIFIED THE REPAIRS BY RUNNING DAILY CHECKS WHICH PASSED AN IMPROVEMENT FOR PLT FLAGGING WAS DEVELOPED AND RELEASED AS A PART OF A SOFTWARE UPDATE TO THE DXH THAT IS DESIGNATED V ON APPROXIMATELY B OF THE WORLDWIDE PRODUCT INSTALLED BASE HAS BEEN UPGRADED AS OF B DUE TO AN INCREASE IN THE NUMBER OF COMPLAINTS RECEIVED RELATED TO UNFLAGGED ERRONEOUSLY HIGH PLT PATIENT RESULTS RELEASED FROM THE LAB BEC HAS DETERMINED THAT A FIELD ACTION WAS WARRANTED THE FIELD ACTION WAS APPROVED B ON AND INCLUDES A SOFTWARE PATCH FOR THE IMPROVEMENT FOR THE PLT FLAGGING COMPONENT OF V WHICH WILL EXPEDITE INSTALLATION OF THIS IMPROVEMENT FOR REST OF THE INSTALLED BASE B BEC INTERNAL IDENTIFIER B | No Findings Available | Medical Evaluation result code
7711954 | 123858115 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPEC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM OR FAILED BATTERY TEST ALARM NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD B | No Audible Alarm | Device problem code
7711954 | 123858115 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPEC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM OR FAILED BATTERY TEST ALARM NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD B | Battery Problem | Device Problem Code
7711954 | 123858115 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPEC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM OR FAILED BATTERY TEST ALARM NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD B | Loss of Power | Device problem code
7711954 | 123858115 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPEC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM OR FAILED BATTERY TEST ALARM NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD B | Mechanical Problem Identified | Medical Evaluation result code
7683894 | 123867644 | N | NO UNEXPECTED LOW BATTERY ALERTS REPLACE BATTERY ALERTS REPLACE BATTERY NOW ALARMS OR POWER LOSS ALARMS NOTED DURING TESTING POWER ERROR DUE TO J CONNECTOR RESISTANCE ISSUE UNIT PASSED DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST | No Findings Available | Medical Evaluation result code
7735612 | 123868761 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7735612 | 123868761 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7735612 | 123868761 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7735612 | 123868761 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7737902 | 123869692 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7737902 | 123869692 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7737902 | 123869692 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7737902 | 123869692 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7709627 | 123871696 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT RECEIVED WITH FADED SERIAL NUMBER LABEL | Component Missing | Device problem code
7709627 | 123871696 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT RECEIVED WITH FADED SERIAL NUMBER LABEL | Mechanical Jam | Device problem code
7709627 | 123871696 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT RECEIVED WITH FADED SERIAL NUMBER LABEL | Missing or Inadequate Safety Measures | Medical Evaluation result code
7709627 | 123871696 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER UNIT RECEIVED WITH FADED SERIAL NUMBER LABEL | Mechanical Problem Identified | Medical Evaluation result code
7957323 | 123875288 | D | MEDTRONIC RECEIVED INFORMATION THAT FOUR DAYS POST IMPLANT OF THIS AORTIC BIOPROSTHETIC VALVE THE PATIENT DEVELOPED COMPLETE HEART BLOCK WHICH REQUIRED THE IMPLANT OF A NONMEDTRONIC PERMANENT PACEMAKER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device problem code
7957323 | 123875288 | D | MEDTRONIC RECEIVED INFORMATION THAT FOUR DAYS POST IMPLANT OF THIS AORTIC BIOPROSTHETIC VALVE THE PATIENT DEVELOPED COMPLETE HEART BLOCK WHICH REQUIRED THE IMPLANT OF A NONMEDTRONIC PERMANENT PACEMAKER NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Complete Heart Block | Patient problem code
7658345 | 123888246 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7658345 | 123888246 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7658765 | 123888403 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7658765 | 123888403 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7813363 | 123891596 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7813363 | 123891596 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7699309 | 123892733 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Component or Accessory Incompatibility | Device problem code
7699309 | 123892733 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Incompatible Component/ Accessory | Medical Evaluation result code
7718307 | 124007334 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7718307 | 124007334 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7718307 | 124007334 | N | PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7721928 | 124178203 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST INSULIN PUMP WAS TESTED WITH NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED B | Display or Visual Feedback Problem | Device problem code
7721928 | 124178203 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST INSULIN PUMP WAS TESTED WITH NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED B | Component Missing | Device problem code
7721928 | 124178203 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST INSULIN PUMP WAS TESTED WITH NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7725463 | 124178381 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7725463 | 124178381 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7725463 | 124178381 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7725463 | 124178381 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7725580 | 124178463 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7725580 | 124178463 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7725580 | 124178463 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7725580 | 124178463 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7663113 | 124180303 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7663113 | 124180303 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7663113 | 124180303 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7663113 | 124180303 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Energy Storage System Problem | Medical Evaluation result code
7662250 | 124368482 | N | PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7953007 | 124688441 | N | THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION THE MALFUNCTION WAS DUPLICATED AND ATTRIBUTED TO A DEPLETED BATTERY THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER | Battery Problem | Device Problem Code
7953007 | 124688441 | N | THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION THE MALFUNCTION WAS DUPLICATED AND ATTRIBUTED TO A DEPLETED BATTERY THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER | Energy Storage System Problem | Medical Evaluation result code
7989126 | 124800347 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED THAT THE POD SITE WAS VERY WET WITH INSULIN PATIENT CHANGED THE POD | Insufficient Information | Device problem code
7989126 | 124800347 | D | IT WAS REPORTED THAT THE PATIENTS BLOOD GLUCOSE LEVELS REACHED MGDL WHILE WEARING THE POD BETWEEN AND HOURS PATIENT NOTICED THAT THE POD SITE WAS VERY WET WITH INSULIN PATIENT CHANGED THE POD | No Known Impact Or Consequence To Patient | Patient Problem Code
7737628 | 124829830 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
7737628 | 124829830 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7737628 | 124829830 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Unintended compatibility | Device Problem Code
8003735 | 124973609 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
8003733 | 124973715 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
8003738 | 124975918 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
8003741 | 124977669 | N | COMPLAINTS RELATING TO THE REPORTED PRODUCTS WERE EVALUATED IT WAS CONCLUDED THAT THE NUMBER OF COMPLAINTS FOR THE PRODUCTS DID NOT BREACH THRESHOLDS INDICATIVE OF A SYSTEMIC ISSUE | No Device Problem Found | Medical Evaluation result code
7734253 | 125005094 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Disconnection | Device problem code
7734253 | 125005094 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Audible Alarm | Device problem code
7734253 | 125005094 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
7734253 | 125005094 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Loss of Power | Device problem code
7734253 | 125005094 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Energy Storage System Problem | Medical Evaluation result code
7735150 | 125015726 | N | DEVICE PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFIC RANGE DEVICE WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED B | No Device Problem Found | Medical Evaluation result code
7735150 | 125015726 | N | DEVICE PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFIC RANGE DEVICE WAS MONITORED AND TESTED WITH NO FAILED BATTERY TEST NOTED B | Battery Problem | Device Problem Code
7720637 | 125016527 | N | COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK | No Device Problem Found | Medical Evaluation result code
7717387 | 125020551 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Inaccurate Delivery | Device Problem Code
7717387 | 125020551 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | Failure to Deliver | Device problem code
8008448 | 125094782 | N | AGEDATE OF BIRTH UNKNOWN INFORMATION NOT PROVIDED GENDERSEX UNKNOWN INFORMATION NOT PROVIDED DATE OF EVENT UNKNOWN INFORMATION NOT PROVIDED IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
8008448 | 125094782 | N | AGEDATE OF BIRTH UNKNOWN INFORMATION NOT PROVIDED GENDERSEX UNKNOWN INFORMATION NOT PROVIDED DATE OF EVENT UNKNOWN INFORMATION NOT PROVIDED IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | No Apparent Adverse Event | Device problem code
8008448 | 125094782 | N | AGEDATE OF BIRTH UNKNOWN INFORMATION NOT PROVIDED GENDERSEX UNKNOWN INFORMATION NOT PROVIDED DATE OF EVENT UNKNOWN INFORMATION NOT PROVIDED IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
8008007 | 125095010 | N | IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Component Missing | Device problem code
8008007 | 125095010 | N | IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | No Apparent Adverse Event | Device problem code
8008007 | 125095010 | N | IF IMPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED IF EXPLANTED GIVE DATE NOT APPLICABLE AS THE LENS WAS INSERTED AND REMOVED B ATTEMPTS HAVE BEEN MADE TO OBTAIN MISSING INFORMATION HOWEVER TO DATE NO RESPONSE HAS BEEN RECEIVED ALL PERTINENT INFORMATION AVAILABLE TO JOHNSON JOHNSON SURGICAL VISION INC HAS BEEN SUBMITTED | Missing or Inadequate Safety Measures | Medical Evaluation result code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Audible Alarm | Device problem code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Loss of Power | Device problem code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Battery Problem | Device Problem Code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Disconnection | Device problem code
7729251 | 125158975 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER LOSS ALARM REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Energy Storage System Problem | Medical Evaluation result code
7730184 | 125174803 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7730184 | 125174803 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7730184 | 125174803 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7737504 | 125175236 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7737504 | 125175236 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7737504 | 125175236 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7737504 | 125175236 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7657602 | 125271557 | N | UNIT RECEIVED WITH OPERATIONAL CURRENTS WITHIN SPEC AND PASSED SELF TEST AND DISPLACEMENT TEST NO LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM NOTED DURING TESTING PUMP ERROR CONFIRMED IN PUMP HISTORY WITH VARIABLE DUE COLD SOLDER JOINT AT U KEYPAD ASSEMBLY | No Audible Alarm | Device problem code
7657602 | 125271557 | N | UNIT RECEIVED WITH OPERATIONAL CURRENTS WITHIN SPEC AND PASSED SELF TEST AND DISPLACEMENT TEST NO LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM NOTED DURING TESTING PUMP ERROR CONFIRMED IN PUMP HISTORY WITH VARIABLE DUE COLD SOLDER JOINT AT U KEYPAD ASSEMBLY | Loss of Power | Device problem code
7657602 | 125271557 | N | UNIT RECEIVED WITH OPERATIONAL CURRENTS WITHIN SPEC AND PASSED SELF TEST AND DISPLACEMENT TEST NO LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM NOTED DURING TESTING PUMP ERROR CONFIRMED IN PUMP HISTORY WITH VARIABLE DUE COLD SOLDER JOINT AT U KEYPAD ASSEMBLY | Component or Accessory Incompatibility | Device problem code
7657602 | 125271557 | N | UNIT RECEIVED WITH OPERATIONAL CURRENTS WITHIN SPEC AND PASSED SELF TEST AND DISPLACEMENT TEST NO LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM NOTED DURING TESTING PUMP ERROR CONFIRMED IN PUMP HISTORY WITH VARIABLE DUE COLD SOLDER JOINT AT U KEYPAD ASSEMBLY | Interoperability Problem Identified | Medical Evaluation result code
7669549 | 125283216 | N | NO MISSING SEGMENTS OR PARTIAL DISPLAY BLANK FLASHING DISPLAY OR DISPLAY ANOMALY NOTED DURING TESTING DEVICE PASSED THE DISPLACEMENT TEST SELF TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT | No Device Problem Found | Medical Evaluation result code
7673045 | 125310360 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THANB CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7673045 | 125310360 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THANB CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7673045 | 125310360 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THANB CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Energy Storage System Problem | Medical Evaluation result code
7727784 | 125338822 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO DARK DISPLAY FLASHING DISPLAY MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO FAILED BATTERY TEST NOTED DURING TESTING OR IN THE DEVICE TRACE DOWNLOAD B | No Device Problem Found | Medical Evaluation result code
7692832 | 125522916 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7692832 | 125522916 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7692832 | 125522916 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Energy Storage System Problem | Medical Evaluation result code
7676783 | 125825004 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Battery Problem | Device Problem Code
7676783 | 125825004 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Loss of Power | Device problem code
7676783 | 125825004 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7676783 | 125825004 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7676783 | 125825004 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST NO UNEXPECTED FAILED BATTERY TEST ALARM NOTED DURING TESTING PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR POWER LOSS ALARM AND REPLACE BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Energy Storage System Problem | Medical Evaluation result code
7725260 | 125826739 | N | B INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPEC RANGE NO UNEXPECTED LOW BATTERY ALERT POWER LOSS ALARM REPLACE BATTERY ALERT REPLACE BATTERY NOW ALARM OR FAILED BATTERY TEST ALARM NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD SOFTWARE ERROR DETECTED ALARM FOUND IN THE PUMP HISTORY DUE TO SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
7707486 | 125828522 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7707486 | 125828522 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7707486 | 125828522 | N | TESTED WITH A TEST PCAP AND THE TEST PCAP DID NOT LOCK IN PLACE PROPERLY DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7697375 | 125828930 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7697375 | 125828930 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7697375 | 125828930 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7686240 | 125830335 | N | INSULIN PUMP PASSED SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT NO REPLACE BATTERY NOW ALARM NOTED IN THE PUMP TRACE DOWNLOAD THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7686240 | 125830335 | N | INSULIN PUMP PASSED SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT NO REPLACE BATTERY NOW ALARM NOTED IN THE PUMP TRACE DOWNLOAD THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7686240 | 125830335 | N | INSULIN PUMP PASSED SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT NO REPLACE BATTERY NOW ALARM NOTED IN THE PUMP TRACE DOWNLOAD THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | No Audible Alarm | Device problem code
7686240 | 125830335 | N | INSULIN PUMP PASSED SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT NO REPLACE BATTERY NOW ALARM NOTED IN THE PUMP TRACE DOWNLOAD THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Loss of Power | Device problem code
7686240 | 125830335 | N | INSULIN PUMP PASSED SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT NO REPLACE BATTERY NOW ALARM NOTED IN THE PUMP TRACE DOWNLOAD THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR LOW BATTERY ALERT POWER LOSS ALARM AND REPLACE BATTERY ALERT WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE UL VLITH WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7730394 | 125846656 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST NO UNEXPECTED PRIMEFILL ANOMALY NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7722860 | 125969532 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7722860 | 125969532 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7722860 | 125969532 | N | THE INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7684521 | 125969921 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7684521 | 125969921 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7684521 | 125969921 | N | THE INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7706667 | 125970461 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7706667 | 125970461 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7706667 | 125970461 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7705612 | 125977861 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7705612 | 125977861 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7705612 | 125977861 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7812924 | 125993997 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7812924 | 125993997 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7812924 | 125993997 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7812130 | 126009152 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Component Missing | Device problem code
7812130 | 126009152 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7812130 | 126009152 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING TEST RESERVOIR DOES NOT LOCK PROPERLY INSIDE THE RESERVOIR TUBE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7730932 | 126075470 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST DEVICE RECEIVED WITH MOTOR ERROR ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST OR VERIFY COMPROMISED FORCED SENSOR SYSTEM ALARM DUE TO MOTOR ERROR ALARM | No Device Problem Found | Medical Evaluation result code
7815418 | 126086148 | N | DEVICE PASSED SELF TEST AND DISPLACEMENT TEST NO MISSING SEGMENTS PARTIAL DISPLAY FLASHING DISPLAY OR SOLID WHITE DISPLAY NOTED B | No Device Problem Found | Medical Evaluation result code
7975231 | 126094898 | N | THE EVALUATION INCLUDED A REVIEW OF SUBMITTED DATA PRODUCT HISTORICAL DATA AND PRODUCT LABELING A REVIEW OF COMPLAINT AND TRENDING DATA DID NOT IDENTIFY ANY ISSUES OR TRENDS THE SUBMITTED DATA WAS REVIEWED AND FOUND THAT THE INITIAL CLOSEDMODE RESULTS GENERATED ON THE CELLDYN RUBY MATCHED THE RESULTS GENERATED ON THE ALTERNATE ANALYZER FOR ALL MAJOR PARAMETERS INCLUDING WBC RBC HGB MCV AND PLT HOWEVER WHEN THE RESULTS FROM THE REPEAT RUN IN OPEN MODE ON THE CELDYN RUBY A DIFFERENCE WAS OBSERVED THE RESULTS WERE HIGHER FOR RBC AND HGB AND LOWER FOR WBC AND PLT WHICH MAY BE INDICATIVE THAT FURTHER MIXING OF THE SPECIMEN WAS NECESSARY A REVIEW OF THE PRODUCT LABELING CONCLUDED THAT THE ISSUE IS SUFFICIENTLY ADDRESSED NO SYSTEMIC ISSUE OR DEFICIENCY OF THE CELLDYN RUBY ANALYZER WAS IDENTIFIED | High Test Results | Device problem code
7975231 | 126094898 | N | THE EVALUATION INCLUDED A REVIEW OF SUBMITTED DATA PRODUCT HISTORICAL DATA AND PRODUCT LABELING A REVIEW OF COMPLAINT AND TRENDING DATA DID NOT IDENTIFY ANY ISSUES OR TRENDS THE SUBMITTED DATA WAS REVIEWED AND FOUND THAT THE INITIAL CLOSEDMODE RESULTS GENERATED ON THE CELLDYN RUBY MATCHED THE RESULTS GENERATED ON THE ALTERNATE ANALYZER FOR ALL MAJOR PARAMETERS INCLUDING WBC RBC HGB MCV AND PLT HOWEVER WHEN THE RESULTS FROM THE REPEAT RUN IN OPEN MODE ON THE CELDYN RUBY A DIFFERENCE WAS OBSERVED THE RESULTS WERE HIGHER FOR RBC AND HGB AND LOWER FOR WBC AND PLT WHICH MAY BE INDICATIVE THAT FURTHER MIXING OF THE SPECIMEN WAS NECESSARY A REVIEW OF THE PRODUCT LABELING CONCLUDED THAT THE ISSUE IS SUFFICIENTLY ADDRESSED NO SYSTEMIC ISSUE OR DEFICIENCY OF THE CELLDYN RUBY ANALYZER WAS IDENTIFIED | Incorrect Interpretation of Results/Data | Medical Evaluation result code
7983171 | 126110339 | N | INSULIN PUMP RECEIVED WITH SERIAL NUMBER LABEL FADING AND MINOR SCRATCHED LCD WINDOW THE AUDIO BEEP FUNCTIONING PROPERLY DURING THE SELF TEST B | Scratched Material | Device problem code
7983171 | 126110339 | N | INSULIN PUMP RECEIVED WITH SERIAL NUMBER LABEL FADING AND MINOR SCRATCHED LCD WINDOW THE AUDIO BEEP FUNCTIONING PROPERLY DURING THE SELF TEST B | No Device Problem Found | Medical Evaluation result code
8022403 | 126131642 | N | THE INITIAL REPORT WAS SENT IN ERROR THERE WAS NO ALLEGATION BY THE PATIENT OF THE PUMP CAUSING A BLOOD GLUCOSE EXCURSION OR A BATTERY LIFE ISSUE PLEASE DISREGARD THE INITIAL REPORT | Component or Accessory Incompatibility | Device Problem Code
8022403 | 126131642 | N | THE INITIAL REPORT WAS SENT IN ERROR THERE WAS NO ALLEGATION BY THE PATIENT OF THE PUMP CAUSING A BLOOD GLUCOSE EXCURSION OR A BATTERY LIFE ISSUE PLEASE DISREGARD THE INITIAL REPORT | Interoperability Problem Identified | Medical Evaluation result code
7728349 | 126212533 | N | B INSULIN PUMP PASSED THE FUNCTIONAL TESTINGS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED NO DELIVERY OR INSULIN FLOW BLOCKED ALARM MISSING SEGMENTS PARTIAL DISPLAY OR BLANK DISPLAY NOTED | No Findings Available | Medical Evaluation result code
7693190 | 126214515 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST THE INSULIN PUMP WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRONIC BOARD DURING VISUAL INSPECTION B | No Display / Image | Device Problem Code
6909210 | 126246503 | N | B INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO MOTOR NOTED DURING TESTING PUMP TESTED OK | No Device Problem Found | Medical Evaluation result code
7673034 | 126274133 | N | DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRICAL BOARD DURING VISUAL INSPECTION THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING | Loss of Power | Device problem code
7673034 | 126274133 | N | DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRICAL BOARD DURING VISUAL INSPECTION THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING | No Display / Image | Device Problem Code
7673034 | 126274133 | N | DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRICAL BOARD DURING VISUAL INSPECTION THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING | Disconnection | Device problem code
7673034 | 126274133 | N | DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRICAL BOARD DURING VISUAL INSPECTION THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARMS NOTED DURING TESTING | Interoperability Problem Identified | Medical Evaluation result code
7673001 | 126274220 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7673001 | 126274220 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7663347 | 126275348 | N | DEVICE PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST DEVICE ALSO PASSED TONE CHECK AND VIBRATE CHECK ON SELF TEST NO AUDIO OR BEEP ANOMALY OR VIBRATE ANOMALY NOTED DURING TESTING THE UNEXPECTED BOLUS DELIVERIES WAS NOT SPECIFIED IN THE CUSTOMER NOTES UNABLE TO VERIFY BOLUS ANOMALY DEVICE WAS MONITORED FOR THREE DAYS NO UNEXPECTED BOLUS DELIVERY NOTED DURING TESTING DEVICE WAS THEN PROGRAMMED WITH MULTIPLE TEST BOLUSES ALL BOLUSES DELIVERED PROPERLY AND WERE RECORDED IN THE DAILY HISTORY SCREEN NO BOLUS ANOMALY OR HISTORY ANOMALY NOTED DURING TESTING DEVICE HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY B | Vibration | Device problem code
7663347 | 126275348 | N | DEVICE PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST DEVICE ALSO PASSED TONE CHECK AND VIBRATE CHECK ON SELF TEST NO AUDIO OR BEEP ANOMALY OR VIBRATE ANOMALY NOTED DURING TESTING THE UNEXPECTED BOLUS DELIVERIES WAS NOT SPECIFIED IN THE CUSTOMER NOTES UNABLE TO VERIFY BOLUS ANOMALY DEVICE WAS MONITORED FOR THREE DAYS NO UNEXPECTED BOLUS DELIVERY NOTED DURING TESTING DEVICE WAS THEN PROGRAMMED WITH MULTIPLE TEST BOLUSES ALL BOLUSES DELIVERED PROPERLY AND WERE RECORDED IN THE DAILY HISTORY SCREEN NO BOLUS ANOMALY OR HISTORY ANOMALY NOTED DURING TESTING DEVICE HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY B | Scratched Material | Device problem code
7663347 | 126275348 | N | DEVICE PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST DEVICE ALSO PASSED TONE CHECK AND VIBRATE CHECK ON SELF TEST NO AUDIO OR BEEP ANOMALY OR VIBRATE ANOMALY NOTED DURING TESTING THE UNEXPECTED BOLUS DELIVERIES WAS NOT SPECIFIED IN THE CUSTOMER NOTES UNABLE TO VERIFY BOLUS ANOMALY DEVICE WAS MONITORED FOR THREE DAYS NO UNEXPECTED BOLUS DELIVERY NOTED DURING TESTING DEVICE WAS THEN PROGRAMMED WITH MULTIPLE TEST BOLUSES ALL BOLUSES DELIVERED PROPERLY AND WERE RECORDED IN THE DAILY HISTORY SCREEN NO BOLUS ANOMALY OR HISTORY ANOMALY NOTED DURING TESTING DEVICE HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY B | No Audible Prompt / Feedback | Device problem code
7663347 | 126275348 | N | DEVICE PASSED THE FUNCTIONAL TEST INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND THE DELIVERY ACCURACY TEST DEVICE ALSO PASSED TONE CHECK AND VIBRATE CHECK ON SELF TEST NO AUDIO OR BEEP ANOMALY OR VIBRATE ANOMALY NOTED DURING TESTING THE UNEXPECTED BOLUS DELIVERIES WAS NOT SPECIFIED IN THE CUSTOMER NOTES UNABLE TO VERIFY BOLUS ANOMALY DEVICE WAS MONITORED FOR THREE DAYS NO UNEXPECTED BOLUS DELIVERY NOTED DURING TESTING DEVICE WAS THEN PROGRAMMED WITH MULTIPLE TEST BOLUSES ALL BOLUSES DELIVERED PROPERLY AND WERE RECORDED IN THE DAILY HISTORY SCREEN NO BOLUS ANOMALY OR HISTORY ANOMALY NOTED DURING TESTING DEVICE HAD MINOR SCRATCHED DISPLAY WINDOW SCRATCHED CASE AND A PILLOWING KEYPAD OVERLAY B | Stress Problem Identified | Medical Evaluation result code
7815428 | 126275387 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST ALL THE BUTTONS FUNCTIONED PROPERLY AND NO MISSING SEGMENTSPARTIAL DISPLAY OR FLASHING WHITE LIGHT ON DISPLAY NOTED DURING TESTING B | Component Missing | Device problem code
7815428 | 126275387 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST ALL THE BUTTONS FUNCTIONED PROPERLY AND NO MISSING SEGMENTSPARTIAL DISPLAY OR FLASHING WHITE LIGHT ON DISPLAY NOTED DURING TESTING B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7675689 | 126287637 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7675689 | 126287637 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7675689 | 126287637 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY UNLOADED VOLTAGE WAS LESS THAN VOLTS FOR TWO HUNDRED AND FORTY CONSECUTIVE MINUTES DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7672972 | 126287818 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Component Missing | Device problem code
7672972 | 126287818 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Mechanical Jam | Device problem code
7672972 | 126287818 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND RESERVOIR UNABLE TO LOCK INTO PLACE DUE TO MISSING RESERVOIR TUBE ORING AND MISSING RESERVOIR RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7667520 | 126288791 | N | INSULIN PUMP SEATS IMMEDIATELY UPON PRIMING PROBLEM ISOLATED TO THE MOTOR OR FORCE SENSOR UNABLE TO PERFORM REWIND TEST PRIME OR SEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TEST DUE TO MOTORFORCE SENSOR ANOMALY | Priming Problem | Device problem code
7667520 | 126288791 | N | INSULIN PUMP SEATS IMMEDIATELY UPON PRIMING PROBLEM ISOLATED TO THE MOTOR OR FORCE SENSOR UNABLE TO PERFORM REWIND TEST PRIME OR SEATING BASIC OCCLUSION OCCLUSION FORCE SENSOR AND DISPLACEMENT TEST DUE TO MOTORFORCE SENSOR ANOMALY | No Findings Available | Medical Evaluation result code
7624922 | 126290693 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR AND LOW BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device problem code
7624922 | 126290693 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR AND LOW BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Mechanical Jam | Device problem code
7624922 | 126290693 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR AND LOW BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Component Missing | Device problem code
7624922 | 126290693 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR AND LOW BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Electrical Power Problem | Device problem code
7624922 | 126290693 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR AND LOW BATTERY ALERT THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALERT DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Energy Storage System Problem | Medical Evaluation result code
7811895 | 126355102 | N | DEVICE ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY MOTOR POSITION ENCODER ERROR ALARM DUE TO MOTOR ERROR ALARMS NO UNEXPECTED CONSTANT AUDIO TONE ANOMALIES NOTED | No Findings Available | Medical Evaluation result code
7574658 | 126379806 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS | Inaccurate Delivery | Device Problem Code
7574658 | 126379806 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS | Component or Accessory Incompatibility | Device Problem Code
7574658 | 126379806 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS | No Findings Available | Medical Evaluation result code
7814987 | 126510758 | N | INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7814987 | 126510758 | N | INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7814987 | 126510758 | N | INSULIN PUMP HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7814983 | 126632257 | N | B INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST INSULIN PUMP WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO CRITICAL BLOCK ALARM NOTED | No Device Problem Found | Medical Evaluation result code
7676535 | 126671491 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS NO FAILED BATTERY TEST ALARMS OR UNEXPECTED BATTERY POWER LOSS ALARMS NOTED DURING TESTING | Battery Problem | Device Problem Code
7676535 | 126671491 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS NO FAILED BATTERY TEST ALARMS OR UNEXPECTED BATTERY POWER LOSS ALARMS NOTED DURING TESTING | Loss of Power | Device problem code
7676535 | 126671491 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS NO FAILED BATTERY TEST ALARMS OR UNEXPECTED BATTERY POWER LOSS ALARMS NOTED DURING TESTING | Energy Storage System Problem | Medical Evaluation result code
7815370 | 126692635 | N | B INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELFTEST INSULIN PUMP WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO BLANK DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7919664 | 126700094 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Defective Device | Device problem code
7729752 | 126711276 | N | B INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST | No Findings Available | Medical Evaluation result code
6953332 | 126897165 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST AND FORCE SENSOR TEST POWER SUPPLY TEST FAILED DURING SELFTEST ALARMED DURING SELF TEST DUE TO THE ELECTRICAL BOARD AND ELECTRICAL BOARD CONNECTOR RESISTANCE | No Findings Available | Medical Evaluation result code
7812581 | 126903475 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Component Missing | Device problem code
7812581 | 126903475 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7812581 | 126903475 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER | Mechanical Problem Identified | Medical Evaluation result code
7812063 | 126903603 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER THE THIN PLASTIC STRIP OVER THE LCD SCREEN IS MISSING LCD SCREEN WAS FUNCTIONING AT THE TIME OF INSPECTION IE DID NOT OBSERVE A BLANK WHITE OR FLASHING DISPLAY | No Display / Image | Device Problem Code
7812063 | 126903603 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER THE THIN PLASTIC STRIP OVER THE LCD SCREEN IS MISSING LCD SCREEN WAS FUNCTIONING AT THE TIME OF INSPECTION IE DID NOT OBSERVE A BLANK WHITE OR FLASHING DISPLAY | Mechanical Jam | Device problem code
7812063 | 126903603 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER THE THIN PLASTIC STRIP OVER THE LCD SCREEN IS MISSING LCD SCREEN WAS FUNCTIONING AT THE TIME OF INSPECTION IE DID NOT OBSERVE A BLANK WHITE OR FLASHING DISPLAY | Component Missing | Device problem code
7812063 | 126903603 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP ORING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER THE THIN PLASTIC STRIP OVER THE LCD SCREEN IS MISSING LCD SCREEN WAS FUNCTIONING AT THE TIME OF INSPECTION IE DID NOT OBSERVE A BLANK WHITE OR FLASHING DISPLAY | Missing or Inadequate Safety Measures | Medical Evaluation result code
7813847 | 126903608 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7813847 | 126903608 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7813847 | 126903608 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE O RING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7814907 | 126905627 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE E HAD MINOR SCRATCHED LCD WINDOW | No Display / Image | Device problem code
7814907 | 126905627 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY DEVICE E HAD MINOR SCRATCHED LCD WINDOW | No Findings Available | Medical Evaluation result code
7811913 | 126911900 | N | NO BUTTON ERROR ALARM NOTED DURING TESTING DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED UP BUTTON DOME SWITCH NO UNLOCKED KEYPAD CONNECTOR NOTED | Component or Accessory Incompatibility | Device Problem Code
7811913 | 126911900 | N | NO BUTTON ERROR ALARM NOTED DURING TESTING DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED UP BUTTON DOME SWITCH NO UNLOCKED KEYPAD CONNECTOR NOTED | Device Alarm System | Device problem code
7811913 | 126911900 | N | NO BUTTON ERROR ALARM NOTED DURING TESTING DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED UP BUTTON DOME SWITCH NO UNLOCKED KEYPAD CONNECTOR NOTED | Interoperability Problem Identified | Medical Evaluation result code
7811869 | 126919073 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DEVICE WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO BLANK DISPLAY NO UNEXPECTED BATTERY POWER LOSS ALARM NOTED DURING TESTING PUMP MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTS IT BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN MINUTES AND NO UNEXPECTED POSTRESET RAM CRC ALARM NOTED B | Loss of Power | Device problem code
7811869 | 126919073 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DEVICE WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO BLANK DISPLAY NO UNEXPECTED BATTERY POWER LOSS ALARM NOTED DURING TESTING PUMP MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTS IT BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN MINUTES AND NO UNEXPECTED POSTRESET RAM CRC ALARM NOTED B | No Display / Image | Device Problem Code
7811869 | 126919073 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DEVICE WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO BLANK DISPLAY NO UNEXPECTED BATTERY POWER LOSS ALARM NOTED DURING TESTING PUMP MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTS IT BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN MINUTES AND NO UNEXPECTED POSTRESET RAM CRC ALARM NOTED B | No Audible Alarm | Device problem code
7811869 | 126919073 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DEVICE WAS MONITORED AND FUNCTIONING PROPERLY NOTED NO BLANK DISPLAY NO UNEXPECTED BATTERY POWER LOSS ALARM NOTED DURING TESTING PUMP MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTS IT BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN MINUTES AND NO UNEXPECTED POSTRESET RAM CRC ALARM NOTED B | Energy Storage System Problem | Medical Evaluation result code
7811971 | 127075296 | N | DEVICE POWER UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED DEVICE WAS RECEIVED WITH NORMAL OPERATING CURRENTS AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED OFF NO POWER LOW BATTERY WEAK BATTERY BATTERY OUT LIMIT AND FAILED BATTERY TEST ALARMS NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7811958 | 127217671 | N | B INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED | No Display / Image | Device problem code
7815497 | 127220301 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO CRITICAL PUMP ERROR OPEN BOOK IMAGE NOTED DURING TESTING B | Component or Accessory Incompatibility | Device problem code
7815497 | 127220301 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO CRITICAL PUMP ERROR OPEN BOOK IMAGE NOTED DURING TESTING B | Interoperability Problem Identified | Medical Evaluation result code
7812999 | 127221572 | N | HARDWARE LOW LEVEL FAILURES ERROR CONFIRMED IN THE DEVICE HISTORY DUE TO COLD SOLDER JOINT ON KEYPAD ASSEMBLY NO AUDIO OR VIBRATE ANOMALY AND NO POSTRESET RAM CRC ALARM NOTED | Vibration | Device problem code
7812999 | 127221572 | N | HARDWARE LOW LEVEL FAILURES ERROR CONFIRMED IN THE DEVICE HISTORY DUE TO COLD SOLDER JOINT ON KEYPAD ASSEMBLY NO AUDIO OR VIBRATE ANOMALY AND NO POSTRESET RAM CRC ALARM NOTED | No Audible Prompt / Feedback | Device problem code
7812999 | 127221572 | N | HARDWARE LOW LEVEL FAILURES ERROR CONFIRMED IN THE DEVICE HISTORY DUE TO COLD SOLDER JOINT ON KEYPAD ASSEMBLY NO AUDIO OR VIBRATE ANOMALY AND NO POSTRESET RAM CRC ALARM NOTED | No Device Problem Found | Medical Evaluation result code
7815172 | 127222431 | N | INSULIN PUMP RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE CONNECTOR NOT PLUGGED IN PROPERLY TO ELECTRICAL BOARD UNABLE TO PERFORM THE DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DUE TO THE BLANK DISPLAY | No Display / Image | Device Problem Code
7815172 | 127222431 | N | INSULIN PUMP RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE CONNECTOR NOT PLUGGED IN PROPERLY TO ELECTRICAL BOARD UNABLE TO PERFORM THE DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DUE TO THE BLANK DISPLAY | Connection Problem | Device problem code
7815172 | 127222431 | N | INSULIN PUMP RECEIVED WITH A BLANK DISPLAY DUE TO BATTERY TUBE CONNECTOR NOT PLUGGED IN PROPERLY TO ELECTRICAL BOARD UNABLE TO PERFORM THE DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST DUE TO THE BLANK DISPLAY | Interoperability Problem Identified | Medical Evaluation result code
7815118 | 127226574 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED CRITICAL PUMP ERROR OPEN BOOK DURING TESTING THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED B | Component or Accessory Incompatibility | Device problem code
7815118 | 127226574 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED CRITICAL PUMP ERROR OPEN BOOK DURING TESTING THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED B | Interoperability Problem Identified | Medical Evaluation result code
7812109 | 127342916 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED CRITICAL PUMP ERROR OPEN BOOK DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7814542 | 127344803 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE POWER UP PROPERLY AFTER BATTERY INSTALLATION DEVICE RECEIVED WITH OPERATING CURRENTS WITHIN SPECIFIC DEVICE PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST POWER ERROR AND LOW BATTERY ALERT DUE TO J CONNECTOR RESISTANCE ISSUE | Component Missing | Device problem code
7814542 | 127344803 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE POWER UP PROPERLY AFTER BATTERY INSTALLATION DEVICE RECEIVED WITH OPERATING CURRENTS WITHIN SPECIFIC DEVICE PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST POWER ERROR AND LOW BATTERY ALERT DUE TO J CONNECTOR RESISTANCE ISSUE | Mechanical Jam | Device problem code
7814542 | 127344803 | N | UNABLE TO PERFORM DISPLACEMENT TEST OR OCCLUSION TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE POWER UP PROPERLY AFTER BATTERY INSTALLATION DEVICE RECEIVED WITH OPERATING CURRENTS WITHIN SPECIFIC DEVICE PASSED SELFTEST REWIND SEATING BASIC OCCLUSION TEST AND FORCE SENSOR TEST POWER ERROR AND LOW BATTERY ALERT DUE TO J CONNECTOR RESISTANCE ISSUE | Mechanical Problem Identified | Medical Evaluation result code
7814588 | 127344880 | N | B INSULIN PUMP PASSED DISPLACEMENT TEST SELF TEST OFF NO POWER TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFIC RANGE INSULIN PUMP WAS MONITORED AND TESTED WITH NO BLANK DISPLAY AND AUDIOBEEP ANOMALY NOTED | No Device Problem Found | Medical Evaluation result code
7726820 | 127424297 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED WITH NORMAL OPERATING CURRENTS DEVICE MONITORED FOR SEVERAL HOURS AND NO BLANK DISPLAY NOTED THE BATTERY TUBE CONNECTOR PLUGGED IN PROPERLY TO ELECTRICAL BOARD DURING VISUAL INSPECTION B | No Device Problem Found | Medical Evaluation result code
7979519 | 127437436 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST TOGGLED ON OR OFF AND INCREASED OR DECREASED VOLUME WITH THE AUDIO FEATURE FUNCTIONING PROPERLY NO AUDIO OF ALARM ANOMALIES NOTED DURING TESTING B | No Audible Prompt / Feedback | Device problem code
7979519 | 127437436 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST TOGGLED ON OR OFF AND INCREASED OR DECREASED VOLUME WITH THE AUDIO FEATURE FUNCTIONING PROPERLY NO AUDIO OF ALARM ANOMALIES NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7813217 | 127537341 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST BUT COMPROMISED FORCED SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCED SENSOR SYSTEM ALARM | No Device Problem Found | Medical Evaluation result code
7814762 | 127538733 | N | INSULIN PUMP RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESCAPE AND ACT BUTTON DOME SWITCH NO CREASE AND PARTIAL UNLOCKED LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT OPERATING CURRENTS UNEXPECTED RESTART ERROR TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST DUE TO BUTTONS NOT RESPONDING | Interoperability Problem Identified | Medical Evaluation result code
7730194 | 127576763 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7730194 | 127576763 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7730194 | 127576763 | N | INSULIN PUMP WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7814636 | 127865532 | N | DEVICE PASSED DISPLACEMENT TEST AND SELF TEST SOFTWARE ERROR DETECTED NOTED IN THE FORMATTED HISTORY FILE DUE TO SOFTWARE ERROR DEVICE PASSED THE SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT NO FAILED BATTERY TEST ALARMS OR BLANK DISPLAY NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | No Device Problem Found | Medical Evaluation result code
7937186 | 127898744 | D | NEW INFORMATION RECEIVED NOTES THAT THE SOURCE OF THE INFECTION WAS UNKNOWN PATIENT WAS STABLE POST PROCEDURE | Infection | Patient problem code
7937186 | 127898744 | D | NEW INFORMATION RECEIVED NOTES THAT THE SOURCE OF THE INFECTION WAS UNKNOWN PATIENT WAS STABLE POST PROCEDURE | Insufficient Information | Device Problem Code
7813455 | 128101185 | N | DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST INCHES NOTED NO MISSING SEGMENT OR PARTIAL DISPLAY ANOMALY DURING TEST B | Display or Visual Feedback Problem | Device problem code
7813455 | 128101185 | N | DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST INCHES NOTED NO MISSING SEGMENT OR PARTIAL DISPLAY ANOMALY DURING TEST B | Component Missing | Device problem code
7813455 | 128101185 | N | DEVICE PASSED THE DISPLACEMENT SELF TEST AND PASSED THE DELIVERY ACCURACY TEST INCHES NOTED NO MISSING SEGMENT OR PARTIAL DISPLAY ANOMALY DURING TEST B | Missing or Inadequate Safety Measures | Medical Evaluation result code
7732060 | 128212192 | N | B INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIMEA TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST | No Device Problem Found | Medical Evaluation result code
7811894 | 128212559 | N | DEVICE PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARM NOTED DURING TEST MOTOR PASSED MOTOR TEST B | No Device Problem Found | Medical Evaluation result code
7737492 | 128291776 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Inappropriate or Unexpected Reset | Device Problem Code
7737492 | 128291776 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Failure to Deliver | Device problem code
7737492 | 128291776 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | No Display / Image | Device problem code
7737492 | 128291776 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Inaccurate Delivery | Device Problem Code
7737492 | 128291776 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO BLANK DISPLAY | Operational Problem Identified | Medical Evaluation result code
7815515 | 128293760 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7815515 | 128293760 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7815515 | 128293760 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7648553 | 128303379 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Component Missing | Device problem code
7648553 | 128303379 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER MISSING RESERVOIR TUBE ORING UNABLE TO PERFORM THE DISPLACEMENT TEST OR LOCK RESERVOIR INTO PLACE DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7811918 | 128305035 | N | B INSULIN PUMP PASSED THE SEAT REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DISPLACEMENT TEST AND MOTOR TEST NO UNEXPECTED MOTOR ERROR ALARM NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7812155 | 128316590 | N | DEVICE PASSED THE DISPLACEMENT TEST SELF TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TESTS AND DELIVERY ACCURACY TEST B | Failure to Deliver | Device problem code
7812155 | 128316590 | N | DEVICE PASSED THE DISPLACEMENT TEST SELF TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TESTS AND DELIVERY ACCURACY TEST B | Inaccurate Delivery | Device Problem Code
7812155 | 128316590 | N | DEVICE PASSED THE DISPLACEMENT TEST SELF TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TESTS AND DELIVERY ACCURACY TEST B | Interoperability Problem Identified | Medical Evaluation result code
7815481 | 128392995 | N | DEVICE PASSED ALL FUNCTIONAL TESTS INCLUDING DISPLACEMENT REWIND PRIMESEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TESTS NO UNEXPECTED CRITICAL PUMP ERRORS OPEN BOOK IMAGE WERE NOTED B | Priming Problem | Device problem code
7815481 | 128392995 | N | DEVICE PASSED ALL FUNCTIONAL TESTS INCLUDING DISPLACEMENT REWIND PRIMESEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TESTS NO UNEXPECTED CRITICAL PUMP ERRORS OPEN BOOK IMAGE WERE NOTED B | No Device Problem Found | Medical Evaluation result code
7698198 | 128451614 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST INSTALLED A TEST RESERVOIR PRIMED THE PUMP SET A BASAL RATE AND ACTIVATED THE SUSPEND DELIVERY FEATURE MONITORED THE INSULIN PUMP DURING SUSPENDED DELIVERY THE PUMP ACTIVATED THE VIBRATE ALERT MINUTES AFTER THE DISPLAY TIMED OUT PROPERLY NO ABSENCE OF VIBRATE ALERT OR ALARM ANOMALY NOTED DURING TESTING B | Vibration | Device problem code
7698198 | 128451614 | N | THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST INSTALLED A TEST RESERVOIR PRIMED THE PUMP SET A BASAL RATE AND ACTIVATED THE SUSPEND DELIVERY FEATURE MONITORED THE INSULIN PUMP DURING SUSPENDED DELIVERY THE PUMP ACTIVATED THE VIBRATE ALERT MINUTES AFTER THE DISPLAY TIMED OUT PROPERLY NO ABSENCE OF VIBRATE ALERT OR ALARM ANOMALY NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7725464 | 128453819 | N | B INSULIN PUMP PASSED FUNCTIONAL TESTINGS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME OR SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DAT TEST | No Device Problem Found | Medical Evaluation result code
7814933 | 128456439 | N | INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED | No Device Problem Found | Medical Evaluation result code
7911483 | 128458969 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | No Device Problem Found | Medical Evaluation result code
7811967 | 128460000 | N | INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO EXCESSIVE NO DELIVERY ALARMS NOTED B | No Device Problem Found | Medical Evaluation result code
7811884 | 128460215 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED DEVICE HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ESC AND ACT BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY MOTOR ERROR ALARM DUE TO UNRESPONSIVE BUTTON MOTOR PASSED MOTOR TEST | Inaccurate Delivery | Device Problem Code
7811884 | 128460215 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED DEVICE HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ESC AND ACT BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY MOTOR ERROR ALARM DUE TO UNRESPONSIVE BUTTON MOTOR PASSED MOTOR TEST | Failure to Deliver | Device problem code
7811884 | 128460215 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED DEVICE HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ESC AND ACT BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY MOTOR ERROR ALARM DUE TO UNRESPONSIVE BUTTON MOTOR PASSED MOTOR TEST | Component or Accessory Incompatibility | Device Problem Code
7811884 | 128460215 | N | DRIVE SUPPORT DISK WAS INSPECTED AND NO ANOMALY WAS NOTED DEVICE HAD UNRESPONSIVE BUTTONS DUE TO FLATTENED CREASED ESC AND ACT BUTTONS DOME SWITCH NO UNLOCKED JLCD KEYPAD CONNECTOR NOTED UNABLE TO PERFORM REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY MOTOR ERROR ALARM DUE TO UNRESPONSIVE BUTTON MOTOR PASSED MOTOR TEST | Interoperability Problem Identified | Medical Evaluation result code
7697356 | 128462681 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7697356 | 128462681 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7853055 | 128464397 | N | THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS CUT PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Jam | Device problem code
7853055 | 128464397 | N | THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS CUT PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION CONFIRMED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Problem Identified | Medical Evaluation result code
7605038 | 128618320 | N | EVALUATED ONE OPEN AND USED RESERVOIR AND CHECKED ORINGS OR STOPPER GROOVE FOR ANOMALIES NONE WERE FOUND RAN BASAL BOLUS LEAKING TEST RESERVOIR FILLED WITH DILUENT AND CONNECTED TO A NEW INFUSION SET INSTALLED RESERVOIR AND CONNECTOR INTO PARADIGM INSULIN PUMP CONCLUSION RESERVOIRS DOES NOT LEAK AND DID NOT CONTINUE DRIPPING INSULIN AFTER TEST B | Leak / Splash | Device problem code
7605038 | 128618320 | N | EVALUATED ONE OPEN AND USED RESERVOIR AND CHECKED ORINGS OR STOPPER GROOVE FOR ANOMALIES NONE WERE FOUND RAN BASAL BOLUS LEAKING TEST RESERVOIR FILLED WITH DILUENT AND CONNECTED TO A NEW INFUSION SET INSTALLED RESERVOIR AND CONNECTOR INTO PARADIGM INSULIN PUMP CONCLUSION RESERVOIRS DOES NOT LEAK AND DID NOT CONTINUE DRIPPING INSULIN AFTER TEST B | Leakage/Seal | Medical Evaluation result code
7706569 | 128621770 | N | THE DEVICE PASSED ALL FUNCTIONAL TESTS INCLUDING DISPLACEMENT REWIND PRIME OR SEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TESTS NO UNEXPECTED CRITICAL PUMP ERRORS WERE NOTED | No Device Problem Found | Medical Evaluation result code
7692996 | 128622623 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Mechanical Jam | Device problem code
7692996 | 128622623 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
7692996 | 128622623 | N | DEVICE HAD A MISSING RETAINER AND MISSING RESERVOIR TUBE ORING RESERVOIR DOES NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7613224 | 128625504 | N | DEVICE RECEIVED WITH A MISSING RETAINER RING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER ALSO THERE ARE SOME VISIBLE SCRATCHES ON THE LCD SCREEN THE USER IS ABLE TO SEE PAST THESE EASILY AND NAVIGATE THROUGH THE MENUS | Scratched Material | Device problem code
7613224 | 128625504 | N | DEVICE RECEIVED WITH A MISSING RETAINER RING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER ALSO THERE ARE SOME VISIBLE SCRATCHES ON THE LCD SCREEN THE USER IS ABLE TO SEE PAST THESE EASILY AND NAVIGATE THROUGH THE MENUS | Mechanical Jam | Device problem code
7613224 | 128625504 | N | DEVICE RECEIVED WITH A MISSING RETAINER RING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER ALSO THERE ARE SOME VISIBLE SCRATCHES ON THE LCD SCREEN THE USER IS ABLE TO SEE PAST THESE EASILY AND NAVIGATE THROUGH THE MENUS | Positioning Failure | Device problem code
7613224 | 128625504 | N | DEVICE RECEIVED WITH A MISSING RETAINER RING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER ALSO THERE ARE SOME VISIBLE SCRATCHES ON THE LCD SCREEN THE USER IS ABLE TO SEE PAST THESE EASILY AND NAVIGATE THROUGH THE MENUS | Component Missing | Device problem code
7613224 | 128625504 | N | DEVICE RECEIVED WITH A MISSING RETAINER RING UNABLE TO PERFORM DISPLACEMENT TEST SINCE PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER ALSO THERE ARE SOME VISIBLE SCRATCHES ON THE LCD SCREEN THE USER IS ABLE TO SEE PAST THESE EASILY AND NAVIGATE THROUGH THE MENUS | Missing or Inadequate Safety Measures | Medical Evaluation result code
7571593 | 128627023 | N | INSULIN PUMP ALARMED MOTOR ERROR DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS | No Device Problem Found | Medical Evaluation result code
7571326 | 128627468 | N | DEVICE RECEIVED WITH ALARM LOOP DURING POWER UP PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TESTS OR VERIFY MOTOR ERROR ALARM DUE TO ALARM MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7571713 | 128628138 | N | DEVICE RECEIVED WITH SOFTWARE ERROR BAD POINTER ALARM LOOP DURING POWER UP PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TESTS OR VERIFY MOTOR ERROR ALARM DUE TO SOFTWARE ERROR BAD POINTER ALARM MOTOR PASSED MOTOR TEST | No Device Problem Found | Medical Evaluation result code
7572735 | 128629246 | N | INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED B | Interoperability Problem Identified | Medical Evaluation result code
7812100 | 128648206 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO CRITICAL PUMP ERROROPEN BOOK IMAGE NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7811919 | 128914855 | N | THE INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION TEST PRIMEA TEST AND DISPLACEMENT TEST PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ALARM IN THE HISTORY FILE UNABLE TO PERFORM THE OCCLUSION TEST AND THE EXCESSIVE NO DELIVERY TEST OR VERIFY NO DELIVERY ALARM DUE TO MOTOR ERROR ALARM THE MOTOR WAS TESTED OUTSIDE TO THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | No Device Problem Found | Medical Evaluation result code
7734143 | 128918316 | N | B INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO CRITICAL PUMP ERROROPEN BOOK IMAGE NOTED DURING TESTING | Insufficient Information | Device Problem Code
7734143 | 128918316 | N | B INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST NO CRITICAL PUMP ERROROPEN BOOK IMAGE NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7813106 | 129021697 | N | CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO FORCE SENSOR | Component or Accessory Incompatibility | Device problem code
7813106 | 129021697 | N | CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO FORCE SENSOR | Interoperability Problem Identified | Medical Evaluation result code
7814609 | 129031307 | N | B INSULIN PUMP PASSED THE SELF TEST AND DISPLACEMENT TEST NO HAL SOFTWARE ERROR DETECTED ALARMS NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7815359 | 129038126 | N | B INSULIN PUMP PASSED THE DISPLACEMENT TEST SELF TEST AND UNEXPECTED RESTART ALARM TEST NO UNEXPECTED RESTART ALARM ANOMALIES NOTED NO UNEXPECTED REWIND ANOMALIES NOTED | No Device Problem Found | Medical Evaluation result code
7812169 | 129104421 | N | B INSULIN PUMP PASSED THE DISPLACEMENT TEST SELF TEST AND UNEXPECTED RESTART ALARM TEST NO UNEXPECTED RESTART ALARM ANOMALIES NOTED | No Device Problem Found | Medical Evaluation result code
7814014 | 129161571 | N | DEVICE PASSED THE DISPLACEMENT TEST AND UNEXPECTED RESTART ERROR TEST NO UNEXPECTED RESTART ERROR ANOMALIES NOTED B | No Device Problem Found | Medical Evaluation result code
7814498 | 129162882 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST HOWEVER PUMP ALARMED COMPROMISED FORCE SENSOR SYSTEM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM THE OCCLUSION TEST PRIMEA TEST AND EXCESSIVE NO DELIVERY TEST DUE TO A ALARM | No Findings Available | Medical Evaluation result code
7813438 | 129184400 | N | DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED SELFTEST UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST NO UNEXPECTED LOW BATTERY WEAK BATTERY BATTERY OUT LIMIT OR FAILED BATTERY TEST ALARMS NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7811999 | 129204320 | N | INSULIN PUMP PASSED FUNCTIONAL TESTINGS INCLUDING STOP CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST SEAT REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DEVICE WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW B | Scratched Material | Device problem code
7811999 | 129204320 | N | INSULIN PUMP PASSED FUNCTIONAL TESTINGS INCLUDING STOP CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST SEAT REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DEVICE WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW B | No Device Problem Found | Medical Evaluation result code
7728510 | 129210753 | N | B THE INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST THE PUMP WAS PROGRAMMED WITH MULTIPLE BOLUSES AND MONITORED ALL BOLUSES DELIVERED PROPERLY AND WERE LISTED IN THE DAILY HISTORY SCREEN THE PUMP THEN WAS PROGRAMMED WITH MULTIPLE BASAL PROFILES AND MONITORED ALL BASAL PROFILES DELIVERED THEIR INDICATED AMOUNTS AND WERE VERIFIED IN THE DAILY AND SUMMARY HISTORY SCREENS THE UNITS LEFT AT THE PUMP DISPLAY MATCHED PROPERLY THE UNITS LEFT AT TEST RESERVOIR NO DELIVERY ANOMALY BOLUS ANOMALY OR BASAL ANOMALY NOTED DURING TESTING ALSO THE PUMP WAS PROGRAMMED WITH A TEST GUARDIAN LINK TRANSMITTER AND A MGDL GLUCOSE SENSOR SIMULATOR THE PUMP COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND DISPLAYED THE CALIBRATE YOUR SENSOR ALARM PROPERLY AFTER COMPLETION OF THE WARM UP THE PUMP CALIBRATED AND DISPLAYED THE PROGRAMMED VALUE OF MGDL PROPERLY ON THE DISPLAY GRAPH NO UNEXPECTED SENSOR ERRORS OR ANOMALIES WERE NOTED DURING TESTING THE PUMP WAS THEN PROGRAMMED FOR AUTO MODE THE DISPLAY SHOWED THE CALIBRATE YOUR SENSOR ALARM PROPERLY AFTER COMPLETION OF THE AUTO MODE WARM UP THE PUMP SENSOR FEATURE AND THE AUTO MODE CONNECTION ARE WORKING PROPERLY NO DELIVERY ANOMALIES NOTED WHILE IN AUTO MODE | Operational Problem Identified | Medical Evaluation result code
7814648 | 129222844 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST AND UNEXPECTED RESTART ALARM TEST NO UNEXPECTED RESTART ALARM ANOMALIES NOTED | No Device Problem Found | Medical Evaluation result code
7813721 | 129226363 | N | DEVICE RECEIVED WITH BLANK DISPLAY PROBLEM ISOLATED TO INTERFACE BOARD | No Display / Image | Device problem code
7813095 | 129228646 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH POWER ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Software Problem Identified | Medical Evaluation result code
7813095 | 129228646 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH POWER ERROR DUE TO NONSLEEP ANOMALY SOFTWARE ERROR | Computer Software Problem | Device Problem Code
7812633 | 129233283 | N | THE INSULIN PUMP WAS RECEIVED WITH POWER ERROR DETECTED DUE TO J CONNECTOR RESISTANCE ISSUE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST | Component or Accessory Incompatibility | Device Problem Code
7812633 | 129233283 | N | THE INSULIN PUMP WAS RECEIVED WITH POWER ERROR DETECTED DUE TO J CONNECTOR RESISTANCE ISSUE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST | Interoperability Problem Identified | Medical Evaluation result code
7815146 | 129238174 | N | THE INSULIN PUMP WAS RECEIVED WITH POWER ERROR DUE TO CONNECTOR RESISTANCE ISSUE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST | Component or Accessory Incompatibility | Device Problem Code
7815146 | 129238174 | N | THE INSULIN PUMP WAS RECEIVED WITH POWER ERROR DUE TO CONNECTOR RESISTANCE ISSUE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST | Electrical Power Problem | Device problem code
7815304 | 129238661 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP RECEIVED WITH POWER ERROR DETECTED DUE TO J CONNECTOR RESISTANCE ISSUE | Electrical Power Problem | Device problem code
7815304 | 129238661 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP RECEIVED WITH POWER ERROR DETECTED DUE TO J CONNECTOR RESISTANCE ISSUE | Component or Accessory Incompatibility | Device Problem Code
7815304 | 129238661 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST PUMP RECEIVED WITH POWER ERROR DETECTED DUE TO J CONNECTOR RESISTANCE ISSUE | Interoperability Problem Identified | Medical Evaluation result code
7577291 | 129241317 | N | INSULIN PUMP RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM HOWEVER INSULIN PUMP PASSED REWIND TEST AND DISPLACEMENT TEST | Failure to Deliver | Device problem code
7577291 | 129241317 | N | INSULIN PUMP RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM HOWEVER INSULIN PUMP PASSED REWIND TEST AND DISPLACEMENT TEST | Component or Accessory Incompatibility | Device Problem Code
7577291 | 129241317 | N | INSULIN PUMP RECEIVED COMPROMISED FORCE SENSOR SYSTEM ALARM DURING THE BASIC OCCLUSION TEST DUE TO LOOSEPROTRUDED DRIVE SUPPORT DISK UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST DUE TO COMPROMISED FORCE SENSOR SYSTEM ALARM HOWEVER INSULIN PUMP PASSED REWIND TEST AND DISPLACEMENT TEST | Interoperability Problem Identified | Medical Evaluation result code
7815517 | 129251096 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT ALL CURRENTS WITHIN SPECIFIC RANGE THE INSULIN PUMP WAS MONITORED FOR SEVERAL HOURS AND NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING B | Loss of Power | Device problem code
7815517 | 129251096 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT ALL CURRENTS WITHIN SPECIFIC RANGE THE INSULIN PUMP WAS MONITORED FOR SEVERAL HOURS AND NO UNEXPECTED POWER LOSS OR REPLACE BATTERY NOW ALARM NOTED DURING TESTING B | Loss of Power | Device Problem Code
7813704 | 129257946 | N | CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO FORCE SENSOR | Component or Accessory Incompatibility | Device problem code
7813704 | 129257946 | N | CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM PROBLEM ISOLATED TO FORCE SENSOR | Interoperability Problem Identified | Medical Evaluation result code
7848390 | 129366983 | N | THE SUSPECT GAS DELIVERY ENGINE GDE WAS RECEIVED AND EVALUATED THE REPORTED PROBLEM WAS DUPLICATED AND THE CAUSE ISOLATED TO BURNT TRANSISTORS AT THE PRINTED CIRCUIT BOARD ASSEMBLY PCBA | Power Source Problem Identified | Medical Evaluation result code
7812012 | 129507868 | N | B INSULIN PUMP PASSED DISPLACEMENT TEST REWIND PRIMESEATING BASIC OCCLUSION FORCE SENSOR OCCLUSION SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND SELF TEST NO UNEXPECTED MAIN ARM FAILED SELFTEST ALARMS NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
7657105 | 129513542 | N | ADVANCED BIONICS CONSIDERS THE INVESTIGATION INTO THIS REPORTABLE EVENT AS CLOSED THE SURGEON REPORTEDLY SUSPECTS THE CHOLESTEATOMA TO BE RELATED TO THE SURGERY THE RECIPIENT REPORTEDLY HEALED THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS CUT PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION REVEALED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK COULD NOT BE OBTAINED AT ANY SPACING THE NO LOCK CONDITION PREVENTED SOME OF THE ELECTRICAL TEST FROM BEING PERFORMED THE DEVICE PASSED SOME OF THE ELECTRICAL TEST PERFORMED THE DEVICE FAILED THE RESIDUAL GAS ANALYSIS TEST THIS DEVICE HAD MOISTURE THAT EXCEEDED THE RESIDUAL GAS ANALYSIS TEST LIMIT THE SOURCE OF THE PROBLEM WAS A FEEDTHRU HERMETICITY ISSUE FROM ONE FEEDTHRU VENDOR A CORRECTIVE ACTION WAS IMPLEMENTED FEEDTHRU ASSEMBLIES FROM THIS VENDOR ARE NO LONGER USED THIS IS THE FINAL REPORT | Moisture Damage | Device Problem Code
7657105 | 129513542 | N | ADVANCED BIONICS CONSIDERS THE INVESTIGATION INTO THIS REPORTABLE EVENT AS CLOSED THE SURGEON REPORTEDLY SUSPECTS THE CHOLESTEATOMA TO BE RELATED TO THE SURGERY THE RECIPIENT REPORTEDLY HEALED THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS CUT PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE PHOTOGRAPHIC IMAGING INSPECTION REVEALED CUT ELECTRODE WIRES THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY SYSTEM LOCK COULD NOT BE OBTAINED AT ANY SPACING THE NO LOCK CONDITION PREVENTED SOME OF THE ELECTRICAL TEST FROM BEING PERFORMED THE DEVICE PASSED SOME OF THE ELECTRICAL TEST PERFORMED THE DEVICE FAILED THE RESIDUAL GAS ANALYSIS TEST THIS DEVICE HAD MOISTURE THAT EXCEEDED THE RESIDUAL GAS ANALYSIS TEST LIMIT THE SOURCE OF THE PROBLEM WAS A FEEDTHRU HERMETICITY ISSUE FROM ONE FEEDTHRU VENDOR A CORRECTIVE ACTION WAS IMPLEMENTED FEEDTHRU ASSEMBLIES FROM THIS VENDOR ARE NO LONGER USED THIS IS THE FINAL REPORT | Storage Problem Identified | Medical Evaluation result code
7881757 | 129522206 | N | DEVICE EVALUATION THE DEVICE WAS VISUALLY INSPECTED IT WAS CONFIRMED THAT THE FAILURE MODE WAS THE PROBE LASER FIBER NOT FULLY CONNECTED TO THE EK ADAPTER ON THE PORTABLE CONNECTOR MODULE THAT RESULTED IN THE EK ADAPTER MELTING | Melted | Device problem code
7815512 | 129629555 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARMS NOTED DURING TESTING NO UNEXPECTED REPLACE BATTERY NOW ALARMS OR LOW BATTERY ALARMS NOTED DURING TESTING B | Loss of Power | Device problem code
7815512 | 129629555 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARMS NOTED DURING TESTING NO UNEXPECTED REPLACE BATTERY NOW ALARMS OR LOW BATTERY ALARMS NOTED DURING TESTING B | Device Alarm System | Device problem code
7815512 | 129629555 | N | INSULIN PUMP PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH NORMAL OPERATING CURRENTS NO UNEXPECTED POWER LOSS ALARMS NOTED DURING TESTING NO UNEXPECTED REPLACE BATTERY NOW ALARMS OR LOW BATTERY ALARMS NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7812941 | 129643860 | N | DEVICE PASSED FUNCTIONAL TESTING INCLUDING STOP CURRENT TEST RUN CURRENT TEST SELF TEST OFF NO POWER TEST UNEXPECTED RESTART ERROR TEST SEAT REWIND BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST B | No Device Problem Found | Medical Evaluation result code
7813009 | 129644781 | N | THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE INSULIN PUMP PASSED THE SELF TEST NO UNEXPECTED POSTRESET RAM CRC ALARM NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
6758192 | 129648575 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT TEST ACTIVE CURRENT TEST AND SELF TEST PRIME ANOMALY NOT CONFIRMED LOW BATTERY ALERT LESS THAN WEEK NOT CONFIRMED B | No Device Problem Found | Medical Evaluation result code
7906470 | 129652663 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Defective Device | Device problem code
7906470 | 129652663 | N | THE RETURNED PRODUCT WAS EVALUATED AND PERFORMED AS DESIGNED NO DEFECT OR DEFICIENCY THAT WOULD RESULT IN THE CANNULA TO FAIL TO INSERT CORRECTLY OR THE PUMP TO FAIL TO DELIVER INSULIN WAS FOUND | Operational Problem Identified | Medical Evaluation result code
7589546 | 129784809 | N | MOTOR ERROR ALARM DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM ANY TESTS DUE TO MOTOR ERROR ALARM | No Findings Available | Medical Evaluation result code
6932817 | 129796155 | N | THE INSULIN PUMP ALARMED MOTOR RELATED ERROR DURING TESTING DUE TO FORCE SENSOR FLEX NOT BEING PLUGGED IN UNABLE TO PERFORM THE DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST AND OCCLUSION TEST DUE TO ALARM NOT BEING ABLE TO BE CLEARED | No Device Problem Found | Medical Evaluation result code
7887963 | 129814331 | N | B SOFTWARE ERRORTHE MOTOR ARM CANNOT COMMUNICATE WITH THE MAIN ARM HARDWARE LOW LEVEL FAILURES AND DEVICE TEST FAILED NOT FOUND DURING TESTING THE PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
7842941 | 129833051 | N | THE DEVICE PASSED THE EXTERNAL VISUAL AND PHOTOGRAPHIC IMAGING INSPECTIONS SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE ELECTRICAL TESTS PERFORMED THIS IS THE FINAL REPORT | No Device Problem Found | Medical Evaluation result code
7811902 | 129899323 | N | DEVICE ALARMED MOTOR ERROR DURING OCCLUSION TEST DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7814604 | 129902095 | N | NO UNEXPECTED INVERSE CRITICAL BLOCK DID NOT ADD TO ZERO ERROR ALARMS DURING TESTING HOWEVER INVERSE CRITICAL BLOCK DID NOT ADD TO ZERO ERROR ALARM IS FOUND IN THE FORMATTED HISTORY FILE DUE TO A SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
7684514 | 129929011 | N | INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARM NOTED MOTOR PASSED MOTOR TEST B | No Device Problem Found | Medical Evaluation result code
7712684 | 129929449 | N | DEVICE RECEIVED WITH CONSTANT MOTOR ARM FAILED SELFTEST ALARM DUE TO A FAULTY Y ON THE RF BOARD UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO MOTOR ARM FAILED SELFTEST ALARM UNABLE TO REPLACE COMPONENT DUE TO LEAD FREE BOARD | No Device Problem Found | Medical Evaluation result code
7612783 | 129929605 | N | DEVICE RECEIVED WITH MOTION SENSOR TEST FAILURE ALARMED DURING REWIND DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM OPERATING CURRENTS SELFTEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST OR VERIFY MOTOR ERROR ALARM DUE TO MOTION SENSOR TEST FAILURE ALARMS | No Device Problem Found | Medical Evaluation result code
7815309 | 129946796 | N | DEVICE PASSED THE DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST AND OCCLUSION TEST NO UNEXPECTED INSULIN FLOW BLOCKED ALARM OR NO DELIVERY ALARM OR DRIVE COUNT OR TIME VALUES OR CHANGES INCORRECT FOR MOVING OR COASTING ERROR NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7811883 | 129946955 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE PASSED SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DEVICE RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Scratched Material | Device problem code
7811883 | 129946955 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE PASSED SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DEVICE RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Component Missing | Device problem code
7811883 | 129946955 | N | UNABLE TO PERFORM DISPLACEMENT TEST DUE TO MISSING RETAINER THE RESERVOIR DOES NOT STAY LOCKED IN DUE TO MISSING RETAINER DEVICE PASSED SELF TEST NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED DEVICE RECEIVED WITH MINOR SCRATCHES ON LCD WINDOW | Missing or Inadequate Safety Measures | Medical Evaluation result code
7815129 | 129948568 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST OCCLUSION TEST SELF TEST AND UNEXPECTED RESTART ERROR TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED DEVICE PASSED ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFIC RANGE AND PASSED OFF NO POWER TEST DEVICE WAS MONITORED AND TESTED WITH NO OFF NO POWER ALERT NOTED B | No Device Problem Found | Medical Evaluation result code
7628577 | 129948656 | N | DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING UNABLE TO PERFORM OCCLUSION TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST AND PRIME TEST | Mechanical Jam | Device problem code
7628577 | 129948656 | N | DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING UNABLE TO PERFORM OCCLUSION TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST AND PRIME TEST | No Device Problem Found | Medical Evaluation result code
7812178 | 130057135 | N | INSULIN PUMP PASSED DISPLACEMENT TEST AND SELF TEST NO FROZEN DISPLAY NOTED DURING TESTING B | No Device Problem Found | Medical Evaluation result code
7815500 | 130060652 | N | DEVICE PASSED DISPLACEMENT TEST UNEXPECTED RESTART ERROR TEST BASIC OCCLUSION OCCLUSION TEST PRIME AND EXCESSIVE NO DELIVERY TEST DEVICE WAS MONITORED AND FUNCTION PROPERLY NOTED NO EXCESSIVE NO DELIVERY ALARM NOTED | No Device Problem Found | Medical Evaluation result code
7815484 | 130061482 | N | NO DELIVERY ALARM OR OCCLUSION DURING THERAPY FOUND DURING TESTING DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST B | No Device Problem Found | Medical Evaluation result code
7812238 | 130061827 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT TEST ACTIVE CURRENT TEST AND SELFTEST NO UNEXPECTED SOFTWARE ERROR DETECTED ALARMS NOTED DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED SOFTWARE ERROR DETECTED ALARM LINE NUMBER FILE NUMBER DUE TO A SOFTWARE ERROR PUMP COMMUNICATED PROPERLY WITH TEST GUARDIAN LINK TRANSMITTER NO AUTO MODE ANOMALY NOTED DURING TESTING B | Computer Software Problem | Device problem code
7812238 | 130061827 | N | THE INSULIN PUMP PASSED THE DISPLACEMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT TEST ACTIVE CURRENT TEST AND SELFTEST NO UNEXPECTED SOFTWARE ERROR DETECTED ALARMS NOTED DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED SOFTWARE ERROR DETECTED ALARM LINE NUMBER FILE NUMBER DUE TO A SOFTWARE ERROR PUMP COMMUNICATED PROPERLY WITH TEST GUARDIAN LINK TRANSMITTER NO AUTO MODE ANOMALY NOTED DURING TESTING B | Software Problem Identified | Medical Evaluation result code
7581373 | 130541325 | N | INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | Mechanical Jam | Device problem code
7581373 | 130541325 | N | INSULIN PUMP PASSED THE REWIND BASIC OCCLUSION PRIME AND DISPLACEMENT TEST INSULIN PUMP RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AND INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | No Device Problem Found | Medical Evaluation result code
7641047 | 130543910 | N | DEVICE PASSED THE DISPLACEMENT TEST DEVICE RECEIVED WITH DETACHED OR LOOSE RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE NOT CONFIRMED B | Mechanical Jam | Device problem code
7641047 | 130543910 | N | DEVICE PASSED THE DISPLACEMENT TEST DEVICE RECEIVED WITH DETACHED OR LOOSE RETAINER RESERVOIR UNABLE TO LOCK INTO PLACE NOT CONFIRMED B | No Device Problem Found | Medical Evaluation result code
7716796 | 130544998 | N | THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS SEVERED NEAR THE ARRAY PRIOR TO RECEIPT THIS IS BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION SYSTEM LOCK WAS VERIFIED THE ELECTRODE CONDITION PREVENTED SOME OF THE ELECTRICAL TESTS FROM BEING PERFORMED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | No Device Problem Found | Medical Evaluation result code
7650155 | 130658813 | N | INSULIN PUMP PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST NO MOTOR ERROR ALARM NOTED THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
7867119 | 130666797 | N | THE DEVICE PASSED THE EXTERNAL VISUAL AND PHOTOGRAPHIC IMAGING INSPECTIONS SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Jam | Device problem code
7867119 | 130666797 | N | THE DEVICE PASSED THE EXTERNAL VISUAL AND PHOTOGRAPHIC IMAGING INSPECTIONS SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | No Device Problem Found | Medical Evaluation result code
7675727 | 130666949 | N | THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS SEVERED AS WELL AS CUT SILASTIC OVERMOLD ON THE TOP AND BOTTOM COVERS THESE ARE BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | Mechanical Jam | Device problem code
7675727 | 130666949 | N | THE EXTERNAL VISUAL INSPECTION REVEALED THE ELECTRODE WAS SEVERED AS WELL AS CUT SILASTIC OVERMOLD ON THE TOP AND BOTTOM COVERS THESE ARE BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION SYSTEM LOCK WAS VERIFIED THE DEVICE PASSED THE ELECTRICAL AND MECHANICAL TESTS PERFORMED THIS DEVICE WAS EXPLANTED FOR MEDICAL REASONS THE DEVICE PASSED THE TESTS PERFORMED THIS IS THE FINAL REPORT | No Device Problem Found | Medical Evaluation result code
7984318 | 130673726 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7984318 | 130673726 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
6026220 | 130685725 | D | ADDITIONAL INFORMATION RECEIVED APPROXIMATELY TWO YEARS LATER INDICATED THAT THE LEAD WAS EXPLANTED AND RETURNED FOR ANALYSIS NO ADDITIONAL INFORMATION WAS PROVIDED REGARDING THE REASON FOR THE EXPLANT | No Information | Patient problem code
6026220 | 130685725 | D | ADDITIONAL INFORMATION RECEIVED APPROXIMATELY TWO YEARS LATER INDICATED THAT THE LEAD WAS EXPLANTED AND RETURNED FOR ANALYSIS NO ADDITIONAL INFORMATION WAS PROVIDED REGARDING THE REASON FOR THE EXPLANT | Insufficient Information | Device Problem Code
8085042 | 130689847 | D | IT WAS REPORTED THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8085042 | 130689847 | D | IT WAS REPORTED THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8085404 | 130689879 | D | IT WAS REPORTED THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS EXPLANTED AND RETURNED FOR ANALYSIS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8085404 | 130689879 | D | IT WAS REPORTED THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR ICD WAS EXPLANTED AND RETURNED FOR ANALYSIS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7436537 | 130689903 | D | ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE SYSTEM WAS REMOVED SUCCESSFULLY DUE TO POSSIBLE INFECTION THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER CRTP WAS EXPLANTED | No Consequences Or Impact To Patient | Patient problem code
7436537 | 130689903 | D | ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE SYSTEM WAS REMOVED SUCCESSFULLY DUE TO POSSIBLE INFECTION THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER CRTP WAS EXPLANTED | Insufficient Information | Device Problem Code
7617959 | 131125080 | N | INSPECTED ONE OPENED AND USED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
7617959 | 131125080 | N | INSPECTED ONE OPENED AND USED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Incompatible Component/ Accessory | Medical Evaluation result code
7628328 | 131302901 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
7628328 | 131302901 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Incompatible Component/ Accessory | Medical Evaluation result code
7815506 | 131303061 | N | B THE INSULIN PUMP PASSED DISPLACEMENT TEST ACTIVE CURRENT TEST SLEEP CURRENT TEST AND SELFTEST NO KEYPAD ANOMALIES NOTED DURING TESTING KEYPAD UNRESPONSIVEBUTTON ERROR ALARM NOT CONFIRMED NO TRACE POINTERS ARE INVALID ALARMS NOTED DURING TESTING TRACE POINTERS ARE INVALID NOT CONFIRMED | No Device Problem Found | Medical Evaluation result code
7612537 | 131371205 | N | INSPECTED ONE OPENED AND USED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
7612537 | 131371205 | N | INSPECTED ONE OPENED AND USED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | No Device Problem Found | Medical Evaluation result code
7628862 | 131382576 | N | DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ERROR CONFIRMED IN THE HISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | Mechanical Jam | Device Problem Code
7628862 | 131382576 | N | DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUS DELIVERY AND MOTOR POSITION ENCODER ERROR CONFIRMED IN THE HISTORY FILE THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING | No Device Problem Found | Medical Evaluation result code
7717592 | 131406245 | N | DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING UNABLE TO PERFORM OCCLUSION TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST AND PRIME TEST | Mechanical Jam | Device problem code
7717592 | 131406245 | N | DEVICE RECEIVED WITH STUCK MOTOR ERROR ALARM LOOP DURING BOLUS DELIVERY THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING TESTING UNABLE TO PERFORM OCCLUSION TEST AND EXCESSIVE NO DELIVERY TEST DUE TO MOTOR ERROR ALARMS DEVICE PASSED DISPLACEMENT TEST REWIND TEST BASIC OCCLUSION TEST AND PRIME TEST | No Findings Available | Medical Evaluation result code
7689323 | 131406265 | N | DEVICE HAD INTERMITTENT BUTTONS RESPONSE DUE TO FLATTENED CREASED ESC ACT AND UP BUTTONS DOME SWITCH NO UNLOCKED LCD KEYPAD CONNECTOR NOTED | Insufficient Information | Device problem code
7735701 | 131406302 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER | Component Missing | Device problem code
7735701 | 131406302 | N | INSULIN PUMP RECEIVED WITH MISSING RETAINER UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER | No Findings Available | Medical Evaluation result code
7702043 | 131406836 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST POWER ERROR DUE TO CONNECTOR RESISTANCE ISSUE | No Device Problem Found | Medical Evaluation result code
7725744 | 131407117 | N | DEVICE RECEIVED WITH MOTOR ERROR ALARM DURING REWIND TEST DUE TO MOTOR ENCODER SIGNAL OUT OF PHASE UNABLE TO PERFORM BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST AND EXCESSIVE NO DELIVERY TEST OR VERIFY REWIND ANOMALY DUE TO MOTOR ERROR ALARM | No Device Problem Found | Medical Evaluation result code
7604780 | 131409132 | N | DEVICE PASSED ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED POWER ERROR DUE TO J CONNECTOR RESISTANCE ISSUE DEVICE FAILED SLEEP CURRENT MEASUREMENT DUE TO POWER ERROR | Electrical Power Problem | Device problem code
7604780 | 131409132 | N | DEVICE PASSED ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE RECEIVED POWER ERROR DUE TO J CONNECTOR RESISTANCE ISSUE DEVICE FAILED SLEEP CURRENT MEASUREMENT DUE TO POWER ERROR | Power Source Problem Identified | Medical Evaluation result code
7672785 | 131654619 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP O RING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER | Component Missing | Device problem code
7672785 | 131654619 | N | DEVICE RECEIVED WITH MISSING RETAINER RING AND RESERVOIR TUBE LIP O RING UNABLE TO PERFORM DISPLACEMENT TEST AND PCAP OR RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
7900687 | 131732637 | N | ONE F FASTCATH INTRODUCER SHEATH AND DILATOR WERE RECEIVED FOR EVALUATION THE SIDEPORT TUBING WAS NO LONGER ATTACHED TO THE HEMOSTASIS BODY NO OTHER VISUAL ANOMALIES WERE NOTED THE PROXIMAL END OF THE SIDEPORT TUBING WAS MICROSCOPICALLY INSPECTED A SMALL AMOUNT OF ADHESIVE WAS NOTED ON THE SIDEPORT TUBING AT THE LOCATION WHERE THE HEMOSTASIS BODY MET THE TUBING NO TEARING OR STRETCHING OF THE TUBING WAS NOTED THE EXTENSION TUBING DETACHED FROM THE SIDEPORT OF THE HEMOSTASIS BODY DUE TO AN INADEQUATE SOLVENT COATING LENGTH ON THE EXTENSION TUBING | Stretched | Device problem code
7900687 | 131732637 | N | ONE F FASTCATH INTRODUCER SHEATH AND DILATOR WERE RECEIVED FOR EVALUATION THE SIDEPORT TUBING WAS NO LONGER ATTACHED TO THE HEMOSTASIS BODY NO OTHER VISUAL ANOMALIES WERE NOTED THE PROXIMAL END OF THE SIDEPORT TUBING WAS MICROSCOPICALLY INSPECTED A SMALL AMOUNT OF ADHESIVE WAS NOTED ON THE SIDEPORT TUBING AT THE LOCATION WHERE THE HEMOSTASIS BODY MET THE TUBING NO TEARING OR STRETCHING OF THE TUBING WAS NOTED THE EXTENSION TUBING DETACHED FROM THE SIDEPORT OF THE HEMOSTASIS BODY DUE TO AN INADEQUATE SOLVENT COATING LENGTH ON THE EXTENSION TUBING | Detachment of Device or device Component | Device problem code
7905864 | 131737275 | N | B WE WERE NOTIFIED THAT THIS LEAD WAS EXPLANTED B DUE TO NOISE | Noise, Audible | Device problem code
7905864 | 131737275 | N | B WE WERE NOTIFIED THAT THIS LEAD WAS EXPLANTED B DUE TO NOISE | Non-Functional Defect | Medical Evaluation result code
3176691 | 131833258 | D | PPF ALLEGES FRACTUREBONE AND ELEVATED METAL IONS | Toxicity | Patient problem code
3176691 | 131833258 | D | PPF ALLEGES FRACTUREBONE AND ELEVATED METAL IONS | Fracture | Patient problem code
3176691 | 131833258 | D | PPF ALLEGES FRACTUREBONE AND ELEVATED METAL IONS | Insufficient Information | Device Problem Code
8079087 | 132055124 | D | NEW INFORMATION RECEIVED ON B INDICATING THAT THE PATIENT WAS STABLE POST PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
8079087 | 132055124 | D | NEW INFORMATION RECEIVED ON B INDICATING THAT THE PATIENT WAS STABLE POST PROCEDURE | Insufficient Information | Device Problem Code
8135936 | 132338152 | D | NEW INFORMATION NOTED THAT THE DEVICE WAS REPROGRAMMED ON B THE PATIENT WAS FINE | No Consequences Or Impact To Patient | Patient problem code
8135936 | 132338152 | D | NEW INFORMATION NOTED THAT THE DEVICE WAS REPROGRAMMED ON B THE PATIENT WAS FINE | Insufficient Information | Device Problem Code
8170443 | 132342514 | D | NEW INFORMATION REPORTED THAT THE LEAD WAS REPOSITIONED ON B THE PATIENT WAS STABLE POSTPROCEDURE | No Consequences Or Impact To Patient | Patient problem code
8170443 | 132342514 | D | NEW INFORMATION REPORTED THAT THE LEAD WAS REPOSITIONED ON B THE PATIENT WAS STABLE POSTPROCEDURE | Insufficient Information | Device Problem Code
8176138 | 132342748 | D | NEW INFORMATION REPORTED THE PATIENT WAS STABLE POSTPROCEDURE | No Consequences Or Impact To Patient | Patient problem code
8176138 | 132342748 | D | NEW INFORMATION REPORTED THE PATIENT WAS STABLE POSTPROCEDURE | Insufficient Information | Device Problem Code
8175890 | 132343099 | D | NEW INFORMATION REPORTED THE DEVICE WAS REMOVED AND REPLACED ON B THE PATIENT WAS STABLE POSTPROCEDURE | No Consequences Or Impact To Patient | Patient problem code
8175890 | 132343099 | D | NEW INFORMATION REPORTED THE DEVICE WAS REMOVED AND REPLACED ON B THE PATIENT WAS STABLE POSTPROCEDURE | Insufficient Information | Device Problem Code
8198213 | 132343525 | D | OLYMPUS WAS INFORMED THAT DURING A POST MARKET SURVEILLANCE STUDY THE SCOPE CULTURED B FOR MICROORGANISMS B AFTER REPROCESSING THERE WERE NO ASSOCIATED PATIENT INFECTIONS REPORTED | No Consequences Or Impact To Patient | Patient problem code
8198213 | 132343525 | D | OLYMPUS WAS INFORMED THAT DURING A POST MARKET SURVEILLANCE STUDY THE SCOPE CULTURED B FOR MICROORGANISMS B AFTER REPROCESSING THERE WERE NO ASSOCIATED PATIENT INFECTIONS REPORTED | Insufficient Information | Device Problem Code
8084522 | 132402008 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN WILL CONTINUE TO MONITOR THE PATIENT THIS RIGHT VENTRICULAR RV LEAD REMAINS IN SERVICE AND THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED | No Consequences Or Impact To Patient | Patient problem code
8084522 | 132402008 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN WILL CONTINUE TO MONITOR THE PATIENT THIS RIGHT VENTRICULAR RV LEAD REMAINS IN SERVICE AND THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED | Insufficient Information | Device Problem Code
6501634 | 132402782 | D | NEW INFORMATION RECEIVED STATED THAT THE LEFT VENTRICULAR LEAD WOULD NOT REMAIN IN A STABLE POSITION IN THE CORONARY SINUS AND THEREFORE IT WAS EXPLANTED SAME DAY THE PATIENT WAS STABLE BEFORE DURING AND AFTER THE PROCEDURE AND THERE WERE NO ADVERSE CONSEQUENCES | No Consequences Or Impact To Patient | Patient problem code
6501634 | 132402782 | D | NEW INFORMATION RECEIVED STATED THAT THE LEFT VENTRICULAR LEAD WOULD NOT REMAIN IN A STABLE POSITION IN THE CORONARY SINUS AND THEREFORE IT WAS EXPLANTED SAME DAY THE PATIENT WAS STABLE BEFORE DURING AND AFTER THE PROCEDURE AND THERE WERE NO ADVERSE CONSEQUENCES | Insufficient Information | Device Problem Code
8124486 | 132427958 | D | ACCORDING TO THE REPORTER DURING COLON SURGERY THE DEVICE NOT SEALING NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
8124486 | 132427958 | D | ACCORDING TO THE REPORTER DURING COLON SURGERY THE DEVICE NOT SEALING NO PATIENT INJURY | Unsealed Device Packaging | Device Problem Code
8201478 | 132567436 | D | THE CUSTOMER OBSERVED FALSELY ELEVATED IGG ANTIBODY TO B IGG ANTIHAV RESULTS WHILE USING ARCHITECT HAVABIGG REAGENTS THE FOLLOWING DATA WAS PROVIDED SID B V REPEAT USING LOT LI B USING LOT LI INITIAL B REPEAT B NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED | No Consequences Or Impact To Patient | Patient problem code
8201478 | 132567436 | D | THE CUSTOMER OBSERVED FALSELY ELEVATED IGG ANTIBODY TO B IGG ANTIHAV RESULTS WHILE USING ARCHITECT HAVABIGG REAGENTS THE FOLLOWING DATA WAS PROVIDED SID B V REPEAT USING LOT LI B USING LOT LI INITIAL B REPEAT B NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED | Insufficient Information | Device Problem Code
7981956 | 132590993 | D | ADDITIONAL INFORMATION WAS RECEIVED CONFIRMING THIS DEVICE WAS EXPLANTED AFTER RECORDING A CODE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7981956 | 132590993 | D | ADDITIONAL INFORMATION WAS RECEIVED CONFIRMING THIS DEVICE WAS EXPLANTED AFTER RECORDING A CODE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7698011 | 132593953 | D | PPF ALLEGES PSEUDOTUMOR METAL WEARMETALLOSIS AND ELEVATED METAL IONS ADDED FACILITY NAME EVENT YEAR EVENT DATE EVENT MONTH PATIENT MIDDLE INITIAL AGE DOB IP PRODUCT CODES LOT NUMBERS DOI DOR PT CODES PE CODES NEW IP STEM AND NEW ASSOCIATED CONTACTS | Toxicity | Patient problem code
7698011 | 132593953 | D | PPF ALLEGES PSEUDOTUMOR METAL WEARMETALLOSIS AND ELEVATED METAL IONS ADDED FACILITY NAME EVENT YEAR EVENT DATE EVENT MONTH PATIENT MIDDLE INITIAL AGE DOB IP PRODUCT CODES LOT NUMBERS DOI DOR PT CODES PE CODES NEW IP STEM AND NEW ASSOCIATED CONTACTS | Appropriate Term/Code Not Available | Device Problem Code
7072350 | 132609841 | D | ACCORDING TO THE REPORTER THE UNITS BIPOLAR WAS ALWAYS ON THERE WAS NO PATIENT INJURY | No Consequences Or Impact To Patient | Patient problem code
7072350 | 132609841 | D | ACCORDING TO THE REPORTER THE UNITS BIPOLAR WAS ALWAYS ON THERE WAS NO PATIENT INJURY | Insufficient Information | Device Problem Code
8085377 | 132627811 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8085377 | 132627811 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7979907 | 132630752 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT INDICATED THIS DEVICE WAS EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7979907 | 132630752 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT INDICATED THIS DEVICE WAS EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7981215 | 132632073 | D | SUBSEQUENT INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND RETURNED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7981215 | 132632073 | D | SUBSEQUENT INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND RETURNED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
3484469 | 132636870 | D | PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
3484469 | 132636870 | D | PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
4536746 | 132639139 | D | PPF ALLEGES METALLOSIS | Toxicity | Patient problem code
4536746 | 132639139 | D | PPF ALLEGES METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
7478031 | 132647705 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THIS ICD WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7478031 | 132647705 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THIS ICD WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7194584 | 132647963 | D | BOSTON SCIENTIFIC RECEIVED ADDITIONAL INFORMATION THAT THIS DEVICE WAS EXPLANTED DUE TO THERAPY UPGRADEDOWNGRADE NO ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7194584 | 132647963 | D | BOSTON SCIENTIFIC RECEIVED ADDITIONAL INFORMATION THAT THIS DEVICE WAS EXPLANTED DUE TO THERAPY UPGRADEDOWNGRADE NO ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
4131477 | 133114927 | D | ADDITIONAL INFORMATION PROVIDED FROM THE FIELD INDICATED SURGICAL INTERVENTION WAS LATER PERFORMED AND THE RV LEAD WAS ABANDONED DUE TO A CONTINUATION OF HIGH PACING IMPEDANCES THE PACEMAKER WAS ALSO EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
4131477 | 133114927 | D | ADDITIONAL INFORMATION PROVIDED FROM THE FIELD INDICATED SURGICAL INTERVENTION WAS LATER PERFORMED AND THE RV LEAD WAS ABANDONED DUE TO A CONTINUATION OF HIGH PACING IMPEDANCES THE PACEMAKER WAS ALSO EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Pacing Problem | Device Problem Code
7304363 | 133234766 | D | IT WAS REPORTED THAT THE PATIENT HAD PAINFUL SCREW HEADS OF MEDIAL ANKLE AFTER INITIAL LEFT TOTAL ANKLE PROCEDURE NO INTERVENTION HAS BEEN REPORTED TO DATE | Pain | Patient problem code
7304363 | 133234766 | D | IT WAS REPORTED THAT THE PATIENT HAD PAINFUL SCREW HEADS OF MEDIAL ANKLE AFTER INITIAL LEFT TOTAL ANKLE PROCEDURE NO INTERVENTION HAS BEEN REPORTED TO DATE | Insufficient Information | Device Problem Code
7992235 | 133260777 | D | ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN STATED THAT THE PATIENT WAS DOING WELL THERE WERE NO FURTHER PATIENT COMPLICATIONS REPORTED IN RELATION TO THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
7992235 | 133260777 | D | ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN STATED THAT THE PATIENT WAS DOING WELL THERE WERE NO FURTHER PATIENT COMPLICATIONS REPORTED IN RELATION TO THIS EVENT | Insufficient Information | Device Problem Code
7349827 | 133269306 | D | IT WAS FURTHER REPORTED THAT THE PATIENT HAS SEVERE PAIN WHILE STANDING AND WALKING PATIENT ALSO REPORTS A LIMP WHILE WALKING AND ALLEGES SEVERE DISCOMFORT | Discomfort | Patient problem code
7349827 | 133269306 | D | IT WAS FURTHER REPORTED THAT THE PATIENT HAS SEVERE PAIN WHILE STANDING AND WALKING PATIENT ALSO REPORTS A LIMP WHILE WALKING AND ALLEGES SEVERE DISCOMFORT | Pain | Patient problem code
7349827 | 133269306 | D | IT WAS FURTHER REPORTED THAT THE PATIENT HAS SEVERE PAIN WHILE STANDING AND WALKING PATIENT ALSO REPORTS A LIMP WHILE WALKING AND ALLEGES SEVERE DISCOMFORT | Insufficient Information | Device Problem Code
7204128 | 133362857 | D | IT WAS REPORTED THAT POST TOTAL HIP ARTHROPLASTY XRAYS SHOWED RADIOLUCENCY OF UNKNOWN MEASUREMENT AROUND THE STEM MEDIALLY TIMEFRAME BETWEEN INITIAL PROCEDURE AND XRAYS ARE UNKNOWN ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | No Code Available | Patient problem code
7204128 | 133362857 | D | IT WAS REPORTED THAT POST TOTAL HIP ARTHROPLASTY XRAYS SHOWED RADIOLUCENCY OF UNKNOWN MEASUREMENT AROUND THE STEM MEDIALLY TIMEFRAME BETWEEN INITIAL PROCEDURE AND XRAYS ARE UNKNOWN ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME | Insufficient Information | Device Problem Code
7990360 | 133366678 | D | ADDITIONAL INFORMATION INDICATED THAT THIS PACEMAKER WAS EXPLANTED DUE TO B THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE PACEMAKER WAS EXPLANTED | No Consequences Or Impact To Patient | Patient problem code
7990360 | 133366678 | D | ADDITIONAL INFORMATION INDICATED THAT THIS PACEMAKER WAS EXPLANTED DUE TO B THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE PACEMAKER WAS EXPLANTED | Insufficient Information | Device Problem Code
7704970 | 133550378 | N | OPEN BOOK OR CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM OPEN BOOK PROBLEM ISOLATED TO FORCE SENSOR | Component or Accessory Incompatibility | Device problem code
7704970 | 133550378 | N | OPEN BOOK OR CRITICAL PUMP ERROR AFTER BATTERY INSTALLATION CONSTANT CRITICAL PUMP ERROR ALARM OPEN BOOK PROBLEM ISOLATED TO FORCE SENSOR | No Device Problem Found | Medical Evaluation result code
7979093 | 133552234 | D | ADDITIONAL INFORMATION WAS RECEIVED INDICATING A REVISION PROCEDURE WAS PERFORMED WHEREIN THE DEVICE WAS EXPLANTED AND REPLACED ADDITIONALLY A NEW RV LEAD WAS IMPLANTED NO ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
7979093 | 133552234 | D | ADDITIONAL INFORMATION WAS RECEIVED INDICATING A REVISION PROCEDURE WAS PERFORMED WHEREIN THE DEVICE WAS EXPLANTED AND REPLACED ADDITIONALLY A NEW RV LEAD WAS IMPLANTED NO ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7990460 | 133554005 | D | ADDITIONAL INFORMATION INDICATED THAT THIS RIGHT VENTRICULAR RV LEAD WAS EXPLANTED DUE TO B THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE RV LEAD WAS EXPLANTED | No Consequences Or Impact To Patient | Patient problem code
7990460 | 133554005 | D | ADDITIONAL INFORMATION INDICATED THAT THIS RIGHT VENTRICULAR RV LEAD WAS EXPLANTED DUE TO B THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED THE RV LEAD WAS EXPLANTED | Insufficient Information | Device Problem Code
8002729 | 133564050 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE FOR THE HIGH RIGHT ATRIAL RA THRESHOLD MEASUREMENTS REMAINS UNKNOWN THE PATIENT WILL CONTINUE TO BE MONITORED NO ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8002729 | 133564050 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE FOR THE HIGH RIGHT ATRIAL RA THRESHOLD MEASUREMENTS REMAINS UNKNOWN THE PATIENT WILL CONTINUE TO BE MONITORED NO ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
7645889 | 133867356 | D | IT WAS FURTHER REPORTED THAT THE PHYSICIAN STATED THERE WAS NO EVIDENCE OF INFECTION HOWEVER THE PATIENT IS REACHING OUT TO GET A SECOND OPINION | No Information | Patient problem code
7645889 | 133867356 | D | IT WAS FURTHER REPORTED THAT THE PHYSICIAN STATED THERE WAS NO EVIDENCE OF INFECTION HOWEVER THE PATIENT IS REACHING OUT TO GET A SECOND OPINION | Insufficient Information | Device Problem Code
7599889 | 133970174 | N | DEVICE PASSED THE SELF TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE INSULIN FLOW BLOCKED ALARM OR OCCLUSION ALARM FUNCTIONED PROPERLY DURING THE BASIC OCCLUSION OCCLUSION AND FORCE SENSOR TESTS NO UNEXPECTED INSULIN FLOW BLOCKED ALARM NOTED DURING TESTING | Partial Blockage | Device problem code
7599889 | 133970174 | N | DEVICE PASSED THE SELF TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND THE DISPLACEMENT TEST THE INSULIN FLOW BLOCKED ALARM OR OCCLUSION ALARM FUNCTIONED PROPERLY DURING THE BASIC OCCLUSION OCCLUSION AND FORCE SENSOR TESTS NO UNEXPECTED INSULIN FLOW BLOCKED ALARM NOTED DURING TESTING | No Device Problem Found | Medical Evaluation result code
2982363 | 133975152 | D | PPF ALLEGES METALLOSIS | Toxicity | Patient problem code
2982363 | 133975152 | D | PPF ALLEGES METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
6100199 | 133976841 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS | Toxicity | Patient problem code
6100199 | 133976841 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
6887901 | 133982331 | D | IN ADDITION TO WHAT WAS PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS UPDATED PART AND LOT NUMBER OF THE HEAD LINER AND STEM PATIENT HARMS DATE OF IMPLANT AND DATE OF REVISION ADDED PATIENTS AGE REVISION HOSPITAL REVISION SURGEON AND LAW FIRM IN THE FACILITY NAME | Toxicity | Patient problem code
6887901 | 133982331 | D | IN ADDITION TO WHAT WAS PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS UPDATED PART AND LOT NUMBER OF THE HEAD LINER AND STEM PATIENT HARMS DATE OF IMPLANT AND DATE OF REVISION ADDED PATIENTS AGE REVISION HOSPITAL REVISION SURGEON AND LAW FIRM IN THE FACILITY NAME | Appropriate Term/Code Not Available | Device Problem Code
7564234 | 133995192 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Component or Accessory Incompatibility | Device problem code
7564234 | 133995192 | N | INSPECTED OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED PER SPECIFICATIONS | Incompatible Component/ Accessory | Medical Evaluation result code
5097906 | 134035964 | D | PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
5097906 | 134035964 | D | PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
6853630 | 134059680 | N | INSULIN PUMP RECEIVED WITH NO BUTTON RESPONSE DUE TO FLATTENED BUTTON DOME SWITCH AND PARTIAL UNLOCKED KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM IDLE CURRENT MEASUREMENT TEST RUN CURRENT MEASUREMENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST DUE TO NO BUTTON RESPONSE | Defective Component | Device problem code
6853630 | 134059680 | N | INSULIN PUMP RECEIVED WITH NO BUTTON RESPONSE DUE TO FLATTENED BUTTON DOME SWITCH AND PARTIAL UNLOCKED KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT UNABLE TO PERFORM IDLE CURRENT MEASUREMENT TEST RUN CURRENT MEASUREMENT TEST SELF TEST OFF NO POWER TEST AND DISPLACEMENT TEST DUE TO NO BUTTON RESPONSE | No Findings Available | Medical Evaluation result code
7639957 | 134064129 | N | B IT WAS REPORTED THAT PATIENT UNDERWENT MESH REVISION ON B AND XENMATRIX SURGICAL GRAFT WAS IMPLANTED DUE TO RECURRENT HERNIA BOWEL OBSTRUCTION AND ABDOMINAL PAIN NO ADDITIONAL INFORMATION WAS PROVIDED | No Device Problem Found | Medical Evaluation result code
2575654 | 134109229 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
2575654 | 134109229 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
5681742 | 134112751 | D | THE PATIENT HARM WAS UPDATED AND ADDED THE STEM TO THE IMPACTED PRODUCT DUE TO ALLEGED ELEVATED METAL IONS LAWYER HOSPITAL NAME LAW FIRM WERE ALSO ADDED AND CHANGE THE REPORT SOURCE TO LEGAL FROM OTHER | Toxicity | Patient problem code
5681742 | 134112751 | D | THE PATIENT HARM WAS UPDATED AND ADDED THE STEM TO THE IMPACTED PRODUCT DUE TO ALLEGED ELEVATED METAL IONS LAWYER HOSPITAL NAME LAW FIRM WERE ALSO ADDED AND CHANGE THE REPORT SOURCE TO LEGAL FROM OTHER | Insufficient Information | Device Problem Code
2872426 | 134115029 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES INFECTION | Infection | Patient problem code
2872426 | 134115029 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES INFECTION | Insufficient Information | Device Problem Code
8005860 | 134202709 | D | IT WAS REPORTED THAT THE BD VACUTAINER BUFFERED SODIUM CITRATE NC BLOOD COLLECTION TUBES OVERFILLED THERE WAS NO REPORT OF EXPOSURE INJURY OR MEDICAL INTERVENTION | No Consequences Or Impact To Patient | Patient Problem Code
8005860 | 134202709 | D | IT WAS REPORTED THAT THE BD VACUTAINER BUFFERED SODIUM CITRATE NC BLOOD COLLECTION TUBES OVERFILLED THERE WAS NO REPORT OF EXPOSURE INJURY OR MEDICAL INTERVENTION | Overfill | Device Problem Code
5854763 | 134205507 | D | PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
5854763 | 134205507 | D | PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
8110279 | 134207895 | D | ADDITIONAL REPORT OF FIRM AND PAINFUL BREAST TREATMENT WITH MONTELUKAST WAS GIVEN THE DEVICE HAS BEEN EXPLANTED | Pain | Patient problem code
8110279 | 134207895 | D | ADDITIONAL REPORT OF FIRM AND PAINFUL BREAST TREATMENT WITH MONTELUKAST WAS GIVEN THE DEVICE HAS BEEN EXPLANTED | Insufficient Information | Device Problem Code
4875832 | 134209653 | D | PPF ALLEGES ELEVATED METAL IONS ADDED AGE LAW FIRM REVISION SURGEON AND HOSPITAL UPDATED PRODUCT DETAILS IN IPS MANUFACTURED AND EXPIRATION DATE CORRECTED PATIENT IDENTIFIER | Toxicity | Patient problem code
4875832 | 134209653 | D | PPF ALLEGES ELEVATED METAL IONS ADDED AGE LAW FIRM REVISION SURGEON AND HOSPITAL UPDATED PRODUCT DETAILS IN IPS MANUFACTURED AND EXPIRATION DATE CORRECTED PATIENT IDENTIFIER | Insufficient Information | Device Problem Code
3627869 | 134220868 | D | PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS ADDED AGE LAWYER SURGEON UPDATED PRODUCT DETAILS AND CORRECTED PATIENT IDENTIFIER ADDED STEM DUE TO ALLEGED ELEVATED METAL IONS | Toxicity | Patient problem code
3627869 | 134220868 | D | PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS ADDED AGE LAWYER SURGEON UPDATED PRODUCT DETAILS AND CORRECTED PATIENT IDENTIFIER ADDED STEM DUE TO ALLEGED ELEVATED METAL IONS | Appropriate Term/Code Not Available | Device Problem Code
6710057 | 134225051 | D | IT WAS REPORTED THAT THE REASON FOR THE REVISION SURGERY IS PATIENTS EXTREMELY LIMITED RANGE OF MOTION AND PAIN NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME | Pain | Patient problem code
6710057 | 134225051 | D | IT WAS REPORTED THAT THE REASON FOR THE REVISION SURGERY IS PATIENTS EXTREMELY LIMITED RANGE OF MOTION AND PAIN NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME | Range Of Motion, Loss Of | Patient problem code
6710057 | 134225051 | D | IT WAS REPORTED THAT THE REASON FOR THE REVISION SURGERY IS PATIENTS EXTREMELY LIMITED RANGE OF MOTION AND PAIN NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME | Insufficient Information | Device Problem Code
2390483 | 134228921 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS BEFORE FIRST REVISION UPDATED LOT NUMBER OF THE STEM AND PATIENT HARMS ADDED REVISION HOSPITAL REVISION SURGEON LAWYER IN THE ASSOCIATED CONTACT AND LAW FIRM IN THE FACILITY NAME | Toxicity | Patient problem code
2390483 | 134228921 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS BEFORE FIRST REVISION UPDATED LOT NUMBER OF THE STEM AND PATIENT HARMS ADDED REVISION HOSPITAL REVISION SURGEON LAWYER IN THE ASSOCIATED CONTACT AND LAW FIRM IN THE FACILITY NAME | Insufficient Information | Device Problem Code
2966226 | 134229968 | D | IN ADDITION TO WHAT PREVIOUSLY ALLEGED PPF ELEVATED METAL IONS ADDED STEM DUE TO NEW ALLEGATION UPDATED PATIENT HARM ADDED EXPIRATION DATE OF THE HEAD LINER STEM AND LAW FIRM IN THE FACILITY NAME | Toxicity | Patient problem code
2966226 | 134229968 | D | IN ADDITION TO WHAT PREVIOUSLY ALLEGED PPF ELEVATED METAL IONS ADDED STEM DUE TO NEW ALLEGATION UPDATED PATIENT HARM ADDED EXPIRATION DATE OF THE HEAD LINER STEM AND LAW FIRM IN THE FACILITY NAME | Insufficient Information | Device Problem Code
7710867 | 134250922 | N | ADVERSE EVENT NOT INCLUDED IN THE SUPPLEMENTAL MDR REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENTDAMAGE DEVICE NOT INCLUDED IN THE SUPPLEMENTAL MDR B | No Device Problem Found | Medical Evaluation result code
7710867 | 134254116 | N | LENS WAS EXCHANGED FOR A SIMILAR POWER AND A LARGER SIZE ON B THE PROBLEM WAS RESOLVED B | Inadequacy of Device Shape and/or Size | Device problem code
7710867 | 134254116 | N | LENS WAS EXCHANGED FOR A SIMILAR POWER AND A LARGER SIZE ON B THE PROBLEM WAS RESOLVED B | No Device Problem Found | Medical Evaluation result code
7980038 | 134387119 | N | THE LEAD WAS RETURNED DUE TO A CARDIAC PERFORATION AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE WITH THE HELIX RETRACTED VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES AFTER CLEANING THE BLOOD OFF FROM THE HELIX AND APPLYING TORQUE DIRECTLY TO THE CONNECTOR PIN THE HELIX COULD BE EXTENDED AND RETRACTED MEETING DEVICE SPECIFICATIONS A TIP STIFFNESS TEST WAS PERFORMED AND THE TEST RESULTS WERE IN SPECIFICATION | No Device Problem Found | Medical Evaluation result code
6895412 | 134561092 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST NO FAILED BATTERY TEST ALARMS NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | No Device Problem Found | Medical Evaluation result code
6656693 | 134634208 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS UPDATED PATIENT HARM EXPIRATION DATE ADDED PATIENTS AGE REVISION HOSPITAL LAW FIRM IN THE FACILITY NAME | Toxicity | Patient problem code
6656693 | 134634208 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METALLOSIS UPDATED PATIENT HARM EXPIRATION DATE ADDED PATIENTS AGE REVISION HOSPITAL LAW FIRM IN THE FACILITY NAME | Appropriate Term/Code Not Available | Device Problem Code
7737334 | 134737772 | N | B IT WAS REPORTED THAT THE PATIENT EXPERIENCED HERNIA ABDOMINAL PAIN AND OBSTRUCTION | Insufficient Information | Device problem code
4577653 | 134785945 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS BEFORE FIRST REVISION UPDATED LOT NUMBER OF THE STEM AND PATIENT HARMS ADDED REVISION HOSPITAL REVISION SURGEON LAWYER IN THE ASSOCIATED CONTACT AND LAW FIRM IN THE FACILITY NAME | Toxicity | Patient problem code
4577653 | 134785945 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES ELEVATED METAL IONS BEFORE FIRST REVISION UPDATED LOT NUMBER OF THE STEM AND PATIENT HARMS ADDED REVISION HOSPITAL REVISION SURGEON LAWYER IN THE ASSOCIATED CONTACT AND LAW FIRM IN THE FACILITY NAME | Insufficient Information | Device Problem Code
2389315 | 134786599 | D | IN ADDITION TO PREVIOUS ALLEGATIONS PPF ALLEGES METALLOSIS | Toxicity | Patient problem code
2389315 | 134786599 | D | IN ADDITION TO PREVIOUS ALLEGATIONS PPF ALLEGES METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Hypersensitivity | Patient problem code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Inflammation | Patient problem code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Pain | Patient problem code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Range Of Motion, Loss Of | Patient problem code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Swelling | Patient problem code
7451953 | 134887786 | D | IT WAS REPORTED THAT PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY AND SUBSEQUENTLY IS CONTINUING TO EXPERIENCE PAIN SWELLING LIMITED RANGE OF MOTION KNEE GIVING OUT POSSIBLE TITANIUM ALLERGY MILD SWELLING AND INFLAMMATION THE PATIENT IS UNABLE TO USE STAIRS | Insufficient Information | Device Problem Code
7578215 | 135013808 | D | PPF ALLEGES METALLOSIS DOI B DOR B RIGHT HIP | Toxicity | Patient problem code
7578215 | 135013808 | D | PPF ALLEGES METALLOSIS DOI B DOR B RIGHT HIP | Appropriate Term/Code Not Available | Device Problem Code
3308963 | 135030341 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | Toxicity | Patient problem code
3308963 | 135030341 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | No Code Available | Patient problem code
3308963 | 135030341 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | Insufficient Information | Device Problem Code
3308950 | 135045764 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | Toxicity | Patient problem code
3308950 | 135045764 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | No Code Available | Patient problem code
3308950 | 135045764 | D | PPF ALLEGES PSEUDOTUMOR AND ELEVATED METAL IONS | Insufficient Information | Device Problem Code
2442396 | 135065642 | D | PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS | Toxicity | Patient problem code
2442396 | 135065642 | D | PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS | Appropriate Term/Code Not Available | Device Problem Code
7737869 | 135365982 | D | NO ADDITIONAL EVENT INFORMATION IS AVAILABLE FOR THIS COMPLAINT SUPPLEMENTAL REQUESTED TO UPDATE REPORT SOURCE SECTION | No Information | Patient problem code
7737869 | 135365982 | D | NO ADDITIONAL EVENT INFORMATION IS AVAILABLE FOR THIS COMPLAINT SUPPLEMENTAL REQUESTED TO UPDATE REPORT SOURCE SECTION | Insufficient Information | Device Problem Code
7905038 | 135534215 | D | DEXCOM WAS MADE AWARE ON B THAT ON B DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7905038 | 135534215 | D | DEXCOM WAS MADE AWARE ON B THAT ON B DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
6836582 | 136984601 | N | INSULIN PUMP PASSED THE FUNCTIONAL TESTS INCLUDING THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST DISPLACEMENT TEST AND DELIVERY ACCURACY TEST AT INCHES | Insufficient Information | Device problem code
7692729 | 137364183 | N | INSPECTED RANDOM OPENEDUSED SENSOR AND PERFORMED CONTINUITY RESISTANCE TEST AND SENSOR FAILED TEST | Insufficient Information | Device problem code
7547118 | 137660508 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7547118 | 137660508 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7547137 | 137660645 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7547137 | 137660645 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7898082 | 137672265 | D | NO DATA WAS AVAILABLE FOR EVALUATION NO PRODUCT TESTING WAS PERFORMED AS THE UNIT WAS UNAVAILABLE BEFORE TESTING CONFIRMATION OF THE ERROR ICON AND A PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7898082 | 137672265 | D | NO DATA WAS AVAILABLE FOR EVALUATION NO PRODUCT TESTING WAS PERFORMED AS THE UNIT WAS UNAVAILABLE BEFORE TESTING CONFIRMATION OF THE ERROR ICON AND A PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7547045 | 137701734 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7547045 | 137701734 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7547033 | 137706732 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7547033 | 137706732 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7547049 | 137714429 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7547049 | 137714429 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8201453 | 137729180 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
8201453 | 137729180 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8201363 | 137731085 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE | No Information | Patient problem code
8201363 | 137731085 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8064173 | 137735317 | D | ADDITIONAL INFORMATION RECEIVED FROM A HEALTHCARE PROFESSIONAL HCP VIA A CLINICAL STUDY INDICATED THE PATIENTS BASELINE WEIGHT WAS B THE CATHETER WAS REVISED IN AN ATTEMPT TO GET BETTER PAIN RELIEF FOR THE PATIENT NO FURTHER COMPLICATIONS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8064173 | 137735317 | D | ADDITIONAL INFORMATION RECEIVED FROM A HEALTHCARE PROFESSIONAL HCP VIA A CLINICAL STUDY INDICATED THE PATIENTS BASELINE WEIGHT WAS B THE CATHETER WAS REVISED IN AN ATTEMPT TO GET BETTER PAIN RELIEF FOR THE PATIENT NO FURTHER COMPLICATIONS WERE REPORTED | Insufficient Information | Device Problem Code
7421186 | 137747082 | D | A DATA INVESTIGATION WILL NOT BE PERFORMED AS SIGNAL LOSS UP TO ONE HOUR IS WITHIN SPECIFICATION THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7421186 | 137747082 | D | A DATA INVESTIGATION WILL NOT BE PERFORMED AS SIGNAL LOSS UP TO ONE HOUR IS WITHIN SPECIFICATION THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Failure to Analyze Signal | Device Problem Code
7849489 | 137929312 | D | IT WAS REPORTED THAT THE CUSTOMER DID NOT EXPERIENCED ISSUE OF DIABETIC KETOACIDOSIS AS PER COMPLAINT NOTES AND CUSTOMER HAD HIGH BLOOD GLUCOSE LEVEL BECAUSE THEY DID NOT BOLUS ENOUGH FOR THE BREAKFAST AND CONTINUED TO EAT HARM CODE HAS BEEN UPDATED TO HYPERGLYCEMIA | Insufficient Information | Device Problem Code
7849489 | 137929312 | D | IT WAS REPORTED THAT THE CUSTOMER DID NOT EXPERIENCED ISSUE OF DIABETIC KETOACIDOSIS AS PER COMPLAINT NOTES AND CUSTOMER HAD HIGH BLOOD GLUCOSE LEVEL BECAUSE THEY DID NOT BOLUS ENOUGH FOR THE BREAKFAST AND CONTINUED TO EAT HARM CODE HAS BEEN UPDATED TO HYPERGLYCEMIA | Hyperglycemia | Patient Problem Code
5696151 | 138017818 | D | NEW INFORMATION RECEIVED STATED THAT THE PATIENT WAS IN STABLE CONDITION POST PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
5696151 | 138017818 | D | NEW INFORMATION RECEIVED STATED THAT THE PATIENT WAS IN STABLE CONDITION POST PROCEDURE | Insufficient Information | Device Problem Code
8034047 | 138074183 | D | IT WAS REPORTED THAT THE CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE HARM CODE HAS BEEN UPDATED IN ANEURYSM AND LOW BLOOD GLUCOSE AND HIGH BLOOD GLUCOSE | Aneurysm | Patient problem code
8034047 | 138074183 | D | IT WAS REPORTED THAT THE CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE HARM CODE HAS BEEN UPDATED IN ANEURYSM AND LOW BLOOD GLUCOSE AND HIGH BLOOD GLUCOSE | Hyperglycemia | Patient Problem Code
8034047 | 138074183 | D | IT WAS REPORTED THAT THE CUSTOMER DECLINED TROUBLESHOOTING FOR HIGH BLOOD GLUCOSE HARM CODE HAS BEEN UPDATED IN ANEURYSM AND LOW BLOOD GLUCOSE AND HIGH BLOOD GLUCOSE | Insufficient Information | Device Problem Code
8266796 | 138132422 | D | IT WAS REPORTED TO PHILIPS THAT THE CHARGE SHOCK FUNCTION WAS TAKING ROUGHLY TO SECONDS TO CHARGE WHILE ON BATTERY POWER THERE WAS NO REPORTED PATIENT INVOLVEMENT | No Patient Involvement | Patient problem code
8266796 | 138132422 | D | IT WAS REPORTED TO PHILIPS THAT THE CHARGE SHOCK FUNCTION WAS TAKING ROUGHLY TO SECONDS TO CHARGE WHILE ON BATTERY POWER THERE WAS NO REPORTED PATIENT INVOLVEMENT | Delayed Charge Time | Device Problem Code
7593165 | 138547525 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7593165 | 138547525 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7489594 | 138548561 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7489594 | 138548561 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7489157 | 138553280 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7489157 | 138553280 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7988056 | 138566721 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7988056 | 138566721 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7588578 | 138572484 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7588578 | 138572484 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7987300 | 138660463 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7987300 | 138660463 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7752663 | 138688622 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7752663 | 138688622 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7750320 | 138712007 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7750320 | 138712007 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
2065322 | 139095091 | D | ALL ATTEMPTS TO THE REPORTER FOR FURTHER INFORMATION REGARDING THE INFECTION HAVE BEEN UNSUCCESSFUL TO DATE | Infection | Patient problem code
2065322 | 139095091 | D | ALL ATTEMPTS TO THE REPORTER FOR FURTHER INFORMATION REGARDING THE INFECTION HAVE BEEN UNSUCCESSFUL TO DATE | Insufficient Information | Device Problem Code
7471449 | 139110328 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7471449 | 139110328 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7750461 | 139119773 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7750461 | 139119773 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7585373 | 139198089 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7585373 | 139198089 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7409124 | 139203235 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7409124 | 139203235 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7421166 | 139215728 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7421166 | 139215728 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
7750145 | 139243480 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7750145 | 139243480 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8194026 | 139256205 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
8194026 | 139256205 | D | NO DATA WAS PROVIDED FOR EVALUATION CONFIRMATION OF THE ISSUE WAS UNDETERMINED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8236240 | 139324483 | D | IT WAS REPORTED TO PHILIPS THAT THE THERAPY KNOB IS NOT WORKING THE DEVICE WAS NOT IN USE ON A PATIENT AS INDICATED BY THE INITIAL REPORTER ANSWERING THE SAFETY QUESTIONS DURING SERVICE ORDER INITIATION | Defective Component | Device Problem Code
8236240 | 139324483 | D | IT WAS REPORTED TO PHILIPS THAT THE THERAPY KNOB IS NOT WORKING THE DEVICE WAS NOT IN USE ON A PATIENT AS INDICATED BY THE INITIAL REPORTER ANSWERING THE SAFETY QUESTIONS DURING SERVICE ORDER INITIATION | No Information | Patient problem code
8243753 | 139330342 | D | IT WAS REPORTED THAT THE DEVICE IS AUTOMATICALLY RESTARTING THERE WAS NO PATIENT INVOLVEMENT THE EFFICIA DFM DEFIBRILLATOR MODEL IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL DEFIBRILLATOR MODEL AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL | No Patient Involvement | Patient problem code
8243753 | 139330342 | D | IT WAS REPORTED THAT THE DEVICE IS AUTOMATICALLY RESTARTING THERE WAS NO PATIENT INVOLVEMENT THE EFFICIA DFM DEFIBRILLATOR MODEL IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL DEFIBRILLATOR MODEL AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL | Self-Activation or Keying | Device Problem Code
7471482 | 139347648 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient problem code
7471482 | 139347648 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8234598 | 139717803 | D | NEW INFORMATION STATES THAT THE PATIENT WAS STABLE | No Consequences Or Impact To Patient | Patient problem code
8234598 | 139717803 | D | NEW INFORMATION STATES THAT THE PATIENT WAS STABLE | Insufficient Information | Device Problem Code
7585375 | 139744556 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | No Information | Patient Problem Code
7585375 | 139744556 | D | DATA WAS PROVIDED FOR EVALUATION THE REPORTED ISSUE WAS NOT CONFIRMED THE PROBABLE CAUSE COULD NOT BE DETERMINED NO ADDITIONAL EVENT OR PATIENT INFORMATION IS AVAILABLE | Insufficient Information | Device Problem Code
8129298 | 139797653 | D | THE ACCOUNT COMMUNICATED THAT THE PATIENTS STROKE WAS NOT RELATED TO THE DEVICE THE PATIENT HAD AN INFECTION AND STROKE AND A CT SCAN WAS DONE TO CONFIRM | Cerebrovascular Accident | Patient problem code
8129298 | 139797653 | D | THE ACCOUNT COMMUNICATED THAT THE PATIENTS STROKE WAS NOT RELATED TO THE DEVICE THE PATIENT HAD AN INFECTION AND STROKE AND A CT SCAN WAS DONE TO CONFIRM | Infection | Patient problem code
8129298 | 139797653 | D | THE ACCOUNT COMMUNICATED THAT THE PATIENTS STROKE WAS NOT RELATED TO THE DEVICE THE PATIENT HAD AN INFECTION AND STROKE AND A CT SCAN WAS DONE TO CONFIRM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8261852 | 139825391 | D | NEW INFORMATION RECEIVED NOTES THE DEVICE WAS SUCCESSFULLY REPROGRAMMED WITH NO PATIENT CONSEQUENCES | No Consequences Or Impact To Patient | Patient problem code
8261852 | 139825391 | D | NEW INFORMATION RECEIVED NOTES THE DEVICE WAS SUCCESSFULLY REPROGRAMMED WITH NO PATIENT CONSEQUENCES | Insufficient Information | Device Problem Code
7431749 | 139882986 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Death | Patient problem code
7431749 | 139882986 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemia | Patient problem code
7431749 | 139882986 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemic Heart Disease | Patient problem code
7431749 | 139882986 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Insufficient Information | Device problem code
7431754 | 139882997 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Death | Patient problem code
7431754 | 139882997 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemia | Patient problem code
7431754 | 139882997 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemic Heart Disease | Patient problem code
7431754 | 139882997 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Insufficient Information | Device Problem Code
7799560 | 139883289 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Death | Patient problem code
7799560 | 139883289 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Myocardial Infarction | Patient problem code
7799560 | 139883289 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7799550 | 139886511 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Death | Patient problem code
7799550 | 139886511 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Myocardial Infarction | Patient problem code
7799550 | 139886511 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7799588 | 139886744 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Death | Patient problem code
7799588 | 139886744 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Myocardial Infarction | Patient problem code
7799588 | 139886744 | D | NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION THE MANNER OF DEATH WAS NATURAL THE PATIENTS DEATH WAS UNRELATED TO THE ICD SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7431750 | 139887763 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Death | Patient problem code
7431750 | 139887763 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemia | Patient problem code
7431750 | 139887763 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Ischemic Heart Disease | Patient problem code
7431750 | 139887763 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS CHECKED FEW WEEKS PRIOR TO PATIENTS DEATH AND WAS FUNCTIONING NORMAL PHYSICIAN BELIEVES THAT THE DEATH WAS RELATED TO PATIENTS ISCHEMIC HEART DISEASE AND WAS NOT RELATED TO DEVICE IN ANY WAY | Insufficient Information | Device Problem Code
8054450 | 139912535 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THIS LEAD WAS EXPLANTED DUE TO AN UNKNOWN REASON UNRELATED TO PRODUCT PERFORMANCE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8054450 | 139912535 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THIS LEAD WAS EXPLANTED DUE TO AN UNKNOWN REASON UNRELATED TO PRODUCT PERFORMANCE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8048516 | 139912819 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS RIGHT VENTRICULAR RV LEAD HAD ELEVATED THRESHOLDS THIS RV LEAD WAS THEN EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | High Capture Threshold | Device Problem Code
8048516 | 139912819 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS RIGHT VENTRICULAR RV LEAD HAD ELEVATED THRESHOLDS THIS RV LEAD WAS THEN EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8048516 | 139912819 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT DURING RECENT DEVICE CHECK THIS PACEMAKERS BATTERY WAS DOWN TO A YEAR AND A HALF BUT HAD SEVEN YEARS LEFT AT LAST CHECK THE PRODUCT WAS EXPLANTED AND WAS REPLACED THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED | No Consequences Or Impact To Patient | Patient problem code
4199825 | 139914176 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Hearing Loss | Patient problem code
4199825 | 139914176 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Scarring | Patient problem code
4199825 | 139914176 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Tinnitus | Patient problem code
4199825 | 139914176 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | No Code Available | Patient problem code
4199825 | 139914176 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Insufficient Information | Device Problem Code
8122982 | 139916734 | D | HEALTHCARE PROFESSIONAL ADDITIONALLY REPORTED CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE HAS BEEN EXPLANTED | Capsular Contracture | Patient problem code
8122982 | 139916734 | D | HEALTHCARE PROFESSIONAL ADDITIONALLY REPORTED CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE HAS BEEN EXPLANTED | Insufficient Information | Device Problem Code
4199933 | 139916855 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Hearing Loss | Patient problem code
4199933 | 139916855 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Scarring | Patient problem code
4199933 | 139916855 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Tinnitus | Patient problem code
4199933 | 139916855 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | No Code Available | Patient problem code
4199933 | 139916855 | D | PFS ALLEGES TINNITUS HEARING LOSS PSEUDOTUMOR TREMOR AND SCARRING | Insufficient Information | Device Problem Code
8361082 | 140009414 | D | OLYMPUS WAS INFORMED THAT THE BIOPSY CHANNEL OF THE SUBJECT DEVICE TESTED POSITIVE FOR BACILLUS DURING ROUTINE SURVEILLANCE CULTURING TEST BY THE FACILITY THE NUMBER OF MICROBES WAS NOT REPORTED THERE WAS NO REPORT OF PATIENT INFECTION ASSOCIATED WITH THIS REPORT | No Consequences Or Impact To Patient | Patient problem code
8361082 | 140009414 | D | OLYMPUS WAS INFORMED THAT THE BIOPSY CHANNEL OF THE SUBJECT DEVICE TESTED POSITIVE FOR BACILLUS DURING ROUTINE SURVEILLANCE CULTURING TEST BY THE FACILITY THE NUMBER OF MICROBES WAS NOT REPORTED THERE WAS NO REPORT OF PATIENT INFECTION ASSOCIATED WITH THIS REPORT | Insufficient Information | Device Problem Code
8343417 | 140013704 | D | ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE ISSUE OCCUR DURING VEIN HARVESTING PROCEDURE WITH ABOUT MINUTES DELAY NO CONSEQUENCES OR IMPACT TO PATIENT PRODUCT WAS CHANGED OUT PROCEDURE WAS COMPLETED SUCCESSFULLY | No Consequences Or Impact To Patient | Patient problem code
8343417 | 140013704 | D | ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE ISSUE OCCUR DURING VEIN HARVESTING PROCEDURE WITH ABOUT MINUTES DELAY NO CONSEQUENCES OR IMPACT TO PATIENT PRODUCT WAS CHANGED OUT PROCEDURE WAS COMPLETED SUCCESSFULLY | Insufficient Information | Device Problem Code
8372453 | 140014529 | D | NEW INFORMATION RECEIVED NOTES THE CAUSE OF DEATH WAS IDIOPATHIC CARDIOMYOPATHY | Cardiomyopathy | Patient problem code
8372453 | 140014529 | D | NEW INFORMATION RECEIVED NOTES THE CAUSE OF DEATH WAS IDIOPATHIC CARDIOMYOPATHY | Death | Patient problem code
8372453 | 140014529 | D | NEW INFORMATION RECEIVED NOTES THE CAUSE OF DEATH WAS IDIOPATHIC CARDIOMYOPATHY | Insufficient Information | Device Problem Code
8076169 | 140016956 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Death | Patient problem code
8076169 | 140016956 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
7574979 | 140017774 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Death | Patient problem code
7574979 | 140017774 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Insufficient Information | Device Problem Code
8369246 | 140022787 | D | IT WAS REPORTED THAT THE CUSTOMER HAD STROKE IN B | Cerebrovascular Accident | Patient problem code
8369246 | 140022787 | D | IT WAS REPORTED THAT THE CUSTOMER HAD STROKE IN B | Insufficient Information | Device Problem Code
8291484 | 140023276 | D | IT WAS REPORTED VIA PHONE CALL THAT THE CUSTOMERS HOSPITALIZATION WAS NOT DIABETES RELATED | No Information | Patient problem code
8291484 | 140023276 | D | IT WAS REPORTED VIA PHONE CALL THAT THE CUSTOMERS HOSPITALIZATION WAS NOT DIABETES RELATED | Insufficient Information | Device Problem Code
8349358 | 140026259 | D | IT WAS FURTHER REPORTED THAT THE PHYSICIAN DID NOT BELIEVE THAT THE DEVICE WAS RELATED TO THE PATIENTS DEATH THE CAUSE OF DEATH IS STILL UNKNOWN | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8349358 | 140026259 | D | IT WAS FURTHER REPORTED THAT THE PHYSICIAN DID NOT BELIEVE THAT THE DEVICE WAS RELATED TO THE PATIENTS DEATH THE CAUSE OF DEATH IS STILL UNKNOWN | Death | Patient problem code
7575034 | 140029077 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Death | Patient problem code
7575034 | 140029077 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Insufficient Information | Device Problem Code
7575048 | 140029120 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Death | Patient problem code
7575048 | 140029120 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Insufficient Information | Device Problem Code
7575012 | 140029657 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Death | Patient problem code
7575012 | 140029657 | D | NEW INFORMATION RECEIVED STATES THAT THE PRIMARY AND SECONDARY CAUSE OF DEATH WAS UNKNOWN THE PHYSICIAN DOES NOT ALLEGE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH SINCE THE DEATH NARRATIVE AND CAUSE ARE UNKNOWN | Insufficient Information | Device Problem Code
8179925 | 140030055 | D | MDRYES SI THIS RIGHT ATRIAL RA LEAD WAS EXPLANTED AND REPLACED DUE TO AN UNKNOWN REASON NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8179925 | 140030055 | D | MDRYES SI THIS RIGHT ATRIAL RA LEAD WAS EXPLANTED AND REPLACED DUE TO AN UNKNOWN REASON NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8172267 | 140030371 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Death | Patient problem code
8172267 | 140030371 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Shock, Septic | Patient problem code
8172267 | 140030371 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8172264 | 140030492 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Death | Patient problem code
8172264 | 140030492 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Shock, Septic | Patient problem code
8172264 | 140030492 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8356295 | 140031215 | D | IT WAS LATER REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SUSPECTED STROKE ELEVEN DAYS POST IMPLANT AND DIED AT HOME NINE DAYS LATER THE PATIENT WAS ALSO NOTED TO HAVE AN UNDERLYING RHYTHM | Cerebrovascular Accident | Patient problem code
8356295 | 140031215 | D | IT WAS LATER REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SUSPECTED STROKE ELEVEN DAYS POST IMPLANT AND DIED AT HOME NINE DAYS LATER THE PATIENT WAS ALSO NOTED TO HAVE AN UNDERLYING RHYTHM | Death | Patient problem code
8356295 | 140031215 | D | IT WAS LATER REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SUSPECTED STROKE ELEVEN DAYS POST IMPLANT AND DIED AT HOME NINE DAYS LATER THE PATIENT WAS ALSO NOTED TO HAVE AN UNDERLYING RHYTHM | Insufficient Information | Device Problem Code
8262615 | 140031233 | D | IT WAS REPORTED THAT THE PATIENTS CAUSE OF DEATH WAS A CORONARY BYPASS GRAFT FAILURE AND NOT RELATED TO THE RA LEAD REVISION PROCEDURE | Death | Patient problem code
8262615 | 140031233 | D | IT WAS REPORTED THAT THE PATIENTS CAUSE OF DEATH WAS A CORONARY BYPASS GRAFT FAILURE AND NOT RELATED TO THE RA LEAD REVISION PROCEDURE | Insufficient Information | Device Problem Code
8076173 | 140032051 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Death | Patient problem code
8076173 | 140032051 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8076187 | 140032556 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Death | Patient problem code
8076187 | 140032556 | D | NEW INFORMATION NOTED THAT THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO COMPLICATION OF DIABETES THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8172263 | 140032755 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Death | Patient problem code
8172263 | 140032755 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Shock, Septic | Patient problem code
8172263 | 140032755 | D | NEW INFORMATION RECEIVED INDICATED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND WAS UNRELATED TO THE PACEMAKER SYSTEM | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8127394 | 140032831 | D | HEALTHCARE PROFESSIONAL ADDITIONALLY REPORTED CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE WAS EXPLANTED | Capsular Contracture | Patient problem code
8127394 | 140032831 | D | HEALTHCARE PROFESSIONAL ADDITIONALLY REPORTED CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE WAS EXPLANTED | Insufficient Information | Device Problem Code
8085377 | 140032995 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8085377 | 140032995 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8272790 | 140036184 | D | THE PHYSICIAN DID NOT ALLEGE THE DEVICE OR PROCEDURE CAUSED OR CONTRIBUTED TO THE INFECTION | Infection | Patient problem code
8272790 | 140036184 | D | THE PHYSICIAN DID NOT ALLEGE THE DEVICE OR PROCEDURE CAUSED OR CONTRIBUTED TO THE INFECTION | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8359178 | 140038494 | D | ADDITIONAL INFORMATION RECEIVED VIA A FOREIGN USER REPORT INDICATED THAT THE PATIENT RECEIVED ANTIBIOTIC TREATMENT FOR THE POSTOPERATIVE INFECTION AFTER THE COMPLEX SPINE SURGERY | Infection | Patient problem code
8359178 | 140038494 | D | ADDITIONAL INFORMATION RECEIVED VIA A FOREIGN USER REPORT INDICATED THAT THE PATIENT RECEIVED ANTIBIOTIC TREATMENT FOR THE POSTOPERATIVE INFECTION AFTER THE COMPLEX SPINE SURGERY | Insufficient Information | Device Problem Code
8335440 | 140038675 | D | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP IT WAS REPORTED THAT THE LEADS AND BATTERY WERE CHANGED AND PATIENT REPROGRAMMED TO NEW SETTINGS PATIENT WAS DOING WELL AND NO OTHER ADVERSE EVENTS REPORTED | No Consequences Or Impact To Patient | Patient problem code
8335440 | 140038675 | D | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP IT WAS REPORTED THAT THE LEADS AND BATTERY WERE CHANGED AND PATIENT REPROGRAMMED TO NEW SETTINGS PATIENT WAS DOING WELL AND NO OTHER ADVERSE EVENTS REPORTED | Insufficient Information | Device Problem Code
8185707 | 140039147 | D | THE ASSOCIATED RIGHT VENTRICULAR RV LEAD WAS LATER EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient Problem Code
8185707 | 140039147 | D | THE ASSOCIATED RIGHT VENTRICULAR RV LEAD WAS LATER EXPLANTED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8122667 | 140042692 | D | HEALTHCARE PROFESSIONAL REPORTED ADDITIONAL EVENT OF CAPSULAR CONTRACTURE BAKER GRADE III | Capsular Contracture | Patient problem code
8122667 | 140042692 | D | HEALTHCARE PROFESSIONAL REPORTED ADDITIONAL EVENT OF CAPSULAR CONTRACTURE BAKER GRADE III | Insufficient Information | Device Problem Code
8357384 | 140043064 | D | HEALTHCARE PROVIDER REPORTED CAPSULAR CONTRACTURE BAKER GRADE III | Capsular Contracture | Patient problem code
8357384 | 140043064 | D | HEALTHCARE PROVIDER REPORTED CAPSULAR CONTRACTURE BAKER GRADE III | Insufficient Information | Device Problem Code
8238211 | 140044363 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS EXPLANTED AND REPLACED PATIENT WAS STABLE | No Consequences Or Impact To Patient | Patient problem code
8238211 | 140044363 | D | NEW INFORMATION RECEIVED NOTES THAT THE DEVICE WAS EXPLANTED AND REPLACED PATIENT WAS STABLE | Insufficient Information | Device Problem Code
8218522 | 140046169 | D | ADDITIONAL INFORMATION RECEIVED FROM THE MANUFACTURERS REPRESENTATIVE REP REPORTED THE PATIENT WAS REIMPLANTED ON THE RIGHT SIDE WITHOUT ISSUE AND WAS DOING FINE NO FURTHER COMPLICATIONS WERE ANTICIPATED | No Consequences Or Impact To Patient | Patient problem code
8218522 | 140046169 | D | ADDITIONAL INFORMATION RECEIVED FROM THE MANUFACTURERS REPRESENTATIVE REP REPORTED THE PATIENT WAS REIMPLANTED ON THE RIGHT SIDE WITHOUT ISSUE AND WAS DOING FINE NO FURTHER COMPLICATIONS WERE ANTICIPATED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8224842 | 140049610 | D | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED VENTRAL HERNIA INCISIONAL HERNIA NEURALGIC PAIN | Hernia | Patient problem code
8224842 | 140049610 | D | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED VENTRAL HERNIA INCISIONAL HERNIA NEURALGIC PAIN | Pain | Patient problem code
8224842 | 140049610 | D | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED VENTRAL HERNIA INCISIONAL HERNIA NEURALGIC PAIN | Insufficient Information | Device Problem Code
7928118 | 140051794 | D | HEALTH PROFESSIONAL ADDITIONALLY REPORTED A RIGHT SIDE CAPSULAR CONTRACTURE BAKER GRADE NOT SPECIFIED AND CREASEFOLDING OF THE IMPLANT DEVICE WAS EXPLANTED | Capsular Contracture | Patient problem code
7928118 | 140051794 | D | HEALTH PROFESSIONAL ADDITIONALLY REPORTED A RIGHT SIDE CAPSULAR CONTRACTURE BAKER GRADE NOT SPECIFIED AND CREASEFOLDING OF THE IMPLANT DEVICE WAS EXPLANTED | Insufficient Information | Device Problem Code
8205963 | 140060469 | D | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED DEHISCENCE AND GRANULATION TISSUE | Wound Dehiscence | Patient problem code
8205963 | 140060469 | D | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED DEHISCENCE AND GRANULATION TISSUE | Insufficient Information | Device Problem Code
8238905 | 140061798 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE SUSPECTED THROMBOSIS WAS TREATED WITH FIBRINOLYSIS THE PATIENT WAS DOING OK BUT THEN DEVELOPED A SUBARACHNOID HEMORRHAGE THE PATIENT EXPIRED ON B NO ADDITIONAL INFORMATION WAS PROVIDED | Death | Patient problem code
8238905 | 140061798 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE SUSPECTED THROMBOSIS WAS TREATED WITH FIBRINOLYSIS THE PATIENT WAS DOING OK BUT THEN DEVELOPED A SUBARACHNOID HEMORRHAGE THE PATIENT EXPIRED ON B NO ADDITIONAL INFORMATION WAS PROVIDED | Hemorrhage, Subarachnoid | Patient problem code
8238905 | 140061798 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE SUSPECTED THROMBOSIS WAS TREATED WITH FIBRINOLYSIS THE PATIENT WAS DOING OK BUT THEN DEVELOPED A SUBARACHNOID HEMORRHAGE THE PATIENT EXPIRED ON B NO ADDITIONAL INFORMATION WAS PROVIDED | Thrombosis | Patient problem code
8238905 | 140061798 | D | ADDITIONAL INFORMATION IT WAS REPORTED THAT THE SUSPECTED THROMBOSIS WAS TREATED WITH FIBRINOLYSIS THE PATIENT WAS DOING OK BUT THEN DEVELOPED A SUBARACHNOID HEMORRHAGE THE PATIENT EXPIRED ON B NO ADDITIONAL INFORMATION WAS PROVIDED | Insufficient Information | Device problem code
8214093 | 140062017 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RV LEAD WAS ALSO PACING TO FAST NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8214093 | 140062017 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RV LEAD WAS ALSO PACING TO FAST NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
5064342 | 140062948 | D | PPF ALLEGES METALLOSIS | Toxicity | Patient problem code
5064342 | 140062948 | D | PPF ALLEGES METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
8346424 | 140063964 | D | HEALTH PROFESSIONAL REPORTED CAPSULAR CONTRACTURE BAKER GRADE IV | Capsular Contracture | Patient problem code
8346424 | 140063964 | D | HEALTH PROFESSIONAL REPORTED CAPSULAR CONTRACTURE BAKER GRADE IV | Insufficient Information | Device Problem Code
2584333 | 140063976 | D | PPF ALLEGED PSEUDOTUMOR AND METALLOSIS DOI B DOR B LEFT HIP | Toxicity | Patient problem code
2584333 | 140063976 | D | PPF ALLEGED PSEUDOTUMOR AND METALLOSIS DOI B DOR B LEFT HIP | No Code Available | Patient problem code
2584333 | 140063976 | D | PPF ALLEGED PSEUDOTUMOR AND METALLOSIS DOI B DOR B LEFT HIP | Appropriate Term/Code Not Available | Device Problem Code
8042324 | 140076426 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT DURING RECENT DEVICE CHECK THIS PACEMAKERS BATTERY WAS DOWN TO A YEAR AND A HALF BUT HAD SEVEN YEARS LEFT AT LAST CHECK THE PRODUCT WAS EXPLANTED AND WAS REPLACED THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED | No Consequences Or Impact To Patient | Patient problem code
8042324 | 140076426 | D | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT DURING RECENT DEVICE CHECK THIS PACEMAKERS BATTERY WAS DOWN TO A YEAR AND A HALF BUT HAD SEVEN YEARS LEFT AT LAST CHECK THE PRODUCT WAS EXPLANTED AND WAS REPLACED THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED | Battery Problem | Device Problem Code
8002790 | 140076948 | D | ADDITIONAL INFORMATION REPORTED INDICATED THAT THE RIGHT VENTRICULAR RV LEAD WAS ABANDONED SURGICALLY AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED THE PRODUCT HAS NOT BEEN RETURNED IF ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE UPDATED | No Consequences Or Impact To Patient | Patient problem code
8002790 | 140076948 | D | ADDITIONAL INFORMATION REPORTED INDICATED THAT THE RIGHT VENTRICULAR RV LEAD WAS ABANDONED SURGICALLY AND REPLACED NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED THE PRODUCT HAS NOT BEEN RETURNED IF ADDITIONAL INFORMATION IS RECEIVED THIS REPORT WILL BE UPDATED | Insufficient Information | Device Problem Code
2244947 | 140078481 | D | IN ADDITION TO PREVIOUS ALLEGATIONS PPF ALLEGES INFECTION | Infection | Patient problem code
2244947 | 140078481 | D | IN ADDITION TO PREVIOUS ALLEGATIONS PPF ALLEGES INFECTION | Insufficient Information | Device Problem Code
8250497 | 140078483 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RA LEAD WAS EXPLANTED AND REPLACED THE PACEMAKER REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8250497 | 140078483 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RA LEAD WAS EXPLANTED AND REPLACED THE PACEMAKER REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
6464797 | 140081566 | D | IN ADDITION TO WHAT WAS PREVIOUSLY REPORTED PFS ALLEGES DIFFICULTY WALKING AND PATIENT STRUGGLE PERFORMING SIMPLE TASK AFTER REVIEW OF THE MEDICAL RECORDS PATIENT WAS REVISED TO ADDRESS TEAR OF THE GLUTEUS MEDIUS TENDON | Ambulation Difficulties | Patient problem code
6464797 | 140081566 | D | IN ADDITION TO WHAT WAS PREVIOUSLY REPORTED PFS ALLEGES DIFFICULTY WALKING AND PATIENT STRUGGLE PERFORMING SIMPLE TASK AFTER REVIEW OF THE MEDICAL RECORDS PATIENT WAS REVISED TO ADDRESS TEAR OF THE GLUTEUS MEDIUS TENDON | Insufficient Information | Device Problem Code
8364896 | 140176582 | D | NEW INFORMATION RECEIVED STATES THAT THE PATIENT WAS ASYMPTOMATIC AND STABLE BEFORE DURING AND AFTER THE CLINIC VISIT | No Consequences Or Impact To Patient | Patient problem code
8364896 | 140176582 | D | NEW INFORMATION RECEIVED STATES THAT THE PATIENT WAS ASYMPTOMATIC AND STABLE BEFORE DURING AND AFTER THE CLINIC VISIT | Insufficient Information | Device Problem Code
8140154 | 140178759 | D | NEW INFORMATION RECEIVED ON B INDICATES THAT THE LEAD WAS REPOSITIONED ON B THE PATIENT WAS STABLE | No Consequences Or Impact To Patient | Patient problem code
8140154 | 140178759 | D | NEW INFORMATION RECEIVED ON B INDICATES THAT THE LEAD WAS REPOSITIONED ON B THE PATIENT WAS STABLE | Insufficient Information | Device Problem Code
8348079 | 140178975 | D | DURING THE SURGERY PATIENT VITALS WERE IN CONCERN AND PATIENT DIED | Death | Patient problem code
8348079 | 140178975 | D | DURING THE SURGERY PATIENT VITALS WERE IN CONCERN AND PATIENT DIED | Insufficient Information | Device Problem Code
8365270 | 140180486 | D | OLYMPUS WAS INFORMED THAT DURING A POST MARKET SURVEILLANCE STUDY THE SCOPE CULTURED POSITIVE FOR STAPHYLOCOCCUS AUREUS AFTER REPROCESSING THERE WERE NO ASSOCIATED PATIENT INFECTIONS REPORTED | No Known Impact Or Consequence To Patient | Patient problem code
8365270 | 140180486 | D | OLYMPUS WAS INFORMED THAT DURING A POST MARKET SURVEILLANCE STUDY THE SCOPE CULTURED POSITIVE FOR STAPHYLOCOCCUS AUREUS AFTER REPROCESSING THERE WERE NO ASSOCIATED PATIENT INFECTIONS REPORTED | Insufficient Information | Device Problem Code
8108180 | 140180970 | D | NEW INFORMATION AVAILABLE ON B STATES THAT THE ISSUE WAS RESOLVED THE PATIENT WAS STABLE BEFORE DURING AND AFTER RESOLUTION | No Consequences Or Impact To Patient | Patient problem code
8108180 | 140180970 | D | NEW INFORMATION AVAILABLE ON B STATES THAT THE ISSUE WAS RESOLVED THE PATIENT WAS STABLE BEFORE DURING AND AFTER RESOLUTION | Insufficient Information | Device Problem Code
8296955 | 140183658 | D | CLARIFICATION WAS RECEIVED THE PATIENT HAD A LITTLE EXTENDED BELLY THEY THOUGHT THE PATIENT MIGHT HAVE WATER RETENTION AND THEY HAD AN ULTRASOUND TO ACCESS REPORTED SYMPTOMS NO FURTHER PATIENT COMPLICATIONS ARE ANTICIPATED OR EXPECTED AS A RESULT OF THIS EVENT | No Consequences Or Impact To Patient | Patient problem code
8296955 | 140183658 | D | CLARIFICATION WAS RECEIVED THE PATIENT HAD A LITTLE EXTENDED BELLY THEY THOUGHT THE PATIENT MIGHT HAVE WATER RETENTION AND THEY HAD AN ULTRASOUND TO ACCESS REPORTED SYMPTOMS NO FURTHER PATIENT COMPLICATIONS ARE ANTICIPATED OR EXPECTED AS A RESULT OF THIS EVENT | Insufficient Information | Device Problem Code
8236002 | 140183755 | D | NEW INFORMATION RECEIVED ON FEB INDICATES THAT THE PATIENT PRESENTED ON B FOR AN EXPLANT PROCEDURE FOR MORE PHRENIC NERVE STIMULATION THE LEAD WAS REPLACED AND THE PATIENT HAD NO COMPLICATION BEFORE DURING OR AFTER THE PROCEDURE | Nerve Stimulation, Undesired | Patient problem code
8236002 | 140183755 | D | NEW INFORMATION RECEIVED ON FEB INDICATES THAT THE PATIENT PRESENTED ON B FOR AN EXPLANT PROCEDURE FOR MORE PHRENIC NERVE STIMULATION THE LEAD WAS REPLACED AND THE PATIENT HAD NO COMPLICATION BEFORE DURING OR AFTER THE PROCEDURE | Insufficient Information | Device Problem Code
7560641 | 140188263 | D | NEW INFORMATION RECEIVED ON JANUARY STATES THAT PATIENT WAS STABLE POST DEVICE CHANGE OUT PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
7560641 | 140188263 | D | NEW INFORMATION RECEIVED ON JANUARY STATES THAT PATIENT WAS STABLE POST DEVICE CHANGE OUT PROCEDURE | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8250136 | 140189289 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE SYSTEM WAS TAKEN OUT OF SERVICE DUE TO INFECTION NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Infection | Patient problem code
8250136 | 140189289 | D | ADDITIONAL INFORMATION WAS RECEIVED THAT THIS DEVICE SYSTEM WAS TAKEN OUT OF SERVICE DUE TO INFECTION NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8260511 | 140189634 | D | ADDITIONAL INFORMATION RECEIVED THAT THE LEAD WAS CAPPED AND REPLACED TO RESOLVE THE EVENT THE PATIENT WAS STABLE WITH NO CONSEQUENCES | No Consequences Or Impact To Patient | Patient problem code
8260511 | 140189634 | D | ADDITIONAL INFORMATION RECEIVED THAT THE LEAD WAS CAPPED AND REPLACED TO RESOLVE THE EVENT THE PATIENT WAS STABLE WITH NO CONSEQUENCES | Insufficient Information | Device Problem Code
8265193 | 140193944 | D | NEW INFORMATION RECEIVED NOTED THE PATIENTS RIGHT ATRIAL AND RIGHT VENTRICULAR LEADS WERE EXPLANTED AND REPLACED ON B THE PATIENT WAS STABLE THROUGHOUT | No Consequences Or Impact To Patient | Patient problem code
8265193 | 140193944 | D | NEW INFORMATION RECEIVED NOTED THE PATIENTS RIGHT ATRIAL AND RIGHT VENTRICULAR LEADS WERE EXPLANTED AND REPLACED ON B THE PATIENT WAS STABLE THROUGHOUT | Insufficient Information | Device Problem Code
8269337 | 140201709 | D | NEW INFORMATION REVEALED THAT A NEW PACER SYSTEM WAS IMPLANTED ON THE PATIENTS RIGHT SIDE ON B THE OLD SYSTEM REMAINED IMPLANTED AS A BACKUP UNTIL THE NEW LEADS MATURE THE PATIENT WAS STABLE | No Consequences Or Impact To Patient | Patient problem code
8269337 | 140201709 | D | NEW INFORMATION REVEALED THAT A NEW PACER SYSTEM WAS IMPLANTED ON THE PATIENTS RIGHT SIDE ON B THE OLD SYSTEM REMAINED IMPLANTED AS A BACKUP UNTIL THE NEW LEADS MATURE THE PATIENT WAS STABLE | Insufficient Information | Device Problem Code
8269325 | 140201777 | D | NEW INFORMATION REVEALED THAT A NEW PACER SYSTEM WAS IMPLANTED ON THE PATIENTS RIGHT SIDE ON B THE OLD SYSTEM REMAINED IMPLANTED AS A BACKUP UNTIL THE NEW LEADS MATURE THE PATIENT WAS STABLE | No Consequences Or Impact To Patient | Patient problem code
8269325 | 140201777 | D | NEW INFORMATION REVEALED THAT A NEW PACER SYSTEM WAS IMPLANTED ON THE PATIENTS RIGHT SIDE ON B THE OLD SYSTEM REMAINED IMPLANTED AS A BACKUP UNTIL THE NEW LEADS MATURE THE PATIENT WAS STABLE | Insufficient Information | Device Problem Code
8334966 | 140201807 | D | IT WAS REPORTED THAT THE CUSTOMER WAS ADMITTED TO THE EMERGENCY ROOM AND GOT INTRAVENOUS BECAUSE SHE WAS DEHYDRATED ON TOP OF BEING TREATED FOR THE HIGH BLOOD GLUCOSE | Dehydrated | Patient problem code
8334966 | 140201807 | D | IT WAS REPORTED THAT THE CUSTOMER WAS ADMITTED TO THE EMERGENCY ROOM AND GOT INTRAVENOUS BECAUSE SHE WAS DEHYDRATED ON TOP OF BEING TREATED FOR THE HIGH BLOOD GLUCOSE | Hyperglycemia | Patient Problem Code
8334966 | 140201807 | D | IT WAS REPORTED THAT THE CUSTOMER WAS ADMITTED TO THE EMERGENCY ROOM AND GOT INTRAVENOUS BECAUSE SHE WAS DEHYDRATED ON TOP OF BEING TREATED FOR THE HIGH BLOOD GLUCOSE | Insufficient Information | Device Problem Code
8250203 | 140204086 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RV LEAD WAS SURGICALLY ABANDONED AND REPLACED THE CRTD REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | No Consequences Or Impact To Patient | Patient problem code
8250203 | 140204086 | D | ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE RV LEAD WAS SURGICALLY ABANDONED AND REPLACED THE CRTD REMAINS IN SERVICE NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED | Insufficient Information | Device Problem Code
8101746 | 140205672 | D | ADDITIONAL INFORMATION RECEIVED FROM A HEALTHCARE PROVIDER HCP VIA A REP INDICATED THAT THE INS WAS TO BE REIMPLANTED AFTER THE INFECTION RESOLVED | Infection | Patient problem code
8101746 | 140205672 | D | ADDITIONAL INFORMATION RECEIVED FROM A HEALTHCARE PROVIDER HCP VIA A REP INDICATED THAT THE INS WAS TO BE REIMPLANTED AFTER THE INFECTION RESOLVED | Insufficient Information | Device Problem Code
6482972 | 140206913 | D | NEW INFORMATION REVEALED THAT THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
6482972 | 140206913 | D | NEW INFORMATION REVEALED THAT THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE | Insufficient Information | Device Problem Code
8214815 | 140211255 | D | PATIENT REPORTED RIGHT SIDE CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE HAS BEEN EXPLANTED | Capsular Contracture | Patient problem code
8214815 | 140211255 | D | PATIENT REPORTED RIGHT SIDE CAPSULAR CONTRACTURE BAKER GRADE UNKNOWN DEVICE HAS BEEN EXPLANTED | Insufficient Information | Device Problem Code
2939102 | 140212480 | D | PPF ALLEGES PSEUDOTUMOR AND METALLOSIS | Toxicity | Patient problem code
2939102 | 140212480 | D | PPF ALLEGES PSEUDOTUMOR AND METALLOSIS | No Code Available | Patient problem code
2939102 | 140212480 | D | PPF ALLEGES PSEUDOTUMOR AND METALLOSIS | Appropriate Term/Code Not Available | Device Problem Code
8271345 | 140218003 | D | FOLLOWUP IDENTIFIED THE ISSUE WAS RESOLVED WITHOUT SURGICAL INTERVENTION STIMULATION WAS RESTORED AND THE IPG IS FUNCTIONING AS INTENDED | No Consequences Or Impact To Patient | Patient problem code
8271345 | 140218003 | D | FOLLOWUP IDENTIFIED THE ISSUE WAS RESOLVED WITHOUT SURGICAL INTERVENTION STIMULATION WAS RESTORED AND THE IPG IS FUNCTIONING AS INTENDED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8257116 | 140218426 | D | DEVICE OF REFERENCE MFR REPORT FOLLOW UP REVEALED THAT THE PATIENTS SYSTEM WAS EXPLANTED DUE TO THE PATIENT NO LONGER REQUIRING IT FOR PAIN MANAGEMENT | Pain | Patient problem code
8257116 | 140218426 | D | DEVICE OF REFERENCE MFR REPORT FOLLOW UP REVEALED THAT THE PATIENTS SYSTEM WAS EXPLANTED DUE TO THE PATIENT NO LONGER REQUIRING IT FOR PAIN MANAGEMENT | Insufficient Information | Device Problem Code
8261779 | 140219021 | D | FOLLOWUP INFORMATION REVEALED THE PATIENT EXPERIENCED DISCOMFORT AT THE IPG SITE THE PATIENT IS HAPPY WITH THE OUTCOME OF THE SURGERY AND THE ISSUE IS RESOLVED | Discomfort | Patient problem code
8261779 | 140219021 | D | FOLLOWUP INFORMATION REVEALED THE PATIENT EXPERIENCED DISCOMFORT AT THE IPG SITE THE PATIENT IS HAPPY WITH THE OUTCOME OF THE SURGERY AND THE ISSUE IS RESOLVED | Adverse Event Without Identified Device or Use Problem | Device Problem Code
8141428 | 140221947 | D | IT WAS REPORTED THAT THE INFECTION HAS BEEN RESOLVED | Infection | Patient problem code
8141428 | 140221947 | D | IT WAS REPORTED THAT THE INFECTION HAS BEEN RESOLVED | Insufficient Information | Device Problem Code
8257134 | 140222079 | D | DEVICE OF REFERENCE MFR REPORT FOLLOW UP REVEALED THAT THE PATIENTS SYSTEM WAS EXPLANTED DUE TO THE PATIENT NO LONGER REQUIRING IT FOR PAIN MANAGEMENT | Pain | Patient problem code
8257134 | 140222079 | D | DEVICE OF REFERENCE MFR REPORT FOLLOW UP REVEALED THAT THE PATIENTS SYSTEM WAS EXPLANTED DUE TO THE PATIENT NO LONGER REQUIRING IT FOR PAIN MANAGEMENT | Insufficient Information | Device Problem Code
8313493 | 140333064 | D | NEW INFORMATION ON STATED THE PATIENT WAS IN STABLE CONDITION THROUGHOUT THE EVENT | No Consequences Or Impact To Patient | Patient problem code
8313493 | 140333064 | D | NEW INFORMATION ON STATED THE PATIENT WAS IN STABLE CONDITION THROUGHOUT THE EVENT | Insufficient Information | Device Problem Code
8255122 | 140333326 | D | NEW INFORMATION RECEIVED INDICATED THAT THE PATIENT PRESENTED ON B FOR EXPLANT PROCEDURE AND THE SYSTEM WAS EXPLANTED THE PATIENT WAS STABLE POST PROCEDURE | No Consequences Or Impact To Patient | Patient problem code
8255122 | 140333326 | D | NEW INFORMATION RECEIVED INDICATED THAT THE PATIENT PRESENTED ON B FOR EXPLANT PROCEDURE AND THE SYSTEM WAS EXPLANTED THE PATIENT WAS STABLE POST PROCEDURE | Insufficient Information | Device Problem Code
7309446 | 140357069 | D | IT WAS REPORTED THAT THE PATIENT HAD THEIR SYSTEM EXPLANTED ON B DUE TO CONTINUED INFECTION | Infection | Patient problem code
7309446 | 140357069 | D | IT WAS REPORTED THAT THE PATIENT HAD THEIR SYSTEM EXPLANTED ON B DUE TO CONTINUED INFECTION | Insufficient Information | Device Problem Code
2232094 | 140358040 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS | Toxicity | Patient problem code
2232094 | 140358040 | D | IN ADDITION TO WHAT WERE PREVIOUSLY ALLEGED PPF ALLEGES METAL WEARMETALLOSIS AND ELEVATED METAL IONS | Appropriate Term/Code Not Available | Device Problem Code
5569741 | 140362532 | D | PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
5569741 | 140362532 | D | PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
2758885 | 140363124 | D | PPF ALLEGES INFECTION ADDED AGE LAWYER SURGEON UPDATED AFFECTED SIDE AND UPDATED PRODUCT DETAILS ADDED STEM SLEEVE CUP AND SCREWS DUE TO ALLEGED INFECTION DOI B DOR B LEFT HIP | Infection | Patient problem code
2758885 | 140363124 | D | PPF ALLEGES INFECTION ADDED AGE LAWYER SURGEON UPDATED AFFECTED SIDE AND UPDATED PRODUCT DETAILS ADDED STEM SLEEVE CUP AND SCREWS DUE TO ALLEGED INFECTION DOI B DOR B LEFT HIP | Insufficient Information | Device Problem Code
4534710 | 140368821 | D | PPF ALLEGES ELEVATED METAL IONS | Toxicity | Patient problem code
4534710 | 140368821 | D | PPF ALLEGES ELEVATED METAL IONS | Insufficient Information | Device Problem Code
8546438 | 142968712 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO DETECT THE ATTACHED MULTIFUNCTION CABLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | Activation, Positioning or Separation Problem | Device Problem Code
8546438 | 142968712 | D | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING THE DEVICE WAS UNABLE TO DETECT THE ATTACHED MULTIFUNCTION CABLE COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION | No Patient Involvement | Patient problem code
6972088 | 143212625 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | Loss of Power | Device problem code
6972088 | 143212625 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | No Device Problem Found | Medical Evaluation result code
6891111 | 143376970 | N | DEVICE RECEIVED ALARMING FAILURE OF FLASH MEMORY FOLLOWED BY AN DEVICE SOFTWARE ERROR ALARM PROBLEM ISOLATED TO MOTHER BOARD UNABLE TO PERFORM OPERATING CURRENTS SELF TEST UNEXPECTED RESTART ERROR TEST REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST DUE TO FAILURE OF FLASH MEMORY AND DEVICE SOFTWARE ERROR ALARM | No Findings Available | Medical Evaluation result code
6829040 | 143514671 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST POWER ERROR DUE TO J CONNECTOR RESISTANCE ISSUE | Electrical Power Problem | Device problem code
6829040 | 143514671 | N | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST POWER ERROR DUE TO J CONNECTOR RESISTANCE ISSUE | Power Source Problem Identified | Medical Evaluation result code
6906325 | 143514778 | N | DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | Electrical Power Problem | Device problem code
6906325 | 143514778 | N | DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING DEVICE FUNCTIONING PROPERLY | No Device Problem Found | Medical Evaluation result code
6891901 | 143549073 | N | INSULIN PUMP MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED INSULIN PUMP PASSED THE DISPLACEMENT TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENT INSULIN PUMP PASSED SELF TEST INSULIN PUMP PASSED OFF NO POWER TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW | Scratched Material | Device problem code
6891901 | 143549073 | N | INSULIN PUMP MONITORED FOR SEVERAL DAYS AND NO BLANK DISPLAY NOTED INSULIN PUMP PASSED THE DISPLACEMENT TEST PUMP RECEIVED WITH NORMAL OPERATING CURRENT INSULIN PUMP PASSED SELF TEST INSULIN PUMP PASSED OFF NO POWER TEST INSULIN PUMP RECEIVED WITH MINOR SCRATCHED LCD WINDOW | No Device Problem Found | Medical Evaluation result code
6936709 | 143680067 | N | NO UNEXPECTED NO DELIVERY ALARM NOTED DURING TESTING DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
6877791 | 143795426 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO BLANK DISPLAY NOTED THE INSULIN PUMP PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST THE INSULIN PUMP WAS MONITORED AND NO UNEXPECTED POWERING OFF OR BLANK DISPLAY NOTED VERIFIED THE BATTERY TUBE POWER CONNECTOR IS PROPERLY CONNECTED TO ELECTRICAL BOARD DURING VISUAL INSPECTION | No Device Problem Found | Medical Evaluation result code
6892017 | 143795561 | N | DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST PRIME TEST AND DISPLACEMENT TEST HOWEVER DEVICE RECEIVED STUCK IN THE MOTOR ERROR LOOP DURING BOLUSBASAL DELIVERY UNABLE TO VERIFY MOTION SENSOR TEST FAILURE ALARM DUE TO MOTOR ERROR ALARM THE MOTOR MAY HAVE HAD AN INTERMITTENT FAILURE THAT WAS NOT DETECTED DURING OUR TESTING THE MOTOR WAS TESTED OUTSIDE THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
6853664 | 144088180 | N | INSULIN PUMP PASSED THE DISPLACEMENT REWIND BASIC OCCLUSION OCCLUSION PRIME AND EXCESSIVE NO DELIVERY TEST NO MOTOR ERROR ALARM NOTED DURING TEST MOTOR PASSED MOTOR TEST ADDITIONAL INFORMATION HAS BEEN RECEIVED WHICH WAS NOT INCLUDED WITH THE INITIAL REPORT THE INFORMATION HAS BEEN PROVIDED WITH THIS REPORT | No Findings Available | Medical Evaluation result code
6755895 | 144088224 | N | THE INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO CRITICAL PUMP ERROR OPEN BOOK IMAGE ALARM NOTED THE INSULIN PUMP PASSED THE DISPLACEMENT TEST REWIND TEST PRIMESEATING TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST ACTIVE CURRENT TEST SLEEP CURRENT TEST AND SELF TEST | No Device Problem Found | Medical Evaluation result code
6757036 | 144241624 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR AND LOW BATTERY THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALARMS WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6757036 | 144241624 | N | DEVICE PASSED THE SELF TEST SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND THE DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR AND LOW BATTERY THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR AND LOW BATTERY ALARMS WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
6840304 | 144241781 | N | INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW THE INSULIN PUMP PASSED THE DISPLACEMENT TEST | Scratched Material | Device problem code
6840304 | 144241781 | N | INSULIN PUMP POWERED UP PROPERLY AFTER BATTERY INSTALLATION NO MISSING SEGMENTS OR PARTIAL DISPLAY NOTED THE INSULIN PUMP WAS RECEIVED WITH MINOR SCRATCHED LCD WINDOW THE INSULIN PUMP PASSED THE DISPLACEMENT TEST | No Device Problem Found | Medical Evaluation result code
6845752 | 144241854 | N | DEVICE PASSED DISPLACEMENT TEST REWIND TEST PRIME TEST BASIC OCCLUSION TEST FORCE SENSOR TEST OCCLUSION TEST SLEEP CURRENT TEST ACTIVE CURRENT TEST AND SELF TEST NO DELIVERY ALARM OR OCCLUSION DURING THERAPY NOT CONFIRMED NO POWER ERROR WAS FOUND IN THE HISTORY FILE AND NO ABNORMAL BEHAVIOR WAS NOTED USING THE POWER MANAGEMENT TOOL AS SHOWN ON THE POWER MANAGEMENT GRAPH POWER ERROR NOT CONFIRMED | No Device Problem Found | Medical Evaluation result code
6879236 | 144241895 | N | INSULIN PUMP PASSED DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6879236 | 144241895 | N | INSULIN PUMP PASSED DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
6987124 | 144308860 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Component Missing | Device problem code
6987124 | 144308860 | N | DEVICE WAS RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING A TEST RESERVOIR WAS INSTALLED AND DID NOT LOCK IN PLACE DUE TO MISSING RETAINER UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER | Missing or Inadequate Safety Measures | Medical Evaluation result code
6981716 | 144309112 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO POWER ERROR ALARM NOTED DURING TESTING OR IN THE DEVICE TRACE DOWNLOAD DEVICE FUNCTIONING PROPERLY DURING TESTING | Electrical Power Problem | Device problem code
6981716 | 144309112 | N | DEVICE PASSED SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT SELF TEST AND DISPLACEMENT TEST THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE AND LOADED VOLTAGE WAS WITHIN SPECIFIC RANGE NO POWER ERROR ALARM NOTED DURING TESTING OR IN THE DEVICE TRACE DOWNLOAD DEVICE FUNCTIONING PROPERLY DURING TESTING | No Device Problem Found | Medical Evaluation result code
6973797 | 144309412 | N | DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6973797 | 144309412 | N | DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
6756817 | 144525750 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Component Missing | Device problem code
6756817 | 144525750 | N | DEVICE RECEIVED WITH MISSING RETAINER AND MISSING RESERVOIR TUBE ORING THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO MISSING RETAINER AND MISSING RESERVOIR TUBE ORING | Missing or Inadequate Safety Measures | Medical Evaluation result code
6972309 | 144525924 | N | DEVICE PASSED THE REWIND TEST PRIME TEST BASIC OCCLUSION TEST OCCLUSION TEST FORCE SENSOR TEST AND DISPLACEMENT TEST NO UNEXPECTED DRIVE COUNT OR TIME VALUES OR CHANGES INCORRECT FOR MOVING OR COASTING ERROR OR POSTRESET RAM CRC ALARM NOTED DURING TESTING THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
6973882 | 144526050 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6973882 | 144526050 | N | DEVICE PASSED THE SLEEP CURRENT MEASUREMENT ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR WAS TRIGGERED WHEN THE LOADED VOLTAGE WAS LESS THAN VOLTS FOR FOUR CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
6891842 | 144527694 | N | DEVICE PASSED REWIND TEST BASIC OCCLUSION TEST OCCLUSION TEST PRIME TEST EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST NO UNEXPECTED PRIME OR FILL ANOMALY NOTED DURING TESTING DEVICE WAS PROGRAMMED WITH TEST MINI LINK AND THE GLUCOSE SENSOR SIMULATOR DEVICE COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR THE SENSOR FEATURE WORKING PROPERLY NO BAD SENSOR OR LOST SENSOR ALARMS NOTED | No Device Problem Found | Medical Evaluation result code
6788660 | 144531943 | N | INSULIN PUMP PASSED DISPLACEMENT TEST SELF TEST FORCE SENSOR TEST PRIME SEATING TEST BASIC OCCLUSION TEST OCCLUSION TEST AND REWIND TEST NO DRIVE COUNT OR TIME VALUES OR CHANGES INCORRECT FOR MOVING OR COASTING TOO SLOW OR TOO FAST AND INTERRUPT SOURCE ERROR NOTED DURING TESTING THE MOTOR WAS TESTED OUTSIDE OF THE DEVICE AND PASSED | No Device Problem Found | Medical Evaluation result code
6973797 | 144550525 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code
6973797 | 144550525 | N | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD THE INSULIN PUMP WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6834362 | 144654908 | N | DEVICE PASSED THE SELFTEST AND THE DISPLACEMENT TEST NO UNEXPECTED SOFTWARE ERROR ALARMS NOTED DURING TESTING HOWEVER THE FORMATTED HISTORY FILE CONFIRMED SOFTWARE ERROR ALARM DUE TO A SOFTWARE ERROR | No Device Problem Found | Medical Evaluation result code
6907048 | 144654985 | N | DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | Electrical Power Problem | Device problem code
6907048 | 144654985 | N | DEVICE PASSED DISPLACEMENT TEST SLEEP CURRENT TEST AND ACTIVE CURRENT TEST THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD DEVICE WAS MONITORED AND FUNCTIONED PROPERLY | No Device Problem Found | Medical Evaluation result code